- 11D-8.013 Blood Alcohol Permit – Analyst.

(1) The application for a permit to determine the alcohol level of a blood sample shall be made on a form provided by the Department and shall include the following information:

(a) Name and address of applicant;

(b) A copy of state license if licensed, or college transcript;

(c) Name and address of employer and laboratory facility where applicant performs analyses;

(d) Identify at least one Agency for which blood analyses are to be performed pursuant to Chapters 316, 322 and 327, F.S.; and,

(e) A complete description of proposed analytical procedure(s) to be used in determining blood alcohol level.

(2) Qualifications for blood analyst permit – To qualify, the applicant must meet all of the following requirements:

(a) Department approval of analytical procedure(s). All proposed analytical procedures will be reviewed and a determination of approval will be made by the Department;

(b) Satisfactory determination of blood alcohol level in five proficiency samples provided by the Department using the proposed analytical procedure. Satisfactory determination shall be made by reporting results for blood alcohol proficiency samples within the acceptable range for the samples. For blood alcohol testing, acceptable ranges shall mean the calculated proficiency sample mean + or - 3 standard deviations iterated twice. The mean and standard deviations will be calculated using the results reported by the analysts and reference laboratories;

(c) Identify at least one Agency for which blood analyses are to be performed pursuant to Chapters 316, 322 and 327, F.S.; and,

(d) Meet one of the following:

1. Possess a clinical laboratory license in clinical chemistry as a technologist, supervisor or director, under Chapter 483, F.S.; or

2. Be a licensed physician pursuant to Chapter 458, F.S.; or

3. Complete a minimum of 60 semester credit hours or equivalent of college, at least 15 semester hours of which must be in college chemistry.

(3) The department shall approve gas chromatographic analytical procedures and enzymatic analytical procedures based on alcohol dehydrogenase which meet the following requirements:

(a) Includes the approved method used and a description of the method, and the equipment, reagents, standards, and controls used;

(b) Uses commercially-prepared standards and controls certified by the manufacturer, or laboratory-prepared standards and controls verified using gas chromatography against certified standards. For commercially-prepared standards and controls, the manufacturer, lot number and expiration date must be documented for each sample or group of samples being analyzed. For laboratory-prepared standards and controls, date, person preparing the solution, method of preparation and verification must be documented;

(c) A statement of the concentration range over which the procedure is calibrated. The calibration curve must be linear over the stated range;

(d) Uses a new or existing calibration curve. The new calibration curve must be generated using at least three (3) standards: one at 0.05 g/100mL or less, one between 0.05 and 0.20 g/100mL (inclusive) and one at 0.20 g/100mL or higher, and must be verified using a minimum of two (2) controls, one at 0.05 g/100mL or less and one at 0.20g/100mL or higher. The existing calibration curve must be verified using a minimum of two (2) controls, one at 0.05 g/100mL or less and one at 0.20g/100mL or higher;

(e) Includes the analysis of an alcohol-free control, and the analysis of a whole blood or serum control. The whole blood or serum control may be used to satisfy the control requirement(s) in paragraph (d);

(f) A gas chromatographic analytical procedure must discriminate between methanol, ethanol, acetone and isopropanol and employ an internal standard technique;

(g) An enzymatic analytical procedure based on alcohol dehydrogenase must use the procedure recommended by the instrument manufacturer/test kit vendor for whole blood alcohol analysis, and the enzyme used must have sufficient selectivity to provide negligible cross-reactivity towards methanol, acetone and isopropanol.

(4) The permit shall be issued by the Department for a specific method and procedure. Any substantial change to the method, analytical procedure, or laboratory facility must receive prior approval by the Department before being used to determine the blood alcohol level of a sample submitted by an agency. The Department shall determine what constitutes a substantial change.

(5) An analyst shall only use a Department-approved procedure to determine the blood alcohol level of samples submitted by an agency. Approval of blood alcohol analysis methods and procedures shall be based on rule requirements in effect at the time they were submitted for approval.

Rulemaking Authority 316.1932(1)(a)2., (f)1., 316.1933(2)(b), 316.1934(3) 322.63(3)(b), 327.352(1)(b)3. FS. Law Implemented 316.1932(1)(b), 316.1933(2)(b), 316.1934(3), 322.63(3)(b), 327.352(1)(b), (e), 327.353(2), 327.354(3) FS. History–New 10-31-93, Amended 4-1-94, 2-1-95, 1-1-97, 11-5-02, 12-9-04.