## 25-30.264 Meter Testing Equipment.

(1)(a) Each utility providing metered water service shall either provide the necessary standard facilities, instruments and other equipment for testing meters in compliance with Rule 25-30.263, F.A.C., or enter into arrangements with other utilities or agencies for the testing of the utility’s meters.

(b) When the utility opts to arrange for its testing to be performed by another utility or agency, that utility shall notify the Commission.

(2)(a) Standard meters may be used by the utility for field tests of meter accuracy provided that they are tested and calibrated to permit the testing of meters within the limits of accuracy set out in Rule 25-30.263, F.A.C.

(b) Testing and calibration of the standard meters shall be done either by the utility with its volumetric or weight standard equipment, or by an approved laboratory.

(c) Testing and calibration of the standard meters shall be done at least once every sixty (60) days while the standard meter is in use.

Rulemaking Authority 350.127(2), 367.121(1) FS. Law Implemented 367.121(1)(a), 367.122(1) FS. History–New 9-12-74, Formerly 25-10.89, 25-10.089, Amended 11-10-86.