- 59A-7.029 General Quality Control Requirements for Non-waived Testing.

(1) The laboratory shall establish and follow written quality control procedures for monitoring and evaluating the quality of the testing process of each method to assure the accuracy and reliability of patient test results and reports in accordance with CLIA requirements. The laboratory shall follow the manufacturers’ instructions and recommendations for instrument or test system operation and test performance if such instructions exceed requirements specified in this rule. In the event of a conflict between these rules and CLIA requirements, the more stringent requirement(s) shall prevail.

(2) The laboratory must utilize test methods, equipment, instrumentation, reagents, materials, and supplies that provide accurate and reliable test results and test reports as required by CLIA.

(a) Methodologies and equipment must be selected and testing must be performed in a manner that provides test results within the laboratory’s stated performance specifications for each test method and reflect procedures that are generally accepted by leading authorities such as the Centers for Disease Control and Prevention (CDC), CLIA recognized accreditation organizations, the American Association of Blood Banks (AABB) or other nationally recognized organizations. Documentation that the test methodologies and equipment meet the requirements of this rule must be maintained by the laboratory and available for review by the Agency.

(b) The laboratory must have equipment, instruments, reagents, materials, and supplies for the type and volume of services provided during the preanalytic, analytic, and postanalytic phases of testing.

(c) All equipment and supplies shall be in good working order, checked and calibrated for the proper performance of tests and services offered in accordance with this rule and CLIA requirements. The laboratory must, at a minimum, follow the manufacturers’ recommendations and instructions for equipment operation and document all such activities required for maintenance and operation of such equipment.

(d) The manufacturers’ instructions and documentation of maintenance and operation of equipment must be maintained by the laboratory and available for review by the Agency.

(e) Out-of-service equipment and supplies shall be clearly labeled to indicate their status.

(f) Expired, substandard or unusable supplies shall be promptly removed from use and clearly labeled to indicate their status. Such supplies shall be isolated from usable supplies until they are removed from the premises.

(g) Procedures must be approved, signed, and dated by the current laboratory director both initially and biennially thereafter.

(3) Quality Control Procedures. In accordance with CLIA requirements and any additional provisions of this rule, the laboratory shall perform control procedures to monitor the ability of the method or test system to give accurate, precise and reliable patient test results.

(a) Quantitative controls shall be of different concentrations that approximate the analytical range of that analyte, e.g., normal and abnormal patient values.

(b) No daily quality control testing is required for those tests listed as Provider-Performed Microscopy tests in 42 CFR 493.19(c)(1-9), provided the laboratory has instituted a quality assessment program containing the elements found in Rule 59A-7.031, F.A.C., Quality Assessment, to verify the accuracy of those tests at least every 6 months.

(c) All control procedures required above shall be documented and available to the Agency upon request.

(d) Use of Equivalent Quality Control (EQC).

1. A laboratory is permitted to use EQC testing pursuant to 42 CFR 493.1256(d) provided that those electronic, procedural or internal controls or combinations thereof are met; and the following requirements are met:

a. The process evaluates each step in the testing process.

b. The process evaluates the potential sources of error.

c. The process evaluation includes specific assessment and documentation of how each step of the testing process is evaluated by the EQC process; and evaluates potential sources of error.

d. The implications of reducing the frequency of the use of external controls and the possibility of providing inaccurate and unreliable test results are evaluated and found acceptable by the clinical consultant and approved in writing by the laboratory director.

e. The choice of EQC options described in 42 CFR 493.1256(d) is consistent with the extent to which the electronic, procedural or internal controls or combinations thereof ensure that the provisions of this rule are met.

f. All EQC studies shall be composed of no less than 20 consecutive different test samples.

g. All EQC evaluations, reevaluations, assessments, actions or other such EQC studies shall be documented and available for review by the Agency.

2. After an acceptable EQC evaluation has been completed in accordance with these rules, the laboratory is permitted to institute EQC in lieu of external quality control requirements of 42 CFR 493.1256(d). However, if any of the following conditions occur, the laboratory shall reinstitute the external quality control provisions of 42 CFR 493.1256(d):

a. A proficiency testing score of less than 80% is obtained for any of the last three proficiency testing events;

b. Personnel competency problems are identified;

c. Major preventive maintenance or replacement of critical parts occurs;

d. Any EQC result that was outside acceptable limits as specified in 42 CFR 493.1256(d); or

e. When there is any indicator that inaccurate, imprecise or unreliable patient testing is being reported.

3. Before EQC can be resumed, the laboratory must repeat the EQC evaluation in accordance with the requirements of this rule. External quality control procedures in accordance with 42 CFR 493.1256(d) shall be performed until the subsequent EQC evaluation meets the requirements of this rule.

4. When an EQC failure occurs, the laboratory’s clinical consultant must examine all patient test results reported during the time that the EQC evaluation was used to determine if there was any clinical impact on the patients tested during that time. Appropriate action shall be taken if such patient impact is found.

Rulemaking Authority 483.051 FS. Law Implemented 483.051 FS. History–New 11-20-94, Amended 6-22-06.