64E-5.628 Permissible Molybdenum 99 Concentration.

(1) A licensee shall not administer a radiopharmaceutical containing more than 0.15 microcurie of molybdenum 99 per millicurie of technetium 99m (5.55 kilobecquerel of molybdenum 99 per 37 megabecquerel of technetium 99m).

(2) A licensee preparing technetium 99m radiopharmaceuticals from molybdenum 99/technetium 99m generators shall measure the molybdenum 99 concentration in each eluate or extract.

(3) A licensee who is required to measure molybdenum concentrations shall retain a record of each measurement for 3 years. The record shall include for each elution or extraction of technetium 99m:

(a) The measured activity of the technetium expressed in millicuries (megabecquerels);

(b) The measured activity of molybdenum expressed in microcuries (kilobecquerels);

(c) The ratio of the measures expressed as microcuries of molybdenum per millicurie of technetium (kilobecquerels of molybdenum per megabecquerel of technetium);

(d) The date of the test; and

(e) The initials of the individual who performed the test.

(4) A licensee shall report immediately to the department each occurrence of molybdenum 99 concentration exceeding the limits specified in subsection 64E-5.628(1), F.A.C.

Specific Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.736.