64E-5.637 Doors, Interlocks, and Warning Systems.

(1) A licensee shall control access to the teletherapy room by a door at each entrance.

(2) A licensee shall equip each entrance to the teletherapy room with an electrical interlock system that shall:

(a) Prevent the operator from turning on the primary beam of radiation unless each treatment room entrance door is closed;

(b) Turn off the beam of radiation immediately when an entrance door is opened; and

(c) Prevent the primary beam of radiation from being turned on following an interlock interruption until all treatment room entrance doors are closed and the beam on-off control is reset at the console.

(3) A licensee shall equip each entrance to the teletherapy room with a conspicuously visible beam condition indicator light.

Specific Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.755.