64E-5.642 Periodic Spot-Checks.

(1) A licensee authorized to use teletherapy units for medical use shall perform output spot-checks on each teletherapy unit at least every month.

(2) Spot-checks shall include the determination of:

(a) Timer constancy and timer linearity over the range of use;

(b) On-off error;

(c) The coincidence of the radiation field and the field indicated by the light beam localizing device;

(d) The accuracy of all distance measuring and localization devices used for medical use;

(e) The output for one typical set of operating conditions; and

(f) The difference between the measurement made in paragraph 64E-5.642(2)(e), F.A.C., and the anticipated output, expressed as a percentage of the anticipated output, which is the value obtained at the last full calibration corrected mathematically for physical decay.

(3) A licensee shall use the dosimetry system described in Rule 64E-5.640, F.A.C., to make the spot-check required in paragraph 64E-5.642(2)(e), F.A.C.

(4) A licensee shall perform spot-checks required by subsection 64E-5.642(1), F.A.C., following procedures established by the teletherapy physicist.

(5) A licensee shall have the teletherapy physicist review the results of each output spot-check within 15 days and promptly notify the licensee in writing of the results of each output spot-check. The licensee shall keep a copy of each written notification for 3 years.

(6) A licensee authorized to use a teletherapy unit for medical use shall perform safety spot-checks of each teletherapy facility monthly.

(7) Safety spot-checks shall assure proper operation of:

(a) Electrical interlocks at each teletherapy room entrance;

(b) Electrical or mechanical stops installed to limit use of the primary beam of radiation, restriction of source housing angulation or elevation, carriage or stand travel, and operation of the beam on-off mechanism;

(c) Beam condition indicator lights on the teletherapy unit, on the control console, and in the facility;

(d) Viewing systems;

(e) Treatment room doors from inside and outside the treatment room; and

(f) Electrically assisted treatment room doors with the teletherapy unit electrical power turned off.

(8) A licensee shall lock the control console in the off position if any door interlock malfunctions. No licensee shall use the unit until the interlock system is repaired unless specifically authorized by the department.

(9) A licensee shall promptly repair any system identified in subsection 64E-5.642(7), F.A.C., that is not operating properly. The teletherapy unit shall not be used until all repairs are completed.

(10) A licensee shall maintain a record of each spot-check required by subsections 64E-5.642(1) and (6), F.A.C., for 3 years. The record shall include:

(a) The date of the spot-check;

(b) The manufacturer’s name, model number, and serial number for both the teletherapy unit and source;

(c) The manufacturer’s name, model number and serial number of the instrument used to measure the output of the teletherapy unit;

(d) The timer linearity and constancy;

(e) The calculated on-off error;

(f) A determination of the coincidence of the radiation field and the field indicated by the light beam localizing device;

(g) The determined accuracy of each distance measuring or localization device;

(h) The difference between the anticipated output and the measured output;

(i) Notations indicating the operability of each entrance door electrical interlock, each electrical or mechanical stop, each beam condition indicator light, the viewing system and doors; and

(j) The signature of the individual who performed the periodic spot-check.

Specific Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.761.