- 64E-5.807 Warning Devices.

(1) All locations designated as high radiation areas, and all entrances to such locations shall be equipped with easily observable warning lights that operate when and only when radiation is being produced.

(2) Except in facilities designed for human exposure, each high radiation area shall have an audible warning device which shall be activated for 15 seconds prior to the possible creation of such high radiation area. Such warning device shall be clearly discernible in all high radiation areas and in any adjacent radiation areas.

(3) Barriers, temporary or otherwise, and pathways leading to high radiation areas shall be identified in accordance with Rules 64E-5.322 and 64E-5.323, F.A.C.

Rulemaking Authority 404.022, 404.051 FS. Law Implemented 404.022(1), 404.051(1), (4) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.909.