- 64E-5.1410 Control of Source Movement.

(1) The mechanism which moves the sources of a panoramic irradiator must require a key to operate. Operation of the mechanism must cause an audible signal to indicate that the sources are leaving the shielded position. Only one key shall be available for use at any time and only operators or facility management can possess it. The key must be attached to a portable radiation survey meter by a chain or cable. The lock must be designed so that the key cannot be removed if the source is in an unshielded position. The door to the radiation room must require the same key. The key must be in the possession of a person entering the radiation room.

(2) The console of a panoramic irradiator must have a source position indicator which indicates when the sources are in the fully shielded position, when they are in transit, and when they are exposed.

(3) The control console of a panoramic irradiator must have a control which promptly returns the sources to the shielded position.

(4) Each control for a panoramic irradiator must be labeled clearly as to its function.

(5) Controls for a panoramic irradiator must be color-coded or illuminated as follows:

(a) Red represents emergency or critical information;

(b) Yellow or orange represents caution, no emergency but some function taking place to be aware of; and,

(c) Green or blue represents normal or safe functioning or information.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1510.