Notice of Proposed Rule

DEPARTMENT OF LAW ENFORCEMENT

Division of Local Law Enforcement Assistance

RULE NOS.:	RULE TITLES:

11D-8.003	Approval of Breath Test Methods and Instruments

11D-8.0035	Approval of Alcohol Reference Solution and Sources

PURPOSE AND EFFECT: 11D-6.001: Updating rule to include Intoxilizer 9000.

11D-8.0035: Updating rule regarding alcohol reference solutions

SUMMARY: Revises rules for inclusion of Intoxilizer 9000 and revisions to alcohol reference solutions.

SUMMARY OF STATEMENT OF ESTIMATED REGULATORY COSTS AND LEGISLATIVE RATIFICATION:

The Agency has determined that this will not have an adverse impact on small business or likely increase directly or indirectly regulatory costs in excess of $200,000 in the aggregate within one year after the implementation of the rule. A SERC has not been prepared by the Agency.

Any person who wishes to provide information regarding a statement of estimated regulatory costs, or provide a proposal for a lower cost regulatory alternative must do so in writing within 21 days of this notice.

RULEMAKING AUTHORITY: 943.03(4), 316.1932(1)(a)2., s. 322.63(3)(a), and s.327.352(1)(b)3.,

LAW IMPLEMENTED: 316.1932(1)(a)2., s. 322.63(3)(a), and s.327.352(1)(b)3.

IF REQUESTED WITHIN 21 DAYS OF THE DATE OF THIS NOTICE, A HEARING WILL BE HELD AT THE DATE,TIME AND PLACE SHOWN BELOW(IF NOT REQUESTED, THIS HEARING WILL NOT BE HELD):

DATE AND TIME: Monday April 22, 2024 at 10:00 a.m.

PLACE: Florida Department of Law Enforcement, Criminal Justice Professionalism, 2331 Phillips Road, Tallahassee, Florida 32308.

Pursuant to the provisions of the Americans with Disabilities Act, any person requiring special accommodations to participate in this workshop/meeting is asked to advise the agency at least 5 days before the workshop/meeting by contacting: Chris Bufano at (850)410-7676, or christopherbufano@fdle.state.fl.us, or write to Florida Department of Law Enforcement, Office of General Counsel, 2331 Phillips Road, Tallahassee, Florida 32308.. If you are hearing or speech impaired, please contact the agency using the Florida Relay Service, 1(800)955-8771 (TDD) or 1(800)955-8770 (Voice).

THE PERSON TO BE CONTACTED REGARDING THE PROPOSED RULE IS: Chris Bufano at (850)410-7676, or christopherbufano@fdle.state.fl.us, or write to Florida Department of Law Enforcement, Office of General Counsel, 2331 Phillips Road, Tallahassee, Florida 32308.

THE FULL TEXT OF THE PROPOSED RULE IS:

- 11D-8.003 Approval of Breath Test Methods and Instruments.

(1) The approved breath test method for evidentiary breath testing is Infrared Spectroscopy, also known as Infrared Light Absorption.

(2) The Department approves breath test methods and new instrumentation to ensure the accuracy and reliability of breath test results. The approved breath test instrument make and model is the CMI, Inc. Intoxilyzer 8000 and Intoxilyzer 9000 using software evaluated by the Department in accordance with Instrument Evaluation Procedures FDLE/ATP Form 34, revised March 2004, effective date July 2015, hereby incorporated by reference, https://www.flrules.org/Gateway/reference.asp?No=Ref-05645. This form may be obtained by contacting the Florida Department of Law Enforcement, Alcohol Testing Program, P.O.Box 1489, Tallahassee, Florida 32302.

(3) A Department inspection performed in accordance with Rule 11D-8.004, F.A.C., validates the approval, accuracy and reliability of an evidentiary breath test instrument.

(4) The Department shall conduct evaluations for approval of new instrumentation under subsection (2) in accordance with Instrument Evaluation Procedures FDLE/ATP Form 34 – Rev. March 2004.

(5) The availability or approval of new instruments, evaluation of software, options or modifications does not negate the approval status of previously approved instruments, or evaluated software, options or modifications. An approved make and model of a breath test instrument remains approved until disapproved by the Department.

Rulemaking Authority 316.1932(1)(a)2., (f)1., 322.63(3)(a), (b), 327.352(1)(b)3. FS. Law Implemented 316.1932(1)(b)2., 316.1934(3), 322.63(3)(b), 327.352(1)(e), 327.354(3) FS. History–New 10-31-93, Amended 1-1-97, 7-29-01, 11-5-02, 12-9-04, 7-29-15, _________.

- 11D-8.0035 Approval of Alcohol Reference Solution and Sources.

(1) The Department shall approve a source of alcohol reference solution for use by agencies in the State of Florida. The source approved by the Department shall be an ISO 17034 or similarly accredited entity that manufactures alcohol reference solutions and meets the following requirements:

(a) The source must prepare alcohol reference solution, and be capable of producing a minimum batch volume of 800 bottles, each containing at least 500 milliliters, to produce the following vapor alcohol concentrations: 0.05 g/210L, 0.08 g/210L, and 0.20g/210L;

(b) The source must have performed and documented tests that demonstrate that the alcohol reference solutions are reliable for at least two years from the date of manufacture.

(2) The Department shall approve each lot of alcohol reference solution prior to distribution for use in Florida.

(a) (2) If an ISO 17034 or similiary accredited provider is not available Tthe Department shall approve each lot prior to distribution in Florida by determineing the alcohol concentration in a minimum of ten (10) sample bottles of each lot of alcohol reference solution using gas chromatography or other scientifically accepted method at the time such lots are submitted for approval. Duplicate analyses will be performed on each sample bottle of alcohol reference solution. If any result falls outside the alcohol reference solution acceptable range, a second set of analyses will be performed. If the second set of results are within the alcohol reference solution acceptable range, the alcohol reference solution shall be approved. If any of the results fall outside the alcohol reference solution acceptable range a second time, the alcohol reference solution shall be disapproved.

(ab) When Tthe Department approved a lot it shall notify the source that the approved lots may be distributed for use in Florida, and shall issue a Certificate of Assurance, FDLE/ATP Form 32 revised March 2001, effective date July 2015, hereby incorporated by reference, https://www.flrules.org/Gateway/reference.asp?No=Ref-05643 This form may be obtained by contacting the Florida Department of Law Enforcement, Alcohol Testing Program, P.O. Box 1489, Tallahassee, Florida 32302.

(3) Alcohol reference solution lots approved by the Department shall be used in agency or Department inspections within two (2) years of the date of manufacture.

Rulemaking Authority 316.1932(1)(a)2., (f)1., 322.63(3)(a), 327.352(1)(b)3., (d) FS. Law Implemented 316.1932(1)(b)2., 316.1934(3), 322.63(3)(b), 327.352(1)(e), 327.354(3) FS. History–New 7-6-99, Amended 7-29-01, 12-9-04, 7-29-15, 9-4-16, __________.

NAME OF PERSON ORIGINATING PROPOSED RULE: Bureau Chief Brett Kirkland

NAME OF AGENCY HEAD WHO APPROVED THE PROPOSED RULE: Governor and Cabinet

DATE PROPOSED RULE APPROVED BY AGENCY HEAD: March 26, 2024

DATE NOTICE OF PROPOSED RULE DEVELOPMENT PUBLISHED IN FAR: October 31, 2023