Title	21—Food	and

Drugs

(This book contains Parts 800 to 1299)

### CHAPTER I—Food and Drug Administration, Department of

Part

### Health and Human Services (Continued)	800

CHAPTER I—FOOD AND DRUG ADMINISTRATION, DEPARTMENT OF HEALTH AND HUMAN SERVICES (CONTINUED)

EDITORIAL  NOTE:  Nomenclature  changes  to  chapter  I  appear  at  59  FR  14366,  Mar.  28,  1994; 68 FR 24879, May 9, 2003; and 69 FR 13717, Mar. 24, 2004.

SUBCHAPTER H—MEDICAL DEVICES

Part	Page

### 800	General  ....................................................................	7

### 801	Labeling   ..................................................................	17

### 803	Medical device reporting  .........................................	49

### 806	Medical devices; reports of corrections and remov-

### als  .........................................................................	66

### 807	Establishment  registration  and  device  listing  for manufacturers and initial importers of devices  ...	70

### 808	Exemptions from Federal preemption of State and   local medical device requirements  .......................	87

### 809	In vitro diagnostic products for human use	97

### 810	Medical device recall authority	106

### 812	Investigational device exemptions	113

### 813	[Reserved]

### 814	Premarket approval of medical devices	132

### 820	Quality system regulation	156

### 821	Medical device tracking requirements	170

### 822	Postmarket surveillance	176

### 830	Unique device identification	184

### 860	Medical device classification procedures	192

### 861	Procedures for performance standards development	205

### 862	Clinical  chemistry  and  clinical  toxicology  devices	209

### 864	Hematology and pathology devices	250

### 866	Immunology and microbiology devices	273

### 868	Anesthesiology devices	327

### 870	Cardiovascular devices	348

### 872	Dental devices	380

### 874	Ear, nose, and throat devices	408

### 876	Gastroenterology-urology devices	423

### 878	General and plastic surgery devices	449

#### 21 CFR Ch. I (4–1–17 Edition)

| Part 880 | General hospital and personal use devices .............. | Page 470 |
| --- | --- | --- |
| 882 | Neurological devices ............................................... | 495 |
| 884 | Obstetrical and gynecological devices .................... | 518 |
| 886 | Ophthalmic devices ................................................. | 544 |
| 888 | Orthopedic devices .................................................. | 571 |
| 890 | Physical medicine devices ....................................... | 598 |
| 892 | Radiology devices .................................................... | 621 |
| 895 | Banned devices ........................................................ | 635 |
| 898 | Performance standard for electrode lead wires and patient cables ....................................................... | 640 |
|  | SUBCHAPTER I—MAMMOGRAPHY QUALITY STANDARDS ACT |  |
| 900 | Mammography ........................................................ | 642 |

SUBCHAPTER  J—RADIOLOGICAL  HEALTH

| 1000 | General .................................................................... | 677 |
| --- | --- | --- |
| 1002 | Records and reports ................................................. | 686 |
| 1003 | Notification of defects or failure to comply ............ | 694 |
| 1004 | Repurchase, repairs, or replacement of electronic |  |
|  | products ............................................................... | 698 |
| 1005 | Importation of electronic products ......................... | 700 |
| 1010 | Performance standards for electronic products: |  |
|  | General ................................................................. | 704 |
| 1020 | Performance standards for ionizing radiation emit- |  |
|  | ting products ........................................................ | 709 |
| 1030 | Performance standards for microwave and radio |  |
|  | frequency emitting products ................................ | 741 |
| 1040 | Performance standards for light-emitting products | 744 |
| 1050 | Performance standards for sonic, infrasonic, and |  |
|  | ultrasonic radiation-emitting products ............... | 768 |

SUBCHAPTER K—TOBACCO PRODUCTS

| 1100 | Tobacco products subject to FDA authority ........... | 772 |
| --- | --- | --- |
| 1105 | General .................................................................... | 773 |
| 1107 1140 | Establishment registration, product listing, and substantial equivalence reports ........................... Cigarettes, smokeless tobacco, and covered to- bacco products ...................................................... | 773 774 |
| 1141 | Cigarette package and advertising warnings .......... | 775 |
| 1143 | Minimum required warning statements .................. | 781 |
| 1150 | User fees .................................................................. | 785 |
| SUBCHAPTER L—REGULATIONS UNDER CERTAIN OTHER ACTS |
|  | ADMINISTERED BY THE FOOD AND DRUG ADMINISTRATION |  |
| 1210 | Regulations under the Federal Import Milk Act .... | 790 |
| 1230 | Regulations under the Federal Caustic Poison Act | 794 |
| 1240 | Control of communicable diseases .......................... | 797 |

#### Food and Drug Administration, HHS

Part	Page

### 1250	Interstate conveyance sanitation	804

### 1251–1269	[Reserved]

### 1270	Human tissue intended for transplantation	809

### 1271	Human cells, tissues, and cellular and tissue-based products	820

### 1272–1299	[Reserved]

SUBCHAPTER H—MEDICAL DEVICES

PART 800—GENERAL

##### Subpart A [Reserved]

Subpart B—Requirements for Specific Medical Devices

Sec.

800.10    Contact lens solutions; sterility.

800.12 Contact lens solutions and tablets; tamper-resistant packaging.

800.20 Patient examination gloves and sur- geons’ gloves; sample plans and  test method  for  leakage  defects;  adulteration.

##### Subpart C—Administrative Practices and Procedures

  - Administrative   detention.    AUTHORITY: 21 U.S.C. 321, 334, 351, 352, 355,

360e, 360i, 360k, 361, 362, 371.

## Subpart A [Reserved]

Subpart B—Requirements for Specific Medical Devices

##### § 800.10 Contact lens solutions; ste- rility.

(a)(1) Informed medical opinion is in agreement that all preparations offered or intended for ophthalmic use, includ- ing contact lens solutions, should be sterile. It is further evident that such preparations purport to be of such pu- rity and quality as to be suitable for safe use in the eye.

    - The Food and Drug Administra- tion concludes that all such prepara- tions, if they are not sterile, fall below their professed standard of purity or quality and may be unsafe. In a state- ment of policy issued on September 1, 1964, the Food and Drug Administra- tion ruled that liquid preparations of- fered or intended for ophthalmic use that are not sterile may be regarded as adulterated within the meaning of sec- tion 501(c) of the Federal Food, Drug, and Cosmetic Act (the act), and, fur- ther, may be deemed misbranded with- in the meaning of section 502(j) of the act. By this regulation, this ruling is applicable to all preparations for oph- thalmic use that are regulated as med- ical devices, i.e., contact lens solu- tions.  By  the  regulation  in  § 200.50  of

this chapter, this ruling  is  applicable to ophthalmic preparations that are regulated as drugs.

    - The containers shall be sterile at the time of filling and closing, and the container or individual carton  shall  be so sealed that the contents cannot be used without destroying the seal. The packaging and labeling of these  solu- tions shall also comply with § 800.12 on tamper-resistant packaging require- ments.

Liquid ophthalmic preparations packed in multiple-dose containers should:

Contain one or more suitable and harmless substances  that  will  inhibit the growth of microorganisms; or

Be so packaged as to volume and type of container and so labeled as to duration of use and with such nec- essary warnings as to afford adequate protection and minimize the hazard of injury resulting from contamination during use.

Eye cups, eye droppers, and other dispensers intended for ophthalmic use should be sterile, and may be regarded as falling below their professed stand- ard of purity or quality if they are not sterile. These articles, which are regu- lated as medical devices unless pack- aged with the drugs with which they are to be used, should be packaged so as to maintain sterility until the pack- age is opened and be labeled, on or within the retail package, so as to af- ford adequate directions and necessary warnings to minimize the hazard of in- jury resulting from contamination dur- ing use.

[47 FR 50455, Nov. 5, 1982]

##### § 800.12   Contact   lens   solutions   and tablets; tamper-resistant packaging.

General. Unless contact lens solu- tions used, for example, to clean, dis- infect, wet, lubricate, rinse, soak, or store contact lenses and salt tablets or other dosage forms to be used to make any such solutions are packaged in tamper-resistant retail packages, there is the opportunity for the malicious adulteration  of  these  products  with

§ 800.12

risks both to individuals who unknow- ingly purchase adulterated products and with loss of  consumer  confidence in the security of the packages of over- the-counter (OTC) health care prod- ucts. The Food and Drug Administra- tion has the authority and responsi- bility under  the  Federal  Food,  Drug, and Cosmetic Act (the act) to establish a uniform national standard for tam- per-resistant packaging of those OTC products vulnerable to malicious adul- teration that will improve the security of OTC packaging and help assure the safety and effectiveness of the products contained therein. A contact lens solu- tion or tablet or other  dosage  form  to be used to make such a solution for re- tail sale that is not packaged in a tam- per-resistant package and labeled in accordance with this section is adulter- ated under section 501 of the act or misbranded under section 502 of the act, or both.

Requirement for tamper-resistant package. Each manufacturer and pack- er who packages for retail sale a prod- uct regulated as a medical device that is a solution intended for use with con- tact lenses, e.g., for cleaning, dis- infecting, wetting, lubricating, rinsing, soaking, or storing contact lenses or tablets or other dosage forms to be used to make any such solution shall package the product in a tamper-resist- ant package, if this product is acces- sible to the public while held for sale. A tamper-resistant package is one hav- ing an indicator or barrier to entry which, if breached or missing, can rea- sonably be expected to provide visible evidence to consumers that tampering has occurred. To reduce the likelihood of substitution of a tamper-resistant feature after tampering, the indicator or barrier to entry is required to be dis- tinctive by design or by the use of an identifying characteristic (e.g., a pat- tern, name, registered trademark, logo, or picture). For purposes of this sec- tion, the term ‘‘distinctive by design’’ means the package cannot be dupli- cated with commonly available mate- rial or through commonly available processes. A tamper-resistant package may involve an immediate-container and closure system or secondary-con- tainer or carton system or any com- bination of systems intended to provide

#### 21 CFR Ch. I (4–1–17 Edition)

a visual indication of package integ- rity. The tamper-resistant  feature shall be designed to and shall remain intact when handled in a reasonable manner during manufacture, distribu- tion, and retail display.

Labeling. Each retail package of a product covered by this section is re- quired to bear a statement that is prominently placed so that consumers are alerted to the tamper-resistant fea- ture of the package. The labeling state- ment is also required to be so placed that it will be unaffected if the tamper- resistant feature of the package is breached or missing. If the tamper-re- sistant feature chosen to meet the re- quirement in paragraph (b) of this sec- tion is one that uses an identifying characteristic, that characteristic is required to be referred to in the label- ing statement. For example, the label- ing statement on a bottle with a shrink band could say ‘‘For your protection, this bottle has an imprinted seal around the neck.’’

Requests for exemptions from pack- aging and labeling requirements. A man- ufacturer or packer may request an ex- emption from the packaging and label- ing requirements of this section. A re- quest for an exemption is  required  to be submitted in the form of a citizen petition under  § 10.30  of  this  chapter and should be clearly identified on the envelope as a ‘‘Request for Exemption from Tamper-resistant Rule.’’ A  peti- tion for an exemption from a require- ment of this section is required to con- tain  the  same  kind  of  information about the  product  as  is  specified  for OTC drugs in § 211.132(d) of this chap- ter.

Products subject to approved pre- market approval applications. Holders of approved premarket approval applica- tions for products subject to this sec- tion are required to submit supple- ments to provide for changes in pack- aging to comply with  the  requirement of paragraph (b)  of this  section unless these changes do not affect the com- position of the container, the torque (tightness) of the container, or the composition of the closure  component in contact with the contents (cap liner or innerseal) as these features are de- scribed in the approved premarket ap- proval   application.   Any   supplemental

#### Food and Drug Administration, HHS	§ 800.20

premarket approval application under this paragraph is required to include data sufficient to show that these changes do not adversely affect the product.

Effective date. Each product sub- ject to this section is required to com- ply with the requirements of this sec- tion on the dates listed below except to the extent that a product’s manufac- turer or packer has obtained an exemp- tion from a packaging or labeling re- quirement:

  - Initial effective date for packaging requirements. (i) The packaging require- ment in paragraph (b) of this section is effective on February 7, 1983 for each contact lens solution packaged for re- tail sale on or after that  date,  except for the requirement in paragraph (b) of this section for a  distinctive  indicator or barrier to entry.

(ii) The packaging requirement in paragraph (b) of this section is effec- tive on May 5, 1983 for each tablet that is to be used to make a contact lens so- lution and that is packaged for retail sale on or after that date.

  - Initial effective date for labeling re- quirements. The requirement in para- graph (b) of this section that the indi- cator or barrier to entry be distinctive by design and the requirement in para- graph (c) of this section for a labeling statement are effective on May 5, 1983 for each product subject to this section packaged for  retail  sale  on  or  after that date, except that the requirement for a specific label reference to any identifying characteristic  is  effective on February 6, 1984 for each affected product subject to this section pack- aged for retail sale on or after that date.

  - Retail level effective date. The tam- per-resistant packaging requirement of paragraph (b) of this section is effec- tive on February 6, 1984 for each prod- uct subject to this section that  is  held for sale at retail level on or after  that date that was packaged for retail sale before May 5, 1983. This does not in- clude the requirement in  paragraph  (b) of this section that the indicator or barrier to entry  be  distinctive  by  de- sign. Products packaged for retail sale after May 5, 1983, are required to be in compliance  with  all  aspects  of  the  reg-

ulations  without  regard  to  the  retail level effective date.

[47 FR 50455, Nov. 5, 1982; 48 FR 1706, Jan. 14,

1983, as amended at 48 FR 16666, Apr. 19, 1983;

48 FR 37625, Aug. 19, 1983; 53 FR 11252, Apr. 6,

1988; 73 FR 34859, June 19, 2008]

EFFECTIVE DATE NOTE: A document pub- lished at 48 FR 41579, Sept. 16, 1983, stayed the effective date of § 800.12(f)(3) until further notice.

##### § 800.20 Patient  examination  gloves and surgeons’ gloves; sample plans and test method for leakage defects; adulteration.

Purpose. The prevalence of human immunodeficiency virus (HIV), which causes acquired immune deficiency syndrome (AIDS), and its risk of trans- mission in the health care context, have caused the Food and Drug Admin- istration (FDA) to look more closely at the quality control of barrier devices, such as surgeons’ gloves and patient examination gloves  (collectively known as medical gloves) to reduce the risk of transmission of HIV and other blood-borne infectious diseases. The Centers for Disease Control (CDC) rec- ommend that health care workers wear medical gloves to reduce the risk of transmission of HIV and other blood- borne infectious deseases. The CDC rec- ommends that health care workers wear medical gloves when touching blood or other body fluids, mucous membranes, or nonintact skin of all pa- tients; when handling items or surfaces soiled with blood or other body fluids; and when performing venipuncture and other vascular access procedures. Among other things, CDC’s rec- ommendation that health care pro- viders wear medical gloves dem- onstrates the proposition that devices labeled as medical gloves purport to be and are represented to be effective bar- riers against the transmission of blood- and fluid-borne pathogens. Therefore, FDA, through this regulation, is defin- ing adulteration for patient examina- tion and surgeons’ gloves as a means of assuring safe and effective devices.

  - For a description of a patient ex- amination glove, see § 880.6250. Finger cots, however,  are  excluded  from  the test method and sample plans in para- graphs (b) and (c) of this section.

#### § 800.20

  - For a description of a surgeons’ glove, see § 878.4460 of this chapter.

(b)(1) General test method. For the pur- poses of this part, FDA’s analysis of gloves for leaks and visual defects will be conducted by a visual examination and by a water leak test method, using 1,000 milliliters (ml) of water.

Units examined. Each medical glove will be analyzed independently. When packaged as pairs, each glove is considered separately, and both gloves will be analyzed.

Identification of defects. For this test, defects include leaks detected when tested in accordance with para- graph (b)(3) of this section. A leak is defined as the appearance of water on the outside of the glove. This emer- gence of water from the glove con- stitutes a watertight barrier failure. Other defects include tears, embedded foreign objects, extrusions of glove ma- terial on the exterior or interior sur- face of the glove, gloves that are fused together so that individual glove sepa- ration is impossible, gloves that adhere to each other and tear when separated, or other visual defects that are likely to affect the barrier integrity.

Factors for counting defects. One defect in one glove is counted as one defect. A defect in both gloves in a pair of gloves is counted as two defects. If multiple defects, as defined in para- graph (b)(1)(ii) of this section, are found in one glove, they are counted as one defect. Visual  defects  and  leaks that are observed in the top 40 millime- ters (mm) of a glove will not be count- ed as a defect for the purposes of this part.

Leak test materials. FDA considers the following to be the minimum mate- rials required for this test :

  - A 60 mm by 380 mm (clear) plastic cylinder with a hook on one end and a mark scored 40 mm from the other end (a cylinder of another size may be used if it accommodates both cuff diameter and any water above the glove capac- ity);

  - Elastic strapping with velcro or other fastening material;

  - Automatic water-dispensing ap- paratus or manual device capable of de- livering 1,000 ml of water;

  - Stand with horizontal rod for hanging  the  hook  end  of  the  plastic

#### 21 CFR Ch. I (4–1–17 Edition)

tube. The  horizontal  support  rod  must be capable of holding the weight of the total number of gloves that will be sus- pended at any one time, e.g.,  five gloves suspended will weigh about 5 kilograms  (kg);

  - Timer capable of measuring two minute  intervals.

Visual defects and leak test proce- dures. Examine the sample and identify code/lot number, size, and brand as ap- propriate. Continue the visual exam- ination using the following procedures:

  - Visual defects  examination.  Inspect the gloves for visual defects by care- fully removing the glove from the wrapper, box, or package. Visually ex- amine each glove for defects. As noted in paragraph (b)(1)(iii) of this section, a visual defect observed in the top 40 mm of a glove will not be counted as a de- fect for the purpose of this part. Vis- ually defective gloves do not require further testing, although they must be included in the total number of defec- tive gloves counted for the sample.

  - Leak test set-up. (A) During this procedure, ensure that the exterior of the glove remains dry. Attach the glove to the plastic fill tube by bring- ing the cuff end to the 40 mm mark and fastening with elastic strapping to make a watertight seal.

(B) Add 1,000 ml of room temperature water (i.e., 20 (deg)C to 30 (deg)C) into the open end of the fill tube. The water should pass freely into the glove. (With some larger sizes of long-cuffed sur- geons’ gloves, the water level may reach only the base of the thumb. With some smaller gloves, the water level may extend several inches up the fill tube.)

  - Leak test examination. Imme- diately after adding the water, examine the glove for water leaks. Do not squeeze the glove; use only minimum manipulation to spread the fingers to check for leaks. Water drops may be blotted to confirm leaking.

If the glove does not leak imme- diately, keep the glove/filling tube as- sembly upright and hang the assembly vertically from the horizontal rod, using the wire hook on the open end of the fill tube (do not support the filled glove while transferring).

Make a second observation for leaks   2   minutes   after   the   water   is

#### Food and Drug Administration, HHS	§ 800.20

added to the glove. Use only minimum manipulation of the fingers to check for leaks.

Record the number of defective gloves.

Sampling, inspection, acceptance, and adulteration. In performing the test for leaks and other visual defects de- scribed in paragraph (b) of this section, FDA will collect and inspect samples of medical gloves, and determine when the gloves are acceptable as set out in paragraphs (c)(1) through (c)(3) of this section.

  - Sample plans. FDA will collect samples from lots of medical gloves in accordance with agency sampling plans. These plans are based on sample sizes, levels of sample inspection, and acceptable quality levels (AQLs) found in the International  Standard Organi- zation’s standard ISO 2859, ‘‘Sampling Procedures For Inspection By At- tributes.’’

  - Sample sizes, inspection levels, and minimum   AQLs.   FDA   will   use   single

normal sampling for lots of 1,200 gloves or less and multiple normal sampling for all larger lots. FDA will use general inspection level II in determining the sample size for any lot size. As shown in the tables following paragraph (c)(3) of this  section,  FDA  considers  a  1.5 AQL to be the minimum  level  of  qual- ity acceptable for surgeons’ gloves and a 2.5  AQL  to  be the minimum level of quality acceptable for patient exam- ination  gloves.

  - Adulteration levels and accept/reject criteria. FDA considers a lot of medical gloves to be adulterated when the num- ber of defective gloves found in the tested sample meets or exceeds the ap- plicable rejection number at the 1.5 AQL for surgeons’ gloves or the 2.5 AQL for patient examination gloves. These acceptance and rejection num- bers are identified in the tables fol- lowing paragraph (c)(3) of this section as follows:

ACCEPT/REJECT CRITERIA AT 1.5 AQL FOR SURGEONS’ GLOVES

| Lot Size | Sample | Sample Size | Number Examined | Number Defective |
| --- | --- | --- | --- | --- |
|  |  |  |  | Accept | Reject |
| 8 to 90 | Single sample |  | 8 | 0 | 1 |
| 91 to 280 | Single sample |  | 32 | 1 | 2 |
| 281 to 500 | Single sample |  | 50 | 2 | 3 |
| 501 to 1,200 | Single sample |  | 80 | 3 | 4 |
| 1,201 to 3,200 | First | 32 | 32 | — | 4 |
|  | Second | 32 | 64 | 1 | 5 |
|  | Third | 32 | 96 | 2 | 6 |
|  | Fourth | 32 | 128 | 3 | 7 |
|  | Fifth | 32 | 160 | 5 | 8 |
|  | Sixth | 32 | 192 | 7 | 9 |
|  | Seventh | 32 | 224 | 9 | 10 |
| 3,201 to 10,000 | First | 50 | 50 | 0 | 4 |
|  | Second | 50 | 100 | 1 | 6 |
|  | Third | 50 | 150 | 3 | 8 |
|  | Fourth | 50 | 200 | 5 | 10 |
|  | Fifth | 50 | 250 | 7 | 11 |
|  | Sixth | 50 | 300 | 10 | 12 |
|  | Seventh | 50 | 350 | 13 | 14 |
| 10,001 to 35,000 | First | 80 | 80 | 0 | 5 |
|  | Second | 80 | 160 | 3 | 8 |
|  | Third | 80 | 240 | 6 | 10 |
|  | Fourth | 80 | 320 | 8 | 13 |
|  | Fifth | 80 | 400 | 11 | 15 |
|  | Sixth | 80 | 480 | 14 | 17 |
|  | Seventh | 80 | 560 | 18 | 19 |
| 35,000 | First | 125 | 125 | 1 | 7 |
|  | Second | 125 | 250 | 4 | 10 |
|  | Third | 125 | 375 | 8 | 13 |
|  | Fourth | 125 | 500 | 12 | 17 |
|  | Fifth | 125 | 625 | 17 | 20 |

#### § 800.20

21 CFR Ch. I (4–1–17 Edition)

ACCEPT/REJECT CRITERIA AT 1.5 AQL FOR SURGEONS’ GLOVES—Continued

| Lot Size | Sample | Sample Size | Number Examined | Number Defective |
| --- | --- | --- | --- | --- |
|  |  |  |  | Accept | Reject |
|  | Sixth Seventh | 125 125 | 750 875 | 21 25 | 23 26 |

ACCEPT/REJECT CRITERIA AT 2.5 AQL FOR PATIENT EXAMINATION GLOVES

| Lot Size | Sample | Sample Size | Number Examined | Number Defective |
| --- | --- | --- | --- | --- |
|  |  |  |  | Accept | Reject |
| 5 to 50 | Single sample |  | 5 | 0 | 1 |
| 51 to 150 | Single sample |  | 20 | 1 | 2 |
| 151 to 280 | Single sample |  | 32 | 2 | 3 |
| 281 to 500 | Single sample |  | 50 | 3 | 4 |
| 501 to 1,200 | Single sample |  | 80 | 5 | 6 |
| 1,201 to 3,200 | First | 32 | 32 | 0 | 4 |
|  | Second | 32 | 64 | 1 | 6 |
|  | Third | 32 | 96 | 3 | 8 |
|  | Fourth | 32 | 128 | 5 | 10 |
|  | Fifth | 32 | 160 | 7 | 11 |
|  | Sixth | 32 | 192 | 10 | 12 |
|  | Seventh | 32 | 224 | 13 | 14 |
| 3,201 to 10,000 | First | 50 | 50 | 0 | 5 |
|  | Second | 50 | 100 | 3 | 8 |
|  | Third | 50 | 150 | 6 | 10 |
|  | Fourth | 50 | 200 | 8 | 13 |
|  | Fifth | 50 | 250 | 11 | 15 |
|  | Sixth | 50 | 300 | 14 | 17 |
|  | Seventh | 50 | 350 | 18 | 19 |
| 10,001 to 35,000 | First | 80 | 80 | 1 | 7 |
|  | Second | 80 | 160 | 4 | 10 |
|  | Third | 80 | 240 | 8 | 13 |
|  | Fourth | 80 | 320 | 12 | 17 |
|  | Fifth | 80 | 400 | 17 | 20 |
|  | Sixth | 80 | 480 | 21 | 23 |
|  | Seventh | 80 | 560 | 25 | 26 |
| 35,000 and above | First | 125 | 125 | 2 | 9 |
|  | Second | 125 | 250 | 7 | 14 |
|  | Third | 125 | 375 | 13 | 19 |
|  | Fourth | 125 | 500 | 19 | 25 |
|  | Fifth | 125 | 625 | 25 | 29 |
|  | Sixth | 125 | 750 | 31 | 33 |
|  | Seventh | 125 | 875 | 37 | 38 |

Compliance. Lots of  gloves  that are sampled, tested, and rejected using procedures in paragraphs (b) and (c) of this section, are considered adulterated within the meaning of section 501(c) of the  act.

  - Detention and seizure. Lots of gloves that are adulterated under sec- tion 501(c) of the act are subject to ad- ministrative  and  judicial  action,   such as detention of imported products and seizure of domestic products.

  - Reconditioning. FDA may author- ize the owner of the product, or the owner’s representative, to attempt to recondition, i.e., bring into compliance with the act, a lot or part of a lot of foreign gloves detained at importation, or a lot or part of a lot of seized domes- tic gloves.

    - Modified sampling, inspection, and acceptance. If FDA authorizes recondi- tioning of a lot or portion of a lot of adulterated gloves, testing to confirm that   the   reconditioned   gloves   meet

#### Food and Drug Administration, HHS	§ 800.55

AQLs  must  be  performed  by  an  inde- pendent  testing  facility.  The  following tightened  sampling  plan  must  be  fol- lowed, as described in ISO 2859 ‘‘Sam- pling  Procedures  for  Inspection  by  At- tributes:’’

      - General inspection level II,

      - Single sampling plans for tight- ened inspection,

      - 1.5 AQL for surgeons’ gloves, and

      - 2.5 AQL for patient examination gloves.

    - Adulteration levels and acceptance criteria for reconditioned gloves. (A) FDA considers a lot or part of a lot of adul- terated gloves, that is reconditioned in

accordance with paragraph (d)(2)(i) of this section, to be acceptable when the number of defective gloves found in the tested sample does not exceed the ac- ceptance number in the appropriate ta- bles in paragraph (d)(2)(ii)(B) of this section for reconditioned surgeons’ gloves or patient examination gloves.

FDA considers a reconditioned lot of medical gloves to be adulterated within the meaning of section 501(c) of the act when the number of defective gloves found in the tested sample meets or exceeds the applicable rejec- tion number in the tables following paragraph (d)(2)(ii)(B) of this section:

ACCEPT/REJECT CRITERIA AT 1.5 AQL FOR RECONDITIONED SURGEONS’ GLOVES

| Lot Size | Sample | Sample Size | Number Defective |
| --- | --- | --- | --- |
|  |  |  | Accept | Reject |
| 13 to 90 | Single sample | 13 | 0 | 1 |
| 91 to 500 | Single sample | 50 | 1 | 2 |
| 501 to 1,200 | Single sample | 80 | 2 | 3 |
| 1,201 to 3,200 | Single sample | 125 | 3 | 4 |
| 3,201 to 10,000 | Single sample | 200 | 5 | 6 |
| 10,001 to 35,000 | Single sample | 315 | 8 | 9 |
| 35,000 and above | Single sample | 500 | 12 | 13 |

ACCEPT/REJECT CRITERIA AT 2.5 AQL FOR RECONDITIONED PATIENT EXAMINATION GLOVES

| Lot Size | Sample | Sample Size | Number Defective |
| --- | --- | --- | --- |
|  |  |  | Accept | Reject |
| 8 to 50 | Single sample | 8 | 0 | 1 |
| 51 to 280 | Single sample | 32 | 1 | 2 |
| 281 to 500 | Single sample | 50 | 2 | 3 |
| 501 to 1,200 | Single sample | 80 | 3 | 4 |
| 1,201 to 3,200 | Single sample | 125 | 5 | 6 |
| 3,201 to 10,000 | Single sample | 200 | 8 | 9 |
| 10,001 to 35,000 | Single sample | 315 | 12 | 13 |
| 35,000 and above | Single sample | 500 | 18 | 19 |

[55 FR 51256, Dec. 12, 1990, as amended at 71

FR 75876, Dec. 19, 2006]

## Subpart C—Administrative Practices and Procedures

##### § 800.55   Administrative detention.

  - General. This section sets forth the procedures for detention of medical devices intended for human use be- lieved to be adulterated or misbranded.

§ 800.55

Administrative detention is intended to protect the public by preventing dis- tribution or use of devices encountered during inspections that may be adul- terated or misbranded, until the Food and Drug Administration (FDA) has had time to consider what action it should take concerning the devices, and to initiate legal action, if appro- priate. Devices that FDA orders de- tained may not be used, moved, al- tered, or tampered with in any manner by any person during the detention pe- riod, except as authorized under para- graph (h) of this section, until FDA terminates the detention order under paragraph (j) of this section, or the de- tention period expires, whichever oc- curs first.

  - Criteria for ordering detention. Ad- ministrative  detention  of  devices  may be ordered in accordance with this sec- tion when an authorized FDA rep- resentative, during an inspection under section 704 of the Federal Food, Drug, and Cosmetic Act (the  act),  has  reason to believe that a device, as defined in section 201(h) of the  act,  is  adulterated or  misbranded.

  - Detention period. The detention is to be for a reasonable period that may not exceed 20 calendar days after the detention order is issued, unless the FDA District Director in whose district the devices are located determines that a greater period is required to seize the devices, to institute injuction pro- ceedings, or to evaluate the need for legal action, in which case the District Director  may  authorize  detention  for

10 additional calendar days. The addi- tional 10-calendar-day detention period may be ordered at the time the deten- tion order is  issued or  at any time thereafter. The entire detention period may not exceed 30 calendar days, ex- cept when the detention period is ex- tended under paragraph (g)(6) of this section. An authorized FDA represent- ative may, in accordance with para- graph (j) of this section, terminate a detention before the expiration of the detention period.

  - Issuance of detention order. (1) The detention order shall be issued in writ- ing, in the form of a detention notice, signed by the authorized FDA rep- resentative who has reason to believe that  the  devices  are  adulterated  or

#### 21 CFR Ch. I (4–1–17 Edition)

misbranded, and issued to the owner, operator, or agent in  charge  of  the place where the devices are located. If the owner or the user of the devices is different from the owner, operator, or agent in charge of the place where the devices are detained, a copy of the de- tention order shall be provided to the owner or user of the devices if the own- er’s or user’s identity can be readily determined.

If detention of devices in a vehicle or other carrier is ordered, a  copy  of the detention order shall be provided to the shipper of record and the owner of the vehicle or other carrier,  if their identities can be readily determined.

The detention order shall include the following information:

  - A statement that the devices iden- tified in the order are detained for the period shown;

  - A brief, general statement of the reasons for the detention;

  - The location of the devices;

  - A statement that these devices are not to be used, moved, altered, or tampered with in any manner during that period, except as permitted under paragraph (h) of this section, without the written permission of an author- ized FDA representative;

  - Identification of the detained de- vices;

  - The detention order number;

  - The date and hour of the deten- tion order;

  - The period of the detention;

  - The text of section 304(g) of the act and paragraph (g) (1) and (2) of this section;

  - A statement that any informal hearing on an appeal of a detention order shall be conducted as a regu- latory hearing under part 16 of this chapter, with certain exceptions de- scribed in paragraph (g)(3) of this sec- tion; and

  - The location and telephone num- ber of the FDA district office and the name of the FDA District Director.

  - Approval of detention order. A de- tention order, before issuance, shall be approved by the FDA  District  Director in whose district the devices are lo- cated. If prior written approval is not feasible, prior oral approval shall be obtained   and   confirmed   by   written

#### Food and Drug Administration, HHS	§ 800.55

memorandum  within  FDA  as  soon  as possible.

  - Labeling or marking a detained de- vice. An FDA representative issuing a detention order under paragraph (d) of this section shall label or mark the de- vices with official FDA tags  that  in- clude the following information:

    - A statement that the devices are detained by the United States Govern- ment in accordance with section 304(g) of the Federal Food, Drug, and Cos- metic Act (21 U.S.C. 334(g)).

    - A statement that the devices shall not be used, moved, altered, or tam- pered with in any manner for the pe- riod shown, without the written per- mission of an authorized  FDA rep- resentative, except as authorized in paragraph (h) of this section.

    - A statement  that  the  violation  of a detention order or the removal or al- teration of the tag is punishable  by fine or imprisonment or both (section 303 of the act, 21 U.S.C. 333).

    - The detention order number, the date and hour of the detention  order, the detention period, and the name of the FDA representative who issued the detention order.

  - Appeal of a detention order. (1) A person who would be entitled to claim the devices, if seized, may appeal a de- tention order. Any appeal shall be sub- mitted in writing to the FDA District Director in whose district the devices are located within 5 working days of receipt of a detention order. If the ap- peal includes a request for an informal hearing, as defined in section 201(x) of the act, the appellant shall request ei- ther that a hearing be held within 5 working days after the  appeal  is  filed or that the hearing be held at a later date, which shall not be later than 20 calendar days after receipt of a deten- tion order.

The appellant of a detention order shall state the ownership or propri- etary interest the appellant has in the detained devices. If the detained de- vices are located at a place other than an establishment owned or operated by the appellant, the appellant shall in- clude documents showing that the ap- pellant would have legitimate author- ity to claim the devices if seized.

Any informal hearing on an ap- peal of a detention order shall be con-

ducted as a regulatory hearing pursu- ant to regulation in accordance with part 16 of this chapter, except that:

The detention order under para- graph (d) of this section, rather than the notice under § 16.22(a) of this chap- ter, provides  notice  of  opportunity  for a hearing under this section and is part of the administrative record of the reg- ulatory hearing under § 16.80(a) of this chapter.

A request for a hearing under this section should be addressed to the FDA District  Director.

The last sentence of § 16.24(e) of this chapter, stating  that  a  hearing may not be required to be held at a time less than 2 working days after re- ceipt of the request for a hearing, does not apply to a hearing under this sec- tion.

Paragraph (g)(4) of this section, rather than § 16.42(a) of this chapter, describes the FDA employees, i.e., Of- fice of Regulatory Affairs Program Di- rectors or other FDA officials senior to an FDA District Director,  who  preside at hearings under this section.

The presiding officer of a regu- latory hearing on an appeal of a deten- tion order, who also shall decide the appeal, shall be an Office of Regulatory Affairs Program Director or  another FDA official senior to an FDA District Director who is permitted by § 16.42(a) of this chapter to preside over the hearing.

If the appellant requests a regu- latory hearing and requests that the hearing be held within 5 working days after the appeal is filed, the presiding officer shall, within 5  working  days, hold the hearing and render a decision affirming or revoking the detention.

If the appellant requests a regu- latory hearing and requests that the hearing be held at a date later than within 5 working days after the appeal is filed, but not later than 20 calendar days after receipt of a detention order, the presiding officer shall hold the hearing at a date agreed upon by FDA and the appellant. The presiding officer shall decide whether to affirm or re- voke the detention within 5 working days after the conclusion of the hear- ing. The detention period extends to the  date  of  the  decision  even  if  the  5-

#### § 800.55

working-day period for making the de- cision extends beyond the otherwise applicable 20-calendar-day or 30-cal- endar-day detention period.

If the appellant appeals the deten- tion order but does not request a regu- latory hearing, the presiding officer shall render a decision on the appeal affirming or revoking the detention within 5 working days  after  the  filing of the appeal.

If the presiding officer affirms a detention order, the devices continue to be detained until FDA terminates the detention under paragraph (j) of this section or the detention period ex- pires, whichever occurs first.

If the presiding officer revokes a detention order, FDA shall terminate the detention under paragraph (j) of this section.

(h)(1) Movement of detained devices. Except as provided in this  paragraph, no person shall move detained devices within or from the place where they have been ordered detained until FDA terminates the detention under para- graph (j) of this section or the deten- tion period expires, whichever occurs first.

If detained devices are not in final form for shipment, the manufacturer may move them within the establish- ment where they are detained to com- plete the work needed to put them in final form. As soon as the devices are moved for this purpose, the individual responsible for their movement shall orally notify the FDA representative who issued the detention order, or an- other responsible district office  offi- cial, of the movement of  the  devices. As soon as the devices are put in final form, they shall be segregated from other devices, and the individual re- sponsible for their movement shall orally notify the FDA representative who issued the detention order, or an- other responsible district office  offi- cial, of their new location. The devices put in final form shall not be moved further without FDA approval.

The FDA representative who issued the detention order, or another responsible district office official, may approve, in writing, the movement of detained devices for any of the fol- lowing purposes:

#### 21 CFR Ch. I (4–1–17 Edition)

To prevent interference with an establishment’s operations or harm to the devices.

To destroy the devices.

To bring the devices into compli- ance.

For any other purpose that the FDA representative who issued the de- tention order, or other responsible dis- trict office official, believes is appro- priate in the case.

If an FDA representative approves the movement of detained  devices under paragraph (h)(3) of this section, the detained devices shall remain seg- regated from other devices and the per- son responsible for their movement shall immediately orally notify the of- ficial who approved the movement of the devices, or another  responsible FDA district office official, of the new location of the detained devices.

Unless otherwise permitted by the FDA representative who is  notified  of, or who approves, the movement of de- vices under this paragraph, the re- quired tags shall accompany the de- vices during and after movement and shall remain with  the  devices  until FDA terminates the detention or the detention period expires, whichever oc- curs  first.

  - Actions involving adulterated or mis- branded devices. If FDA determines that the detained devices, including  any that have  been put  in final  form,  are adulterated or misbranded, or both, it may initiate legal action against the devices or the responsible individuals, or both, or request that the devices be destroyed or otherwise brought into compliance with the act under FDA’s supervision.

Detention termination. If FDA de- cides to terminate a detention or when the detention period expires, whichever occurs first, an FDA representative au- thorized to terminate a detention will issue a detention termination notice releasing the devices to any person who received the original detention order or that person’s  representative and  will remove, or authorize in writing the re- moval of, the required labels or tags.

Recordkeeping requirements. (1) After issuance of a detention order under paragraph (d) of this section, the owner,  operator,  or  agent  in  charge  of

#### Food and Drug Administration, HHS	Pt. 801

any factory, warehouse, other estab- lishment, or consulting laboratory where detained devices are manufac- tured, processed, packed, or held shall have, or establish, and maintain ade- quate records relating to how the de- tained devices may have become adul- terated or misbranded, records on any distribution of the devices before and after the detention period, records on the correlation of any in-process de- tained devices that are put in final form under paragraph (h) of this sec- tion to the completed devices, records of any changes in, or processing of, the devices permitted under the detention order, and records of any other move- ment under paragraph (h) of this sec- tion. Records required under this para- graph shall be provided to the FDA on request for review and copying. Any FDA request for access to records re- quired under this paragraph shall be made at a reasonable time, shall state the reason or purpose for the request, and shall identify to the fullest extent practicable the information or type of information sought in the records to which access is requested.

(2) Records required under this para- graph shall be maintained for a max- imum period of 2 years after the issuance of the detention order or for such other shorter period as FDA di- rects. When FDA terminates the deten- tion or when the detention period ex- pires, whichever occurs first, FDA will advise all persons required under this paragraph to keep records concerning that detention whether further record- keeping is required for  the  remainder of the 2-year, or shorter, period. FDA ordinarily will not require further rec- ordkeeping if the agency  determines that the devices are not adulterated or misbranded or that  recordkeeping  is not necessary to protect the public health, unless the records are required under other regulations in this chapter (e.g., the good manufacturing practice regulation in part 820 of this chapter).

[44 FR 13239, Mar. 9, 1979, as amended at 49

FR  3174,  Jan.  26,  1984;  69  FR  17292,  Apr.  2,

2004;  79  FR  9412,  Feb.  19,  2014;  82  FR  14147,

Mar. 17, 2017]

## PART 801—LABELING

##### Subpart A—General Labeling Provisions

Sec.

801.1 Medical devices; name and place of business of manufacturer, packer or dis- tributor.

  - Definitions.

  - Meaning of intended uses.

  - Medical devices; adequate  directions for use.

  - Medical devices; misleading state- ments.

  - Medical devices; prominence of re- quired label statements; use of symbols in labeling.

  - Medical devices; Spanish-language version of certain required statements.

801.18 Format of dates provided on a med- ical device label.

##### Subpart B—Labeling Requirements for Unique Device Identification

801.20 Label to bear a unique device identi- fier.

801.30 General exceptions from the require- ment for the label of a device to bear a unique  device  identifier.

801.35 Voluntary labeling of a device with a unique  device  identifier.

801.40    Form of a unique device identifier.

801.45 Devices that must be directly marked with  a  unique  device  identifier.

801.50  Labeling  requirements  for   stand- alone   software.

801.55 Request for an exception from or al- ternative to a unique device identifier re- quirement.

801.57 Discontinuation of legacy FDA iden- tification numbers assigned to devices.

##### Subpart C—Labeling Requirements for Over-the-Counter Devices

  - Principal display panel.

  - Statement of identity.

  - Declaration of net quantity of con- tents.

  - Medical devices; warning statements for devices containing or manufactured with chlorofluorocarbons and other class I ozone-depleting substances.

##### Subpart D—Exemptions From Adequate Directions for Use

  - Prescription devices.

  - Retail exemption for prescription de- vices.

801.116 Medical devices having commonly known   directions.

801.119    In vitro diagnostic products.

801.122 Medical devices for processing, re- packing,   or   manufacturing.

§ 801.1

801.125 Medical devices for use  in  teaching, law   enforcement,   research,   and   analysis.

  - Medical devices; expiration of ex- emptions.

  - Exceptions or alternatives to label- ing requirements for medical devices held by the Strategic National Stockpile.

##### Subpart E—Other Exemptions

801.150 Medical  devices;  processing,  label- ing,  or  repacking.

##### Subparts F–G [Reserved]

Subpart H—Special Requirements for Specific Devices

801.405 Labeling of articles intended for lay use in the repairing and/or refitting of dentures.

801.410 Use of impact-resistant lenses in eyeglasses  and  sunglasses.

801.415    Maximum  acceptable  level  of  ozone.

800.417    Chlorofluorocarbon propellants.

801.420 Hearing aid devices; professional and patient    labeling.

801.421 Hearing aid devices; conditions for sale.

801.430 User labeling for  menstrual  tam- pons.

801.433 Warning statements for prescription and restricted device products containing or	manufactured	with chlorofluorocarbons or other ozone-de- pleting  substances.

801.435    User labeling for latex condoms.

801.437 User labeling for devices that con- tain natural rubber.

AUTHORITY: 21 U.S.C. 321, 331, 351, 352, 360d,

360i, 360j, 371, 374.

SOURCE:  41  FR  6896,  Feb.  13,  1976,  unless

otherwise noted.

## Subpart A—General Labeling Provisions

##### § 801.1 Medical devices;  name  and place of business of manufacturer, packer or distributor.

- The label of a device in package form shall specify conspicuously the name and place of business of the man- ufacturer,  packer,  or  distributor.

- The requirement for declaration of the name of the manufacturer, pack- er, or distributor shall be deemed to be satisfied, in the case of a corporation, only by the actual corporate name which may be preceded or followed by the name of the particular division of the corporation. Abbreviations for ‘‘Company,’’  ‘‘Incorporated,’’  etc.,  may

21 CFR Ch. I (4–1–17 Edition)

be used and ‘‘The’’ may be omitted. In the case of an  individual,  partnership, or association, the name under which the business is conducted shall be used.

- Where a device is not manufac- tured by the person whose name ap- pears on the label, the name shall be qualified by a phrase that reveals the connection such person has with such device; such as, ‘‘Manufactured for lll’’, ‘‘Distributed by lllll’’, or any other wording that expresses the facts.

- The statement of the place of business shall include the street ad- dress, city, State, and Zip Code; how- ever, the street address may be omitted if it is shown in a current city direc- tory or telephone directory. The re- quirement for inclusion of the  ZIP Code shall apply only to consumer commodity labels developed or revised after the effective date of this section. In the case of nonconsumer packages, the ZIP Code shall appear on either the label or the labeling (including the in- voice).

- If a person manufactures, packs, or distributes a device at a place other than his principal place of business, the label may state the principal place of business in lieu of the actual place where such device was manufactured or packed or is to be distributed, unless such statement would be misleading.

##### § 801.3   Definitions.

As used in this part:

Automatic identification and data cap- ture (AIDC) means any technology that conveys the unique device identifier or the device identifier of a device in a form that can be entered into an elec- tronic patient record or other com- puter system via an automated proc- ess.

Center Director means the Director of the Center for Devices and Radio- logical Health or the Director of the Center for Biologics Evaluation and Research, depending on which Center has been assigned lead responsibility for the device.

Combination product has the meaning set forth in § 3.2(e) of this chapter.

Convenience kit means two or more different medical devices packaged to- gether for the convenience of the user.

#### Food and Drug Administration, HHS	§ 801.4

Device package means a package that contains a fixed quantity of a par- ticular version or model of a device.

Expiration date means the date by which the label of a device states the device must or should be used.

FDA, we, or us means the Food and Drug Administration.

Finished device means any device or accessory to any device that is suitable for use or capable of functioning.

Global Unique Device Identification Database (GUDID) means the database that serves as a repository of informa- tion to facilitate the identification of medical devices through their distribu- tion and use.

Human cells, tissues, or cellular or tis- sue-based product (HCT/P) regulated as a device  means  an  HCT/P  as  defined  in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

Implantable device means  a  device that is intended to be placed in a sur- gically  or  naturally  formed  cavity  of the human body. A  device  is  regarded as an implantable device for the purpose of this part only if it is intended to re- main implanted continuously for a pe- riod of 30 days or more, unless the Commissioner of Food and Drugs deter- mines otherwise in order to protect human  health.

Label has the meaning set forth in section 201(k) of the Federal Food, Drug, and Cosmetic Act.

Labeler means:

  - Any person who causes a label to be applied to a device with the intent that the device will be commercially distributed without any intended sub- sequent replacement or modification of the label; and

  - Any person who causes the label of a device to be replaced or modified with the intent that the device will be commercially distributed without any subsequent replacement or modifica- tion of the label, except that the addi- tion of the name of, and contact infor- mation for, a person who distributes the device, without making any other changes to the label, is not a modifica- tion for the purposes of determining whether a person is a labeler.

Lot or batch means one finished device or more that consist of a single type, model, class, size, composition, or soft-

ware version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits.

Shipping container means a container used during the shipment or transpor- tation of devices, and whose contents may vary from one shipment to an- other.

Specification means any requirement with which a device must conform.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribution and use  by  meeting  the  requirements  of

§ 830.20 of this chapter. A unique device identifier  is  composed  of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured;

  - For  an  HCT/P  regulated  as  a  de- vice,   the   distinct   identification   code required by § 1271.290(c) of this chapter. Universal  product  code  (UPC)  means the  product  identifier  used  to  identify an  item  sold  at  retail  in  the  United

States.

Version or model means all devices that have specifications, performance, size, and composition, within limits set by the labeler.

[78 FR 55817, Sept. 24, 2013]

§ 801.4   Meaning of intended uses.

The words intended uses or words of similar  import  in  §§ 801.5,  801.119,  and

  - refer to the objective intent of the persons legally responsible for the labeling of devices. The intent is deter- mined by such persons’ expressions or may  be  shown  by  the  circumstances

#### § 801.5

surrounding the distribution of the ar- ticle. This objective intent may, for ex- ample, be shown by labeling claims, ad- vertising matter, or oral or written statements by such persons or their representatives. It may be shown by the circumstances that the article is, with the knowledge of such persons or their representatives, offered and used for a purpose for which it is neither la- beled nor advertised. The intended uses of an article may change after it has been introduced  into interstate com- merce by its manufacturer. If, for ex- ample, a packer, distributor, or seller intends an article for  different  uses than those intended by the person from whom he received the devices,  such packer, distributor, or seller is re- quired to supply adequate labeling in accordance with the new intended uses. But if a manufacturer knows, or has knowledge of facts that would give him notice that a device introduced into interstate commerce by him is to be used for conditions, purposes, or uses other than the ones for which he offers it, he is required to provide adequate labeling for such a device which ac- cords with such other uses to which the article is to be put.

EFFECTIVE  DATE  NOTE: At 82  FR 2217,  Jan.

9, 2017, § 801.4 was revised, eff. Feb. 8, 2017. At

82 FR 9501, Feb. 7, 2017, the effective date was

delayed to Mar. 21, 2017. At 82 FR 14319, Mar. 20, 2017, the effective date was further de- layed to Mar. 19, 2018. For the convenience of the user, the revised text is set forth as fol- lows:

§ 801.4   Meaning of intended uses.

The words intended uses or words of similar import in §§ 801.5, 801.119, 801.122, and 1100.5 of this chapter refer to the objective  intent  of the persons legally responsible for the label- ing of devices. The intent is determined by such persons’ expressions or may be shown by the circumstances surrounding the dis- tribution of  the  article.  This objective  intent may, for example, be shown by labeling claims, advertising matter, or oral or  writ- ten statements by such persons or their rep- resentatives. It may be shown,  for  example, by circumstances in  which  the  article  is, with the knowledge of such persons or their representatives, offered and used for a pur- pose for which it is neither labeled nor ad- vertised. The intended uses of an article may change after it has been introduced into interstate commerce by its manufacturer. If, for example, a packer, distributor, or seller intends   an   article   for   different   uses   than

#### 21 CFR Ch. I (4–1–17 Edition)

those intended by the person from whom he received the device, such packer, distributor, or seller is required to supply adequate label- ing in accordance with  the  new  intended uses. And if the totality of the evidence es- tablishes that a manufacturer objectively in- tends that a device introduced into inter- state commerce by him is to be used for con- ditions, purposes, or uses other than ones for which it has been approved, cleared, granted marketing authorization, or is exempt from premarket  notification  requirements  (if any), he is required, in accordance with sec- tion 502(f) of the Federal Food, Drug, and Cosmetic Act, or, as applicable, duly promul- gated regulations exempting the device from the requirements of section 502(f)(1), to pro- vide for such device adequate labeling that accords with such other intended uses.

##### § 801.5   Medical  devices;  adequate  di- rections for use.

Adequate directions for use means di- rections under which the layman can use a device safely and for the purposes for which it is intended. Section 801.4 defines intended use. Directions for use may be inadequate because, among other reasons, of omission, in whole or in part, or incorrect specification of:

    - Statements of all conditions, pur- poses, or uses for which such device is intended, including conditions, pur- poses, or uses for which it is prescribed, recommended, or suggested in its oral, written, printed, or graphic  adver- tising, and conditions, purposes, or uses for which the device is commonly used; except that such statements shall not refer to conditions, uses, or pur- poses for which the device can be safely used only under the supervision of a practitioner licensed by law and for which it is advertised solely to such practitioner.

    - Quantity of dose, including usual quantities for each of the uses for which it is intended and usual quan- tities for persons of different ages and different physical conditions.

    - Frequency of administration or application.

    - Duration of administration or ap- plication.

    - Time of administration or appli- cation, in relation to time of meals, time of onset of symptoms, or other time factors.

    - Route or method of administra- tion or application.

#### Food and Drug Administration, HHS	§ 801.15

    - Preparation for use, i.e., adjust- ment of temperature, or other manipu- lation or process.

##### § 801.6 Medical devices; misleading statements.

Among representations in the label- ing of a device which render such de- vice misbranded is a false or mis- leading representation with respect to another device or a drug or food or cos- metic.

##### § 801.15  Medical  devices;  prominence of required label statements; use of symbols in labeling.

A word, statement, or other infor- mation required by or under authority of the act to appear on the label  may lack that prominence and conspicuous- ness required by section 502(c) of the act by reason, among other reasons, of:

  - The failure of such word, state- ment, or information to appear on the part or panel of the label which is pre- sented or displayed under customary conditions of purchase;

  - The failure of such word, state- ment, or information to appear on two or more  parts  or  panels  of  the  label, each of which has sufficient space therefor, and each of which is so de- signed as to render it likely to be, under customary conditions of pur- chase, the part or panel displayed;

  - The failure of the label to extend over the area of the container or pack- age available for such extension, so as to provide sufficient label space for the prominent placing of such word, state- ment, or information;

  - Insufficiency of label space for the prominent placing of such word, state- ment, or information, resulting from the use of label space for any word, statement, design, or device which is not required by or under authority of the act to appear on the label;

  - Insufficiency of label space for the placing of such word, statement, or in- formation, resulting from the use  of label space to give materially greater conspicuousness to any other word, statement, or information, or  to any design or device; or

  - Smallness or style of type in which such word, statement, or infor- mation appears, insufficient back- ground contrast, obscuring designs or

vignettes, or crowding with other writ- ten, printed, or graphic matter.

No exemption depending on insuf- ficiency of label space, as prescribed in regulations promulgated under section 502(b) of the act, shall apply if such in- sufficiency is caused by:

  - The use of label space for any word, statement, design, or device which is not required by or under au- thority of the act to appear on the label;

  - The use of label space to give greater conspicuousness to any word, statement, or other  information  than is required by section 502(c) of the act; or

  - The use of label space for any rep- resentation in a foreign language.

(c)(1)(i) All words, statements, and other information required by or under authority of the act to appear on the label or labeling for a device shall ap- pear thereon in one or more of the fol- lowing formats:

    - The English language;

    - In the case of articles distributed solely in Puerto Rico or in a Territory where the predominant language is one other than English, the predominant language may be substituted for English;

    - A symbol accompanied by adja- cent explanatory  English  text, or  text in the predominant language of the Territory, in the case of articles dis- tributed solely in Puerto Rico or in a Territory where the predominant lan- guage is one other than English;

    - A symbol not accompanied by ad- jacent explanatory text that:

      - Is contained in a standard that FDA recognizes under its authority in section 514(c) of the act;

      - Is used according to the specifica- tions for use of the symbol set forth in FDA’s section 514(c) recognition; and

      - Is explained in a paper or elec- tronic symbols glossary that is in- cluded in the labeling for the device and the labeling on or within the pack- age containing the device bears a prominent and conspicuous statement identifying the location of the symbols glossary that is written in English or, in the case of articles distributed sole- ly in Puerto Rico or in a Territory where the predominant language is one

#### § 801.16

other   than   English,   the   predominant language may be used;

    - A symbol not accompanied by ad- jacent explanatory text that:

      - Is established in a standard devel- oped by a standards development orga- nization  (SDO);

      - Is not contained  in  a  standard that is recognized by FDA under its au- thority in section 514(c) of the act or is contained in a standard that is recog- nized by FDA but is not used according to the specifications for use of the sym- bol set forth in FDA’s section 514(c) recognition;

      - Is determined by the manufac- turer to be likely to be read and under- stood by the ordinary individual under customary conditions of purchase and use in compliance with section 502(c) of the  act;

      - Is used according to the specifica- tions for use of the symbol set forth in the SDO-developed standard; and

      - Is explained in a paper or elec- tronic symbols glossary that is in- cluded in the labeling for the device and the labeling on or within the pack- age containing the device bears a prominent and conspicuous statement identifying the location of the symbols glossary that is written in English or, in the case of articles distributed sole- ly in Puerto Rico or in a Territory where the predominant language is one other than English, the predominant language may be used;

    - The symbol statement ‘‘Rx only’’ or ‘‘) only’’ may be used as  provided under § 801.109(b)(1).

    - The use of symbols in device la- beling which  do not  meet the  require- ments of paragraph (c)(1)(i) of this sec- tion renders a device misbranded under section 502(c) of the act.

    - For purposes of paragraph (c)(1)(i) of this section:

An SDO is an organization that is nationally or internationally recog- nized and that follows a process for standard development that is trans- parent, (i.e., open to public scrutiny), where the participation is balanced, where an appeals process is included, where the standard is not in conflict with any statute, regulation, or policy under which FDA operates, and where the standard is national or inter- national in scope.

#### 21 CFR Ch. I (4–1–17 Edition)

The term ‘‘symbols glossary’’ means a compiled listing of:

  - Each SDO-established symbol used in the labeling for the device;

  - The title and designation  number of the SDO-developed standard con- taining the symbol;

  - The title of the symbol and its ref- erence number, if any, in the standard; and

  - The meaning or explanatory text for the symbol as provided in the FDA recognition or, if FDA has not recog- nized the standard or portion of the standard in which the symbol is lo- cated or the symbol is not used accord- ing to the specifications for use of the symbol set forth  in  FDA’s  section 514(c) recognition, the explanatory text as provided in the standard.

If the label contains any represen- tation in a foreign language, all words, statements, and other information re- quired by or under authority of the act to appear on the label shall appear thereon in the foreign language.

If the labeling contains any rep- resentation in a foreign language, all words, statements, and other informa- tion required by or under authority of the act to appear on the label or label- ing shall appear on the labeling in the foreign language.

[41  FR  6896,  Feb.  13,  1976,  as  amended  at  81

FR 38930, June 15, 2016]

##### § 801.16 Medical devices; Spanish-lan- guage version of certain required statements.

If devices restricted to prescription use only are labeled solely in  Spanish for distribution in the  Commonwealth of Puerto Rico where Spanish is the predominant language, such labeling is authorized under § 801.15(c).

##### § 801.18 Format of dates provided on a medical device label.

  - In general. Whenever the label of a medical device includes a printed expi- ration date, date of manufacture, or any other date intended to be brought to the attention of the user of the de- vice, the date must be presented in the following format: The year, using four digits; followed by the month, using two digits; followed by the day, using two digits; each separated by hyphens.

#### Food and Drug Administration, HHS	§ 801.30

For  example,  January  2,  2014,  must  be presented as 2014–01–02.

  - Exceptions. (1) A combination product that properly bears a National Drug Code (NDC) number is not subject to the requirements of paragraph (a) of this  section.

If the device is an electronic prod- uct to which a standard is applicable under subchapter J of this chapter, Ra- diological Health, the date of manufac- ture shall be presented as required by

§ 1010.3(a)(2)(ii) of this chapter.

[78 FR 58818, Sept. 24, 2013]

## Subpart B—Labeling Requirements for Unique Device Identification

##### § 801.20 Label to bear a unique device identifier.

  - In general. (1) The label of every medical device shall bear a unique de- vice identifier (UDI) that meets the re- quirements of this subpart and part 830 of  this  chapter.

(2) Every device package shall bear a UDI that  meets  the  requirements  of this subpart and part 830 of this chap- ter.

  - Exceptions. Exceptions to the  gen- eral  rule  of  paragraph  (a)  of  this  sec- tion are provided by §§ 801.30, 801.45, and 801.128(f)(2), and § 801.55 provides a means to request an exception or alter- native  not  provided  by  those  provi- sions.

[78 FR 55818, Sept. 24, 2013]

##### § 801.30 General exceptions from the requirement for the label of a de- vice to bear a unique device identi- fier.

In general. The following types of devices are excepted from the require- ment of § 801.20; a device within one or more of the following exceptions is not required to bear a unique device identi- fier  (UDI):

  - A finished device manufactured and labeled prior to  the  compliance date established by FDA for § 801.20 re- garding the device. This exception ex- pires with regard to a particular device

3  years  after  the  compliance  date  es- tablished by FDA for the device.

  - A class I device that FDA has by regulation exempted from the good manufacturing  practice  requirements

of part 820 of this chapter, exclusive  of any continuing requirement for record- keeping under §§ 820.180 and 820.198.

  - Individual single-use devices, all of a single version or model, that are distributed together in a single device package, intended to be stored in that device package until removed for use, and which are not intended for indi- vidual commercial distribution. This exception is not available for any implantable device. The device package containing these individual devices is not excepted from the requirement of

§ 801.20, and must bear a UDI.

  - A device used solely for research, teaching, or chemical analysis, and not intended for any clinical use.

  - A custom device within the mean- ing of § 812.3(b) of this chapter.

  - An investigational device  within the meaning of part 812 of this chapter.

  - A veterinary medical device not intended for use in the diagnosis of dis- ease or other conditions in man, in the cure, mitigation, treatment, or preven- tion of disease in man, or intended to affect the structure or any function of the body of man.

  - A device intended for export from the  United  States.

  - A device held by the Strategic Na- tional Stockpile and granted an excep- tion or alternative under § 801.128(f)(2).

  - A device for which FDA has es- tablished  a  performance   standard under section 514(b) of the Federal Food, Drug, and Cosmetic Act and has provided therein an exception from the requirement of § 801.20,  or  for  which FDA has recognized all or part of a per- formance standard under section  514(c) of the Federal Food, Drug,  and  Cos- metic Act and has included an excep- tion from the requirement of § 801.20 within  the  scope  of  that  recognition.

  - A device packaged within the im- mediate container of a combination product or convenience kit, provided that the label of the combination prod- uct or convenience kit bears a UDI.

National Drug Code (NDC) Numbers. If a combination  product  properly bears an NDC number on its label—

  - The combination product is not subject to the requirements of § 801.20.

  - A device constituent of such a combination product whose compo- nents   are   physically,   chemically,   or

§ 801.35

otherwise combined or mixed and pro- duced as a single entity as described by

§ 3.2(e)(1) of this chapter  is  not  subject to the requirements of § 801.20.

  - Each device constituent of such a combination product, other than one described by § 3.2(e)(1) of this chapter, must bear a UDI on its label unless paragraph (a)(11) of this section  ap- plies.

Exception for shipping containers. This rule does not require a UDI to be placed on any shipping container.

The UDI of a class I device is not required to include a production identi- fier.

[78 FR 55818, Sept. 24, 2013]

##### § 801.35   Voluntary labeling of a device with a unique device identifier.

The labeler of a device that is not required to bear a unique device identi- fier (UDI) may voluntarily comply with

§ 801.20. If a labeler voluntarily includes a UDI for a device, the labeler may vol- untarily provide information con- cerning the device under subpart E of part 830 of this chapter.

A device may bear both a Uni- versal Product Code (UPC) and a UDI on its label and packages.

[78 FR 55818, Sept. 24, 2013]

##### § 801.40   Form of a unique device iden- tifier.

Every unique device identifier (UDI) must meet the technical require- ments of § 830.20 of this chapter. The UDI must be presented in two forms:

Easily readable plain-text, and

Automatic identification and data capture (AIDC) technology.

The UDI must include a device identifier segment. Whenever a device label includes a lot or batch number, a serial number, a  manufacturing date, an expiration date, or for a human cell, tissue, or cellular or tissue-based prod- uct (HCT/P) regulated as a device, a distinct identification code as required by § 1271.290(c) of this chapter, the UDI must include a production identifier segment that conveys such informa- tion.

If the AIDC technology is not evi- dent upon visual examination of the label or device package, the label or de-

#### 21 CFR Ch. I (4–1–17 Edition)

vice  package  must  disclose  the  pres- ence of AIDC technology.

A class I device that bears a Uni- versal Product Code (UPC) on its label and device packages is deemed  to  meet all requirements of  subpart  B  of  this part. The UPC will serve as the unique device identifier required by § 801.20.

[78 FR 55818, Sept. 24, 2013]

##### § 801.45 Devices that must be directly marked with a unique device iden- tifier.

In general. A device that must bear a unique device identifier (UDI) on its label must also bear a permanent marking providing the UDI on the de- vice itself if the device is intended to be used more than once and intended to be reprocessed before each use.

UDI for direct marking. The UDI provided through a direct marking on a device may be:

  - Identical to the UDI that appears on the label of the device, or

  - A different  UDI used  to distin- guish the  unpackaged  device  from any device package containing the device.

Form of a UDI when provided as a direct marking. When a device must bear a UDI as a direct marking, the UDI may be provided through either or both of the following:

  - Easily readable plain-text;

  - Automatic identification and data capture (AIDC) technology, or any al- ternative technology, that will provide the UDI of the device on demand.

Exceptions. The requirement of paragraph (a) of this section shall not apply to any device that meets  any  of the  following  criteria:

Any type of direct marking would interfere with the safety or effective- ness of the device;

The device cannot be directly marked because it is not techno- logically feasible;

The device is a single-use device and is subjected to additional proc- essing and manufacturing for the pur- pose of an additional single use.

The device has been previously marked under paragraph (a) of this sec- tion.

Exception to be noted in design his- tory file. A labeler that decides to make use of an exception under paragraph (d of   this   section)   must   document   the

#### Food and Drug Administration, HHS	§ 801.55

basis of that decision in the design his- tory file required by § 820.30(j) of this chapter.

[78 FR 55818, Sept. 24, 2013]

##### § 801.50 Labeling requirements for stand-alone software.

Stand-alone software that is not distributed in packaged  form  (e.g., when downloaded from a Web site) is deemed to meet the UDI labeling re- quirements of this subpart if it com- plies with the requirements of para- graph (b) of this section and  conveys the version number in its production identifier.

Regardless of whether it is or is not distributed in packaged form, stand-alone software regulated as a medical device must provide its unique device identifier through either or both of the following:

An easily readable plain-text statement displayed whenever the soft- ware is started;

An easily readable plain-text statement displayed through a menu command (e.g., an ‘‘About * * *’’ com- mand).

Stand-alone software that is dis- tributed in both packaged form and in a form that is not packaged (e.g., when downloaded from a Web site) may be identified with the same device identi- fier.

[78 FR 55818, Sept. 24, 2013]

##### § 801.55 Request for an exception from or alternative to a unique device identifier requirement.

A  labeler  may  submit  a   request for an exception from or alternative  to the requirement of § 801.20 or any other requirement of this subpart for a speci- fied device or a specified type of device. A written request for an exception or alternative    must:

  - Identify the device or devices that would be subject to the exception or al- ternative;

  - Identify the provisions of this sub- part that are the subject of the request for an exception or alternative;

  - If requesting an exception, explain why you believe the requirements of this subpart are not technologically feasible;

  - If requesting an alternative, de- scribe  the  alternative  and  explain  why

it would provide for more accurate, precise, or rapid device identification than the requirements of  this  subpart or how the alternative would better en- sure the safety or effectiveness of the device that would be subject to the al- ternative;

  - Provide, if known, the number of labelers and the number of devices that would be affected if we grant the re- quested exception or alternative; and

  - Provide other requested informa- tion that the Center Director needs to clarify the scope  and effects of  the re- quested exception or alternative.

A written request for an exception or alternative must be submitted by sending it:

  - If the device is regulated by the Center for Biologics Evaluation and Research (CBER), by email to: cberudirequests@fda.hhs.gov or by cor- respondence to: Food and Drug Admin- istration, Center for Biologics Evalua- tion and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993.

  - In all other cases, by email to: GUDIDSupport@fda.hhs.gov, or by cor- respondence to: UDI Regulatory Policy Support, Center for Devices and Radio- logical Health, Food and Drug Admin- istration, 10903 New Hampshire Ave., Bldg. 66, Rm. 3303, Silver Spring, MD 20993–0002.

The Center Director may grant an exception or alternative, either in re- sponse to a request or on his or her own initiative, if the Center Director deter- mines that an exception is appropriate because the requirements of this sub- part are not technologically  feasible, or that an alternative would provide for more accurate, precise, or rapid de- vice identification than the require- ments of this subpart or would better ensure the safety or effectiveness of the device that would be subject to the alternative. If we grant an exception or alternative, we may include any safe- guards or conditions deemed appro- priate to ensure the adequate identi- fication of the device through its dis- tribution  and  use.  Any  labeler  may

#### § 801.57

make use of an exception or alter- native granted under this section, pro- vided that such use satisfies all safe- guards or  conditions  that are  part  of the exception or alternative.

FDA may initiate and grant an exception or alternative if we deter- mine that the exception or alternative is in the best interest of the public health. Any such exception or alter- native will remain in effect only so long as there remains a public health need for the exception or alternative.

The  Center  Director  may  rescind an exception or alternative  granted under this section if,  after providing an opportunity for an informal hearing as defined in section 201(x) of the Federal Food, Drug,  and  Cosmetic  Act  and under part 16 of this chapter,  the  Cen- ter Director determines that the excep- tion  or  alternative  no  longer  satisfies the  criteria  described  in  this  paragraph

or that any safeguard or condition required under this paragraph (e) has not been met.

[78 FR 58818, Sept. 24, 2013, as amended at 80

FR  18093,  Apr.  3,  2015;  81  FR  11428,  Mar.  4,

2016]

##### § 801.57 Discontinuation of legacy FDA identification numbers assigned to devices.

  - On the date your device must bear a unique device identifier (UDI) on its label, any National  Health-Related Item Code (NHRIC) or National Drug Code (NDC) number assigned to that device is rescinded, and you may no longer provide an NHRIC or NDC num- ber on the label of your device or on any device package.

  - If your device is not required to bear a UDI on its label, any NHRIC or NDC number assigned to that device is rescinded as of September 24, 2018, and beginning on that date, you may no longer provide an NHRIC or NDC num- ber of the label of your device  or  on any device package.

  - A labeler who has been  assigned an FDA labeler code to facilitate use of NHRIC or NDC numbers  may  continue to use that labeler code under a system for the issuance of UDIs, provided that—

    - Such use is consistent with the framework of the issuing agency that operates that system; and

#### 21 CFR Ch. I (4–1–17 Edition)

    - No later than September 24, 2014, the labeler submits, and obtains FDA approval of, a request for continued use of the assigned labeler code. A  request for continued use of an assigned labeler code must be submitted by email to: GUDIDSupport@fda.hhs.gov, or by cor- respondence to: UDI Regulatory Policy Support, Center for Devices and Radio- logical Health, Food and Drug Admin- istration, 10903 New Hampshire Ave., Bldg. 66, Rm. 3303, Silver Spring, MD 20993–0002.

  - Each request for continued use of an assigned labeler code must provide—

    - The name, mailing address, email address, and phone number of the label- er who  is  currently using  the  labeler code;

    - The owner/operator account iden- tification used by the labeler to submit registration and listing information using FDA’s Unified Registration and Listing  System  (FURLS).

    - The FDA labeler code that the la- beler wants to continue using.

[78 FR 55820, Sept. 24, 2013, as amended at 81

FR 11428, Mar. 4, 2016]

## Subpart C—Labeling Require- ments for Over-the-Counter Devices

##### § 801.60   Principal display panel.

The term principal display panel, as it applies to over-the-counter devices in package form and as used in this part, means the part of a label that is most likely to be displayed, presented, shown, or examined under customary conditions of display  for  retail  sale. The principal display panel shall  be large enough to accommodate all the mandatory label information required to be placed thereon by this part with clarity and conspicuousness and with- out obscuring designs, vignettes, or crowding. Where packages bear alter- nate principal display panels, informa- tion required to be placed on the prin- cipal display panel shall be  duplicated on  each  principal  display  panel.  For the purpose of obtaining uniform type size in declaring the quantity of con- tents for all packages of substantially the same size, the term area of the prin- cipal  display  panel  means  the  area  of

Food and Drug Administration, HHS	§ 801.62

the side or surface that bears the prin- cipal display panel, which area shall be:

      - In the case of a rectangular pack- age where one entire side properly can be considered to be the principal dis- play panel side, the product of the height times the width of that side;

      - In the case of a cylindrical or nearly cylindrical container, 40 percent of the product of the height of the con- tainer times the circumference; and

      - In the case of any other shape of container, 40 percent of the total sur- face of the container: Provided, how- ever, That where such container pre- sents an obvious ‘‘principal display panel’’ such as the  top  of  a  triangular or circular package, the area shall con- sist of the entire top surface.

In determining the area of the prin- cipal display panel, exclude tops, bot- toms, flanges at the  tops  and  bottoms of cans, and shoulders and necks of bot- tles or jars. In  the  case  of  cylindrical or nearly cylindrical containers, infor- mation required by this part to appear on the principal display panel shall ap- pear within that 40 percent of the cir- cumference which is most likely to be displayed, presented, shown, or exam- ined under customary conditions of dis- play for retail sale.

##### § 801.61   Statement of identity.

The principal display panel of an over-the-counter device in package form shall bear as one of its principal features a statement of the identity  of the  commodity.

Such statement of identity shall be in terms of the common name of the device followed by an accurate state- ment of the principal intended ac- tion(s) of the device. Such statement shall be placed in direct conjunction with the most prominent display of the name and shall employ terms descrip- tive of the principal intended action(s). The indications for use shall be in- cluded in the directions for use of the device, as required by section 502(f)(1) of the act and by the regulations  in this  part.

The statement of identity shall be presented in bold face type on the prin- cipal display panel, shall be in a size reasonably related to the most promi- nent printed matter on such panel, and

shall be in lines generally parallel  to the base on which the package rests as it is designed to be displayed.

##### § 801.62   Declaration of net quantity of contents.

The label of an over-the-counter device in package form shall bear a declaration of the net quantity of con- tents. This shall be expressed in the terms of weight, measure, numerical count, or a combination of numerical count and weight, measure, or size: Provided, That:

  - In the case of a firmly established general consumer usage and trade cus- tom of declaring the quantity of a de- vice in terms of linear measure or measure of area, such respective term may be used. Such term shall be aug- mented when necessary for accuracy of information by a statement of the weight, measure, or size of the indi- vidual units or of the entire device.

  - If the declaration of contents for a device by numerical  count  does  not give accurate information as to the quantity of the  device  in  the  package, it shall be augmented by such state- ment of weight, measure, or size of the individual units or of the total weight, measure, or size of the device as will give such information; for example, ‘‘100 tongue depressors, adult size’’, ‘‘1 rectal syringe, adult size’’, etc. When- ever the  Commissioner  determines  for a specific packaged device that an ex- isting practice of declaring net quan- tity of contents by weight, measure, numerical count, or a combination of these does not facilitate value comparisions by consumers, he shall by regulation designate the appropriate term or terms to be used for such arti- cle.

Statements of weight of the con- tents shall be expressed in terms of av- oirdupois pound and ounce. A  state- ment of liquid measure of the contents shall be expressed in terms of the U.S. gallon of 231 cubic inches and quart, pint, and fluid-ounce  subdivisions thereof,  and  shall  express   the   volume at 68 F (20 C). See also paragraph (p) of  this  section.

The declaration may contain com- mon or  decimal  fractions.  A  common fraction shall be in terms of halves, quarters, eighths, sixteenths, or thirty-

#### § 801.62

seconds; except that if there exists a firmly established, general consumer usage and trade custom of employing different common fractions in the net quantity declaration of a particular commodity, they may be employed. A common fraction shall be reduced  to its lowest terms; a decimal fraction shall not be carried out to more than two places. A statement that includes small fractions of an ounce shall be deemed to permit smaller variations than one which does not include such fractions.

The declaration shall be located on the principal display panel of the label, and with respect to packages bearing alternate principal panels it shall be duplicated on each principal display panel.

The declaration shall appear as a distinct item on the principal display panel, shall be separated, by at least a space equal to the height of the let- tering used in the declaration, from other printed label information appear- ing above or below the declaration and, by at least a space equal to twice the width of the letter ‘‘N’’ of the style of type used in the quantity of contents statement, from other printed label in- formation appearing to the left or right of the declaration. It shall not include any term qualifying a unit of weight, measure, or count, such as ‘‘giant pint’’ and ‘‘full quart’’, that tends to exag- gerate. It shall be placed on the prin- cipal display panel within the  bottom 30 percent of the area of the label panel in lines generally parallel to  the  base on which the package rests as it is de- signed to be displayed: Provided, That:

  - On packages having a principal display panel of 5 square inches or less the requirement for placement within the bottom 30 percent of the area of the label panel shall not apply when the declaration of net quantity of contents meets the other requirements of this part; and

  - In the case of a device that is mar- keted with both outer and inner retail containers bearing the mandatory label information required by this part and the inner container is not intended to be sold separately, the net quantity of contents placement requirement of this section applicable to such inner con- tainer is waived.

#### 21 CFR Ch. I (4–1–17 Edition)

  - The principal  display  panel  of  a device marketed on a display card to which the immediate container is af- fixed may be considered to be the dis- play panel of the card, and the type size of the net quantity of contents statement is governed by the dimen- sions of the display card.

The declaration shall accurately reveal the quantity of device in the package exclusive of wrappers and other material packed therewith.

The declaration shall appear in conspicuous and easily legible boldface print or type in distinct contrast (by typography,  layout,  color,  embossing, or molding) to other matter on the package; except that a declaration of net quantity blown, embossed, or mold- ed on a glass or plastic surface is per- missible when all label  information  is so formed on the surface. Requirements of conspicuousness and legibility shall include  the  specifications  that:

The ratio of height to width of the letter shall not exceed a differential of

3 units to 1 unit, i.e., no more than 3 times as high as it is wide.

Letter heights pertain to upper case or capital letters. When upper and lower case or all lower case letters are used, it is the lower case letter ‘‘o’’ or its equivalent that shall meet the min- imum standards.

When fractions are used, each component numeral shall  meet  one- half the minimum height standards.

The declaration shall be in letters and numerals in a type size established in relationship to the area of the prin- cipal display panel of the package and shall be uniform for all packages of substantially the same size by com- plying with the following type speci- fications:

  - Not less than one-sixteenth  inch in height on packages the principal dis- play panel of which has an area of 5 square inches or less.

  - Not less than one-eighth inch in height on packages the principal dis- play panel of which has an area of more than 5 but not more than 25 square inches.

  - Not less  than  three-sixteenths inch in height on packages the prin- cipal display panel  of  which  has  an area of more than 25 but not more than 100 square inches.

#### Food and Drug Administration, HHS	§ 801.62

  - Not less than one-fourth inch in height on packages the principal dis- play panel of which has an area of more than 100 square inches, except not less than one-half inch in height if the area is more than 400 square inches.

Where the declaration is blown, em- bossed, or molded on a glass or plastic surface rather than by printing, typ- ing, or coloring, the lettering sizes specified in paragraphs (h)(1) through

of this section shall be increased by one-sixteenth of an inch.

  - On packages containing less  than

4 pounds or 1 gallon and labeled in terms of weight or fluid measure:

- The declaration shall be expressed both in ounces, with identification by weight or by liquid measure and, if ap- plicable (1 pound or 1 pint or more) fol- lowed in parentheses by  a  declaration in pounds for weight units, with any re- mainder in terms of ounces or common or decimal fractions of the pound (see examples set forth in paragraphs (k) (1) and (2) of this section), or in the case of liquid measure, in the largest  whole units (quarts, quarts and pints, or pints, as appropriate) with any remain- der in terms of fluid ounces or common or decimal fractions of the  pint  or quart (see examples set forth in para- graphs (k) (3) and (4) of this section). If the net weight of the package is less than 1 ounce avoirdupois or the net fluid measure is less than 1 fluid ounce, the declaration  shall be  in terms of common or decimal fractions of the re- spective ounce and not in terms of drams.

- The declaration may appear in more than one line. The term ‘‘net weight’’ shall be used when stating the net quantity of contents in terms of weight. Use of the terms ‘‘net’’ or ‘‘net contents’’ in terms of fluid measure or numerical count is optional. It is suffi- cient to distinguish avoirdupois ounce from fluid ounce through association of terms; for example, ‘‘Net wt. 6 oz’’ or ‘‘6 oz net wt.,’’ and ‘‘6 fl oz’’ or ‘‘net contents 6 fl oz.’’

On packages containing 4 pounds or 1 gallon or more and labeled in terms of weight or fluid measure, the declaration shall be expressed in pounds for weight units with any re- mainder in terms of ounces or common or  decimal  fractions  of  the  pound;  in

the case of fluid measure, it shall be expressed  in  the  largest  whole  unit, i.e., gallons, followed by common or decimal fractions of a gallon or by the next smaller whole unit  or  units (quarts or quarts and pints), with any remainder in terms of fluid ounces or common or decimal  fractions  of  the pint or quart; see paragraph (k)(5) of this  section.

Examples: (1) A declaration of 11⁄2 pounds weight shall  be  expressed  as ‘‘net wt.  24 oz  (1  lb 8  oz),’’  or ‘‘Net  wt. 24 oz (11⁄2  lb)’’ or ‘‘Net wt. 24 oz (1.5 lb).’’

- A declaration of three-fourths pound avoirdupois weight shall be ex- pressed as ‘‘Net wt. 12 oz.’’.

- A declaration of 1 quart liquid measure shall be expressed as ‘‘Net contents 32 fl oz (1 qt)’’ or ‘‘32 fl oz (1 qt).’’

- A declaration of 13⁄4 quarts liquid measure shall be expressed as, ‘‘Net contents 56 fl oz (1 qt 1 pt 8  oz)’’  or ‘‘Net contents 56 fl oz (1 qt 1.5 pt),’’ but not in terms of quart and ounce such as ‘‘Net contents 56 fl oz (1 qt 24 oz).’’

- A declaration of 21⁄2 gallons liquid measure shall be expressed as ‘‘Net contents 2 gal 2 qt’’, ‘‘Net contents 2.5 gallons,’’ or ‘‘Net contents 21⁄2 gal’’ but not as ‘‘2 gal 4 pt’’.

For quantities, the following ab- breviations and none other may be em- ployed. Periods and plural forms are optional:

gallon gal	liter l

milliliter ml	cubic centimeter cc

quart qt	yard yd

pint pt	feet or foot ft

ounce oz	inch in

pound lb	meter m

grain gr	centimeter cm

kilogram kg	millimeter mm

gram g	fluid fl

milligram mg	square sq

microgram mcg	weight wt

On packages labeled in terms of linear measure, the declaration shall be expressed both in terms of inches and, if applicable (1 foot or more), the largest whole units (yards, yards and feet, feet). The declaration in terms of the largest whole units shall be in pa- rentheses following the declaration in terms of inches and any remainder shall be in terms of inches or common or decimal fractions of the foot or yard; if applicable, as in the case of ad- hesive tape, the initial declaration in linear  inches  shall  be  preceded  by  a

#### § 801.63

statement  of  the  width.  Examples  of linear measure are ‘‘86 inches (2 yd 1 ft

2  in)’’,  ‘‘90  inches  (21⁄2    yd)’’,  ‘‘30  inches

(2.5 ft)’’, ‘‘3⁄4   inch by 36 in (1 yd)’’, etc.

On packages labeled in terms of area measure, the declaration shall be expressed both in  terms  of  square inches and, if applicable  (1  square  foot or more), the largest whole square unit (square yards, square yards and square feet, square feet). The declaration in terms of the largest whole  units  shall be in parentheses following the dec- laration in terms of square inches and any remainder shall be in terms of square inches or common or decimal fractions of the square foot or square yard; for example,  ‘‘158  sq  inches  (1  sq ft 14 sq in)’’.

Nothing in this section shall pro- hibit supplemental statements at loca- tions other than the principal display panel(s) describing in nondeceptive terms the net quantity of contents, provided that such supplemental state- ments of net quantity of contents shall not include any term qualifying a unit of weight, measure, or count that tends to exaggerate the amount of the device contained in the package; for example, ‘‘giant pint’’ and ‘‘full  quart’’. Dual or combination declarations of net quan- tity of contents as provided for in para- graphs (a) and (i) of this section are not regarded as supplemental net quantity statements and shall be located on the principal display panel.

A separate statement of net quan- tity of contents  in terms of  the metric system of weight or measure is not re- garded as a  supplemental  statement and an accurate statement of the net quantity of contents in terms of the metric system of weight or measure may also appear on the principal dis- play panel or on other panels.

The declaration of net quantity of contents shall express an accurate statement of the quantity of contents of the package. Reasonable variations caused by loss or gain of moisture dur- ing the course of good distribution practice or by unavoidable deviations in good manufacturing practice will be recognized. Variations from stated quantity of contents shall not be un- reasonably large.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 801.63 Medical devices; warning statements for devices containing or	manufactured	with chlorofluorocarbons and other class I ozone-depleting substances.

All over-the-counter devices con- taining or manufactured with chlorofluorocarbons, halons, carbon tetrachloride, methyl chloride, or any other class I substance designated by the Environmental Protection Agency (EPA) shall carry one of the following warnings:

  - The EPA warning statement:

WARNING: Contains [or Manufactured with, if applicable] [insert name of substance], a sub- stance which harms public health and envi- ronment by destroying ozone in the upper at- mosphere.

  - The alternative statement:

NOTE: The indented statement below is re- quired by the Federal government’s Clean Air Act for all products containing or manu- factured with chlorofluorocarbons (CFC’s) [or other class I substance, if applicable]:

WARNING: Contains [or Manufactured with, if applicable] [insert name of substance], a sub- stance which harms public health and envi- ronment by destroying ozone in the upper at- mosphere.

CONSULT WITH YOUR PHYSICIAN, HEALTH PROFESSIONAL, OR SUPPLIER IF YOU HAVE ANY QUESTION ABOUT THE USE OF THIS PRODUCT.

The warning statement shall be clearly legible and conspicuous on the product, its immediate container, its outer packaging, or other labeling in accordance with the requirements of 40 CFR part 82 and appear with such prominence and conspicuousness as to render it likely to be read and under- stood by consumers under normal con- ditions of purchase. This provision does not replace or relieve a person from any requirements imposed under  40 CFR part 82.

[61 FR 20101, May 3, 1996]

## Subpart D—Exemptions From Adequate Directions for Use

##### § 801.109   Prescription devices.

A device which, because of any poten- tiality for harmful effect, or the meth- od of its use, or the collateral measures necessary to its use is not safe except under the supervision of a practitioner

Food and Drug Administration, HHS	§ 801.116

licensed  by  law  to  direct  the  use  of such device, and hence for which ‘‘ade- quate   directions   for   use’’   cannot   be prepared, shall be exempt from section 502(f)(1) of the act if all the following conditions are met:

The device is:

(1)(i) In the possession of a person, or his agents or employees, regularly and lawfully engaged in the manufacture, transportation, storage, or wholesale or retail distribution of such device; or

  - In the possession of a practi- tioner, such as physicians,  dentists, and veterinarians, licensed by law to use or order the use of such device; and

(2) Is to be sold only to or on the pre- scription or other order of such practi- tioner for use in the course of his pro- fessional practice.

The label of the device, other than surgical instruments, bears:

  - The symbol statement ‘‘Rx only’’ or ‘‘) only’’ or the statement ‘‘Caution: Federal law restricts this device to sale by  or  on  the  order  of  a  lll’’, the blank to be filled with the word ‘‘phy- sician’’,  ‘‘dentist’’,  ‘‘veterinarian’’,  or with the descriptive designation of any other practitioner licensed by the law of the State in which the practitioner practices to use or order the use of the device; and

  - The method of its application or use.

Labeling on or within the package from which the device is to be dis- pensed bears information for use, in- cluding indications, effects, routes, methods, and frequency and duration of administration, and any relevant hazards, contraindications, side effects, and precautions under which practi- tioners licensed by law to administer the device can use the device safely and for the purpose for which it is in- tended, including all purposes for which it is advertised or represented: Provided, however, That such informa- tion may be omitted from the dis- pensing package if, but only if, the ar- ticle is a  device for  which directions, hazards, warnings, and other informa- tion are commonly known to practi- tioners licensed by law to use the de- vice. Upon written request, stating rea- sonable grounds therefor, the Commis- sioner will offer an opinion on a pro- posal  to  omit  such  information  from

the dispensing package under this pro- viso.

Any labeling, as defined in section 201(m) of the act, whether or not it is on or within a package from which the device is to be dispensed, distributed by or on behalf of the manufacturer, pack- er, or distributor of the device, that furnishes or purports to furnish infor- mation for use of the device contains adequate information for such use, in- cluding indications, effects, routes, methods, and frequency and duration of administration and any relevant hazards, contraindications, side effects, and precautions, under which practi- tioners licensed by law to employ the device can use the device safely and for the purposes for which it is intended, including all purposes for which it is advertised or represented. This infor- mation will not be required on so- called reminder—piece labeling which calls attention to the name of the de- vice but does not include indications or other use information.

All labeling, except labels and cartons, bearing information for use of the device  also bears  the date  of the issuance or the date of the latest revi- sion of such labeling.

[41  FR  6896,  Feb.  13,  1976,  as  amended  at  81

FR 38930, June 15, 2016]

##### § 801.110   Retail   exemption   for   pre- scription devices.

A device subject to § 801.109 shall be exempt at the time of delivery to the ultimate purchaser or user from sec- tion 502(f)(1) of the act if it is delivered by a licensed practitioner in the course of his professional practice or upon a prescription or other order lawfully issued in the course of his professional practice, with labeling  bearing  the name and address of such licensed prac- titioner and the directions for use  and cautionary statements, if any,  con- tained in such order.

##### § 801.116   Medical devices having com- monly known directions.

A device shall be exempt from sec- tion 502(f)(1) of the act insofar as ade- quate directions for common uses thereof are known to the ordinary indi- vidual.

#### § 801.119

##### § 801.119   In vitro diagnostic products.

A product intended for use in the di- agnosis of disease and which is an in vitro diagnostic product as defined in

§ 809.3(a) of this chapter shall  be deemed to be in compliance with the requirements of this part and section 502(f)(1) of the Federal Food, Drug, and Cosmetic Act if it meets the  require- ments of subpart B of this part and the requirements of § 809.10 of this chapter.

[78 FR 55820, Sept. 24, 2013]

##### § 801.122 Medical devices for  proc- essing, repacking, or manufac- turing.

A device intended for processing, re- packing, or use in the manufacture of another drug or device shall be exempt from section 502(f)(1) of the act if its label bears the  statement  ‘‘Caution: For manufacturing, processing, or re- packing’’.

##### § 801.125 Medical devices for use in teaching, law enforcement, re- search, and analysis.

A device subject to § 801.109 shall be exempt from section 502(f)(1) of this act if shipped or sold to, or in the posses- sion of, persons regularly and lawfully engaged in instruction in pharmacy, chemistry, or medicine not involving clinical use, or engaged in law enforce- ment, or in research not involving clin- ical use, or in chemical analysis, or physical testing, and is to be used only for such instruction, law enforcement, research,  analysis,  or  testing.

##### § 801.127  Medical  devices;  expiration of exemptions.

If a shipment or delivery, or any part thereof, of a device which is ex- empt under the regulations in this sec- tion is made to a person in whose pos- session the article is not exempt, or is made for any purpose other than those specified, such exemption shall expire, with respect to such shipment or deliv- ery or part thereof, at the beginning of that shipment or delivery. The causing of an exemption to expire shall be con- sidered an act which results in such de- vice being misbranded unless it is dis- posed of under circumstances in which it ceases to be a drug or device.

The    exemptions    conferred    by

§§ 801.119,  801.122,  and  801.125  shall  con-

#### 21 CFR Ch. I (4–1–17 Edition)

tinue until the devices are used for the purposes for which they are exempted, or until they are relabeled to comply with section 502(f)(1) of the act. If, how- ever, the device is converted, or manu- factured into a form limited to pre- scription dispensing, no  exemption shall thereafter apply to the article un- less the device is labeled as required by

§ 801.109.

##### § 801.128 Exceptions or alternatives to labeling requirements for medical devices held by the Strategic Na- tional Stockpile.

(a) The appropriate FDA Center Di- rector may grant an exception or alter- native to any provision listed in para- graph (f) of this section and not explic- itly required by statute, for specified lots, batches, or other units of a med- ical device, if the Center Director de- termines that compliance with such la- beling requirement could adversely af- fect the safety, effectiveness, or avail- ability of such devices that are or will be included in the Strategic National Stockpile.

(b)(1)(i) A Strategic National Stock- pile official or any entity that manu- factures (including labeling, packing, relabeling, or repackaging), distrib- utes, or stores devices that are or will be included in the Strategic National Stockpile may submit, with written concurrence from a Strategic National Stockpile official, a written request for an exception or alternative  described in paragraph (a) of this section to the Center Director.

(ii) The Center Director may  grant an exception or alternative  described in paragraph  (a) of this  section on his or her own initiative.

(2) A written request for an exception or  alternative  described  in  paragraph

of this section must:

  - Identify the specified lots, batches, or other units of the medical device that would be subject to the ex- ception or alternative;

  - Identify the labeling provision(s) listed in paragraph (f) of this section that are the subject of the exception or alternative request;

  - Explain why compliance with the labeling provision(s) could ad- versely affect the safety, effectiveness, or  availability  of  the  specified  lots,

#### Food and Drug Administration, HHS	§ 801.150

batches, or other units of a medical de- vice that are or will be held in the Strategic  National  Stockpile;

  - Describe any proposed safeguards or conditions that will be implemented so that the labeling of the device in- cludes appropriate information nec- essary for the safe and effective use of the device, given the anticipated cir- cumstances of use of the device;

  - Provide a draft of the proposed la- beling of the specified lots, batches, or other units of the medical device sub- ject to the exception or alternative; and

  - Provide any other information requested by the Center Director in support of the request.

The Center Director must respond in writing to all requests under this section. The Center Director may im- pose appropriate conditions when granting such an exception or alter- native under this section.

A grant of an exception or alter- native under this section will include any safeguards or conditions deemed appropriate by the Center Director so that the labeling of devices subject to the exception or alternative includes the information necessary for the safe and effective use of the device, given the anticipated circumstances of use.

If the Center Director grants a re- quest for an exception or alternative to the labeling requirements under this section:

  - The Center Director may deter- mine that the submission and grant of a written request under this section satisfies the provisions relating to pre- market  notification  submissions  under

§ 807.81(a)(3) of this chapter.

(2)(i) For a Premarket Approval Ap- plication (PMA)-approved device, the submission and grant of a written re- quest under this section satisfies the provisions  relating  to  submission  of PMA supplements under § 814.39 of this chapter;  however,

(ii) The grant of the request must be identified  in  a  periodic  report  under

§ 814.84 of this chapter.

The Center Director may grant an exception or alternative under this sec- tion to the following provisions of this chapter, to the extent that the require- ments in these provisions are not ex- plicitly required by statute:

(1) § 801.1(d);

(2) Subpart B of this part and part 830 of  this  chapter  in  its  entirety;

(3) § 801.60;

(4) § 801.61;

(5) § 801.62;

(6) § 801.63;

(7) § 801.109; and

Part 801, subpart H.

[72 FR 73601, Dec. 28, 2007, as amended at 78

FR 55820, Sept. 24, 2013]

## Subpart E—Other Exemptions

##### § 801.150   Medical  devices;  processing, labeling, or repacking.

  - Except as provided by paragraphs

  - and (c) of this  section,  a  shipment or other delivery of  a  device  which  is, in accordance with the practice of the trade, to be processed, labeled, or re- packed, in substantial quantity at an establishment other than that where originally processed or packed, shall be exempt, during the time of introduc- tion into and movement in interstate commerce and the time of holding in such establishment, from compliance with the labeling and packaging re- quirements of section 502(b) and (f) of the act if:

    - The person who introduced such shipment or delivery into interstate commerce is the operator of the estab- lishment where such device is to be processed, labeled, or repacked; or

    - In case such person is not such op- erator, such shipment or delivery is made to such establishment under a written agreement, signed by and con- taining the post office  addresses  of such person and such operator, and containing such specifications for the processing, labeling,  or  repacking,  as the case may be, of such device in such establishment as will insure, if such specifications are followed, that such device will not be adulterated or mis- branded within the meaning of the act upon completion of such processing, la- beling, or repacking. Such person and such operator shall each keep a copy of such agreement until 2 years after the final shipment or delivery of such de- vice from such establishment, and shall make such copies available for inspec- tion at any reasonable hour to any offi- cer or employee of the Department who requests them.

§ 801.405

An exemption of a shipment or other delivery of a device under para- graph (a)(1) of this section shall, at the beginning of the act of removing such shipment or delivery, or any part thereof, from such establishment, be- come void ab initio if the device com- prising such shipment, delivery, or part is adulterated or misbranded within the meaning of the act when so re- moved.

An exemption of a shipment or other delivery of a device under para- graph (a)(2) of this section shall be- come void ab initio with respect to the person who introduced such  shipment or delivery  into  interstate commerce upon refusal by such person to make available for inspection a copy of the agreement, as required by such para- graph (a)(2).

An exemption of a shipment or other delivery of a device under para- graph (a)(2) of this section shall expire:

  - At the beginning of the act of re- moving such shipment or delivery, or any part thereof, from such establish- ment if the device comprising such shipment, delivery, or part is adulter- ated or misbranded within the meaning of the act when so removed; or

  - Upon refusal by the operator of the establishment where such device is to be processed, labeled, or repacked, to make available for inspection a copy of the agreement, as required by such clause.

As it is a common industry prac- tice to manufacture and/or assemble, package, and fully label a device as sterile at one establishment and then ship such device in interstate com- merce to another establishment or to a contract sterilizer for sterilization, the Food and Drug Administration will ini- tiate no regulatory action against the device as misbranded or adulterated when the nonsterile device is labeled sterile, provided all the following con- ditions are met:

  - There is in effect a written agree- ment which:

    - Contains the names and post office addresses of the firms involved and is signed by the person authorizing such shipment and the operator or person in charge of the establishment receiving the devices for sterilization.

#### 21 CFR Ch. I (4–1–17 Edition)

    - Provides instructions for main- taining proper records or otherwise ac- counting for the number of units in each shipment to insure that the num- ber of units shipped is the same as the number received and sterilized.

    - Acknowledges that the device is nonsterile and is being shipped for fur- ther processing, and

    - States in detail the sterilization process, the gaseous mixture or other media, the equipment, and the testing method or quality controls to be used by the contract sterilizer to assure that the device will be brought  into full compliance with the Federal Food, Drug, and Cosmetic Act.

  - Each pallet, carton, or other des- ignated unit is conspicuously  marked to show its nonsterile nature when it is introduced into and is moving in inter- state commerce, and while it is being held prior to sterilization. Following sterilization, and until  such  time  as  it is established that the device is sterile and can be released from quarantine, each pallet, carton, or other designated unit is conspicuously marked to show that it has not been released from quarantine, e.g., ‘‘sterilized—awaiting test results’’ or an equivalent designa- tion.

## Subparts F–G [Reserved]

Subpart H—Special Requirements for Specific Devices

##### § 801.405 Labeling of articles intended for lay use in the repairing and/or refitting of dentures.

The American Dental Association and leading dental authorities have ad- vised the Food and Drug Administra- tion of their concern regarding the safety of denture reliners, repair kits, pads, cushions, and other articles mar- keted and labeled for lay use in the re- pairing, refitting, or cushioning of ill- fitting, broken, or  irritating  dentures. It is the opinion of dental authorities and the Food and Drug Administration that to properly repair and properly refit dentures a person must have pro- fessional knowledge and specialized technical skill. Laymen cannot be ex- pected to maintain  the  original vertical dimension of occlusion and the centric relation essential in the proper

Food and Drug Administration, HHS	§ 801.405

repairing or refitting of dentures. The continued wearing of improperly re- paired or refitted dentures may cause acceleration of bone resorption, soft tissue hyperplasia, and other irrep- arable damage to the oral cavity. Such articles designed for lay use should be limited to emergency or temporary sit- uations pending the services of a li- censed dentist.

The Food and Drug Administra- tion therefore regards such articles as unsafe and misbranded under the Fed- eral Food, Drug, and Cosmetic Act, un- less the labeling:

(1)(i) Limits directions for use for denture repair kits to emergency re- pairing pending unavoidable delay in obtaining professional reconstruction of the denture;

(ii) Limits directions for use for den- ture reliners, pads, and cushions to temporary refitting pending unavoid- able delay in obtaining professional re- construction of the denture;

Contains in a conspicuous manner the word ‘‘emergency’’ preceding and modifying each indication-for-use statement for denture repair kits and the word ‘‘temporary’’ preceding and modifying each indication-for-use statement for reliners, pads, and cush- ions; and

Includes a conspicuous warning statement to the effect:

  - For denture repair kits: ‘‘Warn- ing—For emergency repairs only. Long term  use  of  home-repaired  dentures may cause faster bone loss, continuing irritation,  sores,  and  tumors.  This  kit for  emergency  use  only.  See  Dentist Without Delay.’’

  - For denture reliners, pads, and cushions: ‘‘Warning—For temporary use only. Longterm use of this product may lead to faster bone loss, continuing ir- ritation, sores, and tumors. For Use Only Until a Dentist Can Be Seen.’’

Adequate directions for use re- quire full information of the temporary and emergency use recommended in order for the layman to understand the limitations of usefulness, the reasons therefor, and the importance of adher- ing to the warnings. Accordingly, the labeling should contain substantially the following information:

  - For denture repair kits: Special training and tools are needed to repair

dentures to fit properly. Home-repaired dentures may cause irritation to the gums and discomfort and tiredness while eating. Long term use may lead to more troubles, even permanent changes in bones, teeth, and gums, which may make it impossible to wear dentures in the future. For these rea- sons, dentures repaired with this kit should be used only in an emergency until a  dentist  can  be  seen.  Dentures that don’t fit properly cause irritation and injury to the gums and faster bone loss, which is permanent.  Dentures that don’t fit properly cause gum changes that may require surgery for correction. Continuing irritation and injury may lead to cancer in the mouth. You must see your dentist as soon as possible.

  - For denture reliners, pads, and cushions: Use  of  these  preparations or devices may temporarily decrease the discomfort; however, their use will not make the denture fit properly. Special training and tools are needed to repair a denture to fit properly. Dentures that do not fit properly cause irritation and injury to  the  gums  and  faster  bone loss, which is permanent and may re- quire a completely new  denture. Changes in the gums caused by den- tures that do not fit properly may re- quire surgery for correction. Con- tinuing irritation and injury  may  lead to cancer in the mouth. You must see your dentist as soon as possible.

  - If the denture relining or repair- ing material forms a permanent bond with the denture, a warning statement to the following effect should be in- cluded: ‘‘This reliner becomes fixed to the denture and a completely new den- ture may be required because of its use.’’

Labeling claims exaggerating the usefulness or the safety of the material or failing to disclose  all  facts  relevant to the claims of usefulness will be re- garded as false and misleading under sections 201(n) and 502(a) of the Federal Food, Drug, and Cosmetic Act.

Regulatory action may be initi- ated with respect to any article found within the jurisdiction of the act con- trary to the provisions of this policy statement after 90 days following the date of publication of this section  in the  FEDERAL   REGISTER.

#### § 801.410

##### § 801.410   Use	of	impact-resistant lenses in eyeglasses and sunglasses.

Examination of data available on the frequency of eye injuries resulting from the shattering of ordinary crown glass lenses indicates that the use of such lenses constitutes an avoidable hazard to the eye of the wearer.

The consensus of the ophthalmic community is that the number of eye injuries would be substantially reduced by the use in eyeglasses and sunglasses of impact-resistant lenses.

(c)(1) To protect the public more ade- quately from potential eye injury, eye- glasses and sunglasses must be fitted with impact-resistant lenses, except in those cases where the physician or op- tometrist finds that such lenses will not fulfill the visual requirements of the particular patient, directs in writ- ing the use of other lenses, and gives written notification thereof to the pa- tient.

The physician or optometrist shall have the option of ordering glass lenses, plastic lenses, or laminated glass lenses made impact resistant by any method; however, all such lenses shall be capable of withstanding the impact test described in paragraph (d)(2) of this section.

Each finished impact-resistant glass lens for prescription use shall be individually tested for impact resist- ance and shall be capable of with- standing the impact test described in paragraph (d)(2) of this section. Raised multifocal lenses shall be impact re- sistant but need not be tested beyond initial design testing. Prism segment multifocal, slab-off prism, lenticular cataract, iseikonic, depressed segment one-piece multifocal, bioconcave, myodisc and minus lenticular, custom laminate and cemented assembly lenses shall be impact resistant but need not be subjected to impact test- ing. To demonstrate that all other types of impact-resistant lenses, in- cluding impact-resistant laminated glass lenses (i.e., lenses other than those described in the three preceding sentences of this paragraph (c)(3)), are capable of withstanding  the  impact test described in this regulation, the manufacturer of these lenses shall sub- ject to an impact test a statistically significant   sampling   of   lenses   from

#### 21 CFR Ch. I (4–1–17 Edition)

each production batch, and the  lenses so tested shall be representative of the finished forms as worn by the wearer, including finished forms that are of minimal lens thickness and have been subjected to any treatment used to im- part impact resistance. All non- prescription lenses and plastic pre- scription lenses tested on the basis of statistical significance shall be  tested in uncut-finished or finished form.

(d)(1) For the purpose of this regula- tion, the impact test described in para- graph (d)(2) of this section shall be the ‘‘referee test,’’ defined as ‘‘one which will be utilized to determine compli- ance with a regulation.’’ The referee test provides the Food and Drug Ad- ministration with the means of exam- ining a medical device for performance and does not inhibit the manufacturer from using equal or superior test meth- ods. A lens manufacturer shall conduct tests of lenses using the impact test de- scribed in paragraph (d)(2) of this sec- tion or any equal or superior test. Whatever test is used, the lenses shall be capable of withstanding the impact test described in paragraph (d)(2) of this section if the Food and Drug Ad- ministration examines them for per- formance.

(2) In the impact test, a 5⁄8-inch steel ball weighing approximately 0.56 ounce is dropped from a height of 50 inches upon the horizontal upper surface of the lens. The ball  shall strike within a 5⁄8-inch diameter circle located at the geometric center of the lens. The ball may be guided but not restricted in its fall by being dropped through a tube extending to within approximately 4 inches of the lens. To pass the test, the lens must not fracture; for the purpose of this section, a lens will be consid- ered to have fractured if it cracks through its  entire  thickness,  including a laminar layer, if any, and across a complete diameter into two or more separate pieces, or if any lens material visible to the naked eyes becomes de- tached from  the  ocular  surface.  The test shall be conducted with the lens supported by a tube (1-inch inside di- ameter, 11⁄4-inch outside diameter, and approximately 1-inch high) affixed to a rigid iron or steel base plate. The total weight of the base plate and its rigidly attached fixtures shall be not less than

#### Food and Drug Administration, HHS	§ 801.415

27  pounds.  For  lenses  of  small  min- imum diameter, a support tube having an  outside  diameter  of  less  than  11⁄4 inches may be used. The support tube shall be made of  rigid acrylic plastic, steel, or other suitable substance and shall have securely bonded on the top edge a 1⁄8- by 1⁄8-inch  neoprene gasket having a  hardness  of  40  5,  as  deter- mined  by  ASTM  Method  D  1415–88, ‘‘Standard   Test   Method   for   Rubber Property—International   Hardness’’   a minimum   tensile   strength   of   1,200 pounds, as determined by ASTM Meth- od D 412–98A, ‘‘Standard Test Methods for  Vulcanized  Rubber  and  Thermo- plastic   Elastomers—Tension,’’   and   a minimum  ultimate  elongation  of  400 percent, as determined by ASTM Meth- od D 412–68 (Both methods are incor- porated by reference and are available from the American Society for Testing Materials,  100  Barr  Harbor  Dr.,  West Conshohocken, Philadelphia, PA 19428, or available for inspection at the Cen- ter	for	Devices	and	Radiological Health’s Library, 9200 Corporate Blvd., Rockville, MD 20850, or at the National Archives  and  Records  Administration (NARA). For information on the avail- ability  of  this  material  at  NARA,  call 202–741–6030.   The    diameter    or contour  of  the  lens  support  may  be modified as necessary so that the 1⁄8- by 1⁄8-inch  neoprene  gasket  supports  the lens at its periphery.

Copies of invoice(s), shipping doc- ument(s), and records of sale or dis- tribution of all impact resistant lenses, including finished eyeglasses and sun- glasses, shall be kept and maintained for a period of 3 years; however, the names and addresses of individuals pur- chasing nonprescription eyeglasses and sunglasses at the retail level  need  not be kept and maintained by the retailer. The records kept in compliance with this paragraph shall be made available upon request at all reasonable hours by any officer or employee of the Food and Drug Administration or by any other officer or employee acting on be- half of the Secretary of Health and Human Services and such officer or em- ployee shall be  permitted  to  inspect and  copy  such  records,  to  make  such

inventories of stock as he deems nec- essary, and otherwise to check the cor- rectness of such inventories.

In addition, those persons con- ducting tests in accordance with para- graph (d) of this section shall  maintain the results thereof and a description  of the test method and of the test appa- ratus for a period of 3 years. These records shall be made available upon request at any reasonable hour by any officer or employee acting on behalf  of the Secretary of Health and Human Services. The persons conducting tests shall permit the officer or employee to inspect and copy the records, to make such inventories of stock  as  the  officer or employee deems necessary, and oth- erwise to check the correctness of the inventories.

For the purpose of this section, the term ‘‘manufacturer’’ includes an importer for resale. Such importer may have the tests required by paragraph

of this section conducted in the country of origin but must make the results thereof available, upon request, to the Food and Drug Administration, as soon as practicable.

  - All lenses must be impact-resist- ant except when the physician or op- tometrist finds that impact-resistant lenses will not fulfill the visual re- quirements for a particular patient.

  - This statement of policy does not apply to contact lenses.

[41  FR  6896,  Feb.  13,  1976,  as  amended  at  44

FR 20678, Apr. 6, 1979; 47 FR 9397, Mar. 5, 1982;

65 FR 3586, Jan. 24, 2000; 65 FR 44436, July 18,

2000; 69 FR 18803, Apr. 9, 2004]

##### § 801.415   Maximum acceptable level of ozone.

Ozone is a toxic gas with no known useful medical application in specific, adjunctive, or preventive ther- apy. In order for ozone to be effective as a germicide, it must be present in a concentration far greater than that which can be safely tolerated by man and animals.

Although undesirable physio- logical effects on the central nervous system, heart, and vision have been re- ported, the predominant physiological effect of ozone is primary irritation of the mucous membranes. Inhalation of ozone can cause sufficient irritation to the   lungs   to   result   in   pulmonary

#### § 801.417

edema. The onset of pulmonary edema is usually delayed for some hours after exposure; thus, symptomatic response is not a reliable warning of exposure to toxic concentrations of ozone. Since ol- factory fatigue develops readily, the odor of ozone is not a reliable index of atmospheric ozone concentration.

A number of devices currently on the market generate ozone by design or as a  byproduct.  Since  exposure  to ozone  above  a   certain   concentration can be injurious to health, any such de- vice will be considered adulterated and/ or misbranded within the meaning of sections 501 and 502 of the act if it  is used or intended for use under the fol- lowing   conditions:

  - In such a manner that it  gen- erates ozone at a level in excess of 0.05 part per million by volume of air circu- lating through the device or causes an accumulation of ozone in excess of 0.05 part per million by volume of air (when measured under standard conditions at 25 C (77 F) and 760 millimeters of mer- cury) in the atmosphere of enclosed space intended to be occupied by people for extended periods of time,  e.g., houses, apartments, hospitals, and of- fices. This applies to any such device, whether portable or permanent  or  part of any system, which generates  ozone by design or as an inadvertent or inci- dental   product.

  - To generate ozone and release it into the atmosphere in hospitals or other establishments occupied by the ill or infirm.

  - To generate ozone and release it into the atmosphere and does not indi- cate in its labeling the maximum ac- ceptable concentration of ozone which may be generated (not to exceed 0.05 part per million by volume of air circu- lating through the device) as estab- lished herein and the smallest area in which such device can be used so as not to produce an ozone accumulation in excess of 0.05 part per million.

  - In any medical condition for which there is no proof of safety and ef- fectiveness.

  - To generate ozone at a level less than 0.05 part per million by volume of air circulating through the device and it is labeled for use as a germicide or deodorizer.

#### 21 CFR Ch. I (4–1–17 Edition)

This section does not affect the present threshold limit value of 0.10 part per million (0.2 milligram per cubic meter) of ozone exposure for an 8- hour-day exposure of industrial work- ers as recommended by the American Conference of Governmental Industrial Hygienists.

The method and apparatus speci- fied in 40 CFR part 50, or any other equally sensitive and accurate method, may be employed in measuring ozone pursuant  to  this  section.

##### § 801.417   Chlorofluorocarbon	 propel- lants.

The use of chlorofluorocarbon in de- vices as propellants in self-pressurized containers is generally prohibited ex- cept as provided in § 2.125 of this chap- ter.

[43 FR 11318, Mar. 17, 1978]

##### § 801.420   Hearing  aid  devices;  profes- sional and patient labeling.

Definitions for the purposes of this section and § 801.421. (1) Hearing aid means any wearable instrument or de- vice designed for, offered for the pur- pose of, or represented as aiding per- sons with or compensating for, im- paired hearing.

Ear specialist means any licensed physician who specializes in diseases of the ear and is medically trained to identify the symptoms of deafness in the context  of  the total  health  of the patient, and is qualified by special training to diagnose and treat hearing loss. Such physicians are also known as otolaryngologists, otologists, and otorhinolaryngologists.

Dispenser means any person, part- nership, corporation, or association en- gaged in the sale, lease, or rental of hearing aids to any member of the con- suming public  or  any  employee,  agent, sales person, and/or representative of such a person, partnership, corpora- tion,  or  association.

Audiologist means any person qualified by training and experience to specialize in the evaluation and reha- bilitation of individuals whose commu- nication disorders center in whole or in part in the hearing function. In some states audiologists must satisfy spe- cific requirements for licensure.

#### Food and Drug Administration, HHS	§ 801.420

Sale or purchase includes any lease or rental of a hearing aid to a member of the consuming public who is a user or prospective user of a hearing aid.

Used hearing aid means any hear- ing aid that has been worn for any pe- riod of time by a user. However, a hear- ing aid shall not be considered ‘‘used’’ merely because it has been worn by a prospective user as a part of a bona fide hearing aid evaluation conducted to de- termine whether to select that par- ticular hearing aid for that prospective user, if such evaluation has been con- ducted in the presence of the dispenser or a hearing aid health professional se- lected by the dispenser to assist the buyer in making such a determination.

Label requirements for hearing aids. Hearing aids shall be clearly and per- manently marked with:

  - The name of the manufacturer or distributor, the model name or num- ber, the serial number, and the year of manufacture.

  - A ‘‘ + ’’ symbol to indicate the positive connection for battery inser- tion, unless it is physically impossible to insert the battery in the reversed position.

Labeling requirements for hearing aids—(1) General. All labeling informa- tion required by this paragraph  shall be included in a User Instructional Brochure that shall be developed by the manufacturer or distributor, shall accompany the hearing aid, and  shall be provided to the prospective user by the dispenser of the hearing aid in ac- cordance with § 801.421(c). The User In- structional Brochure accompanying each hearing aid shall contain the fol- lowing  information  and   instructions for use, to the extent applicable to the particular requirements and character- istics of the hearing aid:

An illustration(s) of the hearing aid, indicating operating controls, user adjustments, and battery compart- ment.

Information on the function of all controls intended for user adjustment.

A description of any accessory that may accompany the hearing aid, e.g., accessories for use with a tele- vision or telephone.

Specific instructions for:

Use of the hearing aid.

Maintenance and care of the hear- ing aid, including the procedure to fol- low in washing the earmold, when re- placing tubing on those hearing aids that use tubing, and in storing the hearing aid when it will not be used for an extended period of time.

Replacing or recharging the bat- teries, including a generic designation of replacement batteries.

Information on how and where to obtain repair service, including at least one specific address where the user can go, or send the hearing aid to, to ob- tain such repair service.

A description of commonly oc- curring avoidable conditions that could adversely affect or damage the hearing aid, such as dropping, immersing, or exposing the hearing aid to excessive heat.

Identification of any known side effects associated with the use of a hearing aid that may warrant con- sultation with a physician, e.g., skin irritation and accelerated accumula- tion of cerumen (ear wax).

A statement that a hearing aid will not restore normal hearing and will not prevent or improve a hearing impairment resulting from organic conditions.

A statement that in most cases infrequent use of a hearing aid does not permit a user to attain full benefit from it.

A statement that the use of a hearing aid is only part of hearing ha- bilitation and may need to be supple- mented by auditory training and in- struction in lipreading.

The warning statement required by paragraph (c)(2) of this section.

The notice for prospective hear- ing aid users required by paragraph (c)(3) of this section.

The technical data required by paragraph (c)(4) of this section, unless such data is provided in separate label- ing accompanying the device.

Warning statement. The User In- structional Brochure shall contain the following warning statement:

WARNING TO HEARING AID DISPENSERS

A hearing aid dispenser should advise a prospective hearing aid user to consult promptly with a licensed physician (pref- erably an ear specialist) before dispensing a

#### § 801.420

hearing aid if the hearing aid dispenser de- termines through inquiry, actual observa- tion, or review of any other available infor- mation concerning the prospective user, that the prospective user has any of the following conditions:

- Visible congenital or traumatic deform- ity of the ear.

- History of active drainage from the ear within the previous 90 days.

- History of sudden or rapidly progres- sive hearing loss within the previous 90 days.

- Acute or chronic dizziness.

- Unilateral hearing loss of sudden or re- cent onset within the previous 90 days.

- Audiometric air-bone gap equal to or greater than 15 decibels at 500 hertz (Hz), 1,000 Hz, and 2,000 Hz.

- Visible evidence of significant ceru- men accumulation or a foreign body in the ear canal.

- Pain or discomfort in the ear.

Special care should  be exercised  in select- ing and fitting a hearing aid whose max- imum sound pressure level exceeds 132 deci- bels because there may be risk of impairing the remaining hearing  of  the  hearing  aid user. (This  provision  is  required  only  for those hearing aids with a maximum sound pressure capability greater than 132 decibels (dB).)

Notice for prospective hearing aid users. The User Instructional Brochure shall contain the following notice:

IMPORTANT NOTICE FOR PROSPECTIVE HEARING AID USERS

Good health practice requires that a per- son with a hearing loss have a medical eval- uation by a licensed physician (preferably a physician who specializes in diseases of the ear) before purchasing a hearing aid. Li- censed physicians who specialize in diseases of the ear are often referred to as otolaryngologists,	otologists	or otorhinolaryngologists. The purpose of med- ical evaluation is to assure that all medi- cally treatable conditions that may affect hearing are identified and treated before the hearing aid is purchased.

Following the medical evaluation, the phy- sician will give you a written statement that states that your hearing loss has been medi- cally evaluated and that you may be consid- ered a candidate for a hearing aid. The physi- cian will refer you to an audiologist or a hearing aid dispenser, as appropriate, for a hearing  aid  evaluation.

The audiologist or hearing aid dispenser will conduct a hearing aid evaluation to as- sess your ability to hear with and without a hearing aid. The hearing aid evaluation will enable the audiologist or dispenser to select and fit a hearing aid to your individual needs.

#### 21 CFR Ch. I (4–1–17 Edition)

If you have reservations about your ability to adapt to amplification, you should inquire about the availability of a trial-rental or purchase-option program. Many hearing aid dispensers now offer programs that permit you to wear a hearing aid for a period  of time for a nominal fee after which you may decide if you want to purchase the  hearing aid.

Federal law restricts the sale of hearing aids to those individuals who have obtained a medical evaluation from a licensed physi- cian. Federal law permits a fully informed adult to sign a waiver statement  declining the medical evaluation for religious or per- sonal beliefs that preclude consultation with a physician. The exercise of such a waiver is not in your best health interest and its use is strongly discouraged.

CHILDREN WITH HEARING LOSS

In addition to seeing a physician for a med- ical evaluation, a child with a hearing loss should be directed to an audiologist for eval- uation and rehabilitation since hearing loss may cause problems in language develop- ment and the educational and social growth of a child. An audiologist is qualified by training and experience to assist in the eval- uation and rehabilitation of a child with a hearing loss.

Technical  data.  Technical  data useful in selecting, fitting, and check- ing the performance of a hearing aid shall be provided in the User Instruc- tional Brochure or in separate labeling that accompanies the device. The de- termination of technical data values for the hearing aid labeling shall be conducted in accordance with the test procedures of the American National Standard ‘‘Specification of Hearing Aid Characteristics,’’ ANSI S3.22–2003 (Re- vision of ANSI S3.22–1996)  (Includes April 2007 Erratum). The Director of the Office of the Federal Register ap- proves this incorporation by reference in accordance with 5 U.S.C. 552(a) and 1 CFR part 51. Copies are available from the	Standards	Secretariat	of	the Acoustical  Society  of  America,  120 Wall St., New York, NY 10005–3993, or are available for inspection at the Reg- ulations Staff, CDRH (HFZ–215), FDA, 1350 Piccard Dr., rm. 150, Rockville, MD 20850, or at the National Archives and Records Administration (NARA). For information on the availability of this material at NARA, call 202–741–6030

#### Food and Drug Administration, HHS	§ 801.421

. As a minimum, the User Instructional Brochure or such other labeling shall include the appro- priate values or information for the following technical data elements as these elements are defined or used in such standard:

  - Saturation output curve (SSPL 90 curve).

  - Frequency response curve.

  - Average saturation output (HF- Average SSPL 90).

  - Average full-on gain (HF-Average full-on gain).

  - Reference test gain.

  - Frequency range.

  - Total harmonic distortion.

  - Equivalent input noise.

  - Battery current drain.

  - Induction coil sensitivity (tele- phone coil aids only).

  - Input-output curve (ACG aids only).

  - Attack and release times (ACG aids only).

Statement if hearing aid is used or rebuilt. If a hearing aid has been used or rebuilt, this fact shall be declared  on the container in which the hearing aid is packaged and on a tag that is phys- ically attached to such hearing aid. Such fact may also be stated in the User Instructional Brochure.

Statements in User Instructional Brochure other than those required. A User Instructional Brochure may con- tain statements or illustrations in ad- dition to those required by paragraph

  - of this section if the additional statements:

    - Are not false or misleading in any particular, e.g., diminishing the impact of the required statements; and

    - Are not prohibited by this chap- ter or by regulations of the Federal Trade Commission.

[42  FR  9294,  Feb.  15,  1977,  as  amended  at  47

FR  9398,  Mar.  5,  1982;  50  FR  30154,  July  24,

1985; 54 FR 52396, Dec. 21, 1989; 64 FR 59620,

Nov. 3,  1999;  69  FR 18803,  Apr.  9,  2004;  73  FR

31360, June 2, 2008]

##### § 801.421   Hearing  aid  devices;  condi- tions for sale.

Medical  evaluation   requirements—

General. Except as provided in para- graph (a)(2) of this section, a  hearing aid dispenser shall not sell a  hearing aid unless the prospective user has pre-

sented to the hearing aid dispenser a written statement signed by a licensed physician that states that the patient’s hearing loss has been medically evalu- ated and the patient may be considered a candidate for a hearing aid. The med- ical evaluation must have taken place within the preceding 6 months.

Waiver to the medical evaluation re- quirements. If the prospective hearing aid user is 18 years of age or older, the hearing aid dispenser may afford the prospective user an opportunity to waive the medical evaluation require- ment of paragraph (a)(1) of this section provided that the hearing aid dis- penser:

  - Informs the prospective user that the exercise of the waiver is not in the user’s best health interest;

  - Does not in any way actively en- courage the prospective user to waive such a medical evaluation; and

  - Affords the prospective user the opportunity to sign the following statement:

I have been advised by llll llll (Hearing aid dispenser’s name) that the Food and Drug Administration has determined that my best health interest would be served if I had a medical evaluation by a licensed physician (preferably a physician who spe- cializes in diseases of the ear) before pur- chasing a hearing aid. I do not wish a med- ical evaluation before purchasing a hearing aid.

Opportunity to review User Instruc- tional Brochure. Before signing any statement under paragraph (a)(2)(iii) of this section and before the sale of a hearing aid to a prospective user, the hearing aid dispenser shall:

  - Provide the prospective user a copy of the User Instructional Bro- chure for a hearing  aid  that  has  been, or may be selected for the prospective user;

  - Review the content of the User In- structional Brochure with the prospec- tive user orally, or in the predominate method of communication used during the sale;

  - Afford the prospective user an op- portunity to read the User Instruc- tional  Brochure.

Availability of User Instructional Brochure. (1) Upon request by an indi- vidual who is considering purchase of a

#### § 801.430

hearing aid, a dispenser shall, with re- spect to any hearing aid that he dis- penses, provide a copy of the User In- structional Brochure  for  the  hearing aid or the name and address of the manufacturer or distributor  from whom a  User  Instructional  Brochure for the hearing aid may be obtained.

In addition to assuring that a User Instructional Brochure accompanies each hearing aid, a manufacturer or distributor shall with respect to any hearing aid that he manufactures or distributes:

  - Provide sufficient copies of the User Instructional Brochure to sellers for distribution to users and prospec- tive users;

  - Provide a copy of the User In- structional Brochure  to  any  hearing aid professional, user, or prospective user who requests a copy in writing.

Recordkeeping. The dispenser shall retain for 3 years  after  the  dispensing of a hearing aid a copy of any written statement from a physician  required under paragraph (a)(1) of this  section or any written statement waiving med- ical evaluation required  under  para- graph (a)(2)(iii) of this section.

Exemption for group auditory train- ers. Group auditory trainers, defined as a group amplification system pur- chased by a qualified school or institu- tion for the purpose of communicating with and educating individuals with hearing impairments, are exempt from the requirements of this section.

[42 FR 9296, Feb. 15, 1977]

##### § 801.430   User  labeling  for  menstrual tampons.

This section applies to scented or scented deodorized menstrual tampons as identified in § 884.5460 and unscented menstrual    tampons    as    identified    in

§ 884.5470 of this chapter.

Data show that toxic shock syn- drome (TSS), a rare but serious and sometimes fatal disease, is associated with the use of menstrual tampons. To protect the public and to minimize the serious adverse  effects of TSS,  men- strual tampons shall be labeled as set forth in paragraphs (c), (d), and (e) of this section and tested  for  absorbency as set forth in paragraph (f) of this sec- tion.

#### 21 CFR Ch. I (4–1–17 Edition)

If the information specified in paragraph (d) of this section is to be in- cluded as a package insert, the fol- lowing alert statement shall appear prominently and legibly on the pack- age label:

ATTENTION: Tampons are associated with Toxic Shock Syndrome (TSS). TSS is a rare but serious disease that may cause death. Read and save the enclosed information.

The labeling of menstrual tam- pons shall contain the following con- sumer information  prominently and legibly, in such terms as to render the information likely to be read and un- derstood by the ordinary individual under customary conditions of pur- chase and use:

(1)(i) Warning signs of TSS, e.g., sud- den fever (usually 102or more) and vomiting, diarrhea, fainting or near fainting when standing up, dizziness, or a rash that looks like a sunburn;

(ii) What to do if these or other signs of TSS appear, including the need to remove the tampon at once and seek medical attention immediately;

The risk of TSS  to  all  women using tampons during their menstrual period, especially the reported higher risks to women under 30 years of  age and teenage girls, the estimated inci- dence of TSS of 1 to 17 per 100,000 men- struating women  and  girls  per  year, and the risk of death from contracting TSS;

The advisability of using tampons with the minimum absorbency needed to control menstrual flow in order to reduce the risk of contracting TSS;

Avoiding the risk of getting tam- pon-associated TSS by not using tam- pons, and reducing the risk of getting TSS by alternating tampon use with sanitary napkin use during menstrual periods; and

The need to seek medical atten- tion before again using tampons if TSS warning signs have occurred in the past, or if women have any questions about TSS or tampon use.

The statements required by para- graph (e) of this section shall be promi- nently and  legibly placed  on the  pack- age label of menstrual tampons in con- formance with section 502(c) of the Federal  Food,  Drug,  and  Cosmetic  Act

#### Food and Drug Administration, HHS	§ 801.430

(the act) (unless the menstrual tam- pons are exempt under paragraph (g) of this  section).

  - Menstrual tampon package labels shall bear one of the following absorb- ency terms representing the absorb- ency of the production run,  lot,  or batch as measured by the test de- scribed in paragraph (f)(2) of this sec- tion;

| Ranges of absorbency in grams 1 | Corresponding term of ab- sorbency |
| --- | --- |
| 6 and under | Light absorbency |
| 6 to 9 | Regular absorbency |
| 9 to 12 | Super absorbency |
| 12 to 15 | Super plus absorbency |
| 15 to 18 | Ultra absorbency |
| Above 18 | No term |

1These ranges are defined, respectively, as follows: Less than or equal to 6 grams (g); greater than 6 g up to and in- cluding 9 g; greater than 9 g up to and including 12 g; greater than 12 g up to and including 15 g; greater than 15 g up to and including 18 g; and greater than 18 g.

  - The package label shall include an explanation of the ranges of absorb- ency and a description of how con- sumers can use a range of absorbency, and its corresponding absorbency term, to make comparisons of absorbency of tampons to allow selection of the tam- pons with the minimum absorbency needed to control menstrual flow in order to reduce the risk of contracting TSS.

A manufacturer shall measure the absorbency of individual tampons using the test method specified in paragraph (f)(2) of this section and calculate the mean absorbency of a production run, lot, or batch by rounding to the nearest

  - gram.

    - A manufacturer shall design and implement a sampling plan that in- cludes collection of probability  sam- ples of adequate size to yield consistent tolerance intervals such that the prob- ability is 90 percent that at least 90 percent of the absorbencies of indi- vidual tampons within a  brand  and type are within the range of absorb- ency stated on the package label.

    - In the absorbency test, an unlubricated condom, with tensile strength between 17 Mega  Pascals (MPa) and 30 MPa, as measured accord- ing to the procedure in the American Society    for    Testing    and    Materials

(ASTM) D 3492–97, ‘‘Standard Specifica- tion for Rubber Contraceptives (Male Condoms)’’ 1 for determining tensile strength, which is incorporated by ref- erence in accordance with  5  U.S.C. 552(a), is attached to the large end of a glass chamber (or a chamber  made from hard transparent plastic) with a rubber band (see figure 1) and pushed through the small end of the chamber using a smooth, finished rod. The condom  is  pulled  through  until   all slack is removed. The  tip  of  the condom  is  cut  off  and  the  remaining end of the condom is stretched over the end of the tube and secured with a rub- ber  band.  A  preweighed  (to  the  nearest

0.01 gram) tampon is placed within the condom  membrane  so  that  the  center of gravity of the tampon is at the cen- ter of the chamber. An infusion needle (14 gauge) is inserted through the sep- tum created by the condom tip until it contacts  the  end  of  the  tampon.  The outer   chamber   is   filled   with   water pumped from a temperature-controlled waterbath    to    maintain    the    average temperature at 271 C. The water re- turns to the waterbath as shown in fig- ure 2. Syngyna fluid (10 grams sodium chloride,   0.5   gram   Certified   Reagent Acid  Fushsin,  1,000  milliliters  distilled water) is then pumped through the in- fusion needle at a rate of 50 milliliters per  hour.  The  test  shall  be  terminated when the tampon is saturated and the first  drop  of  fluid  exits  the  apparatus. (The  test  result  shall  be  discarded  if fluid  is  detected  in  the  folds  of  the condom   before   the   tampon   is   satu- rated).  The  water  is  then  drained  and the   tampon   is   removed   and   imme- diately   weighed   to   the   nearest   0.01

1The Director of the Federal Register ap- proves this incorporation by reference in ac- cordance with 5 U.S.C. 552(a) and 1 CFR part

You may obtain a copy from the Amer- ican Society for Testing and Materials Inter- national, 100 Barr Harbor Dr., P.O. Box C700, West Conshohocken, PA 19428–2959, 610–832– 9578, . You may inspect a copy at the FDA Main Library, 10903 New Hamp- shire  Ave.,  Bldg.  2,  3d  floor,  Silver  Spring, MD 20993–0002, 301–796–2039, or at the Na- tional Archives and Records Administration (NARA). For information on the availability of this material at NARA, call 202–741–2139.

#### § 801.430

gram. The absorbency of the tampon is determined by subtracting its dry weight from this value. The condom shall be replaced after 10 tests or at the

#### 21 CFR Ch. I (4–1–17 Edition)

end of the day  during which  the condom is used in testing, whichever occurs  first.

#### Food and Drug Administration, HHS	§ 801.430

    - The Food and Drug Administra- tion may permit the use of an absorb- ency  test  method  different  from  the

test method specified in this section if each of the following conditions is met:

  - The  manufacturer  presents  evi- dence, in the form of a citizen petition

#### § 801.433

submitted in accordance with the re- quirements of § 10.30 of this chapter, demonstrating    that    the    alternative test method will yield results that are equivalent to the results yielded by the test method  specified  in  this  section; and

  - FDA approves the method and has published notice of its approval of the alternative test method in the FED- ERAL    REGISTER.

Any menstrual tampon intended to be dispensed by a vending machine is exempt  from  the  requirements  of this  section.

Any menstrual tampon that is not labeled as required by paragraphs (c), (d), and (e) of this section and that is initially  introduced  or  initially  deliv- ered for introduction into commerce after March  1,  1990,  is  misbranded under sections 201(n), 502 (a) and (f) of the   act.

(Information collection requirements con- tained in paragraphs (e) and (f) were ap- proved by the Office of Management and Budget under control number 0910–0257)

[47 FR 26989, June 22, 1982, as amended at 54

FR  43771,  Oct.  26,  1989;  55  FR  17600,  Apr.  26,

1990;  65  FR 3586,  Jan. 24,  2000;  65  FR 44436,

July 18, 2000; 65 FR 62284, Oct. 18, 2000; 69 FR

18803, Apr. 9, 2004; 69 FR 52171, Aug. 25, 2004;

75 FR 20914, Apr. 22, 2010]

##### § 801.433 Warning statements for pre- scription and restricted device products containing or manufac- tured with chlorofluorocarbons or other ozone-depleting substances.

(a)(1) All prescription and restricted device products containing or manufac- tured with chlorofluorocarbons, halons, carbon tetrachloride, methyl chloride, or any other class I substance des- ignated by the Environmental Protec- tion Agency (EPA) shall, except as pro- vided in paragraph (b) of this section, bear the following warning statement:

WARNING: Contains [or Manufactured with, if applicable] [insert name of substance], a sub- stance which harms public health and envi- ronment by destroying ozone in the upper at- mosphere.

(2) The warning statement shall be clearly legible and conspicuous on the product, its immediate container, its outer packaging, or other labeling in accordance with the requirements of 40 CFR   part   82   and   appear   with   such

#### 21 CFR Ch. I (4–1–17 Edition)

prominence and conspicuousness as to render it likely to be read and under- stood by consumers under normal con- ditions of purchase.

(b)(1) For prescription and restricted device   products,   the   following   alter- native warning statement may be used:

NOTE: The indented statement below is re- quired by the Federal government’s Clean Air Act for all products containing or manu- factured with chlorofluorocarbons  (CFC’s) [or name of other class I substance, if appli- cable]:

This product contains  [or is manufactured with, if applicable] [insert name of substance], a substance which harms the environment by destroying ozone in the upper atmosphere.

Your physician has determined that this product is likely to help your personal health. USE THIS PRODUCT AS DIRECTED, UNLESS  INSTRUCTED  TO  DO  OTHERWISE

BY YOUR PHYSICIAN. If you have any ques- tions about alternatives, consult with your physician.

The warning statement shall be clearly legible and conspicuous on the product, its immediate container, its outer packaging, or other labeling in accordance with the requirements of 40 CFR part 82 and appear with such prominence and conspicuousness as to render it likely to be read and under- stood by consumers under normal con- ditions of purchase.

If the  warning  statement in para- graph (b)(1) of this section is used, the following warning statement must be placed on the package labeling in- tended to be read by the physician (physician package insert) after the ‘‘How supplied’’ section, which de- scribes special handling and storage conditions on the physician labeling:

NOTE: The indented statement below is re- quired by the Federal government’s Clean Air Act for all products containing or manu- factured with chlorofluorocarbons  (CFC’s) [or name of other class I substance, if appli- cable]:

WARNING: Contains [or Manufactured with, if applicable] [insert name of substance], a sub- stance which harms public health and envi- ronment by destroying ozone in the upper at- mosphere.

A notice similar to the above WARNING has been placed in the information for the patient [or patient information leaflet, if ap- plicable] of this product under Environ- mental Protection Agency (EPA) regula- tions. The patient’s warning states that the

#### Food and Drug Administration, HHS	§ 801.435

patient should consult his or her physician if there are questions about alternatives.

(c) This section does not replace or relieve a person from any requirements imposed under 40 CFR part 82.

[61 FR 20101, May 3, 1996]

##### § 801.435   User	labeling	for	latex condoms.

This section applies to  the  subset of condoms as identified in § 884.5300 of this chapter, and condoms with spermicidal   lubricant   as   identified   in

§ 884.5310  of  this  chapter,  which  prod- ucts  are  formed  from  latex  films.

Data show that the material in- tegrity of latex condoms degrade over time. To protect the public health and minimize the risk of device failure, latex condoms must bear an expiration date which is supported by testing as described in paragraphs (d) and (h) of this section.

The expiration date, as dem- onstrated by testing procedures re- quired by paragraphs (d) and (h) of this section, must be displayed prominently and legibly on the primary packaging (i.e., individual package), and higher levels of packaging (e.g., boxes of condoms), in order to ensure visibility of the expiration date by consumers.

Except as provided under para- graph (f) of this section, the expiration date must be supported by data dem- onstrating physical and mechanical in- tegrity of the product after three dis- crete and representative lots of the product have been subjected to each of the following conditions:

Storage of unpackaged bulk prod- uct for the maximum amount of  time the manufacturer allows the product to remain unpackaged, followed by stor- age of the packaged product at 70 C (plus or minus 2 C) for 7 days;

Storage of unpackaged bulk prod- uct for the maximum amount of  time the manufacturer allows the product to remain unpackaged, followed by stor- age of the packaged product at a se- lected temperature between 40 and 50

C (plus or minus 2 C) for 90 days; and

Storage of unpackaged bulk prod- uct for the maximum amount of time the manufacturer allows the product to remain unpackaged, followed by stor- age of the packaged product at a mon- itored  or  controlled  temperature  be-

tween 15 and 30 C for the lifetime of the product (real time storage).

If a product fails the physical and mechanical integrity tests commonly used by industry after  the  completion of the accelerated storage tests de- scribed in paragraphs (d)(1) and (d)(2) of this section, the product expiration date must be demonstrated  by  real time storage conditions described in paragraph (d)(3) of this section. If all of the products tested after storage at temperatures as described in para- graphs (d)(1) and (d)(2) of this section pass the manufacturer’s physical and mechanical integrity tests, the manu- facturer may label the product with an expiration date of up to 5  years  from the date of product packaging. If the extrapolated expiration date under paragraphs (d)(1) and (d)(2) of this sec- tion is used, the labeled expiration date must be confirmed by physical and me- chanical integrity  tests  performed  at the end of the stated expiration period as described in paragraph (d)(3) of this section. If the data from tests fol- lowing real time storage described in paragraph (d)(3) of this section fails to confirm the extrapolated expiration date, the manufacturer must, at that time, relabel the product to reflect the actual  shelf  life.

Products that already have estab- lished shelf life data based upon real time storage and testing and have such storage and testing data available for inspection are not required  to  confirm such data using accelerated and inter- mediate aging data described in para- graphs (d)(1) and (d)(2) of this section. If, however, such real time expiration dates were based upon testing of prod- ucts that were  not  first  left unpackaged for the maximum  amount of time as described in paragraph (d)(3) of this section, the real time testing must be confirmed by testing products consistent with the requirements of paragraph (d)(3) of this section. This testing shall be initiated no later  than the effective date of this  regulation. Until the confirmation testing in ac- cordance with paragraph (d)(3) of this section is completed, the product may remain on the market labeled with the expiration date based upon  previous real  time  testing.

#### § 801.437

If a manufacturer uses testing data from one product to support expi- ration dating on any variation of that product, the manufacturer must docu- ment and provide, upon request, an ap- propriate justification for the applica- tion of the testing data to the vari- ation of the tested product.

If a latex condom contains a spermicide, and the expiration date based on spermicidal  stability  testing is different from the expiration date based upon latex integrity testing, the product shall bear only the earlier ex- piration  date.

The time period upon which the expiration date is based shall start with the date of packaging.

As provided in part 820 of this chapter, all testing data must be re- tained in each company’s  files,  and shall be made available upon request for inspection by the Food and Drug Administration.

Any latex condom not labeled with an expiration date as required by paragraph (c) of this section, and ini- tially delivered for introduction into interstate commerce after the effective date of this regulation is misbranded under sections 201(n) and 502(a) and (f) of Federal Food, Drug, and Cosmetic Act (21 U.S.C. 321(n) and 352(a) and (f)).

[62 FR 50501, Sept. 26, 1997]

##### § 801.437   User labeling for devices that contain natural rubber.

Data in the Medical Device Re- porting System and the scientific lit- erature indicate that some individuals are at risk of severe anaphylactic reac- tions to natural latex proteins. This la- beling regulation is intended to mini- mize the risk to  individuals  sensitive to natural latex proteins and protect the public health.

This section applies to all devices composed of or containing, or having packaging   or   components   that   are composed of, or contain, natural rubber that contacts humans. The term ‘‘nat- ural  rubber’’  includes  natural  rubber latex,  dry  natural  rubber,  and  syn- thetic  latex  or  synthetic  rubber  that contains natural rubber in its formula- tion.

  - The term  ‘‘natural rubber  latex’’ means rubber that is produced by the natural  rubber  latex  process  that  in-

#### 21 CFR Ch. I (4–1–17 Edition)

volves the use of natural latex in a con- centrated colloidal suspension. Prod- ucts are formed from natural rubber latex by dipping, extruding, or coating.

  - The term ‘‘dry natural rubber’’ means rubber that is produced by the dry natural rubber process that in- volves the use of coagulated natural latex in the form of dried or milled sheets. Products are formed from dry natural rubber by compression mold- ing, extrusion, or by converting the sheets into a solution for dipping.

  - The term ‘‘contacts humans’’ means that the natural rubber con- tained in a device is intended to con- tact or is likely to contact the user or patient. This includes contact when the device that contains natural rubber is connected to the patient by a liquid path or an enclosed gas path; or the de- vice containing the natural rubber is fully or partially coated with a powder, and such powder may carry natural rubber proteins that may contaminate the environment of the user or patient.

Devices containing natural rubber shall be labeled as set forth in para- graphs (d) through (h) of this section. Each required labeling statement shall be prominently and legibly  displayed in conformance with section 502(c) of the Federal Food, Drug, and Cosmetic Act (the act) (21 U.S.C. 352(c)).

Devices containing natural rubber latex that contacts humans, as de- scribed in paragraph (b) of this section, shall bear the following statement in bold print on the device labeling:

‘‘Caution: This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions.’’

This statement shall appear on all de- vice labels, and other labeling, and shall appear on the principal display panel of the device packaging, the out- side package, container or wrapper, and the immediate device package, container, or wrapper.

Devices containing dry natural rubber that contacts humans, as de- scribed in paragraph (b) of this section, that are not already subject to para- graph (d) of this section, shall bear the following statement in bold print on the device labeling:

‘‘This Product Contains Dry Natural Rubber.’’

#### Food and Drug Administration, HHS	Pt. 803

This statement shall appear on all de- vice labels, and other labeling, and shall appear on the principal display panel of the device packaging, the out- side package, container or wrapper, and the immediate device package, container, or wrapper.

Devices that have packaging con- taining natural rubber latex that con- tacts humans, as described in para- graph (b) of this section, shall bear the following statement in bold print on the device labeling:

‘‘Caution: The Packaging of This Product Contains Natural Rubber Latex Which May Cause Allergic Reac- tions.’’

This statement shall appear on the packaging that contains the natural rubber, and the outside package, con- tainer, or wrapper.

Devices that have packaging con- taining dry natural rubber that con- tacts humans, as described in para- graph (b) of this section, shall bear the following statement in bold print on the device labeling:

‘‘The Packaging of This Product Con- tains Dry Natural Rubber.’’

This statement shall appear on the packaging that contains the natural rubber, and the outside package, con- tainer, or wrapper.

Devices that contain natural rub- ber that contacts humans, as described in paragraph (b) of this section, shall not contain the term ‘‘hypoallergenic’’ on their labeling.

Any affected person may request an exemption or variance from the re- quirements of this  section  by  submit- ting a citizen petition  in  accordance with § 10.30 of this chapter.

Any device subject to this section that is not labeled in accordance with paragraphs (d) through (h) of this sec- tion and that is initially introduced or initially delivered  for introduction into interstate commerce after the effective date of this regulation is misbranded under sections 201(n) and 502(a), (c), and

(f) of the act (21 U.S.C. 321(n)  and 352(a), (c), and (f)).

NOTE TO § 801.437: Paragraphs (f) and (g) are stayed  until  June  27,  1999,  as  those  regula-

tions  relate  to  device  packaging  that  uses ‘‘cold seal’’ adhesives.

[62 FR 51029, Sept. 30, 1997, as amended at 63

FR 46175, Aug. 31, 1998]

## PART 803—MEDICAL DEVICE REPORTING

##### Subpart A—General Provisions

Sec.

803.1    What does this part cover?

803.3 How does FDA define the terms used in this part?

  - What information from the reports do we disclose to the public?

  - Generally, what  are  the  reporting  re- quirements that apply to me?

  - What form should I use to submit re- ports of individual adverse events and where do I obtain these forms?

  - How do I submit initial and supple- mental or followup reports?

  - Do I need to submit reports in English?

  - How will I know if you require more information about my medical device re- port?

  - When I submit a report, does the in- formation in my report constitute an ad- mission that the device caused or con- tributed to the reportable event?

  - What are the requirements for devel- oping, maintaining, and implementing written MDR procedures that apply to me?

  - What are the requirements for estab- lishing and maintaining MDR files or records that apply to me?

  - Are there exemptions, variances, or alternative forms of  adverse event  re- porting requirements?

##### Subpart B—Generally Applicable Require- ments for Individual Adverse Event Re- ports

  - How do I complete and submit an in- dividual adverse event report?

  - Where can I find the reporting codes for adverse events that I use with  med- ical device reports?

  - What are the circumstances in which I am not required to file a report?

  - Where can I find information on how to prepare and submit an MDR in elec- tronic format?

##### Subpart C—User Facility Reporting Requirements

803.30 If I am a user facility, what reporting requirements apply to me?

  - If I am a user facility, what informa- tion must I submit in my individual ad- verse event reports?

§ 803.1

  - If I am a user facility, what must I in- clude when I submit an annual report?

##### Subpart D—Importer Reporting Requirements

803.40 If I am an importer, what reporting requirements  apply  to  me?

803.42 If I am an importer,  what  informa- tion must I submit in my individual ad- verse event reports?

##### Subpart E—Manufacturer Reporting Requirements

803.50 If I am  a  manufacturer,  what  report- ing  requirements  apply  to  me?

  - If I am a manufacturer, what infor- mation must I submit in my individual adverse event reports?

  - If I am a manufacturer, in which cir- cumstances must I submit a 5-day re- port?

803.56 If I am a manufacturer, in what cir- cumstances must I  submit  a  supple- mental or followup report and what  are the  requirements  for  such  reports?

  - Foreign manufacturers.

AUTHORITY: 21 U.S.C. 352, 360, 360i, 360j, 371,

374.

SOURCE:  79  FR  8846,  Feb.  14,  2014,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 803.1   What does this part cover?

    - This part establishes the require- ments for medical device reporting for device user facilities, manufacturers, importers, and distributors. If you are a device user facility, you must report deaths and serious injuries that a de- vice has or may have caused or contrib- uted to, establish and maintain adverse event files, and submit summary an- nual reports. If you are a manufacturer or importer, you must report deaths and serious injuries that your device has or may have caused or contributed to, you must report certain device mal- functions, and you must establish and maintain adverse event files. If you are a manufacturer, you must also submit specified followup. These reports help us to protect the public health by help- ing to ensure that devices are not adul- terated or misbranded and are safe and effective for their intended use. If you are a medical device distributor, you must maintain records (files) of inci- dents, but you are not required to re- port these incidents.

21 CFR Ch. I (4–1–17 Edition)

    - This part supplements and does not supersede other provisions of this chapter, including the provisions  of part 820 of this chapter.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

##### § 803.3   How does FDA define the terms used in this part?

Some of the terms we use in this part are specific to medical device reporting and reflect the language used in the statute (law). Other terms are more general and reflect our interpretation of the law. This section defines the fol- lowing terms as used in this part:

Ambulatory surgical facility (ASF) means a distinct entity that operates for the primary purpose of furnishing same day outpatient surgical services to patients. An ASF may be either an independent entity (i.e., not a part of a provider of services or any other facil- ity) or operated by another medical en- tity (e.g., under the common owner- ship, licensure, or control of an entity). An ASF is subject to this regulation regardless of whether it is licensed by a Federal, State, municipal, or local gov- ernment or regardless of whether it is accredited by a recognized accredita- tion organization. If an adverse event meets the criteria for reporting, the ASF must report that event regardless of the nature or location of the medical service provided by the ASF.

Become aware means that an em- ployee of the entity required to report has acquired information that reason- ably suggests a reportable adverse event has occurred.

  - If you  are a  device  user facility, you are considered to have ‘‘become aware’’ when medical personnel, as de- fined in this section, who are employed by or otherwise formally affiliated with your facility, obtain information about a reportable event.

  - If you are a manufacturer, you are considered to have become aware of an event when any of your employees be- comes aware of a reportable event that is required to be reported within 30 cal- endar days or that is required to be re- ported within 5 work days because we had requested reports in accordance with § 803.53(b). You are also considered

#### Food and Drug Administration, HHS	§ 803.3

to have become aware of an event when any of your employees with manage- ment or supervisory responsibilities over persons with regulatory, sci- entific, or technical responsibilities, or whose duties relate to the collection and reporting of adverse events, be- comes aware, from any information, in- cluding any trend analysis, that a re- portable MDR event or events neces- sitates remedial action to prevent an unreasonable risk of substantial harm to the public health.

  - If you are an importer, you are considered to have become aware of an event when any of your employees be- comes aware of a reportable event that is required to be reported by you with- in 30 days.

Caused or contributed means that a death or serious injury was or may have been attributed to a medical de- vice, or that a medical device was or may have been a factor in a death or serious injury, including events occur- ring as a result of:

  - Failure,

  - Malfunction,

  - Improper or inadequate design,

  - Manufacture,

  - Labeling, or

  - User error.

Device user facility means a hos- pital, ambulatory surgical facility, nursing home, outpatient diagnostic facility, or outpatient treatment facil- ity as defined in this section, which is not a physician’s office, as defined in this section. School nurse offices and employee health units are not device user facilities.

Distributor means any person (other than the manufacturer or im- porter) who furthers the marketing of a device from the original place of manufacture to the person who makes final delivery or sale to the ultimate user, but who does not repackage or otherwise change the container, wrap- per, or labeling of the device or device package. If you repackage or otherwise change the container, wrapper, or la- beling, you are considered a manufac- turer as defined in this section.

Expected life of a device means the time that a device is expected to re- main functional after it is placed into use. Certain implanted devices have specified  ‘‘end  of  life’’  (EOL)  dates.

Other devices are not labeled as to their respective EOL, but are expected to remain operational through activi- ties such as maintenance, repairs, or upgrades, for an estimated period of time.

FDA, we, us, or Agency means the Food and Drug Administration.

Five-day report means a medical device  report  that  must  be   submitted by a manufacturer to us under § 803.53 within 5 work days.

Hospital means a distinct entity that operates for the  primary  purpose of providing diagnostic, therapeutic (such as medical, occupational, speech, physical), surgical, and other patient services for specific and general med- ical conditions. Hospitals include gen- eral, chronic disease, rehabilitative, psychiatric, and other special-purpose facilities. A hospital may be either independent (e.g., not a part of a pro- vider of services or  any  other  facility) or may be operated by another medical entity (e.g.,  under the common  owner- ship, licensure, or control of another entity). A hospital is covered by this regulation regardless of whether it is licensed by a Federal,  State,  municipal or local government or whether it is accredited by a recognized accredita- tion organization. If an adverse event meets the criteria for reporting, the hospital must report that event regard- less of the nature or location of the medical service provided by the hos- pital.

Importer means any person who imports a device into the United States and who furthers the marketing of a device from the original place of manu- facture to  the  person  who makes  final delivery or sale to the ultimate user, but who does not repackage or other- wise change the container, wrapper, or labeling of the device or device pack- age. If you repackage or otherwise change the container, wrapper, or la- beling, you are considered a manufac- turer as defined in this section.

Malfunction means the failure of a device to meet its performance speci- fications or otherwise perform as in- tended. Performance specifications in- clude all claims made in the  labeling for the device. The intended perform- ance of a device refers to the intended use for which the device is labeled or

#### § 803.3

marketed,  as  defined  in  § 801.4  of  this chapter.

Manufacturer means any person who manufactures, prepares, propa- gates, compounds, assembles, or proc- esses a device by chemical, physical, biological, or other procedure. The term includes any person who either:

  - Repackages or otherwise changes the container, wrapper, or labeling of a device in furtherance of the distribu- tion of the device from the original place of manufacture;

  - Initiates specifications for devices that are manufactured by a second party for subsequent distribution by the person initiating the specifica- tions;

  - Manufactures components or ac- cessories that are devices that are ready to be used and are intended to be commercially distributed and intended to be used as is, or are processed by a licensed practitioner or other qualified person to meet the needs of a par- ticular patient; or

  - Is the U.S. agent of a foreign man- ufacturer.

Manufacturer or importer report number. This number uniquely identi- fies each individual adverse event re- port submitted by a manufacturer or importer. This number consists of the following three parts:

The FDA registration number for the manufacturing site of the reported device, or the registration number for the importer. If the  manufacturing site or the importer does not have an estab- lishment registration number, we will assign a temporary MDR reporting number until the site is registered in accordance with part 807 of this chap- ter. We will inform  the  manufacturer or importer of the temporary MDR re- porting number;

The four-digit calendar year in which the report is submitted; and

The five-digit sequence number of the reports submitted during the year, starting with 00001. (For example, the complete  number  will  appear   as   fol- lows:  1234567–2011–00001.)

MDR means medical device re- port.

MDR reportable event (or reportable event) means:

  - An event that user facilities be- come   aware   of   that   reasonably   sug-

#### 21 CFR Ch. I (4–1–17 Edition)

gests that a device has or may have caused or contributed to a death or se- rious injury or

  - An event that manufacturers or importers become aware of that rea- sonably suggests that one of their mar- keted devices:

May have caused or contributed to a death or serious injury, or

Has malfunctioned and that the device or a similar device marketed by the manufacturer or importer would be likely to cause or contribute to a death or serious injury if the malfunction were to recur.

Medical personnel means an indi- vidual who:

  - Is licensed, registered, or certified by a State, territory, or other gov- erning body, to administer health care;

  - Has received a diploma or a degree in a professional or  scientific  dis- cipline;

  - Is an employee responsible for re- ceiving medical complaints or adverse event reports; or

  - Is a supervisor of these persons.

Nursing home means:

  - An independent entity (i.e., not a part of a provider of services or any other facility) or one operated by an- other medical entity (e.g., under the common ownership, licensure, or con- trol of an entity) that operates for the primary purpose of providing:

    - Skilled nursing care and related services for persons who require med- ical or nursing care;

    - Hospice care to the terminally ill; or

    - Services for the rehabilitation of the injured, disabled, or sick.

  - A nursing home is subject to this regulation regardless of whether it is licensed by a Federal, State, munic- ipal, or local government or whether it is accredited by a recognized accredita- tion organization. If an adverse event meets the criteria for reporting, the nursing home must report that event regardless of the nature or location of the medical service provided by the nursing home.

Outpatient	diagnostic	facility

means:

A distinct entity that:

  - Operates for the primary purpose of conducting medical diagnostic tests on patients,

#### Food and Drug Administration, HHS	§ 803.3

  - Does not assume ongoing respon- sibility for patient care, and

  - Provides its services for use by other medical personnel.

Outpatient diagnostic facilities include outpatient facilities providing radiography,	mammography, ultrasonography, electrocardiography, magnetic resonance imaging, comput- erized axial tomography, and in vitro testing. An outpatient diagnostic facil- ity may be either independent (i.e., not a part of a provider of services or any other facility) or operated by another medical entity (e.g., under the common ownership, licensure, or control of an entity). An outpatient diagnostic facil- ity is covered by this regulation re- gardless of whether it is licensed by a Federal, State, municipal, or local gov- ernment or whether it is accredited by a recognized accreditation organiza- tion. If an adverse event meets the cri- teria for reporting, the outpatient di- agnostic facility must  report  that event regardless of the nature or loca- tion of the medical service provided by the outpatient diagnostic facility.

Outpatient treatment facility means a distinct entity that operates for the primary purpose of providing nonsur- gical therapeutic (medical, occupa- tional, or physical) care on an out- patient basis or in a home health care setting. Outpatient treatment facili- ties include ambulance providers, res- cue services, and home health care groups. Examples of services provided by outpatient treatment facilities in- clude	the	following:	Cardiac defibrillation, chemotherapy, radio- therapy, pain control, dialysis, speech or physical therapy, and treatment for substance abuse. An outpatient treat- ment facility may be either inde- pendent (i.e., not a part of a provider of services or any other facility) or oper- ated by another medical entity (e.g., under the common ownership, licen- sure, or control of an entity). An out- patient treatment facility is covered by this regulation regardless of wheth- er it is licensed by a Federal, State, municipal, or local government or whether it is accredited by a recog- nized accreditation organization. If an adverse event meets the criteria for re- porting, the outpatient treatment fa- cility  must  report  that  event  regard-

less of the nature or location of the medical service provided by the out- patient  treatment  facility.

Patient of the facility means any individual who is being diagnosed or treated and/or receiving medical  care at or under the control or authority of the facility. This includes employees of the  facility  or  individuals  affiliated with the facility who, in the course of their duties, suffer a device-related death or serious injury that has or may have been caused or contributed to by a device used at the facility.

Physician’s office means a facility that operates as the office of a physi- cian or other health care professional for the primary purpose of examina- tion, evaluation, and treatment or re- ferral of patients. Examples of physi- cian offices include: Dentist offices, chiropractor offices, optometrist of- fices, nurse practitioner offices, school nurse offices, school clinics, employee health clinics, or freestanding care units. A physician’s office may be inde- pendent, a group practice, or part of a Health Maintenance Organization.

Remedial action means any action other than routine maintenance or servicing of a device where such action is necessary to prevent recurrence of a reportable event.

Serious injury means an injury or illness that:

Is life-threatening,

Results in permanent impairment of a body function or permanent dam- age to a body structure, or

Necessitates medical or surgical intervention to preclude permanent impairment of a body function or per- manent damage to a body structure. Permanent means irreversible impair- ment or damage to a body structure or function, excluding trivial impairment or damage.

User facility report number means the number that uniquely identifies each report submitted by a user facil- ity to manufacturers and to us. This number consists of the following three parts:

  - The user facility’s 10-digit Centers for Medicare and Medicaid Services (CMS) number (if the CMS number has fewer than 10 digits, fill the remaining spaces with zeros);

#### § 803.9

  - The four-digit calendar year in which the report is submitted; and

  - The four-digit sequence number of the reports submitted for the year, starting with 0001. (For example, a complete  user  facility  report  number will appear as follows: 1234560000–2011– 0001. If a user facility has more  than one CMS number, it must  select  one that will be used for all of its MDR re- ports. If a user facility has no CMS number, it should use all zeros in the appropriate space in its  initial  report (e.g., 0000000000–2011–0001). We will as- sign a number for future use and send that number to the user facility. This number is used in our record of the  ini- tial  report,  in  subsequent  reports,  and in any correspondence with the user fa- cility. If a facility has multiple  sites, the  primary  site  may  submit    reports for all sites and use one reporting num- ber for all sites if the primary site pro- vides the name, address, and CMS  num- ber  for  each  respective  site.)

Work day means Monday through Friday, except Federal holidays.

[Reserved]

(aa) Human cell, tissue, or cellular or tissue-based product (HCT/P) regulated as a device means an HCT/P as defined in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

(bb) Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A unique device identifier is composed of:

  - A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

  - A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

    - The lot or batch within which a device was manufactured;

    - The serial number of a specific device;

    - The expiration date of a specific device;

    - The date a specific device was manufactured.

#### 21 CFR Ch. I (4–1–17 Edition)

    - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

##### § 803.9   What information from the re- ports do we disclose to the public?

We may disclose to the public any report, including any FDA record of a telephone report, submitted under this part. Our disclosures are governed by part 20 of this chapter.

Before we disclose a report to the public, we will delete the following:

  - Any information that constitutes trade secret or confidential commer- cial   or   financial   information   under

§ 20.61 of this chapter;

  - Any  personal,  medical,  and  simi- lar information, including the serial number of implanted devices,  which would constitute an invasion of  per- sonal privacy under § 20.63 of this chap- ter. However,  if a patient requests a  re- port,  we  will  disclose  to   that   patient all the information in the report con- cerning   that   patient,   as   provided   in

§ 20.61 of this chapter; and

  - Any names and other identifying information of a third party that vol- untarily submitted an adverse event report.

We may not disclose  the  identity of a device user facility that makes a report under this part except in con- nection with:

  - An action brought to enforce sec- tion 301(q) of the Federal Food, Drug, and Cosmetic Act (21 U.S.C. 331(q)), in- cluding the failure or refusal to furnish material or information required by section 519 of the Federal Food, Drug, and Cosmetic Act (21 U.S.C. 360i));

  - A communication to a manufac- turer of  a  device  that  is  the  subject  of a report required to be submitted by a user facility under § 803.30; or

  - A disclosure to employees of the Department of Health and Human Services, to the Department  of  Justice, or to the duly authorized committees and subcommittees of the Congress.

#### Food and Drug Administration, HHS	§ 803.12

##### § 803.10 Generally, what are the re- porting requirements that apply to me?

    - If you are a device user facility, you must submit reports (described in subpart C of this part), as follows:

      - Submit reports of individual ad- verse events no later than 10 work days after the day that you become aware of a reportable event:

        - Submit reports of device-related deaths to us and to the manufacturer, if known, or

        - Submit reports of device-related serious injuries to the manufacturers or, if the manufacturer is unknown, submit reports to us.

      - Submit  annual  reports  (described in § 803.33) to us.

    - If you are an importer, you must submit reports (described in subpart D of this part), as follows:

      - Submit reports of individual ad- verse events no later than 30 calendar days after the day that you become aware of a reportable event:

        - Submit reports of device-related deaths or serious injuries to us and to the manufacturer or

        - Submit reports of device-related malfunctions to the manufacturer.

      - [Reserved]

    - If you are a manufacturer, you must submit reports (described in sub- part E of this part) to us, as follows:

      - Submit reports of individual ad- verse events no later than 30 calendar days after the day that you become aware of a reportable death, serious in- jury,  or  malfunction.

      - Submit reports of individual ad- verse events no later than 5 work days after the day that you  become  aware of:

        - A reportable event that requires remedial action to prevent an unrea- sonable risk of substantial harm to the public health or

        - A reportable event for which we made a written request.

      - Submit supplemental reports if you obtain information that you did not submit in an initial report.

##### § 803.11 What form should I use to sub- mit reports of individual adverse events and where do I obtain these forms?

If you are a manufacturer or im- porter, you must submit reports of in- dividual adverse events to FDA in an electronic  format  in  accordance  with

§ 803.12(a) and § 803.20, unless granted an exemption under § 803.19.

Importer reports submitted to de- vice manufacturers may be in paper format or an electronic format that in- cludes all required data fields to ensure that the manufacturer has all required information.

If you are a user facility, you must submit reports of individual ad- verse    events    in    accordance    with

§ 803.12(b) and § 803.20.

Form FDA 3500A is available on the or  from  Division of International and Consumer Edu- cation, Center for Devices and Radio- logical Health, Food and Drug Admin- istration, 10903 New Hampshire Ave., Bldg. 66, Rm. 4621, Silver Spring, MD 20993–0002, by email: DICE@fda.hhs.gov, FAX: 301–847–8149, or telephone: 800– 638–2041.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

##### § 803.12 How do I submit initial and supplemental or followup reports?

Manufacturers and importers must submit initial and supplemental or followup reports to FDA in an elec- tronic format that FDA can process, review, and archive.

User  facilities  that  submit  their reports and additional information to FDA  electronically  must  use  an  elec- tronic  format  that  FDA  can  process, review, and archive. User facilities that submit their reports to FDA on paper must submit any written report or ad- ditional   information   required   under this  part  to  FDA,  CDRH,  Medical  De- vice  Reporting,  P.O.  Box  3002,  Rock- ville, MD 20847–3002, using Form FDA 3500A. Each report must be identified (e.g.,  ‘‘User  Facility  Report’’  or  ‘‘An- nual Report’’).

If you are confronted with a pub- lic   health   emergency,   this   can   be

#### § 803.13

brought to FDA’s  attention  by  con- tacting FDA’s Office of Crisis Manage- ment, Emergency Operations Center by telephone, 24-hours a day, at 301–796– 8240 or toll free at 866–300–4374, followed by the submission of an email to: emer- gency.operations@fda.hhs.gov.

NOTE: This action does not satisfy your ob- ligation to report under part 803.

You may submit a voluntary tele- phone report to the MedWatch office at 800–FDA–1088. You may also obtain in- formation  regarding  voluntary  report- ing  from  the  MedWatch  office  at  800– FDA–1088.  You  may  also  find  the  vol- untary Form FDA 3500 and instructions to   complete   it   .

##### § 803.13 Do I need to submit reports in English?

Yes. You must submit all reports re- quired by this part in English.

##### § 803.15 How will I know if you require more information about my medical device report?

We will notify you in writing if we require additional information and will tell you what information we need. We will require additional information if we determine that protection of the public health requires additional or clarifying information for medical de- vice reports submitted to us and in cases when the additional information is beyond the scope of FDA reporting forms or is not readily accessible to us.

In any request under this section, we will state the reason or purpose for the information request, specify the due date for submitting the informa- tion, and clearly identify the reported event(s) related to our request. If we verbally request additional informa- tion, we will confirm the request in writing.

##### § 803.16 When I submit a report, does the information in my report con- stitute an admission that the device caused or contributed to the report- able event?

No. A report or other information submitted by you, and our release of that report or information, is not nec- essarily an admission that the device, or you or your employees, caused or contributed  to  the  reportable  event.

#### 21 CFR Ch. I (4–1–17 Edition)

You do not have to admit and may deny that the report or information submitted under this part constitutes an admission that the device, you, or your employees, caused or contributed to a reportable event.

##### § 803.17 What are the requirements for developing, maintaining, and imple- menting written MDR procedures that apply to me?

If you are a  user  facility,  importer, or manufacturer, you must develop, maintain, and implement written MDR procedures for the following:

Internal systems that provide for:

  - Timely and effective identifica- tion, communication, and evaluation of events that may be subject to MDR re- quirements;

  - A standardized review process or procedure for determining when an event meets the criteria for reporting under this part; and

  - Timely transmission of complete medical device reports to manufactur- ers or to us, or to both if required.

Documentation and record- keeping requirements for:

Information that was evaluated to determine if an event was reportable;

All medical device reports and in- formation submitted to manufacturers and/or us;

Any information that was evalu- ated for the purpose of preparing the submission of annual reports; and

Systems that ensure access to in- formation that facilitates timely fol- lowup and inspection by us.

##### § 803.18 What are the requirements for establishing and maintaining MDR files or records that apply to me?

If you are a user facility, im- porter, or manufacturer, you must es- tablish and maintain MDR event files. You must clearly identify all MDR event files and maintain them to facili- tate timely access.

(b)(1) For purposes of this part, ‘‘MDR event files’’ are written or elec- tronic files maintained by user facili- ties, importers, and manufacturers. MDR event files may incorporate ref- erences to other information (e.g., medical records, patient files, engi- neering reports), in lieu of copying and maintaining duplicates in this file. Your MDR event files must contain:

#### Food and Drug Administration, HHS	§ 803.19

  - Information in your possession or references to  information  related  to the adverse event, including all docu- mentation of your deliberations and decision making processes used to de- termine if a device-related death, seri- ous injury, or malfunction was or was not reportable under this part;

  - Copies of all reports submitted under this part (whether paper or elec- tronic), and of all other information re- lated to the event that you submitted to us or other entities such as an im- porter, distributor, or manufacturer; and

  - Copies of all electronic acknowl- edgments FDA sends you in response to electronic MDR submissions.

(2) If you are a user facility, im- porter, or manufacturer, you must per- mit any authorized FDA employee, at all reasonable times, to access, to copy, and to verify the records required by this part.

(c) If you are a user facility, you must retain an MDR event file relating to an adverse event for a period of 2 years from the date of the event. If you are a manufacturer or importer, you must retain an MDR event file relating to an adverse event for a period of 2 years from the date of the event or a period of time equivalent to the ex- pected life of the device, whichever is greater. If the device is no longer dis- tributed, you still must maintain MDR event files for the time periods de- scribed in this paragraph (c).

(d)(1) If you are a device distributor, you must establish and maintain de- vice complaint records (files). Your records must contain any incident in- formation, including any written, elec- tronic, or oral communication, either received or generated by you, that al- leges deficiencies related to the iden- tity (e.g., labeling), quality, durability, reliability, safety, effectiveness, or per- formance of a device. You must also maintain information about your eval- uation of the allegations, if any, in the incident record. You must clearly iden- tify the records as device incident records and file these records by device name. You may maintain these records in written or electronic format. You must back up any file maintained in electronic format.

You must retain copies of the re- quired device incident records for a pe- riod of 2 years from the date of inclu- sion of the record in the file or for a pe- riod of time equivalent to the expected life of the device, whichever is greater. You must maintain copies of these records for this period even if you no longer distribute the device.

You must maintain the device complaint files established under this section at your principal business es- tablishment. If you are also a manufac- turer, you may maintain the file at the same location as you maintain your complaint file under part 820 of this chapter. You must permit any author- ized FDA employee, at all reasonable times, to access, to copy, and to verify the records required by this part.

(e) If you are a manufacturer, you may maintain MDR event  files  as  part of your complaint file, under part 820 of this chapter, if you prominently iden- tify these records as MDR reportable events. We will not consider your sub- mitted MDR report to comply with this part unless you evaluate an event in accordance with the quality system re- quirements described in part 820 of this chapter. You must document and main- tain in your MDR event files an expla- nation of why you did not submit or could not obtain any information re- quired by this part, as well as the re- sults  of  your  evaluation  of  each  event.

##### § 803.19 Are there exemptions, variances, or alternative forms of adverse event reporting require- ments?

We exempt the following persons from the adverse event reporting re- quirements in this part:

  - A licensed practitioner who pre- scribes or administers devices intended for use in humans and manufactures or imports devices solely for use in diag- nosing and treating persons with whom the practitioner has a ‘‘physician-pa- tient’’ relationship;

  - An individual who manufactures devices intended for use in humans solely for this person’s use in research or teaching and not for sale. This in- cludes any person who is subject to al- ternative reporting requirements under the investigational device exemption regulations  (described  in  part  812  of

#### § 803.20

this chapter), which require  reporting of all adverse device effects; and

  - Dental laboratories or optical lab- oratories.

If you are a manufacturer, im- porter, or user facility, you may re- quest an exemption or variance from any or all of the reporting require- ments in this part, including the re- quirements of § 803.12. You must submit the request to us in writing at the fol- lowing address: MDR Exemption Re- quests, Office of Surveillance and Bio- metrics, 10903 New Hampshire Ave., Bldg. 66, Rm. 3217, Silver Spring, MD 20993–0002. Your request must include information necessary to identify you and the device; a complete statement of the request for exemption, variance, or alternative reporting; and an expla- nation why your request is justified. If you are requesting an exemption from the requirement to submit reports to FDA    in    electronic    format    under

§ 803.12(a), your request should indicate for how long you will require this ex- emption.

If you are a manufacturer, im- porter, or user facility, we may grant in writing an exemption or variance from, or alternative to, any or  all  of the reporting requirements in this part, and may change the frequency of reporting to quarterly, semiannually, annually or other appropriate time pe- riod. We may grant these modifications in response to your request, as de- scribed in paragraph (b) of this section, or at our discretion. When we grant modifications to the reporting require- ments, we may impose other reporting requirements to ensure the protection of public health.

We may revoke or modify in writ- ing an exemption, variance, or alter- native reporting requirement if we de- termine that revocation or  modifica- tion is necessary to protect the public health.

If we grant your request for a re- porting modification, you must submit any reports or information required in our approval of the modification. The conditions of the approval will replace and supersede the regular reporting re- quirement specified  in  this  part  until such time that we revoke or modify the alternative reporting requirements in accordance with paragraph (d) of this

#### 21 CFR Ch. I (4–1–17 Edition)

section or until the date specified in our response granting your variance, at which time the provisions of this part will again apply.

## Subpart B—Generally Applicable Requirements for Individual Adverse Event Reports

##### § 803.20 How do I complete and submit an individual adverse event report?

What form must I complete and sub- mit?

  - If you are a health professional or consumer or other entity, you may submit voluntary reports to FDA re- garding devices or other FDA-regulated products using the Form FDA 3500.

  - To submit a mandatory report in written form, a user facility must use Form FDA 3500A.

  - An electronic submission of a mandatory report from a user facility, importer, or manufacturer must con- tain the information from the applica- ble blocks of Form FDA 3500A. All elec- tronic submissions must include infor- mation about the patient, the event, the device,  and the  ‘‘initial  reporter.’’ An electronic submission from a user facility or importer must include the information from  block  F.  An  elec- tronic submission from a manufacturer must include the information from blocks G and H. If you are a manufac- turer and you receive a report from a user facility or importer, you must in- corporate that information in your electronic submission and include any corrected or missing information.

To whom must I submit reports and when?

  - If you are a user facility, you must submit MDR reports to:

    - The manufacturer and to us no later than 10 work days after the day that you become  aware of information that reasonably suggests that a device has or may have caused or contributed to a death or

    - The manufacturer no later than 10 work days after the day that you be- come aware of information that rea- sonably suggests that a device has or may have caused or contributed to a serious injury. If the manufacturer is not known, you must submit this re- port to us.

Food and Drug Administration, HHS	§ 803.22

  - If you are an importer, you must submit MDR reports to:

    - The manufacturer and to us, no later than 30 calendar days after the day that you become aware of informa- tion that reasonably suggests that a device has or may have caused or con- tributed to a death or serious injury or

    - The manufacturer, no later than

30 calendar days after receiving infor- mation that a device you market has malfunctioned and that this device or a similar device that you market  would be likely to cause or contribute to a death or serious injury if the malfunc- tion were to recur.

  - If you are a manufacturer, you must submit MDR reports to us:

    - No later than 30 calendar days after the day that you become aware of information that reasonably suggests that a device may have caused or con- tributed to a death or serious injury or

    - No later than 30 calendar days after the day that you become aware of information that reasonably suggests a device has malfunctioned and that this device or a similar device that you market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur; or

    - Within 5 work days if required by

§ 803.53.

What kind of information reason- ably suggests that a reportable event has occurred?

  - Any information, including pro- fessional, scientific, or medical facts, observations, or opinions, may reason- ably suggest that a device has caused or may have caused or contributed to an MDR reportable event. An MDR re- portable event is a death, a serious in- jury, or, if you are a manufacturer or importer, a malfunction that would be likely to cause or contribute to a death or serious injury if the malfunction were to recur.

  - If you are a user facility, im- porter, or manufacturer, you do not have to report an adverse event if you have information that would lead a person who is qualified to make a med- ical judgment  reasonably  to  conclude that a device did not cause or con- tribute to a death or serious injury, or that a malfunction would not be likely to cause or contribute to a death or se- rious injury if it were to recur. Persons

qualified to make a medical judgment include physicians, nurses, risk man- agers, and biomedical engineers. You must keep in your MDR event files (de- scribed in § 803.18) the information that the qualified person used to determine whether or not a device-related event was  reportable.

##### § 803.21 Where can I find the reporting codes for adverse events that I use with medical device reports?

The    MedWatch    Medical    Device Reporting     Code     Instruction     Manual contains   adverse   event   codes   for   use with  Form  FDA  3500A.  You  may  obtain the   coding   manual   from   FDA’s   Web site	;	and from   the   Division   of   Small   Manufac- turers,  International  and  Consumer  As- sistance,  Center  for  Devices  and  Radio- logical   Health,   10903   New   Hampshire Ave.,  Bldg.  66,  Rm.  4621,  Silver  Spring, MD 20993–0002, FAX: 301–847–8149, or email

We may sometimes use additional coding of information on the reporting forms or modify the existing codes. If we do make modifications, we will en- sure that we make the new coding in- formation available to all reporters.

##### § 803.22 What are the circumstances in which I am not required to file a re- port?

If you become aware of informa- tion from multiple sources regarding the same patient and same reportable event, you may submit one medical de- vice report.

You are not required to submit a medical device report if:

  - You are a user facility, importer, or manufacturer, and you determine that the information received is erro- neous in that a device-related adverse event did not occur. You must retain documentation of these reports in your MDR files for the time periods speci- fied in § 803.18.

  - You are a manufacturer or im- porter and you did not manufacture or import the device about which you have adverse event information. When

#### § 803.23

you receive reportable event informa- tion in error, you must forward this in- formation to us with a cover letter ex- plaining that you did not manufacture or import the device in question.

##### § 803.23 Where can I find information on how to prepare and submit an MDR in electronic format?

    - You may obtain information on how to prepare and submit  reports  in an electronic format that FDA can process, review, and archive at:

    - We may sometimes update infor- mation on how to prepare and submit reports electronically. If we do make modifications, we will ensure that we alert reporters by updating the eMDR Web page.

## Subpart C—User Facility Reporting Requirements

##### § 803.30 If I am a user facility, what re- porting requirements apply to me?

You must submit reports to the manufacturer or to us, or both, as spec- ified in paragraphs (a)(1) and (a)(2) of this section as follows:

  - Reports of death. You must submit a report to us as soon as practicable but no more than 10 work days after the day that you become aware of in- formation, from any source, that rea- sonably suggests that a device has or may have caused or contributed to the death of a patient of your facility. You must also submit the report to the de- vice manufacturer, if known. You must submit  the  information  required  by

§ 803.32. Reports sent to the  Agency must be submitted in accordance with the requirements of § 803.12(b).

  - Reports of serious injury. You must submit a report to the manufacturer of the device no later than 10 work days after the day that you become aware of information, from any source, that rea- sonably suggests that a device has or may have caused or contributed to a serious injury to a patient of your fa- cility. If  the  manufacturer is not known, you must submit the report to us. You must report information re- quired by § 803.32. Reports sent to the Agency must be submitted in accord-

21 CFR Ch. I (4–1–17 Edition)

ance with the requirements of § 803.12 (b).

What information does FDA con- sider ‘‘reasonably known’’ to me? You must submit all information required in this subpart C that is reasonably known to you. This information in- cludes information found in documents that you possess and any information that becomes available as a result of reasonable followup within your facil- ity. You are not required to evaluate or investigate the event by obtaining or evaluating information that you do not reasonably know.

##### § 803.32 If I am a user facility, what in- formation must I submit in my indi- vidual adverse event reports?

You must include the following infor- mation in  your  report,  if reasonably known     to     you,     as     described     in

§ 803.30(b). These types of information correspond generally to the elements of Form FDA 3500A:

Patient information (Form FDA 3500A, Block A). You must submit the following:

Patient name or other identifier;

Patient age at the time of event, or date of birth;

Patient gender; and

Patient weight.

Adverse event or product problem (Form FDA 3500A, Block B). You must submit the following:

  - Identification of adverse event or product problem;

  - Outcomes attributed to the ad- verse event (e.g., death or serious in- jury). An outcome is considered a seri- ous injury if it is:

    - A life-threatening injury or ill- ness;

    - A disability resulting in perma- nent impairment of a body function or permanent damage to a body structure; or

    - An injury or illness that requires intervention to prevent permanent im- pairment of a body structure or func- tion;

  - Date of event;

  - Date of this report;

#### Food and Drug Administration, HHS	§ 803.33

  - Description of event or problem, including a discussion of how the de- vice was involved, nature of the prob- lem, patient followup or required treat- ment, and any environmental condi- tions that may have influenced the event;

  - Description of relevant tests, in- cluding dates and laboratory data; and

  - Description of other relevant his- tory, including preexisting  medical conditions.

Device information (Form FDA 3500A, Block D). You must submit the following:

Brand name;

Product Code, if known, and Com- mon Device Name;

Manufacturer name, city, and state;

Model number, catalog number, serial number, lot number, or other identifying number; expiration date; and unique device identifier (UDI) that appears on the device label or on the device package;

Operator of the device (health pro- fessional, lay user/patient, other);

Date of device implantation (month, day, year), if applicable;

Date of device explantation (month, day, year), if applicable;

Whether the device is a single-use device that was reprocessed and reused on a patient (Yes, No)?

If the device is a single-use device that was reprocessed and reused on a patient (yes to paragraph (c)(8) of this section), the name and address of the reprocessor;

Whether the device was available for evaluation and whether the device was returned to the manufacturer; if so, the date it was returned to the manufacturer; and

Concomitant medical products and therapy dates. (Do not report prod- ucts that were used to treat the event.)

Initial reporter information (Form FDA 3500A, Block E). You must submit the following:

Name, address, and telephone number of the reporter who initially provided information to you, or to the manufacturer or distributor;

Whether the initial reporter is a health  professional;

Occupation; and

Whether the initial reporter also sent a copy of the report to  us,  if known.

User facility information (Form FDA 3500A, Block F). You must submit the following:

  - An indication  that  this  is  a user facility report (by marking the user fa- cility box on the form);

  - Your user facility number;

  - Your address;

  - Your contact person;

  - Your contact person’s telephone number;

  - Date that you became aware of the event (month, day, year);

  - Type of report (initial or fol- lowup); if it is a followup, you must in- clude the report number of the initial report;

  - Date of your report (month, day, year);

  - Approximate age of device;

  - Event problem codes—patient code and device code (refer to the ‘‘MedWatch Medical Device Reporting Code  Instructions’’);

  - Whether a report was sent to us and the date it was sent (month, day, year);

  - Location where the event oc- curred;

  - Whether the report was sent to the manufacturer and the date it was sent (month, day, year); and

  - Manufacturer name and address, if available.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

##### § 803.33 If I am a user facility, what must I include when I submit an an- nual report?

You must submit to us an annual report on Form FDA 3419. You must submit an annual report by January 1, of each year. You may obtain this form from the following sources:

On	the	Internet or

Division of International and Con- sumer  Education,  Center   for   Devices and  Radiological   Health,   Food   and Drug Administration, 10903 New Hamp- shire Ave., Bldg. 66, Rm. 4621, Silver Spring,    MD    20993–0002,    by    email:

#### § 803.40

DICE@fda.hhs.gov, FAX: 301–847–8149, or telephone: 800–638–2041.

You must clearly identify your annual report as such. You must sub- mit your annual report to FDA, CDRH, Medical Device Reporting, P.O. Box 3002, Rockville, MD 20847–3002. Your an- nual report must include:

  - Your CMS provider number used for medical device reports, or the num- ber assigned by us for reporting pur- poses in accordance with § 803.3;

  - Reporting year;

  - Your name and complete address;

  - Total number of reports  attached or summarized;

  - Date of the annual report and re- port numbers identifying the range of medical device reports that you sub- mitted during the report period (e.g., 1234567890–2011–0001 through 1000);

  - Name, position title, and com- plete address of the individual des- ignated as your contact person respon- sible for reporting to us and whether that person is a new contact for you; and

  - Information for each reportable event that occurred during the annual reporting period including:

Report number;

Name and address of the device manufacturer;

Device brand name and common name;

Product model, catalog, serial, and lot number and unique device iden- tifier (UDI) that appears on the device label or on the device package;

A brief description of the event reported to the manufacturer and/or us; and

Where the report was submitted, i.e., to the manufacturer, importer, or us.

In lieu of submitting the informa- tion in paragraph (b)(7) of this section, you may submit a copy of each medical device report that you submitted to the manufacturers and/or to us during the reporting period.

If you did not submit any medical device reports to manufacturers or us during the time period, you do not need to submit an annual report.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart D—Importer Reporting Requirements

##### § 803.40 If I am an importer, what re- porting requirements apply to me?

  - Reports of deaths or serious injuries. You must submit a report to us, and a copy of this report to  the  manufac- turer, as soon as practicable, but no later than 30 calendar days after the day that you receive or otherwise be- come aware of information from any source, including user facilities, indi- viduals, or medical or scientific lit- erature, whether published or unpub- lished, that reasonably  suggests  that one of your marketed devices may have caused or contributed to a death or se- rious injury. You must submit the in- formation required by § 803.42. Reports sent to the Agency  must  be  submitted in accordance with the requirements of

§ 803.12(a).

  - Reports of malfunctions. You must submit a report to the manufacturer as soon as practicable but no later than 30 calendar days after the day that you receive or otherwise become aware of information from any source, including user facilities, individuals, or through your own research, testing, evaluation, servicing, or maintenance of  one  of your devices, that reasonably suggests that one of your devices has malfunc- tioned and that this device or a similar device that you market would be likely to cause or contribute to a death or se- rious injury if the malfunction were to recur. You must submit the informa- tion required by § 803.42. Reports to manufacturers may be made in accord- ance with § 803.11(b).

##### § 803.42 If I am an importer, what in- formation must I submit in my indi- vidual adverse event reports?

You must include the following infor- mation in your report, if the informa- tion is known or should be known to you, as described in § 803.40. These types of information correspond gen- erally to the format of Form FDA 3500A:

Patient information (Form FDA 3500A, Block A). You must submit the following:

Patient name or other identifier;

Patient age at the time of event, or date of birth;

#### Food and Drug Administration, HHS	§ 803.42

Patient gender; and

Patient weight.

Adverse event or product problem (Form FDA 3500A, Block B). You must submit the following:

  - Identification of adverse event or product problem;

  - Outcomes attributed to the ad- verse event (e.g., death or serious in- jury). An outcome is considered a seri- ous injury if it is:

    - A life-threatening injury or ill- ness;

    - A disability resulting in perma- nent impairment of a body function or permanent damage to a body structure; or

    - An injury or illness that requires intervention to prevent permanent im- pairment of a body structure or func- tion;

  - Date of event;

  - Date of this report;

  - Description of the event or prob- lem, including a discussion of how the device was involved, nature of the problem, patient followup or required treatment, and any environmental con- ditions that may have influenced the event;

  - Description of relevant tests, in- cluding dates and laboratory data; and

  - Description of other relevant pa- tient history, including preexisting medical conditions.

Device information (Form FDA 3500A, Block D). You must submit the following:

Brand name;

Product Code, if known, and Com- mon Device Name;

Manufacturer name, city, and state;

Model number, catalog number, serial number, lot number, or other identifying number; expiration date; and unique device identifier (UDI) that appears on the device label or on the device package;

Operator of the device (health pro- fessional, lay user/patient, other);

Date of device implantation (month, day, year), if applicable;

Date of device explanation (month, day, year), if applicable;

Whether the device is a single-use device that was reprocessed and reused on a patient (Yes, No)?

If the device is a single-use device that was reprocessed and reused on a patient (yes to paragraph (c)(8) of this section), the name and address of the reprocessor;

Whether the device was available for evaluation, and whether the device was returned to the manufacturer, and if so, the date it was returned to the manufacturer; and

Concomitant medical products and therapy dates. (Do not report prod- ucts that were used to treat the event.)

Initial reporter information (Form FDA 3500A, Block E). You must submit the following:

Name, address, and telephone number of the reporter who initially provided information to the manufac- turer, user facility, or distributor;

Whether the initial reporter is a health  professional;

Occupation; and

Whether the initial reporter also sent a copy of the report to  us,  if known.

Importer information (Form FDA 3500A, Block F). You must submit the following:

  - An indication that this is an im- porter report (by marking the importer box on the form);

  - Your importer report number;

  - Your address;

  - Your contact person;

  - Your contact person’s telephone number;

  - Date that you became aware of the event (month, day, year);

  - Type of report  (initial or  fol- lowup). If it is a followup report, you must include the report number of your initial report;

  - Date of your report (month, day, year);

  - Approximate age of device;

  - Event problem codes—patient code and device code (refer to FDA MedWatch Medical Device Reporting Code Instructions);

  - Whether a report was sent to us and the date it was sent (month, day, year);

  - Location where event occurred;

  - Whether a report was sent to the manufacturer and the date it was sent (month, day, year); and

#### § 803.50

  - Manufacturer name and address, if available.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

## Subpart E—Manufacturer Reporting Requirements

##### § 803.50 If I am a manufacturer, what reporting requirements apply to me?

If you are a manufacturer,  you must report to us the information re- quired by § 803.52 in accordance with the requirements of § 803.12(a), no later than 30  calendar  days  after  the  day that you receive or otherwise become aware of information, from any source, that reasonably suggests that a device that  you  market:

May have caused or contributed to a death or serious injury or

Has malfunctioned and this device or a similar device that you market would be likely to cause or contribute to a death or serious injury, if the mal- function were to recur.

What information does FDA con- sider ‘‘reasonably known’’ to me?

  - You must submit all information required in this subpart E that is rea- sonably known to you. We consider the following information to be reasonably known to you:

    - Any information that you can ob- tain by contacting a user facility, im- porter, or other initial reporter;

    - Any information in your posses- sion; or

    - Any information that you can obtain by analysis, testing, or other evaluation of the device.

  - You are responsible for obtaining and submitting to us information that is incomplete or missing from reports submitted by user facilities, importers, and other initial reporters.

  - You are also responsible for con- ducting an investigation of each event and evaluating the cause of the event. If you cannot submit complete infor- mation on a report, you must provide a statement explaining why this infor- mation was incomplete and the steps you took to obtain the information. If you later obtain any required informa- tion that was not available at the time you filed your initial report, you must submit  this  information  in  a  supple-

21 CFR Ch. I (4–1–17 Edition)

mental  report  under  § 803.56  in  accord- ance	with	the	requirements	of

§ 803.12(a).

##### § 803.52 If I am a manufacturer, what information must I submit in my in- dividual adverse event reports?

You must include the following infor- mation in your reports, if known or reasonably known to you,  as  described in § 803.50(b). These types of informa- tion correspond generally  to  the  for- mat of Form FDA 3500A:

Patient information (Form FDA 3500A, Block A). You must submit the following:

Patient name or other identifier;

Patient age at the time of event, or date of birth;

Patient gender; and

Patient weight.

Adverse event or product problem (Form FDA 3500A, Block B). You must submit the following:

  - Identification of adverse event or product problem;

  - Outcomes attributed to the ad- verse event (e.g., death or serious in- jury). An outcome is considered a seri- ous injury if it is:

    - A life-threatening injury or ill- ness;

    - A disability resulting in perma- nent impairment of a body function or permanent damage to a body structure; or

    - An injury or illness that requires intervention to prevent permanent im- pairment of a body structure or func- tion;

  - Date of event;

  - Date of this report;

  - Description of the event or prob- lem, including a discussion of how the device was involved, nature of the problem, patient followup or required treatment, and any environmental con- ditions that may have influenced the event;

  - Description of relevant tests, in- cluding dates and laboratory data; and

  - Other relevant patient history in- cluding preexisting medical conditions.

Device information (Form FDA 3500A, Block D). You must submit the following:

Brand name;

Product Code, if known, and Com- mon Device Name;

#### Food and Drug Administration, HHS	§ 803.52

Manufacturer name, city, and state;

Model number, catalog number, serial number, lot number, or other identifying number; expiration date; and unique device identifier (UDI) that appears on the device label or on the device package;

Operator of the device (health pro- fessional, lay user/patient, other);

Date of device implantation (month, day, year), if applicable;

Date of device explantation (month, day, year), if applicable;

Whether the device is a single-use device that was reprocessed and reused on a patient (Yes, No)?

If the device is a single-use device that was reprocessed and reused on a patient (yes to paragraph (c)(8) of this section), the name and address of the reprocessor;

Whether the device was available for evaluation, and whether the device was returned to the manufacturer, and if so, the date it was returned to the manufacturer; and

Concomitant medical products and therapy dates. (Do not report prod- ucts that were used to treat the event.)

Initial reporter information (Form FDA 3500A, Block E). You must submit the following:

Name, address, and telephone number of the reporter who initially provided information to you, or to the user facility or importer;

Whether the initial reporter is a health  professional;

Occupation; and

Whether the initial reporter also sent a copy of the report to  us,  if known.

Reporting information for all manufacturers (Form FDA 3500A, Block G). You must submit the fol- lowing:

  - Your reporting office’s contact name and address and device manufac- turing  site;

  - Your contact person’s telephone number;

  - Your report sources;

  - Date received by you (month, day, year);

  - PMA/510k Number and whether or not the product is a combination prod- uct;

  - Type of report being submitted (e.g., 5-day, initial, followup); and

  - Your report number.

Device manufacturer information (Form FDA 3500A, Block H). You must submit the following:

  - Type of reportable event (death, serious injury, malfunction, etc.);

  - Type of followup report, if appli- cable (e.g., correction, response to FDA request, etc);

  - If the device was returned to you and evaluated by you, you must in- clude a summary of the evaluation. If you did not perform an evaluation, you must explain why you did not perform an evaluation;

  - Device manufacture date (month, day, year);

  - Whether the device was labeled for single use;

  - Evaluation codes (including event codes, method of evaluation, result, and conclusion codes) (refer to FDA MedWatch Medical Device Reporting Code Instructions);

  - Whether remedial action was taken and the type of action;

  - Whether the use of the device was initial, reuse, or unknown;

  - Whether remedial action was re- ported as a removal or correction under section 519(f) of the Federal Food, Drug, and Cosmetic Act, and if it was, provide the correction/removal re- port number; and

  - Your additional narrative; and/or

  - Corrected data, including:

    - Any information missing on the user facility report or importer report, including any event codes that were not reported, or information corrected on these forms after your verification;

    - For  each  event  code  provided  by the user facility under § 803.32(e)(10) or the importer under § 803.42(e)(10), you must  include  a  statement  of  whether the  type  of  the  event  represented   by the code is addressed in the device la- beling;   and

    - If your report omits any required information, you must  explain  why this information was not provided and the steps taken to obtain this informa- tion.

[79  FR  8846,  Feb.  14,  2014,  as  amended  at  80

FR 10587, Feb. 27, 2015]

#### § 803.53

##### § 803.53 If I am a manufacturer, in which circumstances must I submit a 5-day report?

You must submit a 5-day report to us with    the    information    required    by

§ 803.52 in accordance with the require- ments of § 803.12(a) no later than 5 work days after the day that you  become aware  that:

      - An MDR reportable event neces- sitates remedial action to prevent an unreasonable risk of substantial harm to the public health. You may become aware of the need for remedial action from any information, including any trend analysis or

      - We have made a written request for the submission of a 5-day report. If you receive such a written request from us, you must submit, without fur- ther requests, a 5-day report for all subsequent events of the same nature that involve substantially similar de- vices for the time period specified in the written request. We may extend the time period stated in the original written request if we determine it is in the interest of the public health.

##### § 803.56 If I am a manufacturer, in what circumstances must I submit a supplemental or followup report and what are the requirements for such reports?

If you are a manufacturer, when you obtain information required under this part that you did not provide because it was not known or was not available when you submitted the initial report, you must submit the supplemental in- formation to us within 30 calendar days of the day that you receive this infor- mation. You must submit the supple- mental or followup report in accord- ance     with     the     requirements     of

§ 803.12(a). On a supplemental or fol- lowup  report,  you  must:

Indicate that the report being submitted is a supplemental or fol- lowup report;

Submit the appropriate identi- fication numbers of the report that you are updating with the supplemental in- formation (e.g., your original manufac- turer report number and the user facil- ity or importer report number of any report on which your report was based), if applicable; and

#### 21 CFR Ch. I (4–1–17 Edition)

Include only the new, changed, or corrected information.

##### § 803.58   Foreign manufacturers.

Every foreign manufacturer whose devices are distributed in the United States shall designate a U.S. agent to be responsible for reporting in accordance with § 807.40 of this chapter. The U.S. designated agent accepts re- sponsibility for the duties that such designation entails. Upon the effective date of this regulation, foreign manu- facturers shall inform FDA,  by  letter, of the name and address of the U.S. agent designated under  this  section and § 807.40 of this chapter, and shall update this information as necessary. Such updated information shall be sub- mitted to FDA, within 5 days of  a change in the designated agent infor- mation.

U.S.-designated agents of foreign manufacturers are required to:

  - Report to FDA in accordance with

§§ 803.50, 803.52, 803.53, and 803.56;

  - Conduct, or obtain from the for- eign manufacturer the necessary infor- mation   regarding,   the   investigation and evaluation of the event to comport with the requirements of § 803.50;

  - Forward MDR complaints to the foreign manufacturer and maintain documentation of this requirement;

  - Maintain complaint files in ac- cordance with § 803.18; and

  - Register, list, and submit pre- market notifications  in accordance with part 807 of this chapter.

EFFECTIVE  DATE  NOTE: At 79 FR 8846, Feb.

14, 2014, part 803 was revised. At 79 FR 8855,

Feb.  14,  2014,  § 803.58  was  stayed  indefinitely.

## PART 806—MEDICAL DEVICES; RE- PORTS OF CORRECTIONS AND REMOVALS

##### Subpart A—General Provisions

Sec.

  - Scope.

  - Definitions.

##### Subpart B—Reports and Records

806.10    Reports of corrections and removals.

806.20    Records  of  corrections  and  removals not required to be reported.

806.30    FDA access to records.

  - Public availability of reports.

Food and Drug Administration, HHS	§ 806.2

AUTHORITY: 21 U.S.C. 352, 360, 360i, 360j, 371,

374.

SOURCE:  62  FR  27191,  May  19,  1997,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 806.1   Scope.

    - This part implements the provi- sions of section 519(g) of the Federal Food, Drug, and Cosmetic Act (the act) requiring device manufacturers and importers to report  promptly to the Food and Drug Administration (FDA) certain actions concerning device cor- rections and removals, and to maintain records of all corrections and removals regardless of whether such corrections and removals are required to be re- ported to FDA.

    - The following actions are exempt from the reporting requirements of this part:

      - Actions taken by device manufac- turers or importers to improve the per- formance or  quality  of  a device but that do not reduce a risk to health posed by the device or remedy a viola- tion of the act caused by the device.

      - Market withdrawals as defined in

§ 806.2(h).

      - Routine  servicing  as  defined  in

§ 806.2(k).

      - Stock   recoveries   as   defined   in

§ 806.2(l).

[62 FR 27191, May 19, 1997, as amended at 63

FR 42232, Aug. 7, 1998]

##### § 806.2   Definitions.

As used in this part:

        - Act means the Federal Food, Drug, and Cosmetic Act.

        - Agency or FDA means the Food and Drug Administration.

        - Consignee means any person or firm that has received, purchased, or used a device subject to correction or removal.

        - Correction means the repair, modi- fication, adjustment, relabeling, de- struction, or inspection (including pa- tient monitoring) of a device  without its physical removal from its point of use to some other location.

        - Correction or removal report number means the number that uniquely iden- tifies each report submitted.

        - Human cell, tissue, or cellular or tis- sue-based product (HCT/P) regulated as a

device  means  an  HCT/P  as  defined  in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

        - Importer means, for the purposes of this part, any person who imports a device into the United States.

        - Manufacturer means any person who manufactures, prepares, propa- gates, compounds, assembles, or proc- esses a device by chemical, physical, biological, or other procedures. The term includes any person who:

          - Repackages or otherwise changes the container, wrapper, or labeling of a device in furtherance of the distribu- tion of the device from the original place of manufacture to the person who makes final delivery or sale to the ulti- mate user or consumer;

          - Initiates specifications for devices that are manufactured by a second party for subsequent distribution by the person initiating the specifica- tions; or

          - Manufactures components or ac- cessories which are devices that are ready to be used and are intended to be commercially distributed and are in- tended to be used as is, or are processed by a  licensed  practitioner or  other qualified person to meet the needs of a particular patient.

            - Market withdrawal means a correc- tion or removal of a distributed device that involves a minor violation of the act that would not be subject to legal action by FDA or that involves no vio- lation of the act, e.g., normal stock ro- tation practices.

Removal means the physical re- moval of a device from its point of use to some other location for repair, modification, adjustment, relabeling, destruction, or inspection.

Risk to health means

  - A reasonable probability that use of, or exposure to, the product will cause serious adverse health con- sequences or death; or

  - That use of, or exposure to, the product may cause temporary or medi- cally reversible adverse health con- sequences, or an outcome where the probability of  serious  adverse health consequences is remote.

Routine servicing means any regu- larly scheduled maintenance of a de- vice,   including   the   replacement   of

§ 806.10

parts at the end of their normal life ex- pectancy, e.g., calibration, replace- ment of batteries, and responses to normal wear and tear. Repairs of an unexpected nature, replacement of parts earlier than their normal life ex- pectancy, or identical repairs or re- placements of multiple units of a de- vice are not routine servicing.

Stock recovery means the correc- tion or removal of a device that  has not been marketed or that has not left the direct control of the manufacturer, i.e., the device is located on the prem- ises owned, or under the control of, the manufacturer, and no portion of the lot, model, code, or other relevant unit involved in the corrective or removal action has been released for sale or use.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A UDI is composed of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

[62 FR 27191, May 19, 1997, as amended at 63

FR 42232,  Aug.  7,  1998;  78  FR 55821,  Sept.  24,

2013]

## Subpart B—Reports and Records

##### § 806.10   Reports of corrections and re- movals.

Each device manufacturer or im- porter shall submit a written report to FDA of any correction or removal of a device initiated by such manufacturer or importer if the correction or re- moval was initiated:

21 CFR Ch. I (4–1–17 Edition)

To reduce a risk to  health  posed by the device; or

To remedy a violation of the act caused by the device which  may present a risk to health unless the in- formation has already been provided as set forth in paragraph (f) of this sec- tion or the corrective or removal ac- tion is exempt from the reporting re- quirements under § 806.1(b).

The manufacturer or importer shall submit any report required by paragraph (a) of this section within 10- working days of initiating such correc- tion or removal.

The manufacturer or importer shall include the following information in the report:

The seven digit registration num- ber of the entity responsible for sub- mission of the report of corrective or removal action (if applicable), the month, day, and year that the report is made, and a sequence number (i.e.,  001 for the first report, 002 for the second report, 003 etc.), and the report type designation ‘‘C’’ or ‘‘R’’. For example, the complete number for the first cor- rection report submitted on June 1, 1997, will appear as follows for a firm with the registration number 1234567: 1234567–6/1/97–001–C. The second correc- tion report number submitted by the same firm on July 1, 1997, would be 1234567–7/1/97–002–C etc. For removals, the number will appear as follows: 1234567–6/1/97–001–R and 1234567–7/1/97– 002–R, etc. Firms that do not have a seven digit  registration  number  may use seven zeros followed by the month, date, year, and sequence number (i.e. 0000000–6/1/97–001–C for corrections and 0000000–7/1/97–001–R for removals). Re- ports received without a seven digit registration number will be assigned a seven digit central file number by the district office reviewing the reports.

The name, address, and telephone number of the manufacturer or im- porter, and  the  name,  title,  address, and telephone number of the manufac- turer or importer representative re- sponsible for conducting the device correction or removal.

The brand name and the common name, classification name, or usual name of the device and the  intended use of the device.

#### Food and Drug Administration, HHS	§ 806.20

Marketing status of the device, i.e., any applicable premarket notifica- tion number, premarket approval num- ber, or indication that the device is a preamendments device, and the device listing number. A manufacturer or im- porter that does not have an FDA es- tablishment registration number shall indicate in the report whether it has ever registered with FDA.

The unique device identifier (UDI) that appears on the device label or on the device package, or the device iden- tifier, universal product code (UPC), model, catalog, or code number of the device and the manufacturing lot or se- rial number of the device or other iden- tification number.

The manufacturer’s name, ad- dress, telephone number, and contact person if different from that of the per- son submitting the report.

A description of the event(s) giv- ing rise to the information reported and the corrective or removal actions that have been, and are expected to be taken.

Any illness or injuries that have occurred with use of the device. If ap- plicable, include the medical device re- port numbers.

The total number of devices man- ufactured or distributed subject to the correction or removal and the number in the same batch, lot, or equivalent unit of production subject to the cor- rection or removal.

The date of manufacture or dis- tribution and the device’s expiration date or expected life.

The names, addresses, and tele- phone numbers of all domestic and for- eign consignees of the device and the dates and number of devices distrib- uted to each such consignee.

A copy of all communications re- garding the correction or removal and the names and addresses of all recipi- ents of the communications not pro- vided in accordance with paragraph (c)(11) of this section.

If any required information is not immediately available, a state- ment as to why it is not available and when it will be submitted.

If, after submitting a report under this part, a manufacturer or importer determines that the same correction or removal  should  be  extended  to  addi-

tional lots or batches of the same de- vice, the manufacturer or importer shall within 10-working days of initi- ating the extension of the correction or removal, amend the report by submit- ting an amendment citing the original report number assigned according to paragraph (c)(1) of this section, all of the information required by paragraph (c)(2), and any information required by paragraphs (c)(3) through (c)(12) of this section that is different from the infor- mation submitted in the original re- port. The manufacturer or importer shall also provide a statement in ac- cordance with paragraph (c)(13) of this section for any required information that is not readily available.

A report submitted by a manufac- turer or importer under this section (and any release by FDA of that report or information) does not necessarily re- flect a conclusion by the manufacturer, importer, or FDA that the report or in- formation constitutes an admission that the device caused or contributed to a death or serious injury. A manu- facturer or importer need not admit, and may deny, that the report or infor- mation submitted under this section constitutes an admission that the de- vice caused or contributed to a death or serious injury.

No report of correction or removal is required under this part, if a report of the correction or removal is required and has been submitted under parts 803 or 1004 of this chapter.

[62 FR 27191, May 19, 1997, as amended at 63

FR  42232,  Aug.  7,  1998;  69  FR  11311,  Mar.  10,

2004; 78 FR 55821, Sept. 24, 2013]

##### § 806.20   Records of corrections and re- movals not required to be reported.

Each device manufacturer or im- porter who initiates a correction or re- moval of a device  that  is  not  required to be reported to FDA under § 806.10 shall keep a record  of  such  correction or removal.

Records of corrections and remov- als not required to be reported to FDA under § 806.10 shall contain the fol- lowing   information:

  - The brand name, common or usual name, classification, name and product code if known, and the intended use of the device.

#### § 806.30

  - The unique device identifier (UDI) of the device, or the device identifier, universal product code (UPC), model, catalog, or code number of the device and the manufacturing lot or serial number of the device or other identi- fication number.

  - A description of the event(s) giv- ing rise to the information reported and the corrective or removal action that has been, and is expected to be taken.

  - Justification for not reporting the correction or removal action to FDA, which shall contain  conclusions  and any followups, and be reviewed and evaluated by a designated person.

  - A copy of all communications re- garding the correction or removal.

The manufacturer or importer shall retain records required under this section for a period of 2 years beyond the expected life of  the device,  even if the manufacturer or importer has ceased to manufacture or import the device. Records required to be main- tained under paragraph (b) of this sec- tion must be transferred to the new manufacturer or importer of the device and maintained for the required period of  time.

[62 FR 27191, May 19, 1997, as amended at 63

FR 42233,  Aug.  7,  1998;  78  FR 55821,  Sept.  24,

2013]

##### § 806.30   FDA access to records.

Each device  manufacturer or  im- porter required under this part to maintain records and every person who is in charge or custody of such records shall, upon request of an officer or em- ployee designated by FDA and under section 704(e) of the act, permit such officer or employee at all reasonable times to have access to, and to  copy and verify, such records and reports.

[63 FR 42233, Aug. 7, 1998]

##### § 806.40   Public availability of reports.

Any report submitted under this part is available for  public  disclosure in accordance with part 20 of this chap- ter.

Before public disclosure of a re- port,  FDA  will  delete  from  the  report:

  - Any information that constitutes trade  secret  or  confidential  commer-

#### 21 CFR Ch. I (4–1–17 Edition)

cial   or   financial   information   under

§ 20.61 of this chapter; and

  - Any personnel, medical, or similar information, including the serial num- bers of implanted devices, which would constitute a clearly unwarranted inva- sion of personal privacy under § 20.63 of this chapter or 5 U.S.C. 552(b)(6); pro- vided, that except for the information under § 20.61 of this chapter or 5 U.S.C. 552(b)(4), FDA will disclose to a patient who requests a report all the informa- tion in the report concerning that pa- tient.

## PART 807—ESTABLISHMENT REG- ISTRATION AND DEVICE LISTING FOR MANUFACTURERS AND INI- TIAL IMPORTERS OF DEVICES

##### Subpart A—General Provisions

Sec.

807.3    Definitions.

##### Subpart B—Procedures for Device Establishments

  - Who must register and submit a de- vice list?

  - How to register establishments and list devices.

  - Times for establishment registration and device listing.

  - Information required for device es- tablishment registration and device list- ing.

  - Additional listing information.

807.28     Updating device listing information.

  - Summary of requirements for owners or operators granted a waiver from sub- mitting required information electroni- cally.

  - Notification of registrant.

807.37 Public availability of establishment registration and device listing informa- tion.

  - Misbranding by reference to estab- lishment registration or to registration number.

##### Subpart C—Procedures for Foreign Device Establishments

  - Establishment registration and de- vice listing for foreign establishments importing or offering for import devices into the United States.

  - Identification of importers and per- sons who import or offer for import.

Food and Drug Administration, HHS	§ 807.3

##### Subpart D—Exemptions

807.65    Exemptions     for     device     establish- ments.

##### Subpart E—Premarket Notification Procedures

807.81    When  a  premarket  notification  sub- mission  is  required.

807.85     Exemption  from  premarket  notifica- tion.

807.87    Information  required  in  a  premarket notification     submission.

807.90    Format   of   a   premarket   notification submission.

  - Content and format of a 510(k) sum- mary.

  - Content and format of a 510(k) state- ment.

  - Format of class III certification.

  - Confidentiality of information.

807.97   Misbranding   by   reference   to   pre- market   notification.

  - FDA action on a premarket notifica- tion.

AUTHORITY: 21  U.S.C.  321,  331,  351,  352,  360,

360c, 360e, 360i, 360j, 371, 374, 381, 393; 42 U.S.C.

264, 271.

SOURCE:  42  FR  42526,  Aug.  23,  1977,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 807.3   Definitions.

    - Act means the Federal Food, Drug, and Cosmetic Act.

    - Commercial distribution means any distribution of a device intended for human use which is held or offered for sale but does not include the following:

      - Internal or interplant transfer of a device between establishments within the same parent, subsidiary, and/or af- filiate company;

      - Any distribution of a device in- tended for human use which has in ef- fect an approved exemption for inves- tigational use under section 520(g) of the act and part 812 of this chapter;

      - Any distribution of a device, be- fore the effective date of part 812  of this  chapter,  that  was   not   introduced or delivered for introduction into inter- state commerce for commercial dis- tribution before May  28,  1976,  and  that is classified into class III under section 513(f) of the act: Provided, That the  de- vice is intended solely for investiga- tional use, and under section 501(f)(2)(A) of the act the device is not required    to    have    an    approved    pre-

market   approval   application   as   pro- vided in section 515 of the act; or

      - For foreign establishments, the distribution of any device that is nei- ther imported nor  offered for  import into the United States.

    - Establishment means a place of business under one management at one general physical location at which a device is manufactured, assembled, or otherwise processed.

    - Manufacture, preparation, propaga- tion, compounding, assembly, or proc- essing of a device means the making by chemical, physical, biological, or other procedures of any article that meets the definition of device  in  section 201(h) of the act. These terms include the  following  activities:

      - Repackaging or otherwise chang- ing the container, wrapper, or labeling of any device package in furtherance of the distribution of the device from the original place of manufacture to the person who makes final delivery or sale to the ultimate consumer;

      - Initial importation of devices manufactured in foreign establish- ments; or

      - Initiation of specifications for de- vices that are manufactured by a sec- ond party for subsequent commercial distribution by the person initiating specifications.

    - Official correspondent means the person designated by the owner or op- erator of an establishment as respon- sible for the following:

      - The annual registration of the es- tablishment;

      - Contact with the Food and Drug Administration for device listing;

      - Maintenance and submission of a current list of officers and directors to the Food  and  Drug  Administration upon the request of the Commissioner; and

      - The receipt of pertinent cor- respondence from the Food and Drug Administration directed to and involv- ing the owner or operator and/or any of the firm’s establishments.

    - Owner or operator means the cor- poration, subsidiary, affiliated com- pany, partnership, or proprietor di- rectly responsible for the activities of the registering establishment.

    - Initial importer means any im- porter who furthers the marketing of a

§ 807.3

device from a foreign manufacturer to the person who makes the final deliv- ery or sale of the device to the ulti- mate consumer or user, but does not repackage, or otherwise change the container, wrapper, or labeling of the device or device package.

    - Any term defined in section 201 of the act shall have that meaning.

    - Restricted device means a device for which a requirement restricting sale, distribution, or use has been estab- lished by a  regulation issued under  sec- tion 520(e) of the act, by order as a con- dition of premarket approval under section 515(d)(1)(B)(ii) of the act, or by a performance standard issued in ac- cordance with sections 514(a)(2)(B)(v) and 514(b) of the act.

    - Classification name means the term used by the Food and Drug Administra- tion and its classification panels to de- scribe a device or class of devices for purposes of classifying devices under section 513 of the act.

    - Product code means the code used by FDA to identify the  generic  cat- egory of a device.

    - Representative sampling of advertise- ments means typical advertising mate- rial that gives the promotional claims made for the device.

    - Representative sampling of any other labeling means typical labeling material (excluding labels and package inserts) that gives the promotional claims made for the device.

    - Material change includes any change or modification in the labeling or advertisements that affects the identity or safety and effectiveness of the device. These changes may include, but are not limited to, changes in the common or usual or proprietary name, declared ingredients or components, in- tended use, contraindications, warn- ings, or instructions for use. Changes that are not material may include graphic layouts, grammar, or correc- tion of typographical errors which do not change the content of the labeling, changes in lot number, and, for devices where the biological activity or known composition differs with each lot pro- duced, the labeling containing the ac- tual values for each lot.

    - 510(k) summary (summary of any information respecting safety and ef- fectiveness)   means   a   summary,   sub-

#### 21 CFR Ch. I (4–1–17 Edition)

mitted under section 513(i) of the act, of the safety and effectiveness informa- tion contained in a premarket notifica- tion submission upon which a deter- mination  of  substantial  equivalence can be based. Safety and effectiveness information refers to safety and effec- tiveness data and information sup- porting a finding of substantial equiva- lence, including all adverse safety and effectiveness  information.

    - 510(k) statement means a state- ment, made under section 513(i) of the act, asserting that all information in a premarket notification submission re- garding safety and effectiveness will be made available within 30 days of re- quest by any person if the device de- scribed in the premarket notification submission is determined to be sub- stantially equivalent. The information to be made available will be a duplicate of the premarket notification submis- sion, including any adverse safety and effectiveness information, but  exclud- ing all patient identifiers, and trade se- cret or confidential commercial infor- mation, as defined in § 20.61 of this chapter.

    - Class III certification means a cer- tification that the submitter of the 510(k) has conducted a reasonable search of all known information about the class III device and other similar, legally  marketed  devices.

    - Class III summary means a sum- mary of the types of safety and effec- tiveness problems associated with the type of device being compared and a ci- tation to the information upon which the summary is based. The summary must be comprehensive and describe the problems to which the type of de- vice is susceptible and the causes of such problems.

    - United States agent means a person residing or maintaining a place of busi- ness in the United States whom a for- eign establishment designates as its agent. This definition excludes mail- boxes, answering machines or services, or other places where an individual acting as the foreign establishment’s agent is not physically present.

    - Wholesale distributor means any person (other than the manufacturer or the initial importer) who distributes a device from the original place of manu- facture to the person who makes the

#### Food and Drug Administration, HHS	§ 807.20

final delivery or sale of the device to the ultimate consumer or user.

    - Fiscal year means the FDA fiscal year,  which  runs  from  October  1 through  September  30.

    - FURLS means the Food and Drug Administration’s Unified

Registration and Listing System,

    - FDA premarket submission number means the number assigned by FDA to a premarket device submission, such as a Premarket Approval Application (PMA); Humanitarian Device Exemp- tion (HDE); New Drug Application (NDA); Biologics License Application (BLA); de novo classification petition; or Premarket Notification (510(k)).

    - Importer means, for purposes of this part, a company or individual  in the United States that is an owner, consignee, or recipient, even if not the initial owner, consignee, or  recipient, of the foreign establishment’s device that is imported into the United States. An importer does not include the consumer or patient who ulti- mately purchases, receives, or uses the device, unless the foreign establish- ment ships the device directly to the consumer or patient.

    - Person who imports or offers for im- port means, for purposes of this part, an agent, broker, or other entity, other than a carrier, that the foreign estab- lishment uses  to  facilitate  the  import of its device into the United States.

[42 FR 42526, Aug. 23, 1977, as amended at 43

FR 37997,  Aug. 25,  1978;  57  FR 18066,  Apr. 28,

1992;  58  FR 46522,  Sept. 1,  1993;  59  FR 64295,

Dec. 14, 1994; 60 FR 63606, Dec. 11, 1995; 63 FR

51826, Sept. 29, 1998; 66 FR 59159, Nov. 27, 2001;

77 FR 45940, Aug. 2, 2012]

## Subpart B—Procedures for Device Establishments

##### § 807.20   Who must register and submit a device list?

An owner or operator of an estab- lishment not exempt under section 510(g) of the Federal Food, Drug, and Cosmetic Act or subpart D of this part who is engaged in the manufacture, preparation,	propagation, compounding, assembly, or processing of a device intended for human use shall register and submit listing infor- mation for those devices in commercial distribution,   except   that   registration

and listing information may be sub- mitted by the parent, subsidiary, or af- filiate company for all the domestic or foreign establishments under  the  con- trol of one of these organizations when operations are conducted at more than one  establishment  and  there  exists joint ownership and control among all the establishments. The term ‘‘device’’ includes all in  vitro  diagnostic  prod- ucts and in vitro diagnostic biological products not subject to licensing under section 351 of the Public Health Service Act. An owner or operator of an estab- lishment located in any  State  as  de- fined in section 201(a)(1) of the Federal Food, Drug, and Cosmetic Act  shall register its  name,  places  of  business, and all establishments and list the de- vices whether or not the output of the establishments  or  any  particular   de- vice so listed enters interstate com- merce. The registration and listing re- quirements shall pertain to any person who is engaged in the manufacture, preparation,	propagation, compounding,  assembly,  or  processing of a device intended for human use, in- cluding any person who:

Initiates or develops specifica- tions for a device that is to be manu- factured by a second party;

Sterilizes or otherwise makes a device for or on behalf of a specifica- tions developer or any other person;

Repackages or relabels a device;

Reprocesses a single use device that has previously been used on a pa- tient;

Acts as an initial importer as de- fined in § 807.3(g), except that initial importers may fulfill their listing obli- gation for any device for which  they did not initiate or develop the speci- fications for the device or repackage or relabel the device by submitting the name and address of the manufacturer. Initial importers shall also be prepared to submit, when requested by FDA, the proprietary name, if any, and the com- mon or usual name of each device for which they are the initial importer;

Manufactures components or ac- cessories that are ready to be used for any  intended  health-related  purpose

§ 807.21

and are packaged or labeled for com- mercial distribution for such health-re- lated purpose, e.g. blood filters, hemo- dialysis tubing, or devices which of ne- cessity must be further processed by a licensed practitioner or other qualified person to meet the needs of a par- ticular patient, e.g., a manufacturer of ophthalmic lens blanks.

Registration or listing does not constitute an admission or  agreement or determination that a product is a device within the meaning of section 201(h) of the Federal Food, Drug, and Cosmetic  Act.

Registration and listing require- ments shall not pertain to any person who acts as a wholesale distributor, as defined in § 807.3(t), and who does not manufacture, repackage, process, or relabel a device.

Owners and operators of establish- ments or persons engaged in the recov- ery, screening, testing, processing, storage, or distribution of human cells, tissues, and cellular and tissue-based products, as defined in § 1271.3(d) of this chapter, that are regulated under the Federal Food, Drug, and Cosmetic Act must register and list those  human cells, tissues, and cellular and tissue- based products with the Center for Bio- logics Evaluation and  Research  on Form FDA 3356 following the proce- dures set out in subpart B of part 1271 of this chapter, instead of the proce- dures for registration and listing con- tained in this part, except that  the ad- ditional listing information require- ments of § 807.26 remain applicable.

Owners and operators of establish- ments that manufacture devices li- censed under section 351 of the Public Health Service Act as well as licensed biological products used in the manu- facture of a licensed device must reg- ister and list following the procedures set out in part 607 of this chapter, in- stead of the procedures for registration and listing contained in this part.

[77 FR 45941, Aug. 2, 2012]

##### § 807.21   How    to    register    establish- ments and list devices.

Owners or operators of establish- ments that are subject to the registra- tion and listing requirements of this part must provide the following infor- mation  to  us  using  our  electronic  de-

#### 21 CFR Ch. I (4–1–17 Edition)

vice registration and listing system, except as provided in paragraphs (b), (c), and (d) of this section:

Initial establishment registration information as required by §§ 807.22(a) and 807.25;

Updates to registration informa- tion as required by §§ 807.22(b) and 807.25;

Initial  device  listing  information as required  by  §§ 807.22(a), 807.25,  and 807.28;

Updates to device listing informa- tion as required by §§ 807.22(b), 807.25, and 807.28, including updates to reflect the  discontinuance  or   resumption   of the commercial distribution of a pre- viously-listed device as specified at paragraphs (d) and (e) of § 807.28.

If the information under § 807.21(a) cannot be submitted electronically, a waiver may be requested. Waivers will be granted only if use of electronic means is not reasonable for the person requesting the waiver. To request a waiver, applicants must send  a  letter to the Office of Compliance, Center for Devices and Radiological Health, Food and Drug Administration, 10903 New Hampshire Ave., Bldg. 66, rm. 2621, Sil- ver Spring, MD 20993–0002, that in- cludes the following information:

  - The name and address of the de- vice establishment(s)  to  be  registered, a contact person for the owner or oper- ator of the establishment, and the tele- phone number at which that person can be reached. If the establishment has al- ready registered in the past, the letter should also include the owner or oper- ator number, registration number, and any listing numbers previously as- signed by FDA for devices manufac- tured  at  that  establishment.

  - Information about whether the company is an initial importer as de- fined in § 807.3(g) and, if so, whether it also conducts any other activities or operations relating to devices.

  - A statement that use of the Inter- net is not reasonable for the person re- questing the waiver, and an expla- nation of why such use is not reason- able. The statement must be signed by the owner or operator of the establish- ment, or by a person employed by the owner or operator who is authorized to make the declaration on behalf of the owner or operator.

#### Food and Drug Administration, HHS	§ 807.25

Those owners or operators who have obtained a waiver from filing reg- istration and listing information elec- tronically should refer to § 807.34 for in- formation on how to submit such infor- mation  by  postal  mail.

When additional device listing in- formation (e.g., copies of labeling or advertisements) is requested by FDA as described at § 807.26(e), such informa- tion may be submitted  by  postal  mail or electronically by email,  but  will  not be submitted using the FDA electronic device   registration   and   listing   system.

[77 FR 45941, Aug. 2, 2012]

##### § 807.22   Times  for  establishment  reg- istration and device listing.

Initial registration and listing. An owner or operator of an establishment who has not previously entered into an operation described in § 807.20(a) shall register within 30 days after entering into such an operation and submit de- vice  listing  information  at  that  time.

Registration and listing updates. Owners or operators shall review and update all of their establishment reg- istration and device listing informa- tion that is on file at FDA, docu- menting any changes that were not previously reported as follows:

Annual registration for each fiscal year is required for all establishments. Annual registration shall take place during the period beginning on October

1  and  ending  on  December  31  of  each fiscal   year;

Updates to the registration infor- mation as described in § 807.25(b) shall be made within 30 days of any change to such information;

Every fiscal year, during the pe- riod beginning on October 1 and ending on December 31, owners or operators shall review and update all of their de- vice listing information  that  is  on  file at FDA, reporting any changes or dele- tions to  listings  and  any  new  listings that were not previously reported. The accuracy  of  all  information  on   file must be confirmed each  year  regardless of whether any changes were made  to the owner or operator’s list of devices; and

Changes to listing information may also be made at other times, such as when a device is introduced into commercial     distribution,     when     a

change is made to a previously-listed device, or when a previously-listed de- vice is removed from commercial dis- tribution.

Failure to submit required informa- tion. Failure to submit any  of the re- quired information on time, as speci- fied in paragraphs (a) and (b) of this section, will put the establishment in a ‘‘failed to register’’ or ‘‘failed to list’’ status as applicable. The establishment will not be considered active and the establishment registration and device listing information may not appear on the FDA Web site until such  time  as the owner or operator submits and FDA processes the required informa- tion.

[77 FR 45942, Aug. 2, 2012]

##### § 807.25 Information required for de- vice establishment registration and device listing.

(a) All owners or operators that are subject to the registration and listing requirements of this part shall provide such information  to  us  by  using  the FDA electronic device registration and listing system, unless granted a waiver from electronic submission in accord- ance with § 807.21(b). Electronic submis- sions of registration and listing infor- mation  must  comply  with  part  11  of this chapter, except for the require- ments in § 11.10(b), (c), and (e), and the corresponding requirements in § 11.30 of this chapter. Those  owners  or  opera- tors granted a waiver from electronic submission  should  refer  to  paragraphs

and (g) of this section and  § 807.34 for instructions on how to submit de- vice registration and listing informa- tion.

  - Registration information required to be submitted includes: The name and mailing address of the device es- tablishment; the Web site address of the device establishment, if any; the name, address, phone number, fax num- ber, and email address of the owner or operator; the name, address, phone number, fax number, and email address of the establishment’s official cor- respondent; and all trade names used by the establishment.

  - Owners or operators who have been granted a waiver from electronic filing must submit the establishment registration  information  described  in

#### § 807.26

paragraph (b) of this section, except for the Web site and email address infor- mation, in paper form using the proce- dures set forth in § 807.34.

  - Each owner or operator is re- quired to maintain a listing of all offi- cers, directors, and partners for each establishment registered by the owner or operator and to furnish this infor- mation to FDA upon request.

  - For each establishment, an offi- cial correspondent must be designated by the owner or operator to serve as a point of contact with FDA on matters relating to the registration of device establishments and the listing of de- vice products. Each owner or operator shall also provide FDA with the name of a contact person at the owner or op- erator’s offices who will be responsible for identifying the official cor- respondent for each establishment. The owner or operator contact person will be the official correspondent in the event no one else has been properly designated. The official correspondent is responsible for:

    - Providing FDA with all required registration and listing information electronically unless a waiver from electronic submission has been  granted in accordance with § 807.21(b);

    - Receiving all correspondence from FDA concerning registration and list- ing;

    - Supplying, when requested by FDA, the names of all officers, direc- tors, and partners; and

    - Receiving communications from FDA by email, or by postal mail if the owner or operator has been granted a waiver from the requirement to file registration and listing information electronically.

  - The designation of an official cor- respondent does not in any manner af- fect the liability of the owner or oper- ator of the establishment or any other individual under section 301(p) or any other provision of the Federal Food, Drug, and Cosmetic Act.

  - Device listing  information  must be submitted to FDA electronically un- less a waiver from electronic submis- sion has been granted in accordance with § 807.21(b). Owners or operators who have been granted a waiver must submit the required device listing in- formation,    including    information    re-

#### 21 CFR Ch. I (4–1–17 Edition)

quired by this paragraph, § 807.28, and any listing information requested  by FDA under § 807.26(e), in paper form using    the    procedures    set    forth    in

§ 807.34.  The information  required for each  device  listed  includes:

The current registration number and name of each establishment under the ownership and control of the owner or operator where the device is manu- factured, repackaged, relabeled, or oth- erwise processed, or where specifica- tions are developed.

The product code for each device that is exempt from premarket notifi- cation and approval or which was in commercial distribution prior to May 28, 1976.

The proprietary or brand name(s) under which each device is marketed.

The FDA-assigned premarket sub- mission number of the approved appli- cation, cleared premarket notification, granted de novo classification  petition, or approved humanitarian device ex- emption for each device listed that is subject to sections 505, 510(k), 513(f)(2), 515, or 520(m) of the Federal Food, Drug, and Cosmetic Act, which in- cludes devices that are not  exempt from premarket notification and ap- proval.

Each activity or process that is conducted on or done to the device at each establishment, such as manufac- turing, repacking, relabeling, devel- oping specifications, remanufacturing, single-use device reprocessing, con- tract manufacturing, contract steri- lizing, or manufacturing for export only.

[77 FR 45942, Aug. 2, 2012]

##### § 807.26   Additional listing information.

Each owner or operator shall maintain a historical  file  containing the labeling and advertisements in use on the date of initial listing, and in use after October 10, 1978, but before the date of initial listing, as follows:

For each device subject to section

514 or 515 of the act that is not a re- stricted device, a copy of all  labeling for the device;

For each restricted device, a copy of all labeling and advertisements for the device;

For each device that is neither re- stricted  nor  subject  to  section  514  or

#### Food and Drug Administration, HHS	§ 807.28

515 of the act, a copy of all labels, package inserts, and a representative sampling of any other labeling.

In addition to the requirements set forth in paragraph (a) of this sec- tion, each owner or operator shall maintain in the historical file any la- beling or advertisements in which a material change has been made any- time  after  initial  listing.

Each owner or operator may dis- card labeling and advertisements from the historical file 3 years after the date of the last shipment of a discontinued device by an owner or operator.

Location of the file:

Currently existing systems for maintenance of labeling and adver- tising may be used for the purpose of maintaining the historical file as long as the information included in the sys- tems fulfills the requirements of this section, but only if the labeling and ad- vertisements are retrievable in a time- ly manner.

The contents of the historical file may be physically located in more than one place in the establishment or in more than one establishment provided there exists joint ownership and con- trol among all the establishments maintaining the historical file. If no joint ownership and control exists, the registered establishment must provide the Food and Drug Administration with a letter authorizing the establish- ment outside its control to maintain the historical file.

A copy of the certification and disclosure statements  as  required  by part 54 of this chapter shall be retained and physically located at the establish- ment  maintaining  the  historical  file.

Each owner or operator shall be prepared to submit to the Food and Drug Administration, only upon spe- cific request, the following informa- tion:

  - For a device subject to section 514 or 515 of the act that is not a restricted device, a copy of all labeling for the de- vice.

  - For a device that is a restricted device, a copy of all labeling for the de- vice, a representative sampling of ad- vertisements for the device, and  for good cause, a copy of all  advertise- ments for a particular device.  A  re- quest for all advertisements will, where

feasible, be accompanied by an expla- nation of the basis for such request.

  - For a device that is neither a re- stricted device, nor subject to section 514 of 515 of the act, the label and pack- age insert for the device and a rep- resentative sampling of any other la- beling for the device.

  - For a particular device, a state- ment of the basis upon which the reg- istrant has determined  that  the  device is not subject to section 514 or 515 of the  act.

  - For a particular device, a state- ment of the basis upon which the reg- istrant has determined the device is not a restricted device.

  - For a particular device, a state- ment of the basis for determining that the product is a device rather than a drug.

  - For a device that the owner or op- erator has manufactured for distribu- tion under a label other than its own, the names of all distributors for whom it has been manufactured.

Labeling, advertisements, and other information  to  be  submitted upon request in accordance with para- graph (e) of this section may be sub- mitted by postal mail  or  electronically by email, but will not  be  submitted using the FDA electronic device reg- istration and listing system. Electronic submissions of such information must comply with part 11 of this chapter, ex- cept for the requirements in § 11.10 (a),

through (h), and (k), and the cor- responding requirements in  § 11.30  of this chapter. The information  provided in electronic format must be in a form that we can process, review,  and  ar- chive.

[43 FR 37999, Aug. 25, 1978, as amended at 51

FR  33033,  Sept.  18,  1986;  63  FR  5253,  Feb.  2,

1998. Redesignated and  amended at 77  FR 45943, Aug. 2, 2012]

##### § 807.28   Updating  device  listing  infor- mation.

  - Updating of device listing infor- mation is required if an additional es- tablishment begins to engage in any of the activities described in § 807.3(d) with respect to a listed device, such as manufacturing, developing specifica- tions, repackaging, relabeling, or oth- erwise processing the device. Updating

#### § 807.34

of the listing is also required if an es- tablishment begins performing another activity on or to the device, or ceases to perform an activity on or to the de- vice that had previously been identi- fied on the device listing.

  - An owner or operator shall create a new device listing using the FDA electronic device registration and list- ing system:

    - If introducing into commercial distribution an exempt device identi- fied with a product code that is not currently listed by the owner or oper- ator; or

    - If introducing into commercial distribution a non-exempt device with an FDA premarket submission number that is not currently listed by  the owner or operator.

  - All device listings for foreign es- tablishments must be submitted before the device may be imported or offered for import into the United States.

  - An owner or operator who discon- tinues commercial distribution of a de- vice shall discontinue the device list- ing using the FDA electronic device registration and listing system. A de- vice listing is considered discontinued if:

    - All devices under an exempt prod- uct code have been discontinued or

    - All devices associated with an FDA premarket submission number have been discontinued.

  - If commercial distribution of a discontinued device is resumed, the owner or operator must reactivate the previously-discontinued listing  using the electronic device registration and listing system. Any changes to the list- ing information for the product that is the subject of the listing such as a new establishment, new activity, or new proprietary name must be made using the electronic device registration and listing system at the time the listing is reactivated.

  - FDA will assign one listing num- ber for all devices exempt from pre- market  notification   requirements under a single product code. For prod- ucts not exempt from premarket noti- fication requirements, a single listing number will be assigned by FDA for each FDA premarket submission num- ber.

[77 FR 45943, Aug. 2, 2012]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 807.34 Summary of requirements for owners or operators granted a waiver from submitting required in- formation electronically.

- For initial registration and list- ing, owners or operators who have been granted a waiver from electronic filing using   the   procedures   set   forth   in

§ 807.21(b) must send a letter containing all of the registration and listing infor- mation described in §§ 807.22, 807.25, (and § 807.26 when such information is requested by FDA), at the times de- scribed in § 807.22, to: The Office of Compliance, Center for Devices and Ra- diological Health, Food and Drug Ad- ministration,  10903  New   Hampshire Ave., Bldg. 66, rm. 3521, Silver  Spring, MD 20993–0002.

- As specified in § 807.22(b)(1) and (b)(3), all owners or operators shall up- date their establishment registration and device listings annually during the period beginning on October 1 and end- ing on December 31 of each fiscal year.

- Failure  to  submit  any  of  the  re- quired  information  on  time,  as  speci- fied in § 807.22(a) and (b), will put the establishment  in  a  ‘‘failed  to  register’’ or  ‘‘failed  to  list’’  status  as  applicable. The establishment will not be consid- ered  active  and  the  establishment  reg- istration   and   device   listing   informa- tion  may  not  appear  on  the  FDA  Web site  until  the  required  information  is

submitted to and processed by FDA.

[77 FR 45943, Aug. 2, 2012]

##### § 807.35   Notification of registrant.

The Food and Drug Administra- tion will assign each device establish- ment a registration number after verifying the initial establishment reg- istration information that has been submitted. The owner or operator of the establishment will also be assigned an identifying number. Both numbers will be sent to the official cor- respondent by email, or by postal mail if the owner or operator has been granted a waiver from the requirement to file registration and listing informa- tion  electronically.

Owners or operators of device es- tablishments who also manufacture or process biological products (including devices licensed under section 351 of the Public Health Service Act) or drug products   at   the   same   establishment

#### Food and Drug Administration, HHS	§ 807.40

must also register and list those prod- ucts under part 607 or part 207 of this chapter, as appropriate. Registration and listing for human blood and blood products, devices licensed under sec- tion 351 of the Public Health Service Act, and licensed biological products used in the manufacture of a device li- censed under section 351 of the Public Health Service Act, are subject to part 607 of this chapter; registration and listing for all other drug products (in- cluding other biological products that are also regulated as drug products) are subject to part 207 of this chapter.

Although establishment registra- tion and device listing are required to engage in the device  activities  de- scribed in § 807.20, validation of reg- istration and  the  assignment  of  a  de- vice listing number in itself does not establish that the holder of the reg- istration is legally qualified to deal in such devices and does not represent a determination by the Food and Drug Administration as to the status of any device.

[69 FR 11312, Mar. 10, 2004, as amended at 77

FR 45943, Aug. 2, 2012]

##### § 807.37 Public availability of estab- lishment registration and device listing information.

Establishment registration and device listing  information  is  available for  public  inspection  in  accordance with section 510(f) of the Federal Food, Drug, and Cosmetic Act and will be posted on the FDA Web site, with the exception of the  information  identified in paragraph (b) of this section. Re- quests for information by persons  who do not have access to the Internet should be directed to the Office of Com- pliance, Center for Devices and Radio- logical Health, Food and Drug Admin- istration, 10903 New Hampshire Ave., Bldg. 66,  rm. 3521, Silver Spring, MD 20993–0002. In addition, there will be available for inspection at each of the Food and Drug Administration district offices the same information for firms within the geographical area of such district offices.  Upon  request, verification of a registration number or location of a  registered  establishment will be provided.

The following listing information will not be available for public inspec- tion or posted on the FDA Web site:

  - For contract manufacturers, con- tract sterilizers, and private label man- ufacturers, the proprietary or brand name(s) under which a device is mar- keted and the FDA-assigned premarket submission number, if this information would reveal a confidential business re- lationship;

  - FDA-assigned listing numbers.

[77 FR 45943, Aug. 2, 2012]

##### § 807.39 Misbranding by reference to establishment registration or to registration number.

Registration of a device establish- ment or assignment of a registration number does not in any way denote ap- proval of the establishment or its prod- ucts. Any representation that creates an impression of official approval be- cause of registration or possession of a registration number is misleading and constitutes  misbranding.

## Subpart C—Procedures for Foreign Device Establishments

##### § 807.40  Establishment   registration and device listing for foreign estab- lishments importing or offering for import devices into the United States.

Any establishment within any foreign country engaged in the manu- facture, preparation, propagation, compounding, or processing of a device that is imported or offered for import into the United States  shall  register such establishment and list such de- vices using the FDA electronic device registration and listing system in con- formance with the procedures in this section, § 807.41, and subpart B of this part. The official correspondent for the foreign establishment shall facilitate communication between the foreign es- tablishment’s management and rep- resentatives of FDA for matters relat- ing to the registration of device estab- lishments and the listing of device products.

Each foreign establishment re- quired to register under paragraph (a) of this section shall submit the name, address,   and   phone   number   of   its

§ 807.41

United States agent as part of its ini- tial and updated registration informa- tion in accordance with subpart B of this part. Each foreign establishment shall designate only one United States agent and may designate the United States agent to act as its official cor- respondent.

  - The United States agent shall re- side or maintain a place of business in the United States.

  - Upon request from FDA, the United States agent shall  assist  FDA in communications with the foreign es- tablishment, respond to questions con- cerning the foreign establishment’s products that are imported or offered for import into the United States, and assist FDA in scheduling inspections of the foreign establishment. If the agen- cy is unable to contact the foreign es- tablishment directly or expeditiously, FDA may provide information or docu- ments to the United States agent, and such an action shall be considered to be equivalent to providing the same infor- mation or documents to the foreign es- tablishment.

  - The foreign establishment or the United States agent shall report changes in the United States agent’s name, address, or phone number to FDA within 10-business days of the change.

No device may be imported or of- fered for import into the United States unless it is the subject of a device list- ing as required under subpart B of this part and is manufactured, prepared, propagated, compounded, or processed at a registered foreign establishment; however, this restriction  does  not apply to devices imported or offered for import under the investigational use provisions of part 812 of this chapter.

The device establishment reg- istration and device listing informa- tion shall be in the English language.

[66 FR 59160, Nov. 27, 2001, as amended at 77

FR 45944, Aug. 2, 2012]

##### § 807.41 Identification  of  importers and persons who import or offer for import.

Upon initial registration, annu- ally, and at the time of any changes, each foreign establishment required to register    and    list    as    provided    in

§ 807.40(a)   must,   using   the   FDA   elec-

#### 21 CFR Ch. I (4–1–17 Edition)

tronic device registration and listing system, submit the name, address, tele- phone and fax numbers, email address, and registration number, if any  has been assigned, of any  importer  (defined in § 807.3(x)) of the establishment’s de- vices that is known to the foreign es- tablishment. The foreign establishment must also specify which of the estab- lishment’s listed products each  im- porter receives from the foreign estab- lishment.

Upon initial registration, annu- ally, and at the time of any changes, each foreign establishment required to register    and    list    as    provided    in

§ 807.40(a) must, using the FDA elec- tronic device registration and listing system, submit the name, address, tele- phone and fax numbers, email address, and registration number, if any  has been assigned, of each person who im- ports or offers for import the establish- ment’s devices into the United States. The term ‘‘person who imports or of- fers  for  import,’’  which  is  defined  in

§ 807.3(y), includes agents, brokers, or other parties used by the foreign estab- lishment to facilitate the import of its device into the United States.

For each individual or organiza- tion identified by the foreign establish- ment under paragraphs (a) and (b) of this section, the foreign establishment must submit to FDA electronically the current FDA premarket submission number and any other identifying in- formation that is known to the estab- lishment for each device being im- ported or offered for import by the named individuals or organizations.

[77 FR 45944, Aug. 2, 2012]

## Subpart D—Exemptions

##### § 807.65   Exemptions  for  device  estab- lishments.

The following classes of persons are exempt  from  registration  in  accord- ance with § 807.20 under  the  provisions of section 510(g)(1), (g)(2), and (g)(3) of the  act,  or   because   the   Commissioner of Food and Drugs has found, under section 510(g)(5) of the act, that such registration is not necessary for the protection of the public health. The ex- emptions in  paragraphs (d),  (e), (f),  and

of this section are limited to those classes of persons located in any State

Food and Drug Administration, HHS	§ 807.81

as  defined  in  section  201(a)(1)  of  the act.

  - A manufacturer of raw materials or components to be used in the manu- facture or assembly of a device who would otherwise not be required to reg- ister under the provisions of this part.

  - A manufacturer of devices to be used solely for veterinary purposes.

  - A manufacturer of general pur- pose articles such as chemical reagents or laboratory equipment whose  uses are generally known by persons trained in their use and which are not labeled or promoted for medical uses.

  - Licensed practitioners, including physicians, dentists, and optometrists, who manufacture or otherwise alter de- vices solely for use in their practice.

  - Pharmacies, surgical supply out- lets, or other similar retail establish- ments making final delivery or sale to the ultimate user. This exemption also applies to a pharmacy or other similar retail establishment that purchases a device for subsequent distribution under its own name, e.g., a properly la- beled health aid such as an elastic ban- dage or crutch, indicating ‘‘distributed by’’ or ‘‘manufactured for’’ followed by the name of the pharmacy.

  - Persons who manufacture, pre- pare, propagate, compound, or process devices solely for use in research, teaching, or analysis and do not intro- duce such devices into commercial dis- tribution.

  - [Reserved]

  - Carriers by reason of their re- ceipt, carriage, holding  or delivery of devices in the usual course of business as carriers.

  - Persons who dispense devices to the ultimate consumer or whose major responsibility is to render a service necessary to provide the consumer (i.e., patient, physician, layman, etc.) with a device or the benefits to be derived from the use of a device; for example, a hearing aid dispenser, optician, clinical laboratory, assembler of diagnostic x- ray systems, and personnel from a hos- pital,  clinic,  dental  laboratory, orthotic or  prosthetic retail  facility, whose primary responsibility to the ul- timate consumer is to dispense or pro-

vide a service through the use of a pre- viously manufactured device.

[42 FR 42526, Aug. 23, 1977, as amended at 58

FR 46523,  Sept.  1,  1993;  61  FR 44615,  Aug.  28,

1996; 65 FR 17136, Mar. 31, 2000; 66 FR 59160,

Nov. 27, 2001]

## Subpart E—Premarket Notification Procedures

##### § 807.81   When a premarket notification submission is required.

Except  as  provided  in  paragraph

of this section, each person who is required to register his establishment pursuant to § 807.20 must submit a pre- market notification submission to the Food and Drug Administration at  least 90 days before he proposes to begin the introduction or delivery for  introduc- tion into interstate commerce for com- mercial distribution of  a  device  in- tended for human use  which  meets  any of  the  following  criteria:

  - The device is being  introduced into commercial distribution for  the first time; that is, the device  is  not  of the same type as, or is not substan- tially equivalent to, (i) a device in commercial distribution before May 28, 1976, or (ii) a device introduced for commercial distribution after May 28, 1976, that has subsequently been reclas- sified into class I or II.

  - The device is being introduced into commercial distribution for the first time by a person required to reg- ister, whether or not the device meets the criteria in paragraph (a)(1) of this section.

  - The device is one that the person currently has  in commercial  distribu- tion or is reintroducing into commer- cial distribution, but that is  about  to be significantly changed or modified in design, components, method of manu- facture, or intended use. The following constitute significant changes or modi- fications that require a premarket no- tification:

    - A change or modification in the device that could significantly affect the safety or effectiveness of the de- vice, e.g., a significant change or modi- fication in design, material, chemical composition, energy source, or manu- facturing process.

    - A major change or modification in the intended use of the device.

§ 807.85

(b)(1) A premarket notification under this subpart is not required for a device for which a premarket approval appli- cation under section 515 of the  act,  or for which a petition to reclassify under section 513(f)(2) of the act, is pending before the Food and Drug Administra- tion.

(2) The appropriate FDA Center Di- rector may determine that the submis- sion and grant of a written request for an    exception    or    alternative    under

§ 801.128 or § 809.11 of this chapter satis- fies the requirement in paragraph (a)(3) of  this  section.

In addition to complying with the requirements of this part, owners or operators of device establishments that manufacture radiation-emitting elec- tronic products, as defined in § 1000.3 of this chapter, shall comply with the re- porting requirements  of  part  1002  of this   chapter.

[42 FR 42526, Aug. 23, 1977, as amended at 72

FR 73601, Dec. 28, 2007]

##### § 807.85   Exemption	from	premarket notification.

A custom device is exempt from premarket notification requirements of this subpart if the device is within the meaning of section 520(b) of the Fed- eral Food, Drug, and Cosmetic Act.

  - It is intended for use by a patient named in the order of the physician or dentist (or other specially qualified person); or

  - It is intended solely for use by a physician or dentist (or other specially qualified person) and is not generally available to, or generally used by, other physicians or dentists (or other specially qualified persons).

A distributor who places a device into commercial distribution for  the first time under his own name and a re- packager who places his own name on a device and does not change any other labeling or otherwise affect the device shall be exempted from the premarket notification requirements of this sub- part if:

  - The device was in commercial dis- tribution before May 28, 1976; or

  - A premarket notification submis- sion was filed by another person.

[42 FR 42526, Aug. 23, 1977, as amended at 81

FR 70340, Oct. 12, 2016]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 807.87   Information required in a pre- market notification submission.

Each premarket notification submis- sion shall contain the following infor- mation:

The device name, including both the trade or proprietary name and the common or usual name or classifica- tion name of the device.

The establishment registration number, if applicable, of the owner or operator submitting the premarket no- tification submission.

The class in which the device has been put under section 513 of the act and, if known, its appropriate  panel; or, if the owner or operator determines that the device  has not been classified under such section, a statement of that determination and the basis for the person’s determination that  the  device is not so classified.

Action taken by the person re- quired to register to comply with the requirements of the act under section 514 for performance standards.

Proposed labels, labeling, and ad- vertisements sufficient to describe the device, its intended use, and the direc- tions for its use. Where applicable, pho- tographs or engineering drawings should be supplied.

A statement indicating the device is similar to and/or different from other products of comparable type in commercial distribution, accompanied by data to support the statement. This information may include an identifica- tion of similar products, materials, de- sign considerations, energy expected to be used or delivered by the device, and a description of the operational prin- ciples of the device.

Where a person required to reg- ister intends to introduce into com- mercial distribution a device that has undergone a significant change or modification that could significantly affect the safety or effectiveness of the device, or the device is to be marketed for a new or different indication for use, the premarket notification sub- mission must include appropriate sup- porting data to show that the manufac- turer has considered what con- sequences and effects the change or modification or new use might have on the safety and effectiveness of the de- vice.

#### Food and Drug Administration, HHS	§ 807.90

A 510(k) summary as described in

§ 807.92  or  a  510(k)  statement  as  de- scribed in § 807.93.

A financial certification or disclo- sure statement or both, as required by part 54 of this chapter.

For submissions claiming substan- tial equivalence to a device which has been classified into class III under sec- tion 513(b) of the act:

  - Which was  introduced or delivered for introduction  into interstate  com- merce for commercial distribution  be- fore December 1, 1990; and

  - For which no final regulation re- quiring premarket approval has been issued under section 515(b) of the act, a summary of the types of safety and ef- fectiveness  problems  associated   with the type of devices being compared and a citation  to  the  information  upon which the summary is based (class III summary). The 510(k) submitter  shall also certify that a reasonable  search  of all information known or otherwise available about the class III device and other similar  legally  marketed  devices has been conducted (class III certifi- cation), as described in § 807.94. This in- formation does not  refer  to  informa- tion  that  already  has   been   submitted to the Food and Drug Administration (FDA) under  section  519  of  the  act. FDA may require the submission of the adverse safety and effectiveness data described in the class III summary or citation.

A statement that the submitter believes, to the best of his or her knowledge, that all data and informa- tion submitted in the premarket notifi- cation are truthful and accurate and that no material fact has been omitted.

Any additional information re- garding the device requested by the Commissioner that is necessary for the Commissioner to make a finding as to whether or not the device is substan- tially equivalent to a device in com- mercial distribution. A request for ad- ditional information will advise the owner or operator that there is insuffi- cient information contained in the original premarket notification sub- mission for the Commissioner to make this determination and that the owner or operator may either submit the re- quested data or a new premarket noti- fication containing the requested infor-

mation at least 90 days before the owner or operator intends to market the device, or submit a premarket ap- proval application in accordance with section 515 of the act. If the additional information is not submitted within 30 days following the date of the request, the Commissioner will consider the premarket notification to be with- drawn.

(Information collection requirements in this section were approved by the Office of Man- agement and Budget (OMB)  and  assigned OMB control number 0910–0281)

[42 FR 42526, Aug. 23, 1977, as amended at 57

FR 18066,  Apr. 28,  1992;  59  FR 64295,  Dec. 14,

1994; 63 FR 5253, Feb. 2, 1998]

##### § 807.90   Format of a premarket notifi- cation submission.

Each premarket notification submis- sion pursuant to this part shall be sub- mitted in accordance with this section. Each submission shall:

(a)(1) For devices regulated by the Center for Devices and Radiological Health, be addressed to the Food and Drug Administration,  Center  for  De- vices and  Radiological  Health,  Docu- ment Mail Center, 10903 New  Hamp- shire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993–0002.

For devices regulated by the Cen- ter for Biologics Evaluation and Re- search, be addressed to the Food and Drug Administration, Center for Bio- logics Evaluation and Research, Docu- ment Control Center, 10903 New Hamp- shire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002; or for devices regulated by the Center for Drug Eval- uation and Research, be addressed to the Central  Document  Room,  Center for Drug  Evaluation  and  Research, Food and Drug Administration, 5901–B Ammendale Rd., Beltsville, MD 20705– 1266. Information about devices regu- lated by the Center for Biologics Eval- uation and Research is available by using the Center for Biologics Evalua- tion and Research  electronic  Web- based application on the Internet.

All inquiries regarding a pre- market notification submission should be in writing and sent to one of the ad- dresses above.

Be bound into a volume or vol- umes, where necessary.

#### § 807.92

Be submitted in duplicate on standard size paper, including the original and two copies of the cover letter.

Be submitted separately for each product the manufacturer intends to market.

Designated ‘‘510(k) Notification’’ in the cover letter.

[42 FR 42526, Aug. 23, 1977, as amended at 53

FR  11252,  Apr.  6,  1988;  55  FR  11169,  Mar.  27,

1990; 65 FR 17137, Mar. 31, 2000; 70 FR 14986,

Mar. 24, 2005; 75 FR 20915, Apr. 22, 2010; 80 FR

18094, Apr. 3, 2015]

##### § 807.92   Content and format of a 510(k) summary.

A 510(k) summary shall be in suf- ficient detail to provide an under- standing of the basis for a determina- tion of substantial equivalence. FDA will accept summaries as well as amendments thereto until such time as FDA issues a determination of substan- tial equivalence. All 510(k) summaries shall contain the following informa- tion:

  - The submitter’s name, address, telephone number, a contact person, and the date the summary was pre- pared;

  - The name of the device, including the trade or proprietary name if appli- cable, the common or usual name, and the classification name, if known;

  - An identification of the legally marketed device to which the sub- mitter claims equivalence. A legally marketed device to which a new device may be compared for a determination regarding substantial equivalence is a device that was legally marketed prior to May 28, 1976, or a device which has been reclassified from class III to class II or I (the predicate), or a device which has been found to be substan- tially equivalent through the 510(k) premarket notification process;

  - A description of the device that is the subject of the premarket notifica- tion submission, such as might be found in the labeling or promotional material for the device, including an explanation of how the device func- tions, the scientific concepts that form the basis for the device, and the signifi- cant physical and performance charac- teristics of the device, such as device

#### 21 CFR Ch. I (4–1–17 Edition)

design,   material   used,   and   physical properties;

  - A statement of the intended use of the device that is the subject of the premarket notification submission, in- cluding a general description of the diseases or conditions that the device will diagnose, treat, prevent, cure, or mitigate,  including  a   description, where appropriate, of the patient popu- lation for which the device is intended. If the indication statements are dif- ferent from those of the legally mar- keted device identified in paragraph (a)(3) of this section, the 510(k) sum- mary shall contain an  explanation  as to why the differences  are  not  critical to the intended therapeutic, diag- nostic,  prosthetic,  or  surgical  use  of the device, and why the differences do not affect the safety  and  effectiveness of the device when used as labeled; and

  - If the device has the same techno- logical characteristics (i.e., design, ma- terial, chemical composition, energy source) as the predicate device identi- fied in paragraph (a)(3) of this section, a summary of the technological char- acteristics of  the new  device in  com- parison to those of the predicate de- vice. If the device has different techno- logical characteristics from the predi- cate device, a summary of how the technological characteristics of the de- vice compare to a legally marketed de- vice identified in paragraph (a)(3) of this section.

510(k) summaries for those pre- market submissions in which a deter- mination of substantial equivalence is also based on an assessment of per- formance data shall contain the fol- lowing information:

  - A brief discussion of the nonclin- ical tests submitted, referenced, or re- lied on in the premarket notification submission for a determination of sub- stantial equivalence;

  - A brief discussion of the clinical tests submitted, referenced, or relied on in the premarket notification sub- mission for a determination of substan- tial equivalence. This discussion shall include, where applicable, a description of the subjects upon whom the device was tested, a discussion of the safety or effectiveness data obtained from the testing, with specific reference to ad- verse  effects  and  complications,  and

#### Food and Drug Administration, HHS	§ 807.95

any other information from the clin- ical testing relevant to a determina- tion of substantial equivalence; and

  - The conclusions drawn from the nonclinical and clinical tests that dem- onstrate that the device is as safe, as effective, and performs as well as or better than the legally marketed de- vice identified in paragraph (a)(3) of this  section.

The summary should be in a sepa- rate section of the submission, begin- ning on a new page and ending on a page not shared with any other section of the premarket notification submis- sion, and should be clearly identified as a ‘‘510(k) summary.’’

Any other information reasonably deemed necessary by the agency.

[57 FR 18066, Apr. 28, 1992, as amended at 59

FR 64295, Dec. 14, 1994]

##### § 807.93   Content and format of a 510(k) statement.

(a)(1) A 510(k) statement submitted as part of a  premarket  notification shall state as follows:

I certify that, in my capacity as (the  posi- tion held in company by person required to submit the premarket notification,  pref- erably  the  official  correspondent  in  the firm), of (company name), I will make avail- able all information included in this pre- market notification on safety and effective- ness within 30 days of request by any person if the device described in the premarket no- tification submission is determined to be substantially equivalent.  The information  I agree to make available  will  be  a  duplicate of the premarket notification submission, in- cluding any adverse safety and effectiveness information, but excluding all patient iden- tifiers, and trade secret and confidential commercial information, as  defined  in  21 CFR 20.61.

(2) The statement in paragraph (a)(1) of this section should be signed by the certifier, made on a separate page  of the premarket notification submission, and clearly identified as ‘‘510(k) state- ment.’’

All requests for information in- cluded in paragraph (a) of this section shall be made in writing to the cer- tifier, whose name will be published by FDA on the list of premarket notifica- tion submissions for which substantial equivalence determinations have been made.

The information provided to re- questors will be a duplicate of the pre- market notification submission, in- cluding any adverse information, but excluding all  patient  identifiers,  and trade secret and  confidential  commer- cial information as defined in § 20.61 of this   chapter.

[59 FR 64295, Dec. 14, 1994]

##### § 807.94   Format  of  a  class  III  certifi- cation.

A class III certification submitted as part of a premarket notification shall state as follows:

I certify, in my  capacity  as  (position  held in company), of (company name), that I have conducted a reasonable search of all infor- mation known or otherwise available about the types and causes of safety or effective- ness problems that have been reported  for the (type of device). I further  certify  that  I am aware of the types of problems to which the (type of device)  is  susceptible  and  that, to the best of my knowledge, the following summary of the types and causes of safety or effectiveness problems about the (type of de- vice) is complete and accurate.

The statement in paragraph (a) of this section should be signed by the certifier, clearly identified as ‘‘class III certification,’’ and included at the be- ginning of the section of the premarket notification submission that sets forth the class III summary.

[59 FR 64296, Dec. 14, 1994]

##### § 807.95   Confidentiality of information.

The Food and Drug Administra- tion will disclose publicly whether there exists a premarket notification submission under this part:

  - Where the device is on the mar- ket, i.e., introduced or delivered for in- troduction into interstate commerce for commercial distribution;

  - Where the person submitting the premarket notification submission has disclosed, through advertising or any other manner, his intent to market the device to scientists, market analysts, exporters, or other individuals who are not employees of, or paid consultants to, the establishment and who are not in an advertising or law firm pursuant to commercial arrangements with ap- propriate safeguards for secrecy; or

  - Where the device is not on the market  and  the  intent  to  market  the

#### § 807.95

device has not been so disclosed, except where the submission is subject to an exception under paragraph (b) or (c) of this  section.

The Food and Drug Administra- tion will not disclose publicly the ex- istence of a premarket notification submission for a device that is  not  on the market and where the intent to market the device has not been dis- closed for 90 days from the date of re- ceipt of the submission, if:

  - The person submitting the pre- market notification submission re- quests in the submission that the Food and Drug Administration hold as con- fidential commercial information the intent to market the device and sub- mits a written certification to the Commissioner:

    - That the person considers his in- tent to market the device to be con- fidential commercial information;

    - That neither the person nor, to the best of his knowledge, anyone else, has disclosed through advertising or any other manner, his intent to mar- ket the device to scientists, market an- alysts, exporters, or other individuals, except employees of, or paid consult- ants to, the establishment or individ- uals in an advertising or law firm pur- suant to commercial arrangements with appropriate safeguards for se- crecy;

    - That the person will imme- diately notify the Food and Drug Ad- ministration if he discloses the intent to market the device to anyone, except employees of, or paid consultants to, the establishment or individuals in an advertising or law firm pursuant to commercial arrangements with appro- priate safeguards for secrecy;

    - That the person has taken pre- cautions to protect the confidentiality of the intent to market the device; and

    - That the person understands that the submission to the government of false  information  is  prohibited  by  18

U.S.C. 1001 and 21 U.S.C. 331(q); and

  - The Commissioner agrees that the intent to market the device is con- fidential commercial information.

Where the Commissioner deter- mines that the person  has  complied with the procedures described in para- graph (b) of this section with respect to a device that is not on the market and

#### 21 CFR Ch. I (4–1–17 Edition)

where the intent to market the device has not been disclosed, and the Com- missioner agrees that the intent to market the device is confidential com- mercial information, the Commissioner will not disclose the existence of the submission for 90 days from the date of its receipt by the agency. In  addition, the Commissioner  will  continue  not to disclose the existence of such a submis- sion for the device for an additional time when any of the following occurs:

The Commissioner requests in writing additional information  regard- ing the device pursuant to § 807.87(h), in which case the Commissioner will not disclose the existence  of  the  submis- sion until 90 days after the Food and Drug Administration’s receipt of a complete premarket notification sub- mission;

The Commissioner  determines that the device intended to be  intro- duced is a class III  device  and  cannot be marketed without premarket ap- proval  or  reclassification,  in   which case the Commissioner will not  dis- close the existence of the submission unless a petition for reclassification is submitted under section 513(f)(2) of the act and its existence can be disclosed under § 860.5(d) of this chapter; or

FDA will make a 510(k) summary of the safety and effectiveness data available to the public within 30 days of the issuance of a determination that the device is substantially  equivalent to another device. Accordingly, even when a 510(k) submitter has complied with the conditions set forth in para- graphs (b) and (c) of this section, con- fidentiality for a premarket notifica- tion submission cannot be granted be- yond 30 days after FDA issues a deter- mination of equivalency.

Data or information submitted with, or incorporated by reference in, a premarket notification submission (other than safety and effectiveness data that have not been disclosed  to the public) shall be available for disclo- sure by the Food and Drug Administra- tion when the intent to market the de- vice is no longer confidential in accord- ance with this section, unless exempt from public disclosure in accordance with part 20 of this chapter. Upon final classification, data and information re- lating to safety and effectiveness of a

#### Food and Drug Administration, HHS	Pt. 808

device classified in class I (general con- trols) or class II (performance  stand- ards) shall be available for public dis- closure.  Data  and  information   relating to safety and effectiveness of a device classified in class III (premarket ap- proval) that have not been released  to the public shall be retained as con- fidential unless such data and informa- tion become available for release to the public under § 860.5(d) or other provi- sions  of  this  chapter.

[42 FR 42526, Aug. 23, 1977, as amended at 53

FR  11252,  Apr.  6,  1988;  57  FR  18067,  Apr.  28,

1992; 59 FR 64296, Dec. 14, 1994]

##### § 807.97   Misbranding  by  reference  to premarket notification.

Submission of a premarket notifica- tion in accordance with this subpart, and a subsequent determination by the Commissioner that the device intended for introduction into commercial dis- tribution is  substantially  equivalent  to a device in commercial distribution be- fore May 28, 1976, or is substantially equivalent to a device introduced into commercial distribution after May 28, 1976, that has subsequently been reclas- sified into class I or II, does not in any way denote official approval of the de- vice. Any representation that creates an impression of official approval of a device because of complying with the premarket notification regulations is misleading and constitutes mis- branding.

##### § 807.100   FDA  action  on  a  premarket notification.

After review of a premarket noti- fication, FDA will:

  - Issue an order declaring the device to be substantially equivalent to a le- gally marketed predicate device;

  - Issue an order declaring the device to be not substantially equivalent to any  legally  marketed  predicate  device;

  - Request additional information; or

  - Withhold the decision until a cer- tification or disclosure statement is submitted to FDA under part 54 of this chapter.

  - Advise the applicant that the pre-

alent, the applicant may not proceed to market the device.

FDA will determine that a device is substantially equivalent to a predi- cate device using the following cri- teria:

  - The device has the same intended use as the predicate device; and

  - The device:

    - Has the same technological char- acteristics as the predicate device; or

(ii)(A) Has different technological characteristics, such as a significant change in the materials, design, energy source, or other features of the device from those of the predicate device;

The data submitted establishes that the device is substantially equiva- lent to the predicate device and con- tains information, including clinical data if deemed necessary by the Com- missioner, that demonstrates that the device is as safe and as effective as a legally marketed device; and

Does not raise different questions of safety and effectiveness than the predicate device.

  - The predicate device has not been removed from the market at the initia- tive of the Commissioner of Food and Drugs or has not been determined to be misbranded or adulterated by a judicial order.

[57 FR 58403, Dec. 10, 1992, as amended at 63

FR 5253, Feb. 2, 1998]

## PART 808—EXEMPTIONS FROM FEDERAL PREEMPTION OF STATE AND LOCAL MEDICAL DEVICE REQUIREMENTS

##### Subpart A—General Provisions

Sec.

808.1    Scope.

808.3    Definitions.

808.5    Advisory opinions.

##### Subpart B—Exemption Procedures

808.20    Application.

808.25    Procedures for processing an applica- tion.

808.35    Revocation of an exemption.

##### Subpart C—Listing of Specific State and Local Exemptions

| market notification is not required. | 808.53 | Arizona. |
| --- | --- | --- |
| Until the applicant receives an order | 808.55 | California |
| declaring a device substantially equiv- | 808.57 | Connecticut. |

§ 808.1

808.59   Florida.

808.61    Hawaii.

808.67    Kentucky.

808.69  Maine.

808.71      Massachusetts.

  - Minnesota.

  - Mississippi.

808.77    Nebraska.

  - New Jersey.

  - New Mexico.

  - New York.

808.85   Ohio.

  - Oregon.

  - Pennsylvania.

  - Rhode Island.

808.93  Texas.

  - Washington.

  - West Virginia.

  - District  of  Columbia. AUTHORITY: 21 U.S.C. 360j, 360k, 371.

SOURCE:  43  FR  18665,  May  2,  1978,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 808.1   Scope.

    - This part prescribes procedures for the submission, review, and ap- proval of applications for exemption from Federal preemption of State and local requirements applicable to med- ical devices under section 521 of the act.

    - Section 521(a) of the act contains special provisions governing the regu- lation of devices by States and local- ities. That  section  prescribes  a  general rule that after May 28, 1976, no State or political subdivision of a State may es- tablish or continue in effect any re- quirement with respect to a medical device intended for human use having the force and effect of law (whether es- tablished by statute, ordinance, regula- tion, or court decision), which is dif- ferent from, or in addition to, any re- quirement applicable to such device under any provision of the act  and which relates to the safety or effective- ness of the device or to any other mat- ter included in a requirement applica- ble to the device under the act.

    - Section 521(b) of the act contains a provision whereby the Commissioner of Food and Drugs may, upon applica- tion by a State or political subdivision, allow imposition of a  requirement which is different from, or  in  addition to, any requirement applicable under the act to the device (and which is thereby preempted) by promulgating a

21 CFR Ch. I (4–1–17 Edition)

regulation in accordance with this part exempting the State or local require- ment from preemption. The granting of an exemption does not affect the appli- cability to the device of any require- ments under the act. The Commis- sioner may promulgate an exemption regulation for the preempted require- ment if  he makes  either of  the fol- lowing findings:

That the requirement is more stringent than a requirement under the act applicable to the device; or

That the requirement is required by compelling local conditions and compliance with the  requirement would not cause the device to be in vio- lation of any applicable requirement under the act.

    - State or local requirements are preempted only when the Food  and Drug Administration has established specific   counterpart   regulations   or there are other specific requirements applicable to a particular device under the act, thereby making any existing divergent State or local requirements applicable to the device  different  from, or in addition to, the specific Food and Drug  Administration   requirements. There are other State or local require- ments that affect devices that are not preempted by section 521(a) of the act because they are not ‘‘requirements ap- plicable to a device’’ within the  mean- ing of section 521(a) of the act. The fol- lowing are examples of State or local requirements that are not regarded as preempted by section 521 of the act:

      - Section 521(a) does not preempt State or local requirements of general applicability where the purpose of the requirement relates either to other products in addition to devices (e.g., re- quirements such as general electrical codes, and the Uniform Commercial Code (warranty of fitness)), or to unfair trade practices in which the require- ments are not limited to devices.

      - Section 521(a) does not preempt State or local requirements that are equal to, or substantially identical to, requirements imposed by or under the act.

      - Section 521(a) does not preempt State or local permits, licensing, reg- istration, certification, or other re- quirements relating to the approval or sanction  of  the  practice  of  medicine,

#### Food and Drug Administration, HHS	§ 808.1

dentistry, optometry, pharmacy, nurs- ing, podiatry, or any other of the heal- ing arts or allied medical sciences or related professions or occupations that administer, dispense, or sell devices. However, regulations issued under sec- tion 520(e) or (g) of the act may impose restrictions  on  the  sale,   distribution, or use of a device beyond those pre- scribed in State  or  local  requirements. If there is a conflict between such re- strictions and State or local require- ments, the Federal regulations shall prevail.

      - Section 521(a) does not preempt specifications in contracts entered into by States or localities for procurement of devices.

      - Section 521(a) does not preempt criteria for payment of State or local obligations under Medicaid and similar Federal, State or local health-care pro- grams.

(6)(i) Section 521(a) does not preempt State or local requirements respecting general enforcement, e.g.,  require- ments that State inspection be per- mitted  of  factory  records  concerning all devices, registration, and licensing requirements for manufacturers and others, and prohibition of manufacture of devices in unlicensed establish- ments. However, Federal regulations issued under sections 519 and 520(f) of the act may impose requirements for records and reports and good manufac- turing practices beyond those pre- scribed in State  or  local  requirements. If there is a conflict between such regu- lations and State or local require- ments, the Federal regulations shall prevail.

(ii) Generally, section 521(a) does not preempt a State or local requirement prohibiting the manufacture of adul- terated or misbranded devices. Where, however, such a prohibition has the ef- fect of establishing a substantive re- quirement for a specific device, e.g., a specific labeling requirement, then the prohibition will be preempted if the re- quirement is different from, or in addi- tion to, a Federal requirement estab- lished under the act. In determining whether such a requirement is pre- empted, the determinative factor is how the requirement is interpreted and enforced by the State or local govern- ment  and  not  the  literal  language  of

the statute, which may be identical to a provision in the act.

Section 521(a) does not preempt State or local provisions respecting delegations of authority and related administrative matters relating to de- vices.

Section 521(a) does not preempt a State or local requirement whose sole purpose is raising revenue or charging fees for services, registration, or regu- latory  programs.

Section 521(a) does not preempt State or local requirements of the types that have been developed under the   Atomic   Energy   Act   of   1954   (42

U.S.C. 2011 note), as amended, Sub- chapter C—Electronic Product Radi- ation Control of the Federal Food, Drug, and Cosmetic Act (formerly the Radiation Control for Health and Safe- ty Act of 1968), and other Federal stat- utes, until such time as the Food and Drug Administration issues specific re- quirements under the Federal Food, Drug, and Cosmetic Act applicable to these types of devices.

Part 820 of this chapter (21 CFR part 820) (CGMP requirements) does not preempt remedies created  by States or Territories of the United States, the District of Columbia, or the Commonwealth of Puerto Rico.

    - It is the responsibility of the Food and Drug Administration, subject to review by Federal courts, to determine whether a State or local requirement is equal to, or substantially identical to, requirements imposed by or under the act, or is different from, or in addition to, such requirements, in accordance with the procedures provided by this part. However, it is the responsibility of States and political subdivisions to determine initially whether to seek ex- emptions from preemption. Any State or political subdivision whose require- ments relating to devices are pre- empted by section 521(a) may petition the Commissioner of Food and Drugs for exemption from preemption, in ac- cordance with the procedures provided by this part.

    - The Federal requirement with re- spect to a device applies whether or not a corresponding State or local require- ment is preempted or exempted from preemption. As a result, if a State or local  requirement  that  the  Food  and

#### § 808.3

Drug Administration has exempted from preemption is not as broad in its application as the Federal require- ment, the Federal requirement applies to all circumstances not covered by the State or local requirement.

[43 FR 18665, May 2, 1978, as amended at 45

FR  67336,  Oct.  10,  1980;  61  FR  52654,  Oct.  7,

1996; 73 FR 34859, June 19, 2008]

##### § 808.3   Definitions.

Act means the Federal Food, Drug, and Cosmetic Act.

Compelling local conditions includes any factors, considerations, or cir- cumstances prevailing in, or char- acteristic of, the geographic area or population of the State or political subdivision that justify  exemption from preemption.

More stringent refers to a require- ment of greater restrictiveness or one that is expected to afford to those who may be exposed to a risk of injury from a device a higher degree of protection than is afforded by a requirement ap- plicable to the device under the act.

Political subdivision or locality means any lawfully established local governmental unit within a  State which unit has the authority to estab- lish or continue in effect any require- ment having the force and effect of law with respect to a device intended for human use.

State means a State, American Samoa, the Canal Zone, the Common- wealth of Puerto Rico, the District of Columbia, Guam, Johnston Island, Kingman Reef, Midway Island, the Trust Territory of the Pacific Islands, the Virgin Islands, and Wake Island.

Substantially identical to refers to the fact that a State or local require- ment does not significantly differ in ef- fect from a Federal requirement.

##### § 808.5   Advisory opinions.

Any  State,  political  subdivision, or other interested person may request an advisory opinion from the Commis- sioner with respect to any general mat- ter concerning preemption of State or local device requirements or with re- spect to whether the Food and Drug Administration  regards   particular State or local requirements, or pro- posed requirements, as preempted.

#### 21 CFR Ch. I (4–1–17 Edition)

      - Such an advisory opinion may be requested and may be granted in ac- cordance with § 10.85 of this chapter.

      - The Food and Drug Administra- tion, in its discretion and after con- sultation with the State or political subdivision, may treat a request by a State or political subdivision for an ad- visory opinion as an application for ex- emption      from      preemption      under

§ 808.20.

The Commissioner may issue an advisory opinion  relating to  a State  or local requirement on his own initiative when he makes one of the following de- terminations:

  - A requirement with respect to a device for which an application for ex- emption from preemption has been sub- mitted under § 808.20 is not preempted by section 521(a) of the act because it is: (i) Equal to or substantially iden- tical to a requirement under the act applicable to the device, or (ii) is not a requirement within the meaning of sec- tion 521 of the act and therefore is not preempted;

  - A proposed State or local require- ment with respect to a device is not el- igible for exemption from preemption because the State or local requirement has not been issued in final form. In such a case, the advisory opinion may indicate whether the proposed require- ment would be preempted and, if it would be preempted, whether the Food and Drug Administration would pro- pose to grant an exemption from pre- emption;

  - Issuance of such an advisory opin- ion is in the public interest.

## Subpart B—Exemption Procedures

##### § 808.20   Application.

    - Any State or political subdivision may apply to the Food and Drug Ad- ministration for an exemption from preemption for any requirement that it has enacted and that is preempted. An exemption may only be granted for a requirement that has been enacted, promulgated, or issued in final form by the authorized body or official of the State or political subdivision so as to have the force and effect of law. How- ever, an application for exemption may be submitted before the effective date of the requirement.

Food and Drug Administration, HHS	§ 808.20

    - An  application  for   exemption shall be in the form of a letter to the Commissioner of Food and Drugs and shall be signed by an individual who is authorized to request the exemption on behalf of the State or political subdivi- sion. An original and two copies of the letter  and  any  accompanying   material, as well as any subsequent reports or correspondence concerning an applica- tion, shall be submitted  to  the  Division of  Dockets  Management  (HFA–305), Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852. The outside wrapper of any appli- cation,  report,  or  correspondence should indicate that it concerns an ap- plication for exemption from  preemp- tion  of  device  requirements.

    - For each requirement for which an exemption is sought, the application shall include the following information to the fullest extent possible, or an ex- planation of why such information has not been included:

      - Identification and a current copy of any statute, rule, regulation, or or- dinance of the State or political sub- division considered by the State or po- litical subdivision to be a requirement which is preempted, with a reference to the date of enactment,  promulgation, or issuance in final form. The applica- tion shall also include, where avail- able, copies of any legislative  history or background materials pertinent to enactment, promulgation, or issuance of the requirement, including hearing reports or studies concerning develop- ment or consideration of the require- ment. If the requirement has been sub- ject to any judicial or administrative interpretations, the State or political subdivision shall furnish copies of such judicial or administrative interpreta- tions.

      - A comparison of the  requirement of the  State  or  political  subdivision and any applicable Federal require- ments to show similarities and dif- ferences.

      - Information on the nature of the problem addressed by the requirement of the State or political subdivision.

      - Identification of which (or  both) of the following bases is relied upon for seeking an exemption from preemp- tion:

The requirement is more stringent than a requirement applicable to a de- vice under the act. If the State or po- litical subdivision  relies  upon  this basis for exemption from preemption, the application shall include informa- tion, data,  or  material  showing  how and why the requirement of  the  State or political subdivision is more strin- gent  than  requirements  under  the  act.

The requirement is required by compelling local conditions, and com- pliance with the requirement would not cause the device to be in violation of any applicable requirement under the act. If the State or political sub- division relies upon this basis for ex- emption from preemption, the applica- tion shall include information, data, or material showing why compliance with the requirement of the State or polit- ical subdivision would not cause a de- vice to be in violation of any applicable requirement under the act and why the requirement is required by compelling local conditions. The application shall also explain in detail the compelling local conditions that justify the re- quirement.

      - The title of the chief administra- tive or legal officers of that State or local agency that has primary respon- sibility for administration of the re- quirement.

      - When requested by the Food and Drug Administration, any records con- cerning administration of any require- ment which is the subject of an exemp- tion or an application for an exemption from preemption.

      - Information on how the public health may be benefitted and how interstate commerce may be affected, if an exemption is granted.

      - Any other pertinent information respecting the requirement voluntarily submitted by the applicant.

    - If litigation regarding applica- bility of the requirement is pending, the State or political subdivision may so indicate in its application and re- quest expedited action on such applica- tion.

[43 FR 18665, May 2, 1978; 43 FR 22010, May 23,

1978, as amended at 49 FR 3646, Jan. 30, 1984;

59 FR 14365, Mar. 28, 1994]

#### § 808.25

##### § 808.25   Procedures  for  processing  an application.

Upon receipt of an application for an exemption from preemption sub- mitted in accordance with § 808.20, the Commissioner shall notify the State or political subdivision of the date of such receipt.

If the Commissioner finds that an application does not meet the require- ments of § 808.20, he shall notify the State or political subdivision of the de- ficiencies in  the  application and of the opportunity to correct such defi- ciencies.  A  deficient  application  may be corrected at any time.

After receipt of an application meeting the requirements of § 808.20, the Commissioner shall review such ap- plication and determine whether to grant or deny an exemption from  pre- emption for each requirement which is the subject of the application. The Commissioner shall then issue in the FEDERAL REGISTER a proposed  regula- tion either to grant or to deny an ex- emption from preemption. The Com- missioner shall also issue in the FED- ERAL REGISTER a notice  of  opportunity to request an oral hearing before the Commissioner or the Commissioner’s designee.

A request for an oral hearing may be made by the State or political sub- division or any other interested person. Such request shall be submitted to the Division of Dockets Management  with- in the period of time prescribed in the notice and shall include an explanation of why an oral hearing, rather than submission of written  comments  only, is essential  to the  presentation  of views on the application for exemption from preemption and the proposed reg- ulation.

If a timely request for an oral hearing  is  made,  the  Commissioner shall review such a  request  and  may grant a legislative-type informal oral hearing pursuant to part 15 of this chapter by publishing in the FEDERAL REGISTER a notice of the hearing in ac- cordance with § 15.20  of  this  chapter. The scope of the oral hearing shall be limited to matters relevant to the ap- plication for exemption from  preemp- tion and the proposed  regulation.  Oral or written presentations at the oral hearing  which  are  not  relevant  to  the

#### 21 CFR Ch. I (4–1–17 Edition)

application shall be excluded from the administrative record of the hearing.

If a request for hearing is  not timely made or  a  notice  of  appearance is not filed pursuant to § 15.21 of this chapter, the Commissioner shall con- sider all written  comments  submitted and publish a final rule in accordance with  paragraph  (g)  of  this  section.

(g)(1) The Commissioner shall review all written comments submitted on the proposed rule and the administrative record of the oral hearing, if an oral hearing has been granted, and shall publish in  the  FEDERAL  REGISTER  a final rule in subpart C of this part iden- tifying any requirement in the applica- tion for which exemption from preemp- tion is granted, or conditionally grant- ed, and any requirement in the applica- tion for which exemption from preemp- tion is not granted.

The Commissioner may issue a regulation granting or conditionally granting an application for an exemp- tion from preemption for any require- ment if the Commissioner makes ei- ther of the following findings:

The requirement is more stringent than a requirement applicable to the device under the act;

The requirement is required by compelling local conditions, and com- pliance with the  requirement  would not cause the device to be in violation of any requirement applicable to the device under the act.

The Commissioner may not grant an application for an exemption from preemption for any requirement with respect to a device if the Commissioner determines that the  granting of an  ex- emption would not be in the best inter- est of public health, taking  into  ac- count the potential burden on inter- state  commerce.

(h)  An  advisory  opinion  pursuant  to

§ 808.5 or a regulation pursuant to para- graph (g) of this  section  constitutes final   agency   action.

##### § 808.35   Revocation of an exemption.

An exemption from preemption pursuant to a  regulation  under  this part shall remain effective until the Commissioner revokes such exemption.

The Commissioner may by regula- tion, in accordance with § 808.25, revoke

#### Food and Drug Administration, HHS	§ 808.59

an exemption from preemption for any of the following reasons:

  - An exemption may be revoked upon the effective date of a newly es- tablished requirement under the act which, in the Commissioner’s view, ad- dresses the objectives of an exempt re- quirement and which  is  described, when issued, as preempting a pre- viously exempt State or local require- ment.

  - An exemption may be revoked upon a finding that there has occurred a    change    in    the    bases    listed    in

§ 808.20(c)(4)  upon  which  the  exemption was   granted.

  - An exemption may be revoked if it is determined that a condition placed on the exemption by the regulation under which the exemption was grant- ed has not been met or is no longer being met.

  - An exemption  may  be  revoked  if  a State or local jurisdiction fails to sub- mit records as provided in § 808.20(c)(6).

  - An exemption may be revoked if a State or local jurisdiction to whom the exemption was originally granted re- quests revocation.

  - An exemption may be revoked if it is determined that it is no  longer  in the best interests of the public health to continue the exemption.

An exemption that has been re- voked may be reinstated, upon request from the State  or  political  subdivision, if the Commissioner, in accordance with the procedures in § 808.25, deter- mines that the grounds for revocation are no longer applicable except that the Commissioner may permit abbre- viated submissions of the documents and materials normally required for an application      for      exemption      under

§ 808.20.

## Subpart C—Listing of Specific State and Local Exemptions

##### § 808.53   Arizona.

The following Arizona medical device requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has denied them exemptions from preemption under sec- tion 521(b) of the act:

Arizona  Revised  Statutes,   Chap- ter 17, sections 36–1901.7(s) and 36– 1901.7(t).

Arizona Code of Revised Regula- tions, Title 9, Article 3, sections R9–16– 303 and R9–16–304.

[45 FR 67336, Oct. 10, 1980]

##### § 808.55   California.

The following California medical device requirements are enforceable notwithstanding section 521 of the act because the Food and Drug Adminis- tration exempted them from  preemp- tion under section 521(b) of the act: Business and Professions Code sections 3365 and 3365.6.

The following California medical device requirements are preempted by section 521 of the act, and FDA has de- nied them an exemption from preemp- tion:

  - Sherman Food, Drug, and Cos- metic Law (Division 21 of the  Cali- fornia Health and Safety  Code),  sec- tions 26207, 26607, 26614, 26615, 26618, 26631, 26640, and 26641, to the extent that they apply to devices.

  - Sherman Food, Drug, and Cos- metic Law, section 26463(m) to the ex- tent that it applies to hearing aids.

  - Business and Professions Code sec- tion 2541.3, to the extent that it re- quires adoption of American National Standards Institute standards Z–80.1 and Z–80.2.

[45 FR 67324, Oct. 10, 1980]

##### § 808.57   Connecticut.

The following Connecticut medical device requirements are enforceable notwithstanding section  521(a)  of  the act because the Food and Drug Admin- istration has exempted them from pre- emption under section 521(b) of the act: Connecticut General Statutes, sections 20–403 and 20–404.

[45 FR 67336, Oct. 10, 1980]

##### § 808.59   Florida.

The following Florida medical device requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has denied them an exemption from preemption under section 521(b) of the act:

Florida Statutes, section 468.135(5).

§ 808.61

Florida Administrative Code, sec- tion 10D–48.25(26).

[45 FR 67336, Oct. 10, 1980]

##### § 808.61   Hawaii.

The following Hawaii medical de- vice requirements are enforceable not- withstanding section 521 of the act, be- cause the Food and Drug Administra- tion has exempted them from preemp- tion under section 521(b) of the act: Ha- waii   Revised   Statutes,   chapter   451A,

§ 14.1, subsection (a) with respect to medical examination of a child 10 years of age or under, and subsection (c).

The following Hawaii medical de- vice requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has denied them exemption from preemption: Ha- waii   Revised   Statutes,   chapter   451A,

§ 14.1, subsection (a) to the extent that it requires a written authorization by a physician and does not allow adults to waive this requirement for personal, as well as religious reasons, and sub- section (b).

[50 FR 30699, July 29, 1985; 50 FR 32694, Aug.

14, 1985]

##### § 808.67   Kentucky.

The following Kentucky medical de- vice requirement is preempted by sec- tion 521(a) of the act, and the Food and Drug Administration has denied it an exemption from preemption under sec- tion 521(b) of the act: Kentucky Re- vised Statutes, section 334.200(1).

[45 FR 67336, Oct. 10, 1980]

##### § 808.69   Maine.

The following Maine medical de- vice requirement is enforceable not- withstanding section 521(a) of the act because the Food and Drug Adminis- tration has exempted it from preemp- tion under section 521(b) of the act: Maine Revised Statutes Annotated, Title 32, section 1658–C, on the condi- tion that, in enforcing this require- ment, Maine apply the definition of ‘‘used hearing aid’’ in § 801.420(a)(6) of this chapter.

The following Maine medical de- vice requirement is preempted by sec- tion 521(a) of the act, and the Food and Drug Administration has denied it an exemption from preemption under sec-

#### 21 CFR Ch. I (4–1–17 Edition)

tion 521(b) of the act: Maine Revised Statutes Annotated, Title 32, section 1658–D and the last sentence of section 1658–E.

[45 FR 67336, Oct. 10, 1980]

##### § 808.71   Massachusetts.

The following Massachusetts med- ical device requirements are  enforce- able notwithstanding section 521 of the act because the Food and Drug Admin- istration has exempted them from pre- emption under section 521(b) of the act:

Massachusetts General Laws, Chapter 93, Section 72, to the extent that it requires a hearing test evalua- tion for a child under the age of 18.

Massachusetts General Laws, Chapter 93, Section 74, except as pro- vided in paragraph (6) of the Section, on the condition that, in enforcing this requirement, Massachusetts apply the definition   of   ‘‘used   hearing   aid’’   in

§ 801.420(a)(6) of this chapter.

The following Massachusetts med- ical device requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has de- nied them exemptions from preemption under section 521(b) of the act.

Massachusetts General Laws, Chapter 93, Section 72, except as pro- vided in paragraph (a) of this section.

Massachusetts General Laws, Chapter 93, Section 74, to the extent that it requires that the sales receipt contain a statement that State law re- quires a medical examination and a hearing test evaluation before the sale of a hearing aid.

[45 FR 67326, Oct. 10, 1980]

##### § 808.73   Minnesota.

The following Minnesota medical de- vice requirements are preempted by section 521(a) of the act, and the  Food and Drug  Administration  has  denied them an exemption from preemption under section 521(b) of the act: Min- nesota Statutes, sections 145.43 and 145.44.

[45 FR 67336, Oct. 10, 1980]

##### § 808.74   Mississippi.

The following Mississippi medical de- vice requirement is preempted by sec- tion 521(a) of the act, and the Food and Drug  Administration  has  denied  it  an

#### Food and Drug Administration, HHS	§ 808.85

exemption from preemption under sec- tion 521(b) of the act: Mississippi Code, section  73–14–3(g)(9).

[45 FR 67336, Oct. 10, 1980]

##### § 808.77   Nebraska.

  - The following Nebraska medical device requirement is enforceable not- withstanding section 521(a) of the act because the Food and Drug Adminis- tration has exempted it from preemp- tion under section 521(b) of the act: Ne- braska Revised Statutes, section 71– 4712(2)(c)(vi).

  - The following Nebraska medical device requirement is preempted by section 521(a) of the act, and the Food and Drug Administration has denied  it an exemption from preemption under section 521(b) of the act: Nebraska Re- vised	Statutes,	section	71–

4712(2)(c)(vii).

[45 FR 67336, Oct. 10, 1980]

##### § 808.80   New Jersey.

The following New Jersey medical device requirements are enforceable notwithstanding section 521(a) of the act because the Food and Drug Admin- istration has exempted them from pre- emption under section 521(b) of the act:

  - New Jersey Statutes Annotated, section 45:9A–23 on the condition that, in enforcing this requirement, New Jer- sey apply the definition of ‘‘used hear- ing aid’’ in § 801.420(a)(6) of this chap- ter;

  - New Jersey Statutes Annotated, sections 45:9A–24 and 45:9A–25;

  - Chapter 3, Section 5 of the Rules and Regulations adopted pursuant to New Jersey Statutes Annotated 45:9A–1 et seq. except as provided in paragraph

of this section.

(b) The following New Jersey medical device requirement is preempted by section 521(a) of the act, and the Food and Drug Administration has denied  it an exemption from preemption under section 521(b) of the act: Chapter 3, Section 5 of the Rules and Regulations adopted pursuant to New Jersey Stat- utes Annotated 45:9A–1 et seq. to the ex- tent that it requires testing to be con- ducted in an environment which meets or exceeds the American National Standards  Institute  S3.1  Standard.

[45 FR 67337, Oct. 10, 1980]

##### § 808.81   New Mexico.

The following New Mexico medical device requirement is enforceable not- withstanding section 521(a) of the act because the Food and Drug Adminis- tration has exempted it from preemp- tion under section 521(b) of the act: New Mexico Statutes Annotated, sec- tion  67–36–16(F).

[45 FR 67337, Oct. 10, 1980]

##### § 808.82   New York.

The following New York medical device requirements are enforceable notwithstanding section 521(a) of the act because the Food and Drug Admin- istration has exempted them from pre- emption under section 521(b) of the act:

  - General Business Law, Article 37, sections 784(3) and (4).

  - Official Compilation  of  Codes, Rules and Regulations of the State of New York, Chapter V, Title 19, Sub- chapter G, section 191.10 and section 191.11(a) on the condition that, in en- forcing these requirements, New York apply the definition  of  ‘‘used  hearing aid’’ in § 801.420(a)(6) of this chapter and section 191.11(b), (c), (d), and (e).

The following New York medical device requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has denied them an exemptions from preemption under section 521(b) of the act:

  - General Business Law, Article 37, section 784.1.

  - Official Compilation of Codes, Rules and Regulations of the State of New York, Chapter V, Title 19, Sub- chapter G, sections 191.6, 191.7, 191.8, and 191.9.

[45 FR 67337, Oct. 10, 1980]

##### § 808.85   Ohio.

The following Ohio medical device requirement is enforceable notwith- standing section 521(a) of the act be- cause the Food  and  Drug  Administra- tion has exempted it from preemption under section 521(b) of the act: Ohio Revised Code, section 4747.09, the first two sentences with respect  to  disclo- sure  of  information  to  purchasers  on the condition that, in enforcing these requirements, Ohio apply the  definition of ‘‘used hearing aid’’  in  § 801.420(a)(6) of  this  chapter.

#### § 808.87

The following Ohio medical device requirement is preempted by section 521(a) of the act, and the Food and Drug Administration has denied it an exemption from preemption under sec- tion 521(b) of the act: Ohio Revised Code, section 4747.09, the last two sen- tences with respect to medical exam- ination of children.

[45 FR 67337, Oct. 10, 1980]

##### § 808.87   Oregon.

The following Oregon medical de- vice requriements are enforceable not- withstanding section 521(a) of the act because the Food and Drug Adminis- tration has exempted them from pre- emption under section 521(b) of the act: Oregon      Revised      Statutes,      section

694.036 on the condition that, in enforc- ing this requirement, Oregon apply the definition    of    ‘‘used    hearing    aid’’    in

§ 801.420(a)(6) of this chapter.

The following Oregon medical de- vice requirements are preempted by section 521(a) of the act, and the  Food and Drug  Administration  has  denied them exemptions from  preemption under section 521(b) of the act: Oregon Revised  Statutes,  sections  694.136(6) and (7).

[45 FR 67337, Oct. 10, 1980, as amended at 53

FR 11252, Apr. 6, 1988]

##### § 808.88   Pennsylvania.

The following Pennsylvania med- ical device requirements are  enforce- able notwithstanding section 521(a) of the act because the Food and Drug Ad- ministration has exempted them from preemption under section 521(b) of the act: 35 Purdon’s Statutes 6700, section 504(4) on the condition that, in  enforc- ing this  requirement,  Pennsylvania apply the definition  of  ‘‘used  hearing aid’’ in § 801.420(a)(6) of this chapter; section 506; and, section 507(2).

The following Pennsylvania med- ical  device  requirement  is  preempted by section 521(a)  of  the  act  and  the Food and Drug Administration has de- nied it an exemption from preemption under section 521(b) of the act: 35 Purdon’s Statutes 6700, section 402.

[45 FR 67326, Oct. 10, 1980]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 808.89   Rhode Island.

The following Rhode Island medical device requirements are preempted by section 521(a) of the act, and the Food and Drug Administration has denied them an exemption from preemption under section 521(b) of the act: Rhode Island General Laws, Section 5–49–2.1, and Section 2.2, to the extent that Sec- tion 2.2 requires hearing aid dispensers to keep copies of the  certificates  of need.

[45 FR 67337, Oct. 10, 1980]

##### § 808.93   Texas.

The following Texas medical  de- vice requirement is enforceable not- withstanding section 521(a) of the act because the Food and Drug Adminis- tration has exempted it from preemp- tion under section 521(b) of the act: Vernon’s Civil Statutes, Article 4566, section 14(b) on the condition that, in enforcing  this   requirement,   Texas apply the definition  of  ‘‘used  hearing aid’’ in § 801.420(a)(6) of this chapter.

The following Texas medical de- vice requirement is preempted by sec- tion 521(a) of the act, and the Food and Drug Administration has denied it an exemption from preemption under sec- tion 521(b) of the act: Vernon’s Civil Statutes, Article 4566, section 14(d).

[45 FR 67337, Oct. 10, 1980]

##### § 808.97   Washington.

The following Washington med- ical device requirement is enforceable notwithstanding section 521(a) of the act because the Food and Drug Admin- istration has exempted it from preemp- tion under section 521(b) of the act: Re- vised Code of Washington 18.35.110(2)(e)

(i) and (iii) on the condition that it is enforced in addition to the applicable requirements of this chapter.

The following Washington med- ical device requirements are  preempted by section 521(a) of the act, and the Food and Drug Administration has de- nied them an exemption from preemp- tion under section 521(b) of the act: Re- vised	Code	of	Washington 18.35.110(2)(e)(ii).

[45 FR 67337, Oct. 10, 1980]

#### Food and Drug Administration, HHS	§ 809.4

##### § 808.98   West Virginia.

  - The following West Virginia med- ical device requirements are  enforce- able notwithstanding section 521(a) of the act because the Food and Drug Ad- ministration has exempted them from preemption: West Virginia Code, sec- tions 30–26–14 (b) and (c) and section 30– 26–15(a) on the condition that in enforc- ing section 30–26–15(a) West Virginia apply the definition  of  ‘‘used  hearing aid’’ in § 801.420(a)(6) of this chapter.

  - The following West Virginia med- ical device requirement is preempted by section 521(a) of the act, and the Food and Drug Administration has de- nied it an exemption from preemption under section 521(b) of the act: West Virginia Code, section 30–26–14(a).

[45 FR 67337, Oct. 10, 1980, as amended at 53

FR 35314, Sept. 13, 1988]

##### § 808.101   District of Columbia.

The following District of Colum- bia medical device requirements are enforceable,   notwithstanding   section

521 of the act, because the Food and Drug  Administration  has  exempted them from preemption under section 521(b) of the act:

- Act 2–79, section 5, to the extent that it requires an audiological evalua- tion for children under the age of 18.

- Act 2–79, section 6, on the condi- tion that in enforcing section 6(a)(5), the District of Columbia apply the defi- nition     of     ‘‘used     hearing     aid’’     in

§ 801.420(a)(6) of this chapter.

The following District of Colum- bia medical device requirement is pre- empted by section 521(a) of the act, and the Food and Drug Administration has denied it an exemption from preemp- tion under section 521(b) of the act: Act 2–79, section 5, except as provided in paragraph (a) of this section.

[46 FR 59236, Dec. 4, 1981]

## PART 809—IN VITRO DIAGNOSTIC PRODUCTS FOR HUMAN USE

##### Subpart A—General Provisions

Sec.

  - Definitions.

  - Confidentiality of submitted informa- tion.

##### Subpart B—Labeling

  - Labeling for in vitro diagnostic prod- ucts.

  - Exceptions or alternatives to labeling requirements for in vitro diagnostic products for human use held by the Stra- tegic National Stockpile.

##### Subpart C—Requirements for Manufacturers and Producers

809.20 General requirements for manufac- turers and producers of in vitro  diag- nostic   products.

809.30 Restrictions on the  sale,  distribution and  use  of  analyte  specific  reagents.

  - Restrictions on the sale, distribution, and use of OTC test sample collection systems for drugs of abuse testing.

AUTHORITY: 21 U.S.C. 331, 351, 352, 355, 360b,

360c, 360d, 360h, 360i, 360j, 371, 372, 374, 381.

## Subpart A—General Provisions

##### § 809.3   Definitions.

    - In vitro diagnostic products are those reagents, instruments, and sys- tems intended for use in the diagnosis of disease or other conditions, includ- ing a determination of the state of health,  in  order  to  cure,   mitigate, treat, or prevent  disease  or  its sequelae. Such products are intended for use in the collection, preparation, and examination of specimens taken from the human body. These products are devices as defined in section 201(h) of the Federal Food, Drug, and Cos- metic Act (the act), and may also be bi- ological products subject to section 351 of the Public Health Service Act.

    - A product class is all those prod- ucts intended for use for a particular determination or for a related group of determinations or products with com- mon or related characteristics or those intended for common or related uses. A class may be further divided into sub- classes when appropriate.

    - [Reserved]

    - Act means the Federal Food, Drug, and Cosmetic Act.

[41  FR  6903,  Feb.  13,  1976,  as  amended  at  45

FR 7484, Feb. 1, 1980]

##### § 809.4 Confidentiality of submitted in- formation.

Data and information  submitted under § 809.10(c) that are shown to fall within   the   exemption   established   in

§ 809.10

§ 20.61 of this chapter  shall  be  treated as confidential by the Food and Drug Administration  and  any  person   to whom the data and information are re- ferred. The Food and Drug Administra- tion will determine whether  informa- tion submitted will be treated as con- fidential in accordance with the provi- sions of part 20 of this chapter.

[45 FR 7484, Feb. 1, 1980]

## Subpart B—Labeling

##### § 809.10   Labeling   for   in   vitro   diag- nostic products.

      - The label for an in vitro diag- nostic product shall state the following information, except where such infor- mation is not applicable, or as other- wise specified in a standard for a par- ticular product class or as provided in paragraph (e) of this section. Section 201(k) of the act provides that ‘‘a re- quirement made by or under authority of this act that  any  word,  statement, or other information appear on the label shall not be considered to be com- plied with unless such word, statement, or other information also appears on the outside container or wrapper, if any there be, of the retail package of such article, or is easily legible through the outside container or wrap- per.’’

        - The proprietary name and estab- lished name (common or usual name), if  any.

        - The intended use or uses of the product.

        - For a reagent, a declaration of the established name (common or usual name), if any, and quantity, proportion or concentration of each reactive in- gredient; and for a reagent derived from biological material, the source and a measure of its activity. The quantity, proportion, concentration, or activity shall  be  stated  in the  system generally used and recognized by the intended user, e.g., metric, inter- national units, etc.

        - A statement of warnings or pre- cautions for users as established in the regulations contained in 16 CFR part 1500 and any other warnings appro- priate to the hazard presented by the product; and a statement ‘‘For In Vitro Diagnostic Use’’ and any other limiting statements appropriate to the intended

21 CFR Ch. I (4–1–17 Edition)

use of the product. The limiting state- ment appropriate to the intended use of a prescription in vitro diagnostic product shall bear the symbol state- ment ‘‘Rx only’’ or ‘‘) only’’ or the statement ‘‘Caution: Federal law re- stricts this device to sale by or on the order of a lll’’, the blank to be filled with the word ‘‘physician’’, ‘‘dentist’’, ‘‘veterinarian’’, or with the descriptive designation of any other practitioner licensed by the law of the State in which the practitioner practices to use or order the use of the device.

        - For a reagent, appropriate storage instructions adequate to protect the stability of the product. When applica- ble, these instructions shall include such information as conditions of tem- perature, light, humidity, and other pertinent factors. For products requir- ing manipulation, such as reconstitu- tion and/or mixing before use, appro- priate storage instructions shall be provided for the reconstituted or mixed product which is to be stored in the original container. The basis for such instructions shall be determined by re- liable, meaningful, and specific test methods  such  as  those  described  in

§ 211.166 of this chapter.

        - For a reagent, a means by which the user may be assured that the prod- uct meets appropriate standards of identity, strength, quality and  purity at the time of use. This shall be pro- vided, both for the product as provided and for any resultant reconstituted or mixed product, by including on the label one or more of the following:

          - An expiration date based upon the stated storage instructions.

          - A statement of  an  observable  in- dication of an alteration of the prod- uct, e.g., turbidity, color change, pre- cipitate, beyond its appropriate stand- ards.

          - Instructions for a simple method by which the user can reasonably de- termine that the product meets its ap- propriate standards.

        - For a reagent, a declaration of the net quantity of contents, expressed in terms of weight or volume, numerical count, or any combination of these or other terms which accurately reflect the contents of the package. The use of metric designations is encouraged, wherever appropriate. If more than a

#### Food and Drug Administration, HHS	§ 809.10

single determination may be performed using the product, any statement of the number of tests shall be consistent with instructions for use and amount of material provided.

        - Name and place of business of manufacturer,  packer,  or  distributor.

        - A lot or control number, identi- fied as such, from which it is possible to determine the complete manufac- turing history of the product.

          - If it is a multiple unit product, the lot or control number shall permit tracing the identity of the individual units.

          - For an instrument, the lot or con- trol number shall permit tracing the identity of all functional subassem- blies.

          - For multiple unit products which require the use of included units to- gether as a system, all units should bear the same lot or control number, if appropriate, or other suitable uniform identification should be used.

        - Except that for items in para- graphs (a) (1) through (9) of this sec- tion: (i) In the case of immediate con- tainers too small or  otherwise  unable to accommodate a label with sufficient space to bear all such information and which are packaged within an outer container from which they are removed for use, the information required by paragraphs (a) (2), (3), (4), (5), (6) (ii),

(iii) and (7) of this section may appear in the outer container labeling only.

(ii) In any case in which the presence of this information on the immediate container will interfere with the test, the information may appear on the outside container or wrapper rather than on the immediate container label.

      - Labeling accompanying each product, e.g., a package  insert,  shall state in one place the following infor- mation in the format and order speci- fied below, except where such informa- tion is not applicable, or as specified in a standard for a particular  product class. The labeling for a multiple-pur- pose instrument used for diagnostic purposes, and not committed to spe- cific diagnostic procedures or systems, may bear  only  the  information indi- cated in paragraphs (b) (1), (2), (6), (14), and (15) of this section. The labeling for a reagent intended for use as a replace- ment  in  a  diagnostic  system  may  be

limited to that information necessary to identify the reagent adequately and to describe its proper use in the sys- tem.

The proprietary name and estab- lished name, i.e., common or usual name, if any.

The intended use or uses of the product and the type of procedure, e.g., qualitative  or  quantitative.

Summary and explanation of the test. Include a short history of the methodology, with pertinent references and a balanced statement of the special merits and limitations of this method or product. If the product labeling re- fers to any other procedure, appro- priate literature  citations  shall be  in- cluded and the labeling shall explain the nature of any differences from the original and their effect on the results.

The chemical, physical, physio- logical, or biological principles of the procedure. Explain concisely, with chemical reactions and techniques in- volved, if applicable.

Reagents:

  - A declaration of the established name (common or usual name), if any, and quantity, proportion or concentra- tion or each reactive ingredient; and for biological material, the source and a measure of its activity. The quantity, proportion, concentration or activity shall be stated in the system generally used and  recognized by  the intended user, e.g., metric, international units, etc. A statement indicating the pres- ence of and characterizing any cata- lytic or nonreactive ingredients, e.g., buffers, preservatives, stabilizers.

  - A statement of warnings or pre- cautions for users as established in the regulations contained in 16 CFR part 1500 and any other warnings appro- priate to the hazard presented by the product; and a statement ‘‘For In Vitro Diagnostic Use’’ and any other limiting statements appropriate to the intended use of the product. The limiting state- ment appropriate to the intended use of a prescription in vitro diagnostic product shall bear the symbol state- ment ‘‘Rx only’’ or ‘‘) only’’ or the statement ‘‘Caution: Federal law re- stricts this device to sale by or on the order of a lll’’, the blank to be filled with the word ‘‘physician’’, ‘‘dentist’’, ‘‘veterinarian’’, or with the descriptive

#### § 809.10

designation of any other practitioner licensed by the law of the State in which the practitioner practices to use or order the use of the device.

  - Adequate instructions for recon- stitution, mixing, dilution, etc.

  - Appropriate storage instructions adequate to protect the stability of the product. When applicable, these in- structions shall include such informa- tion as  conditions  of  temperature, light, humidity, and other pertinent factors. For products requiring manip- ulation, such as reconstitution and/or mixing before  use,  appropriate  storage instructions shall be provided for the reconstituted or mixed product. The basis for such instructions shall be de- termined by reliable, meaningful, and specific test methods such as those de- scribed in § 211.166 of this chapter.

  - A statement of any purification or treatment required for use.

  - Physical, biological, or chemical indications of instability or deteriora- tion.

Instruments:

  - Use or function.

  - Installation procedures and spe- cial  requirements.

  - Principles of operation.

  - Performance characteristics and specifications.

  - Operating instructions.

  - Calibration procedures including materials and/or equipment to be used.

  - Operational precautions and limitations.

  - Hazards.

  - Service and maintenance infor- mation.

Specimen collection and prepara- tion for analysis, including a descrip- tion of:

Special precautions regarding specimen collection including special preparation of the patient as it bears on the validity of the test.

Additives, preservatives, etc., necessary to maintain the integrity of the specimen.

Known interfering substances.

Recommended storage, handling or shipping instructions for the protec- tion and maintenance of stability of the specimen.

Procedure: A step-by-step outline of recommended procedures from re- ception  of  the  specimen  to  obtaining

#### 21 CFR Ch. I (4–1–17 Edition)

results. List any points that may be useful in improving precision and accu- racy.

A list of all materials provided, e.g., reagents, instruments and equip- ment, with instructions for their use.

A list of all materials required but not provided. Include such details as sizes, numbers, types, and quality.

A description of the amounts of reagents necessary, times required for specific steps, proper temperatures, wavelengths, etc.

A statement describing the sta- bility of the final reaction material to be measured and the time within which it shall be measured to assure accurate results.

Details of calibration: Identify reference material. Describe prepara- tion of reference sample(s), use of blanks, preparation of the standard curve, etc. The description of the range of calibration should include the high- est and the lowest values  measurable by the procedure.

Details of kinds of quality con- trol procedures and materials required. If there is need for both positive and negative controls, this should be stat- ed. State what are considered satisfac- tory limits of performance.

Results: Explain the procedure for calculating the value of the unknown. Give an explanation for each compo- nent of the formula used for the cal- culation of the unknown. Include a sample calculation,  step-by-step,  ex- plaining the answer. The values  shall be expressed to the appropriate number of significant figures. If the test pro- vides other than quantitative results, provide an adequate description of ex- pected results.

Limitation of the procedure: In- clude a statement of limitations of the procedure. State known extrinsic fac- tors or interfering substances affecting results. If further testing, either more specific or more sensitive, is indicated in all cases where certain results are obtained, the need for the additional test shall be stated.

Expected values: State the range(s) of expected values as obtained with the product from studies of var- ious populations. Indicate how the range(s) was established and identify

#### Food and Drug Administration, HHS	§ 809.10

the population(s) on which it was es- tablished.

Specific performance characteris- tics: Include, as appropriate, informa- tion describing such things as accu- racy, precision, specificity, and sensi- tivity. These shall be related to a gen- erally accepted method using biologi- cal specimens from normal and abnor- mal populations. Include a statement summarizing the data upon which the specific performance characteristics are based.

Bibliography: Include pertinent references keyed to the text.

Name and place of business of manufacturer,  packer,  or  distributor.

Date of issuance of the last revi- sion of the labeling identified as such.

      - A shipment or other delivery of an in vitro diagnostic product shall be ex- empt from the requirements of para- graphs (a) and (b) of this section and from a standard promulgated  under part 861 provided that the following conditions are met:

        - In the case of a shipment or deliv- ery for an investigation subject to part 812, if there has been compliance with part 812; or

        - In the case of a shipment or deliv- ery for an investigation that is not subject to part 812 (see § 812.2(c)), if the following conditions are met:

          - For a product in the laboratory re- search phase of development, and not represented as an effective in vitro di- agnostic product, all labeling bears the statement, prominently placed: ‘‘For Research Use Only. Not for use in diag- nostic procedures.’’

          - For a product being shipped or de- livered for product testing prior to full commercial marketing (for example, for use on specimens derived from hu- mans to compare the usefulness of the product with other products or proce- dures which are in current use or rec- ognized as useful), all labeling bears the statement, prominently placed: ‘‘For Investigational Use Only. The performance characteristics of this product have not been established.’’

      - The labeling of general purpose laboratory reagents (e.g., hydrochloric acid) and equipment (e.g., test tubes and pipettes) whose uses are generally known by persons trained in their use need not bear the directions for use re-

quired by § 809.10(a) and (b), if their la- beling meets the requirements of this paragraph.

The label of a reagent shall bear the following information:

  - The proprietary name and estab- lished name (common or usual name), if any, of the reagent.

  - A declaration of the established name (common or usual name), if any, and quantity, proportion or concentra- tion of the reagent ingredient (e.g., hy- drochloric acid: Formula weight 36.46, assay 37.9 percent, specific gravity 1.192 at 60 F);  and  for  a  reagent  derived from  biological  material,   the   source and where applicable a measure of its activity. The quantity, proportion, concentration or activity shall  be  stat- ed in the system generally used and recognized by the intended user, e.g., metric,   international   units,   etc.

  - A statement of the purity and quality of the reagent, including a quantitative declaration of any impuri- ties present. The requirement for this information may be met by a state- ment of conformity with a generally recognized and generally available standard which contains the same in- formation, e.g., those established by the American Chemical Society, U.S. Pharmacopeia, National Formulary, National Research Council.

  - A statement of warnings or pre- cautions for users as established in the regulations contained in 16 CFR part 1500 and any other warnings appro- priate to the hazard presented by the product; and a statement ‘‘For Labora- tory  Use.’’

  - Appropriate storage  instructions adequate to protect the stability of the product. When applicable, these in- structions shall include such informa- tion as  conditions  of  temperature, light, humidity, and other pertinent factors. The basis for such information shall be determined by reliable, mean- ingful, and specific test  methods  such as those described in § 211.166 of this chapter.

  - A declaration of the net quantity of contents, expressed in terms of weight or volume, numerical count, or any combination of these or other terms which accurately reflect the con- tents of the package. The use of metric

#### § 809.10

designations   is   encouraged,   wherever appropriate.

  - Name and place of business of manufacturer,  packer,  or  distributor.

  - A lot or control number, identi- fied as such, from which it is possible to determine the complete manufac- turing history of the product.

  - In the case of immediate con- tainers too small or  otherwise  unable to accommodate a label with sufficient space to bear all such information, and which are packaged within an outer container from which they are removed for use, the information required by paragraphs (d)(1)(ii), (iii), (iv), (v), and

(vi) of this section may appear in the outer container labeling only.

The label of general purpose lab- oratory equipment, e.g., a beaker or a pipette, shall bear a statement ade- quately describing the product, its composition, and physical characteris- tics if necessary for its proper use.

(e)(1) The labeling for analyte spe- cific reagents (e.g., monoclonal anti- bodies, deoxyribonucleic acid (DNA) probes, viral antigens, ligands) shall bear the following information:

  - The proprietary name and estab- lished name (common or usual name), if any, of the reagent;

  - A declaration of the established name (common or usual name), if any;

  - The quantity, proportion, or con- centration of the reagent ingredient; and for a reagent derived from biologi- cal material, the source and where ap- plicable, a measure of its activity. The quantity, proportion, concentration, or activity shall  be  stated  in the  system generally used and recognized by the intended user, e.g., metric, inter- national units, etc.;

  - A statement of the purity and quality of the reagent, including a quantitative declaration of any impuri- ties present and method of analysis or characterization. The requirement for this information may be met by a statement of conformity with a gen- erally recognized and generally avail- able standard that contains the same information, e.g., those established by the American Chemical Society, U.S. Pharmacopeia,   National   Formulary, and National Research Council. The la- beling may also include information concerning    chemical/molecular    com-

#### 21 CFR Ch. I (4–1–17 Edition)

position, nucleic acid sequence, binding affinity, cross-reactivities, and inter- action with substances of known clin- ical significance;

  - A statement of warnings or pre- cautions for users as established in the regulations contained in 16 CFR part 1500 and any other warnings appro- priate to the hazard presented by the product;

  - The date of manufacture and ap- propriate storage instructions ade- quate to protect the stability of the product. When applicable, these in- structions shall include such informa- tion as  conditions  of  temperature, light, humidity, date of expiration, and other pertinent factors. The basis for such instructions shall be determined by reliable, meaningful,  and  specific test methods, such as  those  described in § 211.166 of this chapter;

  - A declaration of the net quan- tity of contents, expressed in terms of weight or volume, numerical count, or any combination of these or other terms that accurately reflect the con- tents of the package. The use of metric designations is encouraged, wherever appropriate;

  - The name and place of business of manufacturer, packer, or dis- tributor;

  - A lot or control number, identi- fied as such, from which it is possible to determine the complete manufac- turing history of the product;

  - For class I exempt ASR’s, the statement: ‘‘Analyte Specific Reagent. Analytical and performance character- istics are not established’’; and

  - For class II and III ASR’s, the statement: ‘‘Analyte Specific Reagent. Except as a component of the approved/ cleared test (Name of approved/cleared test), analytical and performance char- acteristics of this ASR are not estab- lished.’’

(2) In the case of immediate con- tainers too small or  otherwise  unable to accommodate a label with sufficient space to bear all such information, and which are packaged within an outer container from which they are removed for use, the information required by paragraphs (e)(1) through (e)(6) of this section may appear in the outer con- tainer  labeling  only.

#### Food and Drug Administration, HHS	§ 809.10

The labeling for over-the-counter (OTC) test sample collection systems for drugs of abuse testing shall  bear the following information in language appropriate for the intended users:

  - Adequate instructions for speci- men collection and handling, and for preparation and mailing of the speci- men to the laboratory for testing.

  - An identification system to en- sure that specimens are not mixed up or otherwise misidentified at the lab- oratory, and that user anonymity is maintained.

  - The intended use or uses of the product, including what drugs are to be identified in the specimen, a quan- titative description of the performance characteristics for those drugs (e.g., sensitivity and specificity) in terms understandable to lay users, and the detection period.

  - A statement that confirmatory testing will be conducted on all sam- ples that initially test positive.

  - A statement of warnings or pre- cautions for users as established in the regulations contained in 16 CFR part 1500 and any other warnings appro- priate to the hazard presented by the product.

  - Adequate instructions on how to obtain test results from a person who can explain their meaning, including the probability of false positive and false negative results, as well as how to contact a trained health professional if additional information on interpreta- tion of test results from the laboratory or followup counseling is desired.

  - Name and place of business of the manufacturer, packer, or distributor.

(g)(1) The applicant may provide the labeling information referenced in this section in the form of:

    - A symbol accompanied by explan- atory text adjacent to the symbol;

    - A symbol not accompanied by ad- jacent explanatory text that:

      - Is contained in a standard that FDA recognizes under its authority in section 514(c) of the act;

      - Is used according to the specifica- tions for use of the symbol set forth in FDA’s section 514(c) recognition; and

      - Is explained in a paper or elec- tronic symbols glossary that is in- cluded in the labeling for the device and the labeling on or within the pack-

age containing the device bears a prominent and conspicuous statement identifying the location of the symbols glossary that is written in English or, in the case of articles distributed sole- ly in Puerto Rico or in a Territory where the predominant language is one other than English, the predominant language may be used;

    - A symbol not accompanied by adjacent explanatory text that:

Is established in a standard devel- oped by a standards development orga- nization  (SDO);

Is not contained in a standard that is recognized by FDA under its au- thority in section 514(c) of the act or is contained in a standard that is recog- nized by FDA but is not used according to the specifications for use of the sym- bol set forth in FDA’s section 514(c) recognition;

Is determined by the manufac- turer to be likely to be read and under- stood by the ordinary individual under customary conditions of purchase and use in compliance with section 502(c) of the  act;

Is used according to the specifica- tions for use of the symbol set forth in the SDO-developed standard; and

Is explained in a paper or elec- tronic symbols glossary that is in- cluded in the labeling for the device and the labeling on or within the pack- age containing the device bears a prominent and conspicuous statement identifying the location of the symbols glossary that is written in English or, in the case of articles distributed sole- ly in Puerto Rico or in a Territory where the predominant language is one other than English, the predominant language may be used; or

    - The symbol statement ‘‘Rx only’’ or ‘‘) only’’ used as provided under paragraphs (a)(4) and (b)(5)(ii) of this section.

The use of symbols in device la- beling which  do not  meet the  require- ments of paragraph (g)(1) of this sec- tion renders a device misbranded under section 502(c) of the act.

For purposes of paragraph (g)(1) of this  section:

  - An SDO is an organization that is nationally or internationally recog- nized  and  that  follows  a  process  for

#### § 809.11

standard development that is trans- parent, (i.e., open to public scrutiny), where the participation is balanced, where an appeals process is included, where the standard is not in conflict with any statute, regulation, or policy under which FDA operates, and where the standard is national or inter- national in scope.

  - The term ‘‘symbols glossary’’ means a compiled listing of:

    - Each SDO-established symbol used in the labeling for the device;

    - The title and designation number of the SDO-developed standard con- taining the symbol;

    - The title of the symbol and its reference number, if any, in the stand- ard; and

    - The meaning or explanatory text for the symbol as provided in the FDA recognition or, if FDA has not recog- nized the standard or portion of the standard in which the symbol is lo- cated or the symbol is not used accord- ing to the specifications for use of the symbol set forth  in  FDA’s  section 514(c) recognition, the explanatory text as provided in the standard.

[41  FR  6903,  Feb.  13,  1976,  as  amended  at  45

FR 3750, Jan. 18, 1980; 45 FR 7484, Feb. 1, 1980;

47 FR 41107, Sept. 17, 1982; 47 FR 51109, Nov.

12, 1982; 48 FR 34470, July 29, 1983; 62 FR 62259,

Nov. 21, 1997; 65 FR 18234, Apr. 7, 2000; 81 FR

38931, June 15, 2016]

##### § 809.11 Exceptions or alternatives to labeling requirements for in vitro diagnostic products for human use held by the Strategic National Stockpile.

      - The appropriate FDA Center Di- rector may grant an exception or alter- native to any provision listed in para- graph (f) of this section and not explic- itly required by statute, for specified lots, batches, or other units of an in vitro diagnostic product for  human use, if the Center Director determines that compliance with such labeling re- quirement could adversely affect the safety, effectiveness, or availability of such products  that are  or will  be  in- cluded in the Strategic National Stockpile.

(b)(1)(i) A Strategic National Stock- pile official or any entity that manu- factures (including labeling, packing, relabeling, or repackaging), distrib- utes,  or  stores  an  in  vitro  diagnostic

#### 21 CFR Ch. I (4–1–17 Edition)

product for human use that is or will be included in the Strategic National Stockpile may submit, with written concurrence from a Strategic National Stockpile official, a written request for an exception or alternative  described in paragraph (a) of this section to the Center Director.

(ii) The Center Director may  grant an exception or alternative  described in paragraph  (a) of this  section on his or her own initiative.

(2) A written request for an exception or  alternative  described  in  paragraph

of this section must:

  - Identify the specified lots, batches, or other units of an in vitro diagnostic product for human use that would be subject to the exception or al- ternative;

  - Identify the labeling provision(s) listed in paragraph (f) of this section that are the subject of the exception or alternative request;

  - Explain why compliance with such labeling provision(s) could ad- versely affect the safety, effectiveness, or availability of the specified lots, batches, or other units of the in vitro diagnostic product for human use that are or will be held in the Strategic Na- tional  Stockpile;

  - Describe any proposed safeguards or conditions that will be implemented so that the labeling of the product in- cludes appropriate information nec- essary for the safe and effective use of the product, given the anticipated cir- cumstances of use of the product;

  - Provide a draft of the proposed la- beling of the specified lots, batches, or other units of the in vitro diagnostic products for human use subject to the exception or alternative; and

  - Provide any other information requested by the Center Director in support of the request.

The Center Director must respond in writing to all requests under this section. The Center Director may im- pose appropriate conditions or safe- guards when granting such an excep- tion or alternative under this section.

A grant of an exception or alter- native under this section will include any safeguards or conditions deemed appropriate by the Center Director to ensure that the labeling of the product subject to the exception or alternative

#### Food and Drug Administration, HHS	§ 809.30

includes the information necessary for the safe and effective use of the prod- uct, given the anticipated cir- cumstances of use.

If the Center Director grants a re- quest for an exception or alternative to the labeling requirements under this section:

  - The Center Director may deter- mine that the submission and grant of a written request under this section satisfies the provisions relating to pre- market  notification  submissions  under

§ 807.81(a)(3) of this chapter.

(2)(i) For a Premarket Approval Ap- plication (PMA)-approved in vitro diag- nostic product for human use, the sub- mission and grant of a written request under this section satisfies the  provi- sions relating to submission of PMA supplements under § 814.39 of this chap- ter;  however,

(ii) The grant of the request must be identified  in  a  periodic  report  under

§ 814.84 of this chapter.

The Center Director may grant an exception or alternative under this sec- tion to the following provisions of this part, to the extent that the require- ments in these provisions are not ex- plicitly required by statute:

(1) § 809.10(a)(1) through (a)(6) and (a)(9);

(2) § 809.10(b);

(3) § 809.10(c)(2);

(4)    § 809.10(d)(1)(i)    through    (d)(1)(v),

(d)(1)(viii), and (d)(2); and

(5)    § 809.10(e)(1)(i)    through    (e)(1)(vi)

and (e)(1)(ix) through (e)(1)(xi).

[72 FR 73601, Dec. 28, 2007]

## Subpart C—Requirements for Manufacturers and Producers

##### § 809.20 General requirements for manufacturers and producers of in vitro diagnostic products.

[Reserved]

Compliance with good manufac- turing practices. In vitro diagnostic products shall be manufactured in ac- cordance with the good manufacturing practices requirements found in  part 820 of this chapter and, if applicable, with § 610.44 of this chapter.

[41  FR  6903,  Feb.  13,  1976,  as  amended  at  42

FR 42530, Aug. 23, 1977; 43 FR 31527, July 21,

1978; 66 FR 31165, June 11, 2001]

##### § 809.30 Restrictions on the sale, dis- tribution and use of analyte specific reagents.

Analyte specific reagents (ASR’s) (§ 864.4020 of this chapter) are restricted devices under section 520(e) of the Fed- eral Food, Drugs,  and  Cosmetic  Act (the act) subject to the restrictions set forth in this section.

ASR’s may only be sold to:

  - In vitro diagnostic manufacturers;

  - Clinical laboratories regulated under the Clinical Laboratory Improve- ment Amendments of 1988 (CLIA), as qualified to perform high  complexity testing under 42 CFR part 493 or clin- ical laboratories regulated under VHA Directive 1106 (available from Depart- ment of Veterans Affairs, Veterans Health     Administration,     Washington,

DC 20420); and

  - Organizations that use the re- agents to make tests for purposes other than providing diagnostic information to patients and practitioners, e.g., fo- rensic, academic, research, and other nonclinical laboratories.

ASR’s must be labeled in accord- ance with § 809.10(e).

Advertising and promotional ma- terials for ASR’s:

  - Shall include the identity and pu- rity (including source and method of acquisition) of the analyte specific rea- gent and the identity of the analyte;

  - Shall include the statement for class I exempt ASR’s: ‘‘Analyte Spe- cific Reagent. Analytical and perform- ance characteristics are not estab- lished’’;

  - Shall include the statement for class II or III ASR’s: ‘‘Analyte Specific Reagent. Except as a component of the approved/cleared test (name of ap- proved/cleared test), analytical and performance characteristics are not es- tablished’’; and

  - Shall not make any statement re- garding analytical or clinical perform- ance.

The laboratory that develops an in-house test using the ASR shall in- form the ordering person of the test re- sult by appending to the test report the statement: ‘‘This test was developed and its performance characteristics de- termined by (Laboratory Name). It has not been cleared  or approved by  the

U.S. Food and Drug Administration.’’

§ 809.40

This statement would not be applicable or required when test results are gen- erated using the test that was cleared or approved in conjunction with review of the class II or III ASR.

Ordering in-house tests that are developed using analyte specific re- agents is limited under section 520(e) of the act to physicians and other persons authorized by applicable State law to order such tests.

The restrictions in paragraphs (c) through (f) of this section do not apply when reagents that otherwise meet the analyte specific reagent definition are sold to:

  - In vitro diagnostic manufacturers; or

  - Organizations that use the re- agents to make tests for purposes other than providing diagnostic information to patients and practitioners, e.g., fo- rensic, academic, research, and other nonclinical laboratories.

[62 FR 62259, Nov. 21, 1997]

##### § 809.40 Restrictions on the sale, dis- tribution, and use of OTC test sam- ple collection systems for drugs of abuse testing.

    - Over-the-counter (OTC) test sam- ple collection  systems  for  drugs  of abuse testing (§ 864.3260 of this chapter) are restricted devices under section 520(e) of the Act subject to the restric- tions  set  forth  in  this  section.

    - Sample testing shall be performed in a laboratory using screening tests that have been approved, cleared, or otherwise recognized by the Food and Drug Administration as accurate and reliable for the testing of such speci- mens for identifying drugs of abuse or their metabolites.

    - The laboratory performing the test(s) shall have, and shall be recog- nized as having, adequate capability to reliably perform the necessary screen- ing and confirmatory tests, including adequate capability to perform integ- rity checks of the biological specimens for possible adulteration.

    - All OTC test sample collection systems for drugs of abuse testing shall be labeled in accordance with § 809.10(f) and shall provide an  adequate  system to  communicate  the  proper  interpreta-

#### 21 CFR Ch. I (4–1–17 Edition)

tion of test results from the laboratory to the lay purchaser.

[65 FR 18234, Apr. 7, 2000]

## PART 810—MEDICAL DEVICE RECALL AUTHORITY

##### Subpart A—General Provisions

Sec.

  - Scope.

  - Definitions.

  - Computation of time.

  - Service of orders.

##### Subpart B—Mandatory Medical Device Recall Procedures

  - Cease   distribution   and   notification order.

  - Regulatory hearing.

  - Written request for review of cease distribution and notification order.

  - Mandatory recall order.

  - Cease distribution and notification or mandatory recall strategy.

  - Communications concerning a cease distribution and notification or manda- tory recall order.

  - Cease distribution and notification or mandatory recall order status reports.

  - Termination of a cease distribution and notification or mandatory recall order.

  - Public notice.

AUTHORITY: 21  U.S.C.  321,  331,  332,  333,  334,

351, 352, 355, 360h, 360i, 371, 374, 375.

SOURCE:  61  FR  59018,  Nov.  20,  1996,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 810.1   Scope.

Part 810 describes the  procedures that the Food  and  Drug  Administra- tion will follow in exercising its med- ical device recall authority under sec- tion 518(e) of the Federal Food, Drug, and  Cosmetic  Act.

##### § 810.2   Definitions.

As used in this part:

Act means the Federal Food, Drug, and Cosmetic Act.

Agency or FDA means the Food and Drug Administration.

Cease distribution and notification strategy or mandatory recall strategy means a planned, specific course of ac- tion to be taken by the person named in a cease distribution and notification

Food and Drug Administration, HHS	§ 810.10

order or in a mandatory recall order, which addresses the extent of the noti- fication or recall, the need for public warnings, and the extent of effective- ness checks to be conducted.

Consignee means any person or firm that has received, purchased, or used a device that is subject to a cease distribution and notification order or a mandatory recall order. Consignee does not mean lay individuals or patients, i.e.,  nonhealth  professionals.

Correction means repair, modifica- tion, adjustment, relabeling, destruc- tion, or inspection (including patient monitoring) of a device, without its physical removal from its point of use to some other location.

Device user facility means a hos- pital, ambulatory surgical facility, nursing home, or outpatient treatment or diagnostic facility that is not a phy- sician’s office.

Health professionals means practi- tioners, including physicians, nurses, pharmacists, dentists, respiratory therapists, physical therapists, tech- nologists, or any other practitioners or allied health professionals that have a role in using a device for human use.

Human cell, tissue, or cellular or tis- sue-based product (HCT/P) regulated as a device  means  an  HCT/P  as  defined  in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

Reasonable probability means that it is more likely than not that an event will occur.

Serious, adverse health consequence means any significant adverse experi- ence, including those that may be ei- ther life-threatening or involve perma- nent or long-term injuries, but exclud- ing injuries that are nonlife-threat- ening and that are temporary and rea- sonably reversible.

Recall means the correction or re- moval of a device for human use where FDA finds that there is a reasonable probability that the device would cause serious, adverse health consequences or death.

Removal means the physical re- moval of a device from its point of use to some other location for repair, modification, adjustment, relabeling, destruction, or inspection.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A unique  device  identifier  is  composed  of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

[61 FR 59018, Nov. 20, 1996, as amended at 78

FR 55821, Sept. 24, 2013]

##### § 810.3   Computation of time.

In computing any period of time pre- scribed or allowed by this part, the day of the act or event from which the des- ignated period of time begins to run shall not be included. The computation of time is based only on working days.

##### § 810.4   Service of orders.

Orders issued under this part will be served in person by a designated em- ployee of FDA, or by certified or reg- istered mail or similar mail delivery service with a return receipt record re- flecting receipt, to the named person or designated agent at the named person’s or designated agent’s last known ad- dress in FDA’s records.

## Subpart B—Mandatory Medical Device Recall Procedures

##### § 810.10   Cease  distribution  and  notifi- cation order.

If, after providing the appropriate person with an opportunity to consult with the agency, FDA finds that there is a reasonable probability that a de- vice  intended  for  human  use  would

§ 810.10

cause serious, adverse health con- sequences or death, the agency may issue a cease distribution and notifica- tion order requiring the person named in the order to immediately:

  - Cease distribution of the device;

  - Notify health professionals and device user facilities of the order; and

  - Instruct these professionals and device user facilities to cease use of the device.

FDA will include the following in- formation in the order:

  - The requirements of the order re- lating to cessation of distribution and notification of health professionals and device user facilities;

  - Pertinent descriptive information to enable accurate and immediate iden- tification of  the  device  subject  to the order, including, where known:

    - The brand name of the device;

    - The common name, classification name, or usual name of the device;

    - The model, catalog, or product code numbers of the device;

    - The manufacturing lot numbers or serial numbers of the device or other identification numbers; and

    - The unique device identifier (UDI) that appears on the device label or on the device package; and

  - A statement of the grounds for FDA’s finding that there is a reason- able probability that the device would cause serious, adverse health con- sequences or death.

FDA may also include in the order a model letter for notifying health pro- fessionals and device user facilities of the order and a requirement that noti- fication of health professionals and de- vice user facilities be completed within a specified timeframe. The model letter will include the key elements of infor- mation that the agency in its discre- tion has determined, based on the cir- cumstances  surrounding  the  issuance of each order, are necessary to inform health professionals and device user fa- cilities about the order.

FDA may also require that the person named in the cease distribution and notification order submit any  or all of the following information to the agency by a time specified in the order:

  - The total number of units of the device produced and the timespan of the production;

#### 21 CFR Ch. I (4–1–17 Edition)

  - The total number of units of the device estimated to be in distribution channels;

  - The total number of units of the device estimated to be distributed to health professionals and device user fa- cilities;

  - The total number of units of the device estimated to be in the hands of home users;

  - Distribution information, includ- ing the names and addresses of all con- signees;

  - A copy of any written communica- tion used by the person named in the order to notify health professionals and device user facilities;

  - A proposed strategy for complying with the cease distribution and notifi- cation order;

  - Progress reports to be made at specified intervals, showing the names and addresses of health professionals and device user facilities that  have been notified, names of specific individ- uals contacted within device user fa- cilities, and the dates of such contacts; and

  - The name, address, and telephone number of the person who should be contacted concerning  implementation of the order.

FDA will provide the person named in a cease distribution and noti- fication order with an opportunity for a regulatory hearing on the actions re- quired by the cease distribution and notification order and on whether the order should be modified, or vacated, or amended to require a mandatory recall of the device.

FDA will also provide the person named in the cease distribution and no- tification order with an opportunity, in lieu of a regulatory hearing, to submit a written request to FDA asking that the order be modified, or vacated, or amended.

FDA will include in the cease dis- tribution and notification order the name, address, and telephone number of an agency employee to whom any re- quest for a regulatory hearing or agen- cy review is to be addressed.

[61 FR 59018, Nov. 20, 1996, as amended at 78

FR 55821, Sept. 24, 2013]

#### Food and Drug Administration, HHS	§ 810.12

##### § 810.11   Regulatory hearing.

Any request for a regulatory hear- ing shall be submitted in writing to the agency employee identified  in  the order within the timeframe specified by FDA. Under § 16.22(b) of this chapter, this timeframe ordinarily will not be fewer than 3 working days after receipt of the cease distribution and notifica- tion   order.   However,   as   provided   in

§ 16.60(h) of this chapter, the Commis- sioner of Food and Drugs or presiding officer may waive, suspend, or modify any provision of part 16 under § 10.19 of this chapter,  including  those  per- taining  to  the  timing  of   the   hearing. As provided in § 16.26(a), the Commis- sioner or presiding officer may deny a request for a hearing, in whole or in part, if he or she determines that no genuine and substantial issue of fact is raised by the material submitted in the request.

If a request for a regulatory hear- ing is granted, the regulatory hearing shall be limited to:

  - Reviewing the actions required by the cease distribution and notification order, determining if FDA should af- firm, modify, or vacate the order, and addressing an appropriate cease dis- tribution and notification strategy; and

  - Determining whether FDA should amend the cease distribution and noti- fication order to require a recall of the device that was the subject  of the order. The hearing may also address the actions that might be required by a recall order, including an appropriate recall strategy, if FDA later orders a recall.

If a request by  the  person  named in a cease distribution and notification order for a regulatory hearing is  grant- ed, the regulatory hearing will be con- ducted in accordance with the proce- dures set out in section 201(x)  of  the act (21 U.S.C. 321(x)) and part 16 of this chapter, except that the order issued under § 810.10, rather than a notice under § 16.22(a) of this chapter, provides the notice of opportunity for a hearing and is part of  the  administrative  record of      the      regulatory      hearing      under

§ 16.80(a) of this chapter. As provided in

§ 16.60(h) of this chapter, the Commis- sioner of  Food and  Drugs or presiding officer  may  waive,  suspend,  or  modify

any provision of part 16 under § 10.19 of this chapter. As provided in § 16.26(b), after the hearing commences, the pre- siding officer may issue a summary de- cision on any issue if the presiding offi- cer determines that there is no genuine and substantial issue of fact respecting that  issue.

If the person named in the cease distribution and notification order does not request a regulatory hearing with- in the timeframe specified by FDA in the cease distribution and notification order, that person will be deemed to have waived his or her right to request a hearing.

The presiding officer will ordi- narily hold any regulatory hearing re- quested under paragraph (a) of this sec- tion no fewer than 2 working days after receipt of the request for a  hearing, under § 16.24(e) of this chapter, and no later than  10  working  days  after  the date of issuance of the cease distribu- tion and notification order.  However, FDA and the person named in the order may agree to a later date or the pre- siding officer may determine that the hearing should be held in fewer than 2 days.    Moreover,    as    provided    for    in

§ 16.60(h) of this chapter, the Commis- sioner of Food and Drugs or presiding officer may waive, suspend, or modify any provision of part 16 under § 10.19 of this chapter,  including  those  per- taining to the timing  of  the  hearing. After the presiding officer prepares a written report of the hearing and the agency issues a final decision based  on the report, the presiding officer shall provide the requestor written notifica- tion of the final decision to affirm, modify,  or  vacate  the  order  or  to amend the order to require a  recall  of the device within 15 working days of conducting  a  regulatory  hearing.

##### § 810.12 Written request for review of cease distribution and notification order.

In lieu  of requesting a  regulatory hearing under § 810.11, the person named in a cease distribution and noti- fication order may submit a written re- quest to FDA asking that the  order be modified or vacated. Such person shall address the written request to the agency employee identified  in  the order   and   shall   submit   the   request

#### § 810.13

within the timeframe specified in the order, unless FDA and the person named in the order agree to a later date.

A written request for review of a cease distribution and notification order shall identify each ground upon which the requestor relies in asking that the order be modified or vacated, as well as addressing an appropriate cease distribution and notification strategy, and shall address whether the order should be amended to require a recall of the device that was the sub- ject of the order and the actions re- quired by such a recall order, including an appropriate recall strategy.

The agency official who issued the cease distribution  and  notification order shall provide the requestor writ- ten notification of the agency’s deci- sion to affirm, modify, or vacate the order or amend the order to require a recall of  the device  within 15  working days of receipt of the written request. The agency official  shall  include  in this  written  notification:

  - A statement of the grounds for the decision to affirm, modify, vacate, or amend the order; and

  - The requirements of any modified or amended order.

##### § 810.13   Mandatory recall order.

    - If the person named in a cease dis- tribution and notification  order  does not request a regulatory hearing or submit a request for agency review  of the order, or, if the Commissioner of Food and Drugs or the presiding officer denies a request for a hearing, or, if after conducting a regulatory hearing under § 810.11 or completing agency re- view of a cease distribution and notifi- cation order under § 810.12, FDA deter- mines that the order should be amend- ed to require a recall of the device with respect to which the order was issued, FDA shall amend the order to require such a recall. FDA shall  amend  the order to require such a recall within 15 working days of issuance of a cease dis- tribution and notification order if a regulatory hearing or agency review of the order is not requested, or within 15 working days of denying a request for a hearing, or within 15 working days of completing  a  regulatory  hearing  under

§ 810.11, or within 15 working days of re-

#### 21 CFR Ch. I (4–1–17 Edition)

ceipt of  a  written  request  for  review  of a cease  distribution  and  notification order under § 810.12.

    - In a mandatory recall order, FDA may:

      - Specify that the recall is to ex- tend to the wholesale, retail, or user level;

      - Specify a timetable in accordance with which the recall is to begin and be completed;

      - Require the person named in the order to submit to the agency a pro- posed recall strategy, as described in

§ 810.14, and periodic reports describing the  progress  of  the   mandatory   recall, as described in § 810.16; and

      - Provide the person named in the order with a model recall notification letter that includes the  key  elements of information that FDA has deter- mined are necessary to inform health professionals and device user facilities.

    - FDA will not include in a manda- tory recall order a requirement for:

      - Recall of a device from individ- uals; or

      - Recall of a device from device user facilities, if FDA determines that the risk of recalling the device from the fa- cilities presents a greater health risk than the health risk of not  recalling the device from use, unless the device can be replaced immediately with an equivalent device.

    - FDA will include in a mandatory recall order provisions for  notification to individuals subject to the risks asso- ciated with use of the device. If a sig- nificant number of such individuals cannot be identified, FDA may notify such individuals under section 705(b) of the  act.

##### § 810.14   Cease  distribution  and  notifi- cation or mandatory recall strategy.

General. The person named in a cease distribution  and  notification order issued under § 810.10 shall comply with the order, which FDA will fashion as appropriate for the individual cir- cumstances of the case. The person named in a cease distribution and noti- fication order modified under § 810.11(e) or § 810.12(c) or a mandatory recall order issued under § 810.13 shall develop a    strategy    for    complying    with    the

#### Food and Drug Administration, HHS	§ 810.15

order that is appropriate for the indi- vidual circumstances and that takes into account the following factors:

  - The nature of the serious, adverse health consequences related to the de- vice;

  - The ease of identifying the device;

  - The extent to which the risk pre- sented by the device is obvious to a health professional or device user facil- ity; and

  - The extent to which the device is used by health professionals and device user  facilities.

Submission and review. (1) The per- son named  in  the  cease  distribution and  notification  order  modified  under

§ 810.11(e) or § 810.12(c) or mandatory re- call order shall submit a copy of the proposed strategy to the agency within the timeframe specified in the order.

(2) The agency will review the pro- posed strategy and make  any  changes to the strategy that it deems necessary within 7 working days of receipt of the proposed strategy. The person  named in the order shall act in accordance with a strategy determined by  FDA  to be appropriate.

Elements of the strategy. A pro- posed strategy shall meet all of the fol- lowing requirements:

(1)(i) The person named in the order shall specify the level in the chain of distribution to which the cease dis- tribution and notification order or mandatory recall order is to extend as follows:

Consumer or user level, e.g., health professionals, consignee, or de- vice user facility level, including any intermediate wholesale or retail  level; or

Retail level, to the level imme- diately preceding the consumer or user level, and including any intermediate level; or

Wholesale level.

    - The person named in the order shall not recall a device from individ- uals; and

    - The person named in the order shall not  recall  a  device  from  device user facilities if FDA notifies the per- son not to do so because of a risk deter- mination under § 810.13(c)(2).

The person named in a recall order shall ensure that the strategy provides for notice to individuals sub-

ject to the risks associated with use of the recalled device. The notice may be provided through the  individuals’ health professionals if FDA determines that such consultation is appropriate and would be the most effective meth- od of notifying patients.

Effectiveness checks by the person named in the order are required to verify that all health professionals, de- vice user facilities, consignees, and in- dividuals, as appropriate, have been no- tified of the cease distribution and no- tification order or mandatory recall order and of the need to take appro- priate action. The person named in the cease distribution  and  notification order or the mandatory recall order shall specify in the strategy the meth- od(s) to be used in addition to written communications as required by § 810.15, i.e., personal  visits,  telephone  calls,  or a combination thereof to contact all health professionals,  device user  facili- ties, consignees, and individuals, as ap- propriate. The agency may conduct ad- ditional audit checks where appro- priate.

##### § 810.15 Communications concerning a cease distribution and notification or mandatory recall order.

General. The person named in a cease distribution  and  notification order issued under § 810.10 or a manda- tory recall order issued under § 810.13 is responsible  for  promptly   notifying each health professional, device user facility, consignee, or individual, as ap- propriate, of the order. In accordance with § 810.10(c) or § 810.13(b)(4), FDA may provide the person named in the cease distribution and notification or mandatory recall order with a model letter for notifying each health profes- sional, device user facility,  consignee, or individual, as appropriate, of the order. However, if FDA does not pro- vide the person named in the cease dis- tribution and notification or  manda- tory recall order with a model  letter, the person named in a cease distribu- tion and notification order  issued under § 810.10, or a mandatory recall order issued under § 810.13, is respon- sible for providing  such  notification. The purpose of the  communication  is to convey:

#### § 810.16

  - That FDA has found that there is a reasonable probability that use of the device would cause a serious, adverse health consequence or death;

  - That the person named in the order has ceased distribution of the de- vice;

  - That health professionals and de- vice user facilities should cease use  of the device immediately;

  - Where appropriate, that the de- vice is subject to a mandatory recall order; and

  - Specific instructions on what should be done with the device.

Implementation. The person named in a cease distribution and notification order, or a mandatory recall order, shall notify the appropriate person(s) of the order by verified written com- munication, e.g., telegram, mailgram, or fax. The written communication and any envelope in which it is sent or en- closed shall be conspicuously marked, preferably in bold red ink: ‘‘URGENT— [DEVICE CEASE DISTRIBUTION AND NOTIFICATION ORDER] or [MANDA- TORY DEVICE RECALL ORDER].’’ Telephone calls or other personal con- tacts may be made in addition to, but not as a substitute for, the verified written communication, and shall be documented in an appropriate manner.

Contents. The person named in the order shall ensure that the notice of a cease distribution and notification order or mandatory recall order:

  - Is brief and to the point;

  - Identifies clearly the device, size, lot number(s), code(s), or serial num- ber(s), and any other pertinent descrip- tive information to facilitate accurate and immediate identification of the de- vice;

  - Explains concisely the serious, ad- verse health consequences that may occur if use of the device were contin- ued;

  - Provides specific instructions on what should be done with the device;

  - Provides a ready means for the re- cipient of the  communication to  con- firm receipt of the communication and to notify the person named in the order of the actions taken in response to the communication. Such means may in- clude, but are not limited to, the re- turn of a postage-paid, self-addressed

#### 21 CFR Ch. I (4–1–17 Edition)

post card or a toll-free call to the per- son named in the order; and

  - Does not contain irrelevant quali- fications, promotional materials, or any other statement that may detract from the message.

Followup communications. The per- son named in the cease distribution and notification order or mandatory recall order shall ensure that followup communications are sent to all who fail to respond to the initial commu- nication.

Responsibility of the recipient. Health professionals, device user facili- ties, and consignees who receive a com- munication concerning a cease dis- tribution and notification order or a mandatory recall order should imme- diately follow  the  instructions  set forth in the communication. Where ap- propriate, these recipients should im- mediately notify  their  consignees  of the order in accordance with para- graphs (b) and (c) of this section.

##### § 810.16 Cease distribution and  notifi- cation or mandatory recall order status reports.

The person named in a cease dis- tribution and notification order issued under § 810.10 or a mandatory recall order issued under § 810.13 shall submit periodic status reports  to FDA  to en- able the agency to assess the person’s progress in complying with the order. The frequency of such reports and the agency official to whom  such  reports shall be submitted will be specified  in the  order.

Unless otherwise specified in the order, each status report shall contain the following information:

  - The number and type of health professionals, device user facilities, consignees, or  individuals  notified about the order and the date and meth- od of notification;

  - The number and type of health professionals, device user facilities, consignees, or individuals who have re- sponded to the communication and the quantity of the device on hand at these locations at the time they received the communication;

  - The number and type of health professionals, device user facilities, consignees, or individuals who have not responded to the communication;

#### Food and Drug Administration, HHS	Pt. 812

  - The number of devices returned or corrected by each health professional, device user facility, consignee, or indi- vidual contacted, and the quantity of products accounted for;

  - The number and results of effec- tiveness checks that have been made; and

  - Estimated timeframes for comple- tion of the requirements of the cease distribution and notification order or mandatory recall order.

The person named in the cease distribution and notification order or recall order  may discontinue  the sub- mission of status reports  when  the agency terminates the order in accord- ance with § 810.17.

##### § 810.17 Termination of a cease dis- tribution and notification or man- datory recall order.

The person named in a cease dis- tribution and notification order issued under § 810.10 or a mandatory recall order issued under § 810.13 may request termination of  the  order  by  submitting a written request to FDA. The person submitting a request shall  certify  that he or she has  complied  in  full  with  all of the requirements of the order and shall include a copy of the most cur- rent status report submitted  to  the agency under § 810.16. A request for ter- mination of  a  recall  order  shall  include a description of the disposition of the recalled   device.

FDA may terminate a cease dis- tribution and notification order issued under § 810.10 or a mandatory recall order issued under § 810.13 when the agency determines that the person named in the order:

  - Has taken all reasonable efforts to ensure and to verify that all health professionals, device user facilities, consignees, and, where appropriate, in- dividuals have been notified of the cease distribution and  notification order, and to  verify  that  they  have been instructed to cease use of the de- vice and to take other appropriate ac- tion; or

  - Has removed the device from the market or has corrected the device so that use of the device would not cause serious, adverse health consequences or death.

FDA will provide written notifica- tion to the person named in the order when a request for termination of a cease distribution and notification order or a mandatory recall order has been granted or denied. FDA will re- spond to a written request for termi- nation of a cease distribution and noti- fication or recall order within 30 work- ing days of its receipt.

##### § 810.18   Public notice.

The agency  will  make  available  to the public in the weekly FDA Enforce- ment Report a descriptive  listing  of each new mandatory recall  issued under § 810.13. The agency will delay public notification of orders when the agency determines that such notifica- tion may cause unnecessary and harm- ful anxiety in individuals and that  ini- tial consultation between individuals and their health professionals is essen- tial.

## PART 812—INVESTIGATIONAL DEVICE EXEMPTIONS

##### Subpart A—General Provisions

Sec.

  - Scope.

  - Applicability.

  - Definitions.

812.5    Labeling of investigational devices.

812.7   Prohibition  of  promotion  and  other practices.

812.10   Waivers.

  - Import and export requirements.

  - Address for IDE correspondence.

##### Subpart B—Application and Administrative Action

  - Application.

812.25   Investigational plan.

812.27    Report of prior investigations.

812.30     FDA action on applications.

  - Supplemental applications.

  - Treatment  use  of  an  investigational device.

812.38   Confidentiality  of  data  and  informa- tion.

##### Subpart C—Responsibilities of Sponsors

812.40    General responsibilities of sponsors.

  - FDA and IRB approval.

  - Selecting investigators and monitors.

  - Informing investigators.

  - Monitoring investigations.

  - Emergency  research  under  § 50.24  of this   chapter.

§ 812.1

##### Subpart D—IRB Review and Approval

812.60    IRB   composition,   duties,   and   func- tions.

812.62   IRB approval.

  - IRB’s continuing review.

  - [Reserved]

  - Significant   risk   device   determina- tions.

##### Subpart E—Responsibilities of Investigators

812.100    General responsibilities of investiga- tors.

812.110     Specific responsibilities of investiga- tors.

812.119     Disqualification  of  a  clinical  investi- gator.

##### Subpart F [Reserved]

Subpart G—Records and Reports

812.140   Records.

812.145    Inspections.

  - Reports.

AUTHORITY: 21  U.S.C.  331,  351,  352,  353,  355,

360, 360c–360f, 360h–360j, 371, 372, 374, 379e, 381,

382, 383; 42 U.S.C. 216, 241, 262, 263b–263n.

SOURCE: 45  FR 3751,  Jan. 18,  1980,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 812.1   Scope.

    - The purpose of this part is to en- courage, to the extent consistent with the protection of public health and safety and with ethical standards, the discovery and development of useful devices intended for human use, and to that end to maintain optimum freedom for scientific investigators in their pur- suit of this purpose. This part provides procedures for the conduct of clinical investigations of devices. An approved investigational device exemption (IDE) permits a device that otherwise would be required to comply with a perform- ance standard or to have premarket ap- proval to be shipped lawfully for the purpose of conducting investigations of that  device.  An  IDE  approved  under

§ 812.30   or   considered   approved   under

§ 812.2(b) exempts a device from the re- quirements of the following sections of the  Federal  Food,  Drug,  and  Cosmetic Act (the act) and regulations issued thereunder: Misbranding under section 502 of the act, registration, listing, and premarket notification under  section 510,  performance  standards  under  sec-

21 CFR Ch. I (4–1–17 Edition)

tion 514, premarket approval under sec- tion 515, a banned device regulation under section 516, records and reports under section 519, restricted device re- quirements under section 520(e), good manufacturing practice requirements under section 520(f) except for the re- quirements found in § 820.30, if applica- ble (unless the sponsor states an inten- tion to  comply  with  these  require- ments	under	§ 812.20(b)(3)	or

§ 812.140(b)(4)(v)) and color additive re- quirements  under  section  721.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 59

FR  14366,  Mar.  28,  1994;  61  FR  52654,  Oct.  7,

1996]

##### § 812.2   Applicability.

General. This  part applies  to all clinical investigations of devices to de- termine safety and effectiveness, ex- cept as provided in paragraph (c) of this section.

Abbreviated requirements. The fol- lowing categories of investigations are considered to have approved applica- tions for  IDE’s,  unless  FDA  has  noti- fied a sponsor under § 812.20(a) that ap- proval  of  an  application  is  required:

  - An investigation of a device other than a significant risk device, if the de- vice is not a banned device and the sponsor:

    - Labels the device in accordance with § 812.5;

    - Obtains IRB approval of the in- vestigation after presenting the re- viewing IRB with a brief explanation of why the device is not a significant risk device, and maintains such approval;

    - Ensures that each investigator participating in an investigation of the device obtains from each subject under the investigator’s care, informed con- sent under part 50 and documents it, unless documentation is waived by an IRB under § 56.109(c).

    - Complies  with  the  requirements of § 812.46 with respect to monitoring investigations;

    - Maintains the records required under § 812.140(b) (4) and (5) and makes the  reports  required  under  § 812.150(b)

through (3) and (5) through (10);

#### Food and Drug Administration, HHS	§ 812.3

    - Ensures that participating inves- tigators  maintain  the  records  required by § 812.140(a)(3)(i) and make the re- ports required under § 812.150(a) (1), (2), (5), and (7); and

    - Complies  with  the  prohibitions in § 812.7 against promotion and other practices.

An investigation of a device other than one subject to paragraph (e) of this section, if the investigation was begun on or before July 16, 1980, and to be completed,  and  is completed,  on  or before January 19, 1981.

Exempted investigations. This part, with the exception of § 812.119, does not apply to investigations of the following categories  of  devices:

  - A device, other than a transi- tional device, in commercial distribu- tion immediately before May 28, 1976, when used or investigated in accord- ance with the indications in labeling in effect  at  that  time.

  - A device, other than a transi- tional device, introduced into commer- cial distribution on or after  May  28, 1976, that FDA has determined to be substantially equivalent to a device in commercial distribution  immediately before May 28, 1976, and that is used or investigated in accordance with the in- dications in the labeling FDA reviewed under subpart E of part 807 in deter- mining   substantial   equivalence.

  - A diagnostic device, if the sponsor complies with applicable requirements in § 809.10(c) and if the testing:

    - Is noninvasive,

    - Does not require an invasive sam- pling procedure that presents signifi- cant  risk,

    - Does not by design or intention introduce energy into a subject, and

    - Is not used as a diagnostic proce- dure without confirmation of the diag- nosis by another, medically established diagnostic product or procedure.

  - A device undergoing consumer preference testing,  testing of  a  modi- fication, or testing of a combination of two or more devices in commercial dis- tribution, if the testing is not for the purpose of determining safety or effec- tiveness and does not put subjects at risk.

  - A device intended solely for vet- erinary use.

  - A device shipped solely for re- search on or with laboratory animals and     labeled     in     accordance     with

§ 812.5(c).

  - A  custom  device  as  defined  in

§ 812.3(b), unless the device is being used to determine safety or effective- ness for commercial distribution.

Limit on certain exemptions. In the case of class II or class III device de- scribed in paragraph (c)(1) or (2) of this section, this part applies beginning on the date stipulated in an FDA regula- tion or order that calls for the submis- sion of premarket approval applica- tions for an unapproved class III de- vice, or establishes a performance standard for a class II device.

Investigations subject to IND’s. A sponsor that, on July 16, 1980, has an effective investigational new drug ap- plication (IND) for an investigation of a device shall continue to comply with the requirements of part 312 until 90 days after that date. To continue the investigation after that date, a sponsor shall comply with paragraph (b)(1) of this section, if the device is not a sig- nificant risk device, or shall have ob- tained FDA approval under § 812.30 of an IDE application for the investiga- tion of the device.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 46

FR 8956,  Jan. 27,  1981;  46  FR 14340,  Feb. 27,

1981; 53 FR 11252, Apr. 6, 1988; 62 FR 4165, Jan.

29, 1997; 62 FR 12096, Mar. 14, 1997]

##### § 812.3   Definitions.

Act means  the  Federal  Food, Drug, and Cosmetic Act (sections 201– 901, 52 Stat. 1040 et seq., as amended (21

U.S.C. 301–392)).

A custom device means a device within the meaning of section 520(b) of the Federal Food, Drug, and Cosmetic Act.

FDA means the Food and Drug Administration.

Implant means a device that is placed into a surgically or naturally formed cavity of the human body if it is intended to remain there for a period of 30 days or more. FDA may, in order to protect public health,  determine that devices placed in subjects for shorter periods are also ‘‘implants’’ for purposes of this part.

Institution means a person, other than an individual, who engages in the

#### § 812.3

conduct of research on subjects or  in the delivery of medical services to indi- viduals as a primary activity or as an adjunct to providing residential or cus- todial care to humans. The term in- cludes, for example, a hospital, retire- ment home, confinement facility, aca- demic establishment, and device manu- facturer. The term has the same mean- ing as ‘‘facility’’ in section 520(g) of the act.

Institutional review board (IRB) means any board, committee, or other group formally designated by an insti- tution to review biomedical research involving subjects and established, op- erated, and functioning in conformance with part 56. The term has the same meaning as ‘‘institutional review com- mittee’’ in section 520(g) of the act.

Investigational device means a de- vice, including a transitional device, that is the object of an investigation.

Investigation means a clinical in- vestigation or research involving  one or more subjects to determine the safe- ty or effectiveness of a device.

Investigator means an individual who actually conducts a clinical inves- tigation, i.e., under whose immediate direction the test article is adminis- tered or dispensed to, or used involv- ing, a subject, or, in the event of an in- vestigation conducted by a team of in- dividuals, is the responsible leader of that team.

Monitor, when used as a noun, means an individual designated by a sponsor or contract research organiza- tion to oversee the progress of an in- vestigation. The monitor may be an employee of a sponsor or a consultant to the sponsor, or an employee of or consultant to a contract research orga- nization. Monitor, when used as a verb, means to oversee an investigation.

Noninvasive, when applied to a di- agnostic device or  procedure,  means one that does not by design or inten- tion: (1) Penetrate or pierce the skin or mucous membranes of the body, the oc- ular cavity, or the urethra, or (2) enter the ear beyond the external auditory canal, the nose beyond the nares, the mouth beyond the pharynx, the anal canal beyond the rectum, or the vagina beyond the cervical os. For purposes of this part, blood sampling that involves simple     venipuncture     is     considered

#### 21 CFR Ch. I (4–1–17 Edition)

noninvasive, and the use of surplus samples of body fluids or tissues that are left over from samples taken for noninvestigational purposes is also considered noninvasive.

Person includes any individual, partnership, corporation, association, scientific or academic establishment, Government agency or organizational unit of a Government agency, and any other legal entity.

Significant risk device means  an investigational device that:

Is intended as an implant and pre- sents a potential for serious risk to the health, safety, or welfare of a subject;

Is purported or represented to be for a use in supporting or sustaining human life and presents a potential for serious risk to the health, safety, or welfare of a subject;

Is for a use of substantial impor- tance in diagnosing, curing, miti- gating, or treating disease, or other- wise preventing impairment of human health and presents a potential for se- rious risk to the health, safety, or wel- fare of a subject; or

Otherwise presents a potential for serious risk to the health, safety, or welfare of a subject.

Sponsor means a person who initi- ates, but who does not actually con- duct, the investigation, that is, the in- vestigational device is administered, dispensed, or used under the immediate direction of another individual. A per- son other than an individual that uses one or more of its own employees to conduct an investigation that it has initiated is a sponsor, not a sponsor-in- vestigator, and the employees are in- vestigators.

Sponsor-investigator means an indi- vidual who both initiates and actually conducts, alone or with others, an in- vestigation, that is, under whose im- mediate direction the investigational device is administered, dispensed, or used. The term does not include any person other than an individual. The obligations of a sponsor-investigator under this part include those of an in- vestigator and those of a sponsor.

Subject means a human who par- ticipates in an investigation, either as an individual on whom or on whose specimen an investigational device is used or as a control. A subject may be

#### Food and Drug Administration, HHS	§ 812.10

in normal health or may have a med- ical condition or disease.

Termination means a discontinu- ance, by sponsor or by withdrawal of IRB or FDA approval, of an investiga- tion before completion.

Transitional device means a device subject to section 520(l) of the act, that is, a  device  that  FDA  considered  to  be a new drug or an antibiotic drug before May 28, 1976.

Unanticipated adverse device effect means any serious adverse effect on health or safety or any life-threatening problem or death caused by, or associ- ated with, a device, if that effect, prob- lem, or death was not previously iden- tified in nature, severity, or degree of incidence in the investigational plan or application (including a supplementary plan or application), or any other un- anticipated serious problem associated with a device that relates to the rights, safety, or welfare of subjects.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 46

FR 8956,  Jan. 27,  1981;  48  FR 15622,  Apr. 12,

1983; 81 FR 70340, Oct. 12, 2016]

##### § 812.5   Labeling  of  investigational  de- vices.

Contents. An  investigational  de- vice or its  immediate  package  shall bear a label with the following infor- mation: the name and  place  of  business of the manufacturer, packer, or dis- tributor (in  accordance  with  § 801.1), the quantity of  contents,  if  appro- priate, and the following statement: ‘‘CAUTION—Investigational	device. Limited by  Federal  (or  United  States) law  to  investigational  use.’’   The   label or other labeling shall describe all rel- evant contraindications, hazards, ad- verse effects, interfering substances or devices,  warnings,  and  precautions.

Prohibitions. The labeling of an in- vestigational device shall not bear any statement that is false or misleading in any particular and shall not represent that the device is safe or effective for the purposes for which it is being in- vestigated.

Animal research. An investiga- tional device shipped solely for re- search on or with laboratory animals shall bear on its label the following statement: ‘‘CAUTION—Device for in- vestigational use in laboratory animals

or other tests that do not involve human  subjects.’’

The appropriate FDA Center Di- rector, according to the procedures set forth in § 801.128 or § 809.11 of this chap- ter, may grant an exception or  alter- native  to  the  provisions  in  paragraphs

and (c) of this section, to the extent that these provisions are not explicitly required by statute, for specified lots, batches, or other units of a device that are or will be included in the Strategic National  Stockpile.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 45

FR  58842,  Sept.  5,  1980;  72  FR  73602,  Dec.  28,

2007]

##### § 812.7   Prohibition  of  promotion  and other practices.

A sponsor, investigator, or any per- son acting for or on behalf of a sponsor or investigator shall not:

Promote or test market an inves- tigational device, until after FDA has approved the device for commercial distribution.

Commercialize an investigational device by charging the subjects or in- vestigators for a device a price larger than that necessary to recover costs of manufacture, research, development, and handling.

Unduly prolong an investigation. If data developed by the investigation indicate in the case of a class III device that premarket approval cannot be jus- tified or in the case of a class II device that it will not comply with an appli- cable performance standard or an amendment to that standard, the spon- sor shall promptly terminate the inves- tigation.

Represent that an investigational device is safe or effective for the pur- poses for which it is being investigated.

##### § 812.10   Waivers.

Request. A sponsor may request FDA to waive any requirement of this part. A waiver request, with supporting documentation, may be submitted sep- arately or as part of an  application  to the address in § 812.19.

FDA action. FDA may by letter grant a waiver of any requirement that FDA finds is not required by the  act and is  unnecessary  to  protect  the rights, safety, or welfare of human sub- jects.

#### § 812.18

Effect of request. Any requirement shall continue to apply  unless  and until FDA waives it.

##### § 812.18   Import   and   export   require- ments.

Imports. In addition to complying with other requirements of this part, a person who imports or offers for impor- tation an investigational device sub- ject to this part shall be the agent of the foreign exporter with respect to in- vestigations of the device and shall act as the sponsor of the clinical investiga- tion, or ensure that another person acts as the agent of the foreign ex- porter and the sponsor of the investiga- tion.

Exports. A person exporting an in- vestigational   device   subject   to    this part shall obtain FDA’s  prior  approval, as required by section 801(e) of the act or comply with section 802 of the act.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 62

FR 26229, May 13, 1997]

##### § 812.19   Address  for  IDE  correspond- ence.

If you are sending an application, supplemental application, report, re- quest for waiver, request for import or export approval, or other correspond- ence relating  to  matters  covered  by this part, you must send the submis- sion to the appropriate address as fol- lows:

  - For devices regulated by the Cen- ter for  Devices  and  Radiological Health, send it to Food and Drug Ad- ministration, Center for Devices and Radiological Health, Document Mail Center, 10903 New Hampshire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993–0002.

  - For devices regulated by the Cen- ter for Biologics Evaluation and Re- search, send it to the Food and Drug Administration, Center for Biologics Evaluation and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver  Spring, MD 20993–0002.

  - For devices regulated by the Cen- ter for Drug Evaluation and Research, send it to Central Document Control Room, Center for Drug Evaluation and Research, Food and Drug Administra- tion, 5901–B Ammendale Rd., Beltsville, MD 20705–1266.

#### 21 CFR Ch. I (4–1–17 Edition)

You must state on the outside wrapper of each submission what the submission is, for example, an ‘‘IDE ap- plication,’’ a ‘‘supplemental IDE appli- cation,’’ or a ‘‘correspondence con- cerning an IDE (or an IDE applica- tion).’’

[71 FR 42048, July 25, 2006, as amended at 75

FR  20915,  Apr.  22,  2010;  80  FR  18094,  Apr.  3,

2015]

## Subpart B—Application and Administrative Action

##### § 812.20   Application.

Submission. (1) A sponsor shall submit an application to FDA if the sponsor intends  to  use  a  significant risk device in an  investigation,  intends to conduct an investigation that in- volves an exception from informed con- sent under § 50.24 of this chapter, or if FDA notifies the sponsor that an appli- cation  is  required  for  an  investigation.

A sponsor shall not begin an in- vestigation for which FDA’s  approval of an application is required until FDA has approved the application.

A sponsor shall submit three  cop- ies of a signed ‘‘Application for an In- vestigational Device Exemption’’ (IDE application), together with accom- panying  materials,  by  registered   mail or by hand to the address in § 812.19. Subsequent correspondence  concerning an application or a supplemental appli- cation shall be submitted by registered mail or by hand.

(4)(i) A sponsor shall submit a sepa- rate IDE for any clinical investigation involving an exception from informed consent under § 50.24 of this chapter. Such a clinical investigation is not permitted to proceed without the prior written  authorization  of  FDA.  FDA shall provide a written determination 30 days  after FDA  receives the  IDE or earlier.

(ii) If the investigation involves an exception from informed consent under

§ 50.24 of this chapter, the sponsor shall prominently  identify  on   the   cover sheet  that  the  investigation   is   subject to the requirements in § 50.24 of this chapter.

Contents. An IDE application shall include, in the following order:

  - The name and address of the spon- sor.

Food and Drug Administration, HHS	§ 812.25

  - A complete report of prior inves- tigations of the device and an accurate summary of those sections of the inves- tigational plan described in § 812.25(a) through (e) or, in lieu of the summary, the complete plan. The sponsor shall submit to FDA a complete investiga- tional plan and a complete report of prior investigations of the device if no IRB has reviewed them, if FDA  has found an IRB’s review inadequate, or if FDA requests them.

  - A description of the methods, fa- cilities, and controls used for the man- ufacture, processing, packing, storage, and, where appropriate, installation of the device, in sufficient detail so that a person generally familiar with good manufacturing practices can make a knowledgeable judgment about the quality control used in the manufac- ture of the device.

  - An example of the agreements to be entered into by all investigators to comply with investigator obligations under this part, and a list of the names and addresses of all investigators who have signed the agreement.

  - A certification that all investiga- tors who will participate in the inves- tigation have signed the agreement, that the list of investigators  includes all the investigators participating in the investigation, and that no inves- tigators will be added to the investiga- tion until they have signed the agree- ment.

  - A list of the name, address, and chairperson of each IRB that has been or will be asked to review the inves- tigation and a certification of the ac- tion concerning  the  investigation taken by each such IRB.

  - The name and address of any in- stitution at which a part of the inves- tigation may be conducted that has not been identified in accordance with paragraph (b)(6) of this section.

  - If the device is to be sold, the amount to be charged and an expla- nation of why sale does not constitute commercialization of the device.

  - A claim for categorical exclusion under § 25.30 or § 25.34 or an environ- mental assessment under § 25.40.

  - Copies of all labeling for the de- vice.

  - Copies of all forms and informa- tional materials to be provided to sub- jects to obtain informed consent.

  - Any other relevant information FDA requests for review of the applica- tion.

Additional information. FDA may request additional information con- cerning an investigation or revision in the investigational plan. The sponsor may treat such a request as a dis- approval of the application for pur- poses of requesting a hearing under part 16.

Information previously submitted. Information previously submitted to the Center for Devices and Radio- logical Health, the Center for Biologics Evaluation and Research, or the Center for Drug Evaluation and Research, as applicable, in accordance with this chapter ordinarily need not be resub- mitted, but may be incorporated by reference.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 46

FR 8956,  Jan. 27,  1981;  50  FR 16669,  Apr. 26,

1985;  53  FR  11252,  Apr.  6,  1988;  61  FR  51530,

Oct. 2, 1996; 62 FR 40600, July 29, 1997; 64 FR

10942, Mar. 8, 1999; 73 FR 49942, Aug. 25, 2008]

##### § 812.25   Investigational plan.

The investigational plan shall in- clude, in the following order:

Purpose. The name and intended use of the device and the objectives and duration of the investigation.

Protocol. A written  protocol  de- scribing the methodology to be used and an analysis  of the  protocol dem- onstrating that the investigation is scientifically sound.

Risk analysis. A description and analysis of all increased risks to which subjects will be exposed by the inves- tigation; the manner in which these risks will be minimized; a justification for the investigation; and a description of the patient population, including the number, age, sex, and condition.

Description of device. A description of each important component, ingre- dient, property, and principle of oper- ation of the device and of each antici- pated change in the device during the course of the investigation.

Monitoring procedures. The spon- sor’s written procedures for monitoring the investigation and the name and ad- dress of any monitor.

#### § 812.27

Labeling. Copies of all labeling for the device.

Consent materials. Copies of all forms and informational materials to be provided to subjects to obtain in- formed consent.

IRB information. A list of the names, locations, and chairpersons of all IRB’s that have been or will be asked to review the investigation, and a certification of any action taken by any of those IRB’s with respect to the investigation.

Other institutions. The name and address of each institution at which a part of the investigation may be con- ducted that has not been identified in paragraph (h) of this section.

Additional records and reports. A de- scription of records and reports that will be maintained on the investigation in addition to those prescribed in sub- part G.

##### § 812.27   Report of prior investigations.

General. The report of prior inves- tigations shall include reports of all prior clinical, animal, and laboratory testing of the device and shall be com- prehensive and adequate to justify the proposed investigation.

Specific contents. The report also shall include:

  - A bibliography of all publications, whether adverse or supportive, that are relevant to an evaluation of the safety or effectiveness of the device, copies of all published and unpublished adverse information, and, if requested by an IRB or FDA, copies of other significant publications.

  - A summary of all other unpub- lished information (whether adverse or supportive) in the possession of, or rea- sonably obtainable by, the sponsor that is relevant to an evaluation of the safe- ty or effectiveness of the device.

  - If information on nonclinical lab- oratory studies is provided, a  state- ment that all such studies have been conducted in compliance with applica- ble requirements in the good labora- tory practice regulations in part 58, or if any such study was not conducted in compliance with such regulations, a brief statement of the reason for the noncompliance. Failure or inability to comply with this requirement does not

#### 21 CFR Ch. I (4–1–17 Edition)

justify  failure  to  provide  information on a relevant nonclinical test study.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 50

FR 7518, Feb. 22, 1985]

##### § 812.30   FDA action on applications.

Approval or disapproval. FDA will notify the sponsor in writing of the date it receives an application. FDA may approve an investigation as pro- posed, approve it with  modifications, or disapprove it. An investigation may not begin until:

  - Thirty  days   after   FDA   receives the application at the address in § 812.19 for the investigation of a device other than a banned device, unless FDA noti- fies the sponsor that the investigation may not begin; or

  - FDA approves, by order, an IDE for the investigation.

Grounds for disapproval or with- drawal. FDA may disapprove or with- draw approval of an application if FDA finds that:

There has been a failure to comply with any requirement of this part  or the act, any other applicable regula- tion or statute, or any condition of ap- proval imposed by an IRB or FDA.

The application or a report con- tains an untrue statement of a mate- rial fact, or omits material informa- tion required by this part.

The sponsor fails to respond to a request for additional information within the time prescribed by FDA.

There is reason to believe that the risks to the subjects are not out- weighed by the anticipated benefits to the subjects and the importance of the knowledge to be gained, or informed consent is inadequate, or the investiga- tion is scientifically unsound, or there is reason to believe that the device as used is ineffective.

It is otherwise unreasonable to begin or to continue the investigation owing to the way in which the device is used or the inadequacy of:

  - The report of prior investigations or the investigational plan;

  - The methods, facilities, and con- trols used for the manufacturing, proc- essing, packaging, storage, and, where appropriate, installation of the device; or

  - Monitoring and review of the in- vestigation.

#### Food and Drug Administration, HHS	§ 812.35

Notice of disapproval or withdrawal. If FDA disapproves an application or proposes to withdraw approval of an application, FDA will notify the spon- sor in writing.

  - A disapproval order will contain a complete statement of the reasons for disapproval and a statement that the sponsor has an opportunity to request a hearing under part 16.

  - A notice of a proposed withdrawal of approval will contain a complete statement of the reasons for with- drawal and a statement that the spon- sor has an opportunity to request a hearing under part 16. FDA will provide the opportunity for hearing before withdrawal of approval, unless FDA de- termines in the notice that continu- ation of testing under the exemption will result in an unreasonble risk  to the public health and orders with- drawal of approval before any hearing.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 45

FR 58842, Sept. 5, 1980]

##### § 812.35   Supplemental applications.

    - Changes in investigational plan—(1) Changes requiring prior approval.  Except as described in paragraphs  (a)(2) through (a)(4) of this section, a sponsor must obtain approval of  a  supple- mental application  under  § 812.30(a), and  IRB  approval  when   appropriate (see §§ 56.110 and 56.111 of this chapter), prior to implementing a change to an investigational plan. If a sponsor in- tends to conduct an investigation that involves an exception to informed con- sent under § 50.24 of this chapter, the sponsor shall submit a separate inves- tigational device exemption (IDE) ap- plication in accordance with § 812.20(a).

Changes effected for emergency use. The requirements of paragraph (a)(1) of this section regarding FDA approval of a supplement do not apply in the case of a deviation from the investigational plan to protect the life or physical well-being of a subject in an  emer- gency. Such deviation shall be reported to FDA within 5-working days after the sponsor learns of it (see § 812.150(a)(4)).

Changes effected with notice to FDA within 5 days. A sponsor may make cer- tain changes without prior approval of a supplemental application under para- graph (a)(1) of this section if the spon- sor    determines    that    these    changes

meet the criteria described in para- graphs (a)(3)(i) and (a)(3)(ii) of this sec- tion, on the basis of credible informa- tion defined in paragraph (a)(3)(iii) of this section, and the sponsor provides notice to FDA within 5-working days of making these changes.

Developmental changes. The re- quirements in paragraph (a)(1) of this section regarding FDA approval of a supplement do not apply to develop- mental changes in the device (includ- ing manufacturing changes) that do not constitute a significant change in design or basic principles of operation and that are made in response to infor- mation gathered during the course of an investigation.

Changes to clinical protocol. The requirements in paragraph (a)(1) of this section regarding FDA approval of a supplement do not apply to changes to clinical protocols that do not affect:

  - The validity of the data or infor- mation resulting from the completion of the approved protocol, or the rela- tionship of likely patient risk to ben- efit relied upon to approve the pro- tocol;

  - The scientific soundness of the in- vestigational plan; or

  - The rights, safety, or welfare of the human subjects involved in the in- vestigation.

Definition of credible information.

  - Credible information to support de- velopmental changes in the device (in- cluding manufacturing changes) in- cludes data generated under the design control procedures of § 820.30, pre- clinical/animal testing, peer reviewed published literature, or other reliable information such as clinical informa- tion gathered during a trial or mar- keting.

  - Credible information to support changes to clinical protocols is defined as the sponsor’s documentation sup- porting the conclusion that a change does not have a significant impact on the study design or planned statistical analysis, and that the change does not affect the rights, safety, or welfare of the subjects. Documentation shall in- clude information such as peer re- viewed published literature, the rec- ommendation of the clinical investi- gator(s), and/or the data gathered dur- ing the clinical trial or marketing.

#### § 812.36

Notice of IDE change. Changes meeting the criteria in paragraphs (a)(3)(i) and (a)(3)(ii) of this section that are supported by credible informa- tion as defined in paragraph (a)(3)(iii) of this section may be made without prior FDA approval if the sponsor sub- mits a notice of the change to the IDE not later than 5-working days after making the change. Changes to devices are deemed to occur on the date the de- vice, manufactured incorporating the design or manufacturing change, is dis- tributed to the investigator(s). Changes to a clinical protocol are deemed to occur when a clinical investigator is notified by the sponsor that the change should be implemented in the protocol or, for sponsor-investigator studies, when a sponsor-investigator incor- porates the change in  the  protocol. Such notices shall be identified as a ‘‘notice  of  IDE  change.’’

For a developmental or manufac- turing change to the device, the notice shall include a summary of the rel- evant information gathered during the course of the investigation upon which the change was based; a description of the change to the device or manufac- turing process (cross-referenced to the appropriate sections of the original de- vice description or manufacturing process); and, if design controls were used to assess the change, a statement that no new risks were identified by appropriate risk analysis and that the verification and validation testing, as appropriate, demonstrated that the de- sign outputs met the design input re- quirements. If another method of as- sessment was used, the notice shall in- clude a summary of the information which served as the credible informa- tion supporting the change.

For a protocol change, the notice shall include a description  of  the change (cross-referenced to the appro- priate sections of the original pro- tocol); an assessment supporting the conclusion that  the  change does not have a significant impact on the study design or planned statistical analysis; and a summary of the information that served as the credible information sup- porting the  sponsor’s  determination that the change does not affect  the rights, safety, or welfare of the sub- jects.

#### 21 CFR Ch. I (4–1–17 Edition)

Changes submitted in annual report. The requirements of paragraph (a)(1) of this section do not apply to minor changes to the purpose of the study, risk analysis, monitoring procedures, labeling, informed consent materials, and IRB information that do not affect:

  - The validity of the data or infor- mation resulting from the completion of the approved protocol, or the rela- tionship of likely patient risk to ben- efit relied upon to approve the pro- tocol;

  - The scientific soundness of the investigational plan; or

  - The rights, safety, or welfare of the human subjects involved in the in- vestigation. Such changes shall be re- ported in the annual progress report for the IDE, under § 812.150(b)(5).

    - IRB approval for new facilities. A sponsor shall submit to FDA a certifi- cation of any IRB approval of an inves- tigation  or  a  part  of  an   investigation not included in the IDE application.  If the investigation is otherwise un- changed, the supplemental application shall consist of an updating of the in- formation required by § 812.20(b) and (c) and a description  of  any  modifications in the investigational plan required by the IRB as a condition of approval. A certification of IRB approval need  not be included in the initial submission  of the supplemental application, and such certification is not a precondition for agency  consideration  of  the  applica- tion. Nevertheless, a sponsor  may not begin a part of an investigation at a fa- cility until the IRB has approved the investigation, FDA has received the certification of IRB approval, and FDA, under § 812.30(a), has approved the sup- plemental application relating to  that part       of       the       investigation       (see

§ 56.103(a)).

[50 FR 25909, June 24, 1985; 50 FR 28932, July

17, 1985, as amended at 61 FR 51531, Oct. 2,

1996; 63 FR 64625, Nov. 23, 1998]

##### § 812.36   Treatment use of an investiga- tional device.

General. A device that is not ap- proved for marketing may be under clinical investigation for a serious or immediately life-threatening disease or

#### Food and Drug Administration, HHS	§ 812.36

condition in patients for whom no com- parable or satisfactory alternative de- vice or other therapy is available. Dur- ing the clinical trial or prior to final action on the marketing application, it may be appropriate to use the device in the treatment of patients not in the trial under the provisions of a treat- ment investigational device exemption (IDE). The purpose of this section is to facilitate the availability of promising new devices to desperately ill patients as early in the device development process as possible, before general mar- keting begins, and to obtain additional data on the device’s safety and effec- tiveness. In the case of a serious dis- ease, a device ordinarily may be made available for treatment use under this section after all clinical trials  have been completed. In the case of an im- mediately life-threatening disease, a device may be made available for treat- ment use under this section prior  to the completion of all clinical trials. For the purpose of this section, an ‘‘immediately life-threatening’’ disease means a stage of a disease in which there is a reasonable likelihood that death will occur within a matter of months or in which premature death is likely without early treatment. For purposes of this section, ‘‘treatment use’’of a device includes the use of a de- vice for diagnostic purposes.

Criteria. FDA shall consider the use of an investigational device under a treatment IDE if:

  - The device is intended to treat or diagnose a serious or immediately life- threatening disease or condition;

  - There is no comparable or satis- factory alternative device or other therapy available to treat or diagnose that stage of the  disease  or  condition in the intended patient population;

  - The device is under investigation in a controlled clinical trial for the same use under  an approved  IDE, or such clinical trials have been com- pleted; and

  - The sponsor of the investigation is actively pursuing marketing approval/ clearance of the investigational device with due diligence.

Applications for treatment use. (1) A treatment IDE application  shall  in- clude, in the following order:

The name, address, and telephone number of the sponsor of the treatment IDE;

The intended use of the device, the criteria for patient selection, and a written protocol describing the treat- ment use;

An explanation of the rationale for use of the device, including, as ap- propriate, either a list of the available regimens that ordinarily should be tried before using the investigational device or an explanation of why the use of the investigational device is pref- erable to the use of available marketed treatments;

A description of clinical proce- dures, laboratory tests,  or other meas- ures that will be used to evaluate the effects of the device and to minimize risk;

Written procedures for moni- toring the treatment use and the name and address of the monitor;

Instructions for use for the de- vice and all other labeling as required under § 812.5(a) and (b);

Information that is relevant to the safety and effectiveness of the de- vice for the intended treatment use. In- formation from other IDE’s may be in- corporated by reference to support the treatment  use;

A statement of the sponsor’s commitment to meet all applicable re- sponsibilities under this part and  part 56 of this chapter and to ensure compli- ance of all participating investigators with the informed consent require- ments of part 50 of this chapter;

An example of the agreement to be signed by all investigators partici- pating in the treatment IDE and cer- tification that no investigator will be added to the treatment IDE before the agreement is signed; and

If the device is to be sold, the price to be charged and a statement in- dicating that the price is based on manufacturing and handling costs only.

A licensed practitioner who re- ceives an investigational device for treatment use under  a  treatment  IDE is an ‘‘investigator’’ under the IDE and is responsible for meeting all applica- ble investigator responsibilities under this part and parts 50 and 56 of this chapter.

#### § 812.38

FDA action on treatment IDE appli- cations—(1) Approval of treatment IDE’s. Treatment use may begin 30 days after FDA receives the treatment IDE sub- mission   at   the   address   specified   in

§ 812.19, unless FDA notifies the sponsor in writing earlier than the 30 days that the treatment use may or may  not begin. FDA may approve the treatment use as proposed or approve it with modifications.

  - Disapproval or withdrawal of ap- proval of treatment IDE’s. FDA may dis- approve or withdraw approval of a treatment IDE if:

    - The criteria specified in § 812.36(b) are not met or the treatment IDE does not contain the  information  required in § 812.36(c);

    - FDA determines that any of the grounds for disapproval or  withdrawal of approval listed in § 812.30(b)(1) through (b)(5) apply;

    - The device is intended for a seri- ous disease or condition and there is insufficient evidence of safety and ef- fectiveness to support such use;

    - The device is intended for an im- mediately life-threatening disease or condition and the available scientific evidence, taken as a whole, fails to pro- vide a reasonable basis for concluding that the device:

      - May be effective for its intended use in its intended population; or

      - Would not expose the patients to whom the device is to be administered to an unreasonable and significant ad- ditional risk of illness or injury;

    - There is reasonable evidence that the treatment use is impeding enroll- ment in, or otherwise interfering with the conduct or completion of, a con- trolled investigation of the same or an- other investigational device;

    - The device has received mar- keting approval/clearance or a com- parable device or therapy becomes available to treat or diagnose the same indication in the same patient popu- lation for which the investigational de- vice is being used;

    - The sponsor of the controlled clinical trial is not pursuing marketing approval/clearance with due diligence;

    - Approval of the IDE for the con- trolled clinical investigation of the de- vice has been withdrawn; or

#### 21 CFR Ch. I (4–1–17 Edition)

    - The clinical investigator(s) named in the treatment IDE are not qualified by reason of their scientific training and/or experience to use the investigational device for the intended treatment  use.

  - Notice of disapproval or withdrawal. If FDA disapproves or proposes to with- draw approval of  a  treatment  IDE, FDA will follow the procedures  set forth in § 812.30(c).

Safeguards. Treatment use of an investigational device is conditioned upon the sponsor and investigators complying with the safeguards of the IDE process and the regulations gov- erning informed consent (part 50 of this chapter) and  institutional  review boards (part 56 of this chapter).

Reporting requirements. The spon- sor of a treatment IDE shall submit progress reports  on  a  semi-annual  basis to  all  reviewing  IRB’s  and  FDA   until the filing of a marketing application. These reports shall be based on the pe- riod  of  time  since  initial  approval   of the  treatment  IDE  and  shall  include the  number  of  patients  treated  with the  device  under  the  treatment  IDE, the names of the investigators partici- pating in the  treatment  IDE,  and  a brief description of the  sponsor’s  ef- forts to pursue marketing approval/ clearance of  the  device.  Upon  filing  of a marketing application, progress re- ports shall be submitted annually in accordance with § 812.150(b)(5). The sponsor of a treatment IDE is respon- sible for submitting all other reports required under § 812.150.

[62 FR 48947, Sept. 18, 1997]

##### § 812.38   Confidentiality of data and in- formation.

Existence of IDE. FDA will not dis- close the existence of an IDE unless its existence has previously been publicly disclosed or acknowledged, until FDA approves an application for premarket approval of the device subject to the IDE; or a notice of completion of a product development protocol for the device has become effective.

Availability of summaries or data.

FDA will make publicly available, upon request, a detailed summary of information concerning the safety and effectiveness of the device that was the

#### Food and Drug Administration, HHS	§ 812.43

basis for an order approving, dis- approving, or withdrawing approval of an application for an IDE for a banned device. The summary shall include in- formation on any adverse effect on health caused by the device.

If a device is a banned device or if the existence  of  an  IDE has been pub- licly disclosed or acknowledged,  data or information contained in the file  is not available for public disclosure be- fore approval of an application for pre- market approval or the  effective  date of a notice of completion of a product development protocol except as pro- vided in this section. FDA may, in its discretion, disclose a summary of se- lected portions of the safety and effec- tiveness data, that is,  clinical,  animal, or laboratory studies and tests of the device, for public consideration of a specific pending issue.

If the existence of an IDE file has not been publicly disclosed or acknowl- edged, no data or information in  the file are available for public disclosure except as provided in paragraphs (b)(1) and (c) of this section.

Notwithstanding paragraph  (b)(2) of this  section,  FDA  will  make  avail- able to the public, upon request, the in- formation in the IDE  that  was  required to be filed in  Docket  Number  95S–0158 in the Division of Dockets Management (HFA–305), Food and Drug Administra- tion, 5630 Fishers Lane, rm. 1061, Rock- ville, MD 20852, for  investigations in- volving an exception from informed consent under § 50.24  of  this chapter. Persons wishing to request this infor- mation  shall  submit  a  request  under the  Freedom  of  Information  Act.

Reports of adverse effects. Upon re- quest or on its own initiative, FDA shall disclose to an individual on whom an investigational device has been used a copy of a report of adverse device ef- fects relating to that use.

Other rules. Except as otherwise provided in this section, the avail- ability for public disclosure of data and information in an IDE file shall be han- dled in accordance with § 814.9.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 53

FR  11253,  Apr.  6,  1988;  61  FR  51531,  Oct.  2,

1996]

## Subpart C—Responsibilities of Sponsors

##### § 812.40   General	responsibilities	of sponsors.

Sponsors are responsible for selecting qualified investigators and providing them with the information they need to conduct the investigation properly, ensuring proper monitoring of the in- vestigation, ensuring that IRB review and approval are obtained, submitting an IDE application to FDA, and ensur- ing that any reviewing IRB and  FDA are promptly informed of significant new information about an investiga- tion. Additional responsibilities of sponsors are described in subparts B and G.

##### § 812.42   FDA and IRB approval.

A sponsor shall not begin an inves- tigation or part of an investigation until an IRB and FDA have both ap- proved the application or supplemental application relating to the investiga- tion or part of an investigation.

[46 FR 8957, Jan. 27, 1981]

##### § 812.43   Selecting	investigators	and monitors.

Selecting investigators. A sponsor shall select investigators qualified by training and experience to investigate the device.

Control of device. A sponsor shall ship investigational devices only to qualified investigators participating in the investigation.

Obtaining agreements. A sponsor shall obtain from each participating investigator a signed agreement that includes:

  - The investigator’s curriculum vitae.

  - Where applicable, a statement of the investigator’s relevant experience, including the dates, location, extent, and type of experience.

  - If the investigator was involved in an investigation or other research that was terminated, an explanation of the circumstances that led to termination.

  - A statement of the investigator’s commitment to:

    - Conduct the investigation in ac- cordance with the agreement, the in- vestigational plan, this part and other

§ 812.45

applicable FDA regulations, and condi- tions of  approval  imposed  by  the re- viewing IRB or FDA;

    - Supervise all testing of the device involving human subjects; and

    - Ensure that the requirements for obtaining informed consent are met.

  - Sufficient accurate financial dis- closure information  to allow  the spon- sor to submit a complete and accurate certification or disclosure statement as required under part 54 of this chapter. The sponsor shall obtain a  commit- ment from the clinical investigator to promptly update this  information  if any relevant changes occur during the course of the investigation and for 1 year following completion of the study. This information shall not be sub- mitted in an investigational device ex- emption application, but shall be sub- mitted in any marketing application involving the device.

Selecting monitors. A sponsor shall select monitors qualified  by  training and experience to monitor the inves- tigational study in  accordance  with this part and other applicable FDA reg- ulations.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 63

FR 5253, Feb. 2, 1998]

##### § 812.45   Informing investigators.

A sponsor shall supply all investiga- tors participating in the investigation with copies of the investigational plan and the report of prior investigations of the device.

##### § 812.46   Monitoring investigations.

Securing compliance. A sponsor who discovers  that  an investigator  is not complying with the signed agree- ment, the investigational plan, the re- quirements of this part or other appli- cable FDA regulations, or any condi- tions of  approval  imposed  by  the re- viewing IRB or FDA shall promptly ei- ther secure compliance, or discontinue shipments of the device to the investi- gator and terminate the investigator’s participation in the investigation. A sponsor shall also require such an in- vestigator to dispose of or return the device, unless this action would jeop- ardize the rights, safety, or welfare of a subject.

Unanticipated adverse device effects.

A sponsor shall   immediately  con-

#### 21 CFR Ch. I (4–1–17 Edition)

duct an evaluation of any unantici- pated adverse device effect.

A sponsor who determines that an unanticipated adverse device effect presents an unreasonable risk to sub- jects shall terminate all  investigations or parts of investigations  presenting that risk as soon as possible. Termi- nation shall occur not later than 5 working days after the sponsor makes this determination and not later  than 15 working days after the sponsor first received notice of the effect.

Resumption of terminated studies. If the device is a significant risk device, a sponsor may not resume a termi- nated investigation without IRB and FDA approval. If the device is not a significant risk device, a sponsor may not resume a terminated investigation without IRB approval and, if the inves- tigation was terminated under para- graph (b)(2) of this section, FDA ap- proval.

##### § 812.47   Emergency    research    under

§ 50.24 of this chapter.

The sponsor shall monitor the progress of all investigations involving an exception from informed consent under § 50.24 of this chapter. When the sponsor receives from the IRB informa- tion concerning the public disclosures under § 50.24(a)(7)(ii) and (a)(7)(iii) of this chapter, the sponsor shall prompt- ly submit to the IDE file and to Docket Number 95S–0158 in the Division of Dockets Management (HFA–305), Food and Drug  Administration,  5630 Fishers Lane, rm. 1061, Rockville, MD 20852, copies of the information that was dis- closed, identified by the IDE number.

The sponsor also shall monitor such investigations to determine when an IRB determines that it cannot ap- prove the research because it does not meet  the  criteria  in  the  exception  in

§ 50.24(a) of this chapter or because of other relevant ethical concerns. The sponsor promptly shall provide this in- formation in writing to FDA, inves- tigators who are asked  to  participate in this or a substantially equivalent clinical investigation, and other IRB’s that are asked to review this or a sub- stantially   equivalent   investigation.

[61 FR 51531, Oct. 2, 1996, as amended at 64 FR

10943, Mar. 8, 1999]

#### Food and Drug Administration, HHS	§ 812.110

Subpart D—IRB Review and Approval

##### § 812.60   IRB  composition,  duties,  and functions.

An IRB reviewing and approving in- vestigations under this part shall com- ply with the requirements of part 56 in all respects, including its composition, duties, and functions.

[46 FR 8957, Jan. 27, 1981]

##### § 812.62   IRB approval.

  - An IRB shall review and have au- thority to approve, require modifica- tions in (to secure approval), or dis- approve all investigations covered by this part.

  - If no IRB exists or if FDA finds that an  IRB’s  review  is  inadequate,  a sponsor may submit an application to FDA.

[46 FR 8957, Jan. 27, 1981]

##### § 812.64   IRB’s continuing review.

The IRB shall conduct its continuing review of an investigation in accord- ance with part 56.

[46 FR 8957, Jan. 27, 1981]

##### § 812.65   [Reserved]

§ 812.66   Significant  risk  device  deter- minations.

If an IRB determines that an inves- tigation, presented for approval under

§ 812.2(b)(1)(ii), involves a significant risk device, it shall so notify the inves- tigator and, where appropriate, the sponsor. A sponsor may not begin the investigation    except    as    provided    in

§ 812.30(a).

[46 FR 8957, Jan. 27, 1981]

## Subpart E—Responsibilities of Investigators

##### § 812.100   General responsibilities of in- vestigators.

An investigator is responsible for en- suring that an investigation is con- ducted according to the signed agree- ment, the investigational plan and ap- plicable FDA regulations, for pro- tecting the rights, safety, and welfare of subjects under the investigator’s care,  and  for  the  control  of  devices

under investigation.  An  investigator also is responsible for ensuring that in- formed consent is obtained in accord- ance with part 50 of this chapter. Addi- tional responsibilities of investigators are described in subpart G.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 46

FR 8957, Jan. 27, 1981]

##### § 812.110   Specific responsibilities of in- vestigators.

Awaiting approval. An investigator may determine whether potential sub- jects would be interested in partici- pating in an investigation, but shall not request the written informed con- sent of any subject to participate, and shall not allow any subject to partici- pate before obtaining IRB and FDA ap- proval.

Compliance. An investigator shall conduct an investigation in accordance with the signed agreement with the sponsor, the investigational plan, this part and other applicable FDA regula- tions, and any conditions of approval imposed by an IRB or FDA.

Supervising device use. An investi- gator shall permit an investigational device to be used only with subjects under the investigator’s supervision. An investigator shall not supply an in- vestigational device to any person not authorized under  this part  to  receive it.

Financial disclosure. A clinical in- vestigator shall disclose to the sponsor sufficient accurate financial informa- tion to allow the applicant to submit complete and accurate certification or disclosure statements required under part 54 of this chapter. The investi- gator shall promptly update this infor- mation if any relevant changes occur during the course of the investigation and for 1 year following completion of the  study.

Disposing of device. Upon comple- tion or termination of a clinical inves- tigation or the investigator’s part of an investigation, or at the sponsor’s re- quest, an investigator shall return to the sponsor any remaining supply of the device or otherwise dispose of the device as the sponsor directs.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 63

FR 5253, Feb. 2, 1998]

§ 812.119

##### § 812.119   Disqualification  of  a  clinical investigator.

  - If FDA has information indicating that an investigator (including a spon- sor-investigator) has repeatedly or de- liberately failed to comply with the re- quirements of this part, part 50, or part 56 of this chapter, or has repeatedly or deliberately  submitted  to  FDA  or  to the sponsor false information in any required report, the Center for Devices and Radiological Health, the Center for Biologics Evaluation and Research, or the Center for Drug Evaluation and Re- search will furnish the investigator written notice of the matter com- plained of and offer the investigator an opportunity to explain the matter in writing, or, at the option of the inves- tigator, in an informal conference.  If an explanation  is offered  and accepted by the  applicable  Center,  the  Center will discontinue the disqualification proceeding. If an explanation is offered but not accepted by the applicable Cen- ter, the investigator will be given an opportunity for a regulatory hearing under part 16 of this chapter on the question of whether the investigator is eligible to receive test articles under this part and eligible to conduct any clinical investigation that supports an application for a research or marketing permit  for  products  regulated  by  FDA.

  - After evaluating all available in- formation, including any explanation presented by the investigator, if the Commissioner determines that the in- vestigator has repeatedly or delib- erately failed to comply with the re- quirements of this part, part 50, or part 56 of this chapter, or has repeatedly or deliberately  submitted  to  FDA  or  to the sponsor false information in any required report, the Commissioner will notify the investigator, the sponsor of any investigation in which the investi- gator has been named as a participant, and the reviewing investigational re- view boards (IRBs) that the  investi- gator is not eligible to receive test ar- ticles under this part.  The  notification to the investigator, sponsor and IRBs will provide a statement of the  basis for such determination. The  notifica- tion also will explain that an investi- gator determined to be ineligible to re- ceive test articles under this  part  will be ineligible to conduct any clinical in-

#### 21 CFR Ch. I (4–1–17 Edition)

vestigation that supports an  applica- tion for a research or marketing per- mit for products regulated by FDA, in- cluding drugs, biologics, devices, new animal drugs, foods, including dietary supplements, that bear a nutrient con- tent claim or a health claim, infant formulas, food and color additives, and tobacco products.

  - Each application or submission to FDA under the provisions of this chap- ter containing data reported by an in- vestigator who has been determined to be ineligible to receive FDA-regulated test articles is subject to examination to determine whether the investigator has submitted unreliable data that are essential to the continuation of an in- vestigation or essential to the clear- ance or approval of a marketing appli- cation, or essential to the continued marketing of an FDA-regulated prod- uct.

  - If the Commissioner determines, after the unreliable data submitted by the investigator are eliminated from consideration, that the data remaining are inadequate to support a conclusion that it is reasonably safe to  continue the  investigation,  the  Commissioner will notify the sponsor, who shall have an opportunity for a regulatory hear- ing under part 16 of this chapter. If a danger to the public health exists, how- ever, the Commissioner shall termi- nate the investigational device exemp- tion (IDE) immediately and notify the sponsor and the reviewing IRBs of the termination. In such case, the sponsor shall have an opportunity for a regu- latory  hearing  before  FDA  under  part

16 of this chapter on the question of whether the IDE should be reinstated. The determination that an  investiga- tion may not be considered in support of a research or  marketing  application or a notification or petition submission does not, however, relieve the sponsor of any obligation under any other ap- plicable regulation to submit to FDA the results of the investigation.

  - If the Commissioner determines, after the unreliable data submitted by the investigator are eliminated from consideration, that the continued clearance or approval of the product for which the data were submitted can- not be justified, the Commissioner will

#### Food and Drug Administration, HHS	§ 812.140

proceed to rescind clearance or with- draw approval of the product in accord- ance with the applicable provisions of the  relevant  statutes.

  - An investigator who has been de- termined to be ineligible under para- graph (b) of this section may be rein- stated as eligible when the Commis- sioner determines that the investigator has presented adequate assurances that the investigator will employ all test articles, and will conduct any clinical investigation that supports an applica- tion for a research or marketing per- mit for products regulated by FDA, solely in compliance with the applica- ble provisions of this chapter.

[77 FR 25360, Apr. 30, 2012]

## Subpart F [Reserved]

Subpart G—Records and Reports

##### § 812.140   Records.

Investigator records. A partici- pating investigator shall maintain the following accurate, complete, and cur- rent records relating to the investiga- tor’s participation in an investigation:

All correspondence with another investigator, an IRB, the sponsor, a monitor, or FDA, including required reports.

Records of receipt, use or disposi- tion of a device that relate to:

  - The type and quantity of the de- vice, the dates  of its  receipt, and  the batch number or code mark.

  - The names of all persons who re- ceived, used, or disposed of each device.

  - Why and how many units of the device have been returned to the spon- sor, repaired, or otherwise disposed of.

Records of each subject’s case his- tory and exposure to the device. Case histories include the case report forms and supporting data including, for ex- ample, signed and dated consent forms and medical records including, for ex- ample, progress notes of the physician, the individual’s hospital chart(s), and the nurses’ notes. Such records shall include:

  - Documents evidencing informed consent and, for any use of a device by the investigator without  informed con- sent, any written concurrence of a li- censed  physician  and  a  brief  descrip-

tion of  the  circumstances  justifying the failure to obtain informed consent. The case history for each individual shall document that informed consent was obtained prior to participation in the study.

  - All relevant observations, includ- ing records concerning adverse device effects (whether anticipated or unan- ticipated), information and data on the condition of each subject upon enter- ing, and during the course of, the in- vestigation,  including  information about relevant previous medical  his- tory and the results of all diagnostic tests.

  - A record of the exposure of each subject to the investigational device, including the date and time of each use, and any other therapy.

The protocol, with documents showing the dates of and reasons for each deviation from the protocol.

Any other records that FDA re- quires to be maintained by regulation or by specific requirement for a cat- egory of investigations or a particular investigation.

Sponsor records. A sponsor shall maintain the following accurate, com- plete, and current records relating to an investigation:

  - All correspondence with another sponsor, a monitor, an investigator, an IRB, or FDA, including required re- ports.

  - Records of shipment and disposi- tion. Records of shipment shall include the name and address of the consignee, type and quantity of device, date of shipment, and batch number or code mark. Records of disposition shall de- scribe the batch number or code marks of any devices returned to the sponsor, repaired, or disposed of in other ways by the investigator or another person, and the reasons for and method of dis- posal.

  - Signed investigator agreements including the financial disclosure in- formation required to be  collected under § 812.43(c)(5) in accordance with part 54 of this chapter.

§ 812.145

  - For each investigation subject to

§ 812.2(b)(1) of a device other than a sig- nificant risk device, the records de- scribed in paragraph (b)(5) of this sec- tion and the following records, consoli- dated in one location and available for FDA  inspection  and  copying:

    - The name and intended use of the device and the objectives of the inves- tigation;

    - A brief explanation of why the de- vice is not a significant risk device:

    - The name and address of each in- vestigator:

    - The name and address of each IRB that has reviewed the investiga- tion:

    - A statement of the extent to which the good manufacturing practice regulation in part 820 will be followed in manufacturing the device; and

    - Any other information required by FDA.

  - Records concerning adverse device effects (whether anticipated or unan- ticipated) and complaints and

  - Any other records that FDA re- quires to be maintained by regulation or by specific requirement for a cat- egory of investigation or a particular investigation.

IRB records. An IRB shall main- tain records in accordance with part 56 of this chapter.

Retention period. An investigator or sponsor shall maintain the records required by this subpart during the in- vestigation and for a period of 2 years after the latter of the following two dates: The date on which the investiga- tion is terminated or completed, or the date that the records are no longer re- quired for purposes of supporting a pre- market approval application or a no- tice of completion of a product devel- opment protocol.

Records custody. An  investigator or sponsor may withdraw from the re- sponsibility to maintain records for the period required in paragraph (d) of this section and transfer custody of the records to any other person who will accept responsibility for  them  under this  part,  including  the   requirements of § 812.145. Notice of a transfer shall be

#### 21 CFR Ch. I (4–1–17 Edition)

given to FDA not later than 10 working days after transfer occurs.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 45

FR 58843,  Sept. 5,  1980;  46  FR 8957,  Jan. 27,

1981; 61 FR 57280, Nov. 5, 1996; 63 FR 5253, Feb.

2, 1998]

##### § 812.145   Inspections.

- Entry and inspection. A sponsor or an investigator who has authority to grant access shall permit authorized FDA employees, at reasonable times and in a reasonable manner, to enter and inspect any establishment where devices are held (including any estab- lishment where devices are manufac- tured, processed, packed, installed, used, or implanted or where records of results from use of devices are kept).

- Records inspection. A sponsor, IRB, or investigator, or any other person acting on behalf of such a person with respect to an investigation, shall per- mit authorized FDA employees, at rea- sonable times and in a reasonable man- ner, to inspect and copy all records re- lating to an investigation.

- Records identifying subjects. An in- vestigator shall permit authorized FDA employees to inspect and copy records that identify subjects, upon notice that FDA has reason to suspect that ade- quate informed consent was not ob- tained, or that reports required to be submitted by the investigator to the sponsor or IRB have not been sub- mitted or are incomplete, inaccurate, false, or misleading.

##### § 812.150   Reports.

Investigator reports. An investi- gator shall prepare and submit the fol- lowing complete, accurate, and timely reports:

Unanticipated adverse device effects. An investigator shall submit to the sponsor and to the reviewing IRB a re- port of any unanticipated adverse de- vice effect occurring during an inves- tigation as soon as possible, but in no event later than 10 working days after the investigator first learns of the ef- fect.

Withdrawal of IRB approval. An in- vestigator shall report to the sponsor, within 5 working days, a withdrawal of approval by the reviewing IRB of the investigator’s  part  of  an  investigation.

#### Food and Drug Administration, HHS	§ 812.150

Progress. An investigator shall submit progress reports on the inves- tigation to the sponsor, the monitor, and the reviewing IRB at regular inter- vals, but in no event less often than yearly.

Deviations from the investigational plan. An investigator shall notify the sponsor  and  the  reviewing  IRB  (see

§ 56.108(a) (3) and (4)) of any deviation from the  investigational  plan  to  pro- tect the life or physical well-being of a subject in an emergency.  Such  notice shall be given as  soon  as  possible,  but in no event later than 5 working days after  the  emergency  occurred.   Except in  such  an  emergency,  prior   approval by the sponsor is  required  for  changes in or deviations from  a  plan,  and  if these changes or deviations may affect the scientific soundness of the plan  or the rights, safety, or welfare of human subjects, FDA and IRB  in  accordance with § 812.35(a) also is required.

Informed consent.  If an investi- gator uses a device without obtaining informed  consent,  the   investigator shall report such use to the sponsor and the reviewing IRB within 5  work- ing days after the use occurs.

Final report. An investigator shall, within 3 months after termination or completion of the investigation or the investigator’s part of the investigation, submit a final report to the sponsor and the reviewing IRB.

Other. An investigator shall, upon request by a reviewing IRB or FDA, provide accurate, complete, and cur- rent information about any aspect of the investigation.

Sponsor reports. A sponsor shall prepare and submit the following com- plete, accurate, and timely reports:

Unanticipated adverse  device  effects. A sponsor who conducts  an  evaluation of an unanticipated adverse device ef- fect under § 812.46(b) shall report the re- sults of such evaluation  to  FDA  and  to all reviewing IRB’s and participating investigators within 10 working days after  the  sponsor  first   receives   notice of the effect. Thereafter  the  sponsor shall submit such additional reports concerning  the  effect  as  FDA  requests.

Withdrawal of IRB approval. A sponsor shall notify FDA and all re- viewing IRB’s and participating inves- tigators of any withdrawal of approval

of an investigation or a part of an in- vestigation by a reviewing IRB within

5   working   days   after   receipt   of   the withdrawal of approval.

Withdrawal of FDA approval. A sponsor shall notify all reviewing IRB’s and participating investigators of any withdrawal of FDA approval of the in- vestigation, and shall do so within 5 working days after receipt of notice of the withdrawal of approval.

Current investigator list. A sponsor shall submit to FDA, at 6-month inter- vals, a current list of the names and addresses of all investigators partici- pating in the investigation. The spon- sor shall submit the first such list 6 months after FDA approval.

Progress reports. At regular inter- vals, and at  least  yearly,  a  sponsor shall submit progress reports to all re- viewing IRB’s. In the case of a signifi- cant risk device, a sponsor shall also submit progress reports to FDA. A sponsor of a treatment IDE shall sub- mit semi-annual progress reports to all reviewing IRB’s and FDA in accordance with § 812.36(f) and annual reports in ac- cordance with this section.

Recall and device disposition. A sponsor shall notify FDA and all re- viewing IRB’s of any request that an investigator return, repair, or other- wise dispose of any units of a device. Such notice shall occur within 30 work- ing days after the request is made and shall state why  the request was made.

Final report. In the case of a sig- nificant risk device, the sponsor shall notify FDA within 30 working days  of the completion or termination of the investigation and shall submit a final report to FDA and all reviewing the IRB’s and participating investigators within 6 months after completion or termination. In the  case of a device that  is  not  a  significant  risk   device, the sponsor shall submit a final report to all reviewing IRB’s within 6 months after termination  or  completion.

Informed consent. A sponsor shall submit to FDA a copy of any report by an investigator under paragraph (a)(5) of this section of use of a device with- out obtaining informed consent, within

5  working  days  of  receipt  of  notice  of such use.

#### Pt. 814

Significant risk device determina- tions. If an IRB determines that a de- vice is a significant risk device,  and the sponsor had proposed that the IRB consider the device not to be a signifi- cant risk device, the sponsor shall sub- mit to FDA a report of the IRB’s deter- mination within 5 working days after the sponsor first learns of the IRB’s de- termination.

Other. A sponsor shall, upon re- quest by a reviewing IRB or FDA, pro- vide accurate, complete, and current information about any aspect of the in- vestigation.

[45  FR 3751,  Jan. 18,  1980,  as  amended  at 45

FR  58843,  Sept.  5,  1980;  48  FR  15622,  Apr.  12,

1983; 62 FR 48948, Sept. 18, 1997]

## PART 813 [RESERVED]

PART 814—PREMARKET APPROVAL OF MEDICAL DEVICES

##### Subpart A—General

21 CFR Ch. I (4–1–17 Edition)

Subparts F–G [Reserved]

Subpart H—Humanitarian Use Devices

814.100  Purpose and scope.

814.102    Designation of HUD status.

814.104    Original applications.

814.106    HDE   amendments   and   resubmitted HDE’s.

814.108    Supplemental applications.

814.110   New indications for use.

814.112    Filing an HDE.

814.114    Timeframes for reviewing an HDE.

814.116    Procedures for review of an HDE.

814.118    Denial  of  approval  or  withdrawal  of approval  of  an  HDE.

814.120   Temporary suspension of approval of an  HDE.

814.122    Confidentiality  of  data  and  informa- tion.

814.124   Institutional Review Board require- ments.

  - Postapproval   requirements   and   re- ports.

AUTHORITY: 21 U.S.C. 351, 352, 353, 360, 360c–

360j, 371, 372, 373, 374, 375, 379, 379e, 381.

SOURCE: 51 FR 26364, July 22, 1986, unless

otherwise noted.

Sec.

  - Scope.

  - Purpose.

  - Definitions.

814.9 Confidentiality of data and informa- tion in a premarket approval application (PMA) file.

814.15 Research	conducted	outside	the United   States.

814.17   Service of orders.

  - Product  development  protocol  (PDP).

##### Subpart B—Premarket Approval Application (PMA)

  - Application.

814.37    PMA   amendments   and   resubmitted PMAs.

  - PMA supplements.

##### Subpart C—FDA Action on a PMA

  - Time frames for reviewing a PMA.

814.42   Filing a PMA.

  - Procedures for review of a PMA.

  - Denial of approval of a PMA.

  - Withdrawal of approval of a PMA.

  - Temporary suspension of approval of a  PMA.

##### Subpart D—Administrative Review

[Reserved]

Subpart E—Postapproval Requirements

814.80   General.

814.82    Postapproval requirements.

814.84    Reports.

## Subpart A—General

##### § 814.1   Scope.

    - This section implements sections

515 and 515A of the act by providing procedures for the premarket  approval of medical devices intended for human use.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - This part applies to any class III medical device, unless exempt under section 520(g) of the act, that:

      - Was not on the market (intro- duced or delivered for introduction into commerce for commercial distribution) before May 28, 1976, and is not substan- tially equivalent to a device on the market before May 28, 1976, or to a de- vice first marketed on, or after that date, which has been classified into class I or class II; or

      - Is required to have an approved premarket approval application (PMA) or a declared completed product devel- opment protocol under a regulation issued under section 515(b) of the act; or

      - Was regulated by FDA as a new drug or antibiotic drug before May 28,

Food and Drug Administration, HHS	§ 814.3

1976,  and  therefore  is  governed  by  sec- tion 520(1) of the act.

    - This part amends the conditions to approval for any PMA approved be- fore the effective date of this part. Any condition to approval for an approved PMA that is inconsistent with this part is revoked. Any condition to approval for an approved PMA that is consistent with this part remains in effect.

[51 FR 26364, July 22, 1986, as amended at 79

FR 1740, Jan. 10, 2014]

##### § 814.2   Purpose.

The purpose of this part is to estab- lish an efficient and thorough device review process—

To facilitate the approval  of PMA’s for devices  that  have  been shown to be safe and effective and that otherwise meet the statutory criteria for approval; and

To ensure the disapproval of PMA’s for devices that have not been shown to be safe and effective  or that do not otherwise meet the statutory criteria for approval. This part shall be construed in light of these objectives.

##### § 814.3   Definitions.

For the purposes of this part:

Act means  the  Federal  Food, Drug, and Cosmetic Act (sections 201– 902, 52 Stat. 1040 et seq., as amended (21

U.S.C. 321–392)).

FDA means the Food and Drug Administration.

IDE means an approved or consid- ered approved investigational device exemption under section 520(g) of the act and parts 812 and 813.

Master file means a reference source that a person submits to FDA. A master file may contain detailed infor- mation on a specific manufacturing fa- cility, process, methodology, or compo- nent used in the manufacture, proc- essing, or packaging of a medical de- vice.

PMA means any premarket ap- proval application for a class III med- ical device, including all information submitted with or incorporated by ref- erence therein. ‘‘PMA’’ includes a new drug application for a device under sec- tion 520(1) of the act.

PMA amendment means informa- tion  an  applicant  submits  to  FDA  to

modify  a  pending  PMA  or  a  pending PMA  supplement.

PMA supplement means a supple- mental application to an approved PMA for approval of a change or modi- fication in a class III medical device, including all information submitted with or incorporated by reference therein.

Person includes any individual, partnership, corporation, association, scientific or academic establishment, Government agency, or organizational unit thereof, or any other legal entity.

Statement of material fact means a representation that tends to show that the safety or effectiveness of a device is more probable than it would be in the absence of such a representation. A false affirmation or silence or an omis- sion that would lead a reasonable per- son to draw a particular conclusion as to the safety or effectiveness of a de- vice also may be a false statement of material fact, even if the statement was not intended by the person making it to be misleading or to have any pro- bative effect.

30-day PMA supplement means a supplemental application  to an  ap- proved     PMA     in     accordance     with

§ 814.39(e).

Reasonable probability means that it is more likely than not that an event will occur.

Serious, adverse health consequences means any significant adverse experi- ence, including those which may be ei- ther life-threatening or involve perma- nent or long term injuries, but exclud- ing injuries that are nonlife-threat- ening and that are temporary and rea- sonably reversible.

HDE means  a premarket approval application submitted pursuant to this subpart seeking a humanitarian device exemption from the effectiveness re- quirements of sections 514 and 515  of the act as authorized by section 520(m)(2) of the act.

HUD (humanitarian use device) means a medical device intended to benefit patients in the treatment or di- agnosis of a disease or condition that affects or is manifested in fewer than 4,000 individuals in the United States per year.

Newly acquired information means data,  analyses,  or  other  information

#### § 814.9

not previously submitted to the  agen- cy, which may include (but are not lim- ited to) data derived from new clinical studies, reports of adverse events, or new analyses of previously submitted data (e.g., meta-analyses) if the stud- ies, events or analyses reveal risks of a different type or greater severity or frequency than previously included in submissions  to  FDA.

Human cell, tissue, or cellular or tis- sue-based product (HCT/P) regulated as a device  means  an  HCT/P  as  defined  in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A unique  device  identifier  is  composed  of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

Universal product code (UPC) means the product identifier used to identify an item sold at retail in the United States.

Pediatric patients means patients who are 21 years of age or younger (that is, from birth through the twen- ty-first year of life, up to but not in- cluding the twenty-second birthday) at the time of the diagnosis or treatment.

Readily available means available in the public domain through com- monly  used  public  resources  for  con-

#### 21 CFR Ch. I (4–1–17 Edition)

ducting biomedical, regulatory, and medical product research.

[51 FR 26364, July 22, 1986, as amended at 61

FR  15190,  Apr.  5,  1996;  61  FR  33244,  June  26,

1996; 73 FR 49610, Aug. 22, 2008; 78 FR 55821,

Sept. 24, 2013; 79 FR 1740, Jan. 10, 2014]

##### § 814.9 Confidentiality of data and in- formation in a premarket approval application (PMA) file.

- A ‘‘PMA file’’ includes all data and information submitted with or in- corporated by reference in the PMA, any IDE incorporated into the PMA, any PMA supplement, any report under

§ 814.82, any master file, or any other related submission. Any record  in  the PMA file will be available for public disclosure in accordance with the pro- visions of this section and part 20. The confidentiality  of   information   in   a color  additive  petition  submitted   as part of a PMA is governed by § 71.15.

- The existence of a PMA file may not be disclosed by FDA before an ap- proval order is issued to the applicant unless it previously has been publicly disclosed or acknowledged.

- If the existence of a PMA file has not been publicly disclosed or acknowl- edged, data or information in the PMA file are not available for public disclo- sure.

(d)(1) If the existence of a PMA  file has been publicly disclosed or acknowl- edged before an order approving, or an order denying approval of the PMA is issued, data or information  contained in the file are not available for public disclosure before such order  issues. FDA may, however, disclose a  sum- mary of portions of the safety and ef- fectiveness data before an approval order or an order denying approval of the PMA issues if disclosure is relevant to public consideration of a specific pending issue.

(2) Notwithstanding paragraph (d)(1) of this  section,  FDA  will  make  avail- able to the public upon request the in- formation in the IDE  that  was  required to be filed in  Docket  Number  95S–0158 in the Division of Dockets Management (HFA–305), Food and Drug Administra- tion, 12420 Parklawn Dr., rm. 1–23, Rockville, MD 20857, for investigations involving an exception from informed consent  under   § 50.24   of   this   chapter.

#### Food and Drug Administration, HHS	§ 814.9

Persons wishing to request this infor- mation shall submit a request under the Freedom of Information Act.

  - Upon issuance of an order approv- ing, or an order denying approval of any PMA, FDA will make available to the public the fact of the existence of the PMA and a detailed summary of in- formation submitted to FDA respect- ing the safety and effectiveness of the device that is the subject of the PMA and that is the basis for the order.

  - After FDA issues an order approv- ing, or an order denying approval of any PMA, the following data and infor- mation in the PMA file are imme- diately available for public disclosure:

    - All safety and  effectiveness  data and information previously disclosed to the public, as such  disclosure  is  defined in § 20.81.

    - Any protocol for a test or study unless the protocol is shown to con- stitute trade secret or confidential commercial or  financial  information under § 20.61.

    - Any adverse reaction report, prod- uct experience report, consumer com- plaint, and other similar data and in- formation, after deletion of:

      - Any information that constitutes trade secret or confidential commer- cial   or   financial   information   under

§ 20.61;  and

      - Any personnel, medical, and simi- lar information  disclosure  of  which would constitute a clearly unwarranted invasion   of   personal   privacy   under

§ 20.63; provided, however, that except for the  information  that  constitutes trade secret or  confidential  commer- cial     or     financial     information     under

§ 20.61, FDA will disclose to a patient who requests a report all the informa- tion in the report concerning that pa- tient.

    - A list of components previously disclosed to the public, as such disclo- sure is defined in § 20.81.

    - An assay method or other analyt- ical method, unless it does not serve any regulatory purpose and is shown to fall within the exemption in § 20.61 for trade secret or confidential  commer- cial  or  financial  information.

    - All correspondence and written summaries of oral discussions  relating to the PMA file, in accordance with the provisions of §§ 20.103 and 20.104.

  - All safety and effectiveness data and other information not previously disclosed to the public are available for public disclosure if any one of the fol- lowing events occurs and the data and information do not constitute trade se- cret or confidential commercial or fi- nancial information under § 20.61:

    - The   PMA   has   been   abandoned. FDA will consider a PMA abandoned if: (i)(A)  The  applicant  fails  to  respond to a request for additional information within  180  days  after  the  date  FDA

issues the request or

(B) Other circumstances indicate that further work is not being under- taken with respect to it, and

(ii) The applicant fails to commu- nicate with FDA within 7 days  after the date on which FDA notifies the ap- plicant that the PMA appears to have been abandoned.

    - An order denying approval of the PMA has issued, and all legal appeals have been exhausted.

    - An order withdrawing approval of the PMA has issued, and all legal ap- peals have been exhausted.

    - The device has been reclassified.

    - The device has been found to be substantially equivalent to a class I or class II device.

    - The PMA is  considered  volun- tarily withdrawn under § 814.44(g).

  - The following data and  informa- tion in a PMA file are not available for public disclosure unless they have been previously disclosed to the  public,  as such disclosure is defined in § 20.81, or they relate  to  a  device  for  which  a PMA has been abandoned and they no longer represent a trade secret or con- fidential commercial or financial infor- mation as defined in § 20.61:

    - Manufacturing methods or proc- esses, including quality control proce- dures.

    - Production, sales, distribution, and similar data and information, ex- cept that any compilation of such data and information aggregated and pre- pared in a way that does not reveal data or information which are not available for public disclosure under this provision is available for public disclosure.

#### § 814.15

    - Quantitative or semiquantitative formulas.

[51 FR 26364, July 22, 1986, as amended at 61

FR 51531, Oct. 2, 1996]

##### § 814.15   Research   conducted   outside the United States.

(a) A study conducted outside the United States  submitted  in  support  of a PMA and conducted under an IDE shall comply with part 812. A study conducted outside the United States submitted in support of a PMA and not conducted under an IDE shall comply with the provisions in paragraph (b) or

of this section, as applicable.

  - Research begun on or after effective date. FDA will accept studies sub- mitted in support of a PMA which have been conducted outside the United States and begun on or after November 19, 1986, if the data are valid and the in- vestigator has conducted the studies in conformance with the ‘‘Declaration of Helsinki’’ or the laws  and  regulations of the country in which the research is conducted, whichever accords greater protection  to  the  human  subjects.  If the standards of the country are  used, the applicant shall state in detail any differences between those standards and the ‘‘Declaration of Helsinki’’ and explain why they offer greater protec- tion to the human subjects.

  - Research begun before effective date. FDA will accept studies sub- mitted in support of a PMA which have been conducted outside the United States and begun before November 19, 1986, if FDA is satisfied that the data are scientifically valid and that the rights, safety, and welfare of human subjects have not been violated.

  - As sole basis for marketing approval. A PMA based solely on foreign clinical data and otherwise meeting  the  cri- teria for approval under this part may be approved if:

    - The foreign data are applicable to the U.S. population and U.S. medical practice;

    - The studies have been performed by clinical investigators of recognized competence; and

    - The data may be considered valid without the need for an on-site inspec- tion by FDA or, if FDA considers such an inspection to be necessary, FDA can

#### 21 CFR Ch. I (4–1–17 Edition)

validate the data through an on-site in- spection or other appropriate means.

  - Consultation between FDA and ap- plicants. Applicants are encouraged to meet  with  FDA  officials  in  a  ‘‘pre- submission’’  meeting  when  approval based  solely  on  foreign  data  will  be sought.

(Approved  by  the  Office  of  Management  and Budget under control number 0910–0231)

[51 FR 26364, July 22, 1986; 51 FR 40415, Nov.

7,  1986,  as  amended  at  51  FR  43344,  Dec.  2,

1986]

##### § 814.17   Service of orders.

Orders issued under this part will be served in person by a designated officer or employee of FDA on, or by reg- istered mail to, the applicant or the designated agent at the applicant’s or designated agent’s last known address in FDA’s records.

##### § 814.19   Product development protocol (PDP).

A class III device for which a product development protocol has been de- clared completed by FDA under this chapter will be considered to have an approved PMA.

## Subpart B—Premarket Approval Application (PMA)

##### § 814.20   Application.

The applicant or an authorized representative shall sign the PMA. If the applicant does not reside or have a place of business within the United States, the PMA shall be countersigned by an authorized representative resid- ing or maintaining a  place  of  business in the United States and shall  identify the representative’s name and address.

Unless the applicant justifies an omission in accordance with paragraph

of this section, a PMA shall include:

The name and address of the ap- plicant.

A table of contents that specifies the volume and page number for each item referred to in the table. A PMA shall include separate sections on non- clinical laboratory studies and on clin- ical investigations involving human subjects. A PMA shall be submitted in six copies each bound in one or more numbered volumes of reasonable size.

Food and Drug Administration, HHS	§ 814.20

The applicant shall include informa- tion that it believes to be trade secret or confidential commercial or financial information in all copies of the PMA and identify in at least one copy the in- formation that it believes to be trade secret or confidential commercial or fi- nancial information.

A summary in sufficient detail that the reader may gain a general un- derstanding of the data and informa- tion in the application. The summary shall contain the following informa- tion:

  - Indications for use. A general de- scription of the disease or condition the device will diagnose, treat, prevent, cure, or mitigate, including a descrip- tion of the patient population for which the device is intended.

  - Device description. An explanation of how the device functions, the basic scientific concepts that form the basis for the device, and the significant physical and performance characteris- tics of the device. A brief description of the manufacturing process should be included if it will significantly enhance the reader’s understanding of the de- vice. The generic name of the device as well as any proprietary name or trade name should be included.

  - Alternative practices and proce- dures. A description of existing alter- native practices or procedures for diag- nosing, treating, preventing, curing, or mitigating the disease or condition for which the device is intended.

  - Marketing history. A brief descrip- tion of the foreign and U.S. marketing history, if  any,  of  the  device,  including a list of all countries in which the  de- vice has been marketed and a list of all countries in which the device has been withdrawn from marketing for any rea- son related to the safety or effective- ness of the device.  The  description shall include the history of the mar- keting of the device by the applicant and, if known, the history of the mar- keting of the device by any other per- son.

  - Summary of studies. An abstract of any information or report described in the PMA under paragraph (b)(8)(ii) of this section and a summary of the re- sults of technical data submitted under paragraph (b)(6) of this section. Such summary shall include a description of

the objective of the study, a descrip- tion of the experimental design of the study, a brief description of how the data were collected and analyzed, and a brief description of the results, wheth- er positive, negative, or inconclusive. This section shall include the fol- lowing:

A summary of the nonclinical laboratory studies submitted in the ap- plication;

A summary of the clinical inves- tigations involving human subjects submitted in the application including a discussion of subject selection and exclusion criteria, study population, study period, safety and effectiveness data, adverse reactions and complica- tions, patient discontinuation, patient complaints, device failures and replace- ments, results of statistical analyses of the clinical investigations, contra- indications and precautions for use of the device, and other information from the clinical investigations as appro- priate (any investigation conducted under an IDE shall be identified as such).

  - Conclusions drawn from the stud- ies. A discussion  demonstrating  that the data and information in the appli- cation constitute valid scientific evi- dence within the meaning of § 860.7 and provide reasonable assurance that the device is safe and effective for its in- tended use. A concluding discussion shall present benefit and risk consider- ations related to the device including a discussion of any adverse effects of the device on health and any proposed ad- ditional studies or surveillance the ap- plicant intends to conduct following approval of the PMA.

A complete description of:

The device, including pictorial representations;

Each of the functional compo- nents or ingredients of the device if the device consists of more than one phys- ical component or ingredient;

The properties of the device rel- evant to the diagnosis, treatment, pre- vention, cure, or mitigation of a dis- ease or condition;

The principles of operation of the device; and

The methods used in, and the fa- cilities and controls used for, the man- ufacture,  processing,  packing,  storage,

#### § 814.20

and, where appropriate, installation of the device, in sufficient detail so that a person generally familiar with current good manufacturing practice can make a knowledgeable judgment about the quality control used in the manufac- ture of the device.

Reference to any performance standard under section 514 of the act or under section 534 of Subchapter C— Electronic   Product   Radiation    Control of the Federal Food, Drug,  and  Cos- metic Act  (formerly  the  Radiation Control for Health and Safety Act  of 1968) in effect or  proposed  at  the  time of the submission and to any voluntary standard that is  relevant  to  any  aspect of the safety or effectiveness of the de- vice and that is known to  or  that should reasonably be known to the ap- plicant.  The  applicant  shall—

Provide adequate information to demonstrate how the device meets, or justify any deviation from, any per- formance standard established under section 514 of the act or under section

534 of Subchapter C—Electronic Prod- uct Radiation Control of the Federal Food, Drug, and Cosmetic Act (for- merly the Radiation Control for Health and Safety Act of 1968), and

Explain any deviation from a vol- untary standard.

The following technical sections which shall contain data and informa- tion in sufficient detail to permit FDA to determine whether to approve or deny approval of the application:

A section containing results of the nonclinical laboratory studies with the device including microbiological, toxi- cological, immunological, biocompat- ibility, stress, wear, shelf life, and other laboratory or animal tests as ap- propriate. Information on nonclinical laboratory studies shall include a statement that each such study was conducted in compliance with part 58, or, if the study was not conducted in compliance with such regulations, a brief statement of the reason for the noncompliance.

A section containing results of the clinical investigations involving human subjects with the device includ- ing clinical protocols, number of inves- tigators and subjects per investigator, subject selection and exclusion cri- teria,  study  population,  study  period,

#### 21 CFR Ch. I (4–1–17 Edition)

safety and effectiveness data, adverse reactions and complications, patient discontinuation, patient complaints, device failures and replacements, tab- ulations of data from all individual subject report forms and copies of such forms for each subject who died during a clinical investigation or who did not complete the investigation, results of statistical analyses of the clinical in- vestigations, device failures and re- placements, contraindications and pre- cautions for use of the device, and any other appropriate information from the clinical investigations. Any investiga- tion conducted under an IDE shall be identified as such. Information on clin- ical investigations involving human subjects shall include the following:

  - A statement with respect to each study that it either was conducted in compliance with the institutional re- view board regulations in part 56, or was not subject to the regulations under § 56.104 or § 56.105, and that it was conducted in compliance with the in- formed consent regulations in part 50; or if the study was not conducted in compliance with those regulations, a brief statement of the reason for the noncompliance.

  - A statement that each study was conducted in compliance with part 812 or part 813 concerning sponsors of clin- ical investigations and clinical inves- tigators, or if the study was not con- ducted in compliance with those regu- lations, a brief statement of the reason for the noncompliance.

For a PMA supported solely by data from one investigation, a jus- tification showing that data and other information from a single investigator are sufficient to demonstrate the safe- ty and effectiveness of the device and to ensure reproducibility of test re- sults.

(8)(i) A bibliography of all published reports not submitted under paragraph (b)(6) of this section, whether adverse or supportive, known to or that should reasonably be known to the applicant and that concern the safety or effec- tiveness of the device.

An identification, discussion, and analysis of any other data, informa- tion, or report relevant to an evalua- tion of the safety and effectiveness of

#### Food and Drug Administration, HHS	§ 814.20

the device known to or that should rea- sonably be known to the applicant from any source, foreign or domestic, including information derived from in- vestigations other than those proposed in the application and from commer- cial marketing experience.

Copies of such published reports or unpublished information in the pos- session of or reasonably obtainable by the applicant if an FDA advisory com- mittee or FDA requests.

One or more samples of the device and its components, if  requested  by FDA. If it is impractical to submit a requested sample of the device, the ap- plicant shall  name  the  location  at which FDA may examine  and  test  one or more devices.

Copies of all proposed labeling for the device. Such labeling may in- clude, e.g., instructions for installation and any information, literature, or ad- vertising that constitutes labeling under section 201(m) of the act.

An environmental assessment under § 25.20(n) prepared in the applica- ble format in § 25.40, unless the action qualifies  for  exclusion  under  § 25.30  or

§ 25.34. If the applicant believes that the action qualifies for exclusion, the PMA shall under § 25.15(a) and (d) pro- vide information that establishes to FDA’s satisfaction that the action re- quested is included within the excluded category and meets the criteria for the applicable  exclusion.

A financial certification or dis- closure statement or both as  required by part 54 of this chapter.

Information concerning uses in pe- diatric patients. The application must include the following information, if readily available:

A description of any pediatric sub- populations (neonates, infants, chil- dren, adolescents) that suffer from the disease or condition that the device is intended to  treat,  diagnose,  or  cure; and

The number of affected pediatric patients.

Such other information as FDA may request. If necessary, FDA will ob- tain the concurrence of the appropriate FDA advisory committee before re- questing additional information.

Pertinent information in FDA files specifically referred to by an ap-

plicant may be incorporated into a PMA by reference. Information in a master file or other information sub- mitted to FDA by a person other than the applicant will not be considered part of a PMA unless such reference is authorized in writing by the person who submitted the information or the master file. If a master file is not ref- erenced within 5 years after the date that it is submitted to FDA, FDA will return the master file to the person who submitted it.

If the applicant believes that cer- tain information required under para- graph (b) of this section to be in a PMA is not applicable to the device that is the subject of the PMA, and omits any such information from its PMA, the ap- plicant shall submit a statement that identifies the omitted information and justifies the omission. The statement shall be submitted as a separate sec- tion in the PMA and identified in the table of contents. If the justification for the omission is not accepted by the agency, FDA will so notify the appli- cant.

The applicant shall periodically update its pending application with new safety and effectiveness informa- tion learned about the device from on- going or completed studies that may reasonably affect an evaluation of the safety or effectiveness of the device or that may reasonably affect the state- ment of contraindications, warnings, precautions, and adverse reactions in the draft labeling. The update report shall be consistent with the data re- porting provisions of the protocol. The applicant shall submit three copies of any update report and shall include in the report the number assigned by FDA to the PMA. These updates are consid- ered to be amendments to the PMA. The time frame for review of a  PMA will not be extended due to the submis- sion of an update report unless the up- date   is   a   major   amendment   under

§ 814.37(c)(1). The applicant shall submit these reports—

3 months after the filing date,

Following receipt of an approv- able letter, and

At any other time as requested by FDA.

If a color additive subject to sec- tion 721 of the act is used in or on the

#### § 814.37

device and has not previously been list- ed for such use, then, in lieu of submit- ting a color  additive  petition  under part 71, at the option of the applicant, the information required to be sub- mitted under part 71 may be submitted as part of the PMA. When submitted as part of the PMA, the  information  shall be submitted in three copies each bound in one or more numbered vol- umes of reasonable size. A PMA for a device that contains a color additive that is subject to section 721 of the act will not be approved until the color ad- ditive is listed for use in or on the de- vice.

Additional   information   on   FDA policies   and   procedures,   as   well   as links  to  PMA  guidance  documents,  is available   on   the   Internet   .

If you are sending a PMA, PMA amendment, PMA supplement, or cor- respondence with respect to a PMA, you must send the submission to the appropriate address as follows:

  - For devices regulated by the Cen- ter  for   Devices   and   Radiological Health, Food and Drug Administration, Center for Devices and Radiological Health, Document Mail Center, 10903 New  Hampshire  Ave.,  Bldg.  66,  rm. G609, Silver Spring, MD 20993–0002.

  - For devices regulated by the Cen- ter for Biologics Evaluation and Re- search, send it to: Food and Drug Ad- ministration, Center for Biologics Evaluation and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002.

  - For devices regulated by the Cen- ter for Drug Evaluation and Research, send it to: Central Document Control Room, Center for Drug Evaluation and Research, Food and Drug Administra-

#### 21 CFR Ch. I (4–1–17 Edition)

tion, 5901–B Ammendale Rd., Beltsville, MD 20705–1266.

[51 FR 26364, July 22, 1986; 51 FR 40415, Nov.

7,  1986,  as  amended  at  51  FR  43344,  Dec.  2,

1986; 55 FR 11169, Mar. 27, 1990; 62 FR 40600,

July 29, 1997; 63 FR 5253, Feb. 2, 1998; 65 FR

17137, Mar. 31, 2000; 65 FR 56480, Sept. 19, 2000;

67 FR 9587, Mar. 4, 2002; 71 FR 42048, July 25,

2006;  72  FR  17399,  Apr.  9,  2007;  73  FR  34859,

June 19, 2008; 74 FR 14478, Mar. 31, 2009; 75 FR

20915, Apr. 22, 2010; 78 FR 18233, Mar. 26, 2013;

79 FR 1740, Jan. 10, 2014; 80 FR 18094, Apr. 3,

2015]

##### § 814.37   PMA  amendments  and  resub- mitted PMAs.

(a) An applicant may amend a pend- ing PMA or PMA supplement to revise existing information or provide addi- tional information.

(b)(1) FDA may request the applicant to amend a PMA or PMA supplement with any information regarding the de- vice that is necessary for FDA or the appropriate advisory committee to complete the review of the PMA or PMA supplement.

(2) FDA may request the applicant to amend a PMA or PMA supplement with information concerning pediatric uses as required under §§ 814.20(b)(13) and 814.39(c)(2).

A PMA amendment submitted to FDA shall include the PMA or PMA supplement number assigned to the original submission and, if submitted on the applicant’s own initiative, the reason for submitting the amendment. FDA may extend the time required for its review of the PMA, or PMA supple- ment, as follows:

  - If the applicant on its own initia- tive or at FDA’s request submits a major PMA amendment (e.g., an amendment that contains significant new data from a previously unreported study, significant updated data from a previously reported study, detailed new analyses of previously submitted data, or significant required information previously omitted), the review period may be extended up to 180 days.

  - If an applicant declines to submit a major amendment requested by FDA, the review period may be extended for the number of days that elapse between the date of such request and the date that FDA receives the written response declining to submit the requested amendment.

#### Food and Drug Administration, HHS	§ 814.39

An applicant may on its own ini- tiative withdraw a PMA or PMA sup- plement. If FDA requests an applicant to submit a PMA amendment and a written response to FDA’s request is not received within 180 days of the date of the request, FDA will consider the pending PMA or PMA supplement to be withdrawn voluntarily by the appli- cant.

An applicant may resubmit a PMA or PMA supplement after with- drawing it or after it is considered withdrawn under paragraph (d) of this section, or after FDA has refused to ac- cept it for filing, or has  denied  ap- proval of the PMA or PMA supplement. A resubmitted PMA or PMA  supple- ment shall comply with the require- ments of § 814.20 or § 814.39, respec- tively, and shall include the PMA num- ber assigned to the original submission and the applicant’s reasons for resub- mission of the PMA or PMA supple- ment.

[51 FR 26364, July 22, 1986, as amended at 79

FR 1740, Jan. 10, 2014]

##### § 814.39   PMA supplements.

After FDA’s approval of a  PMA, an applicant shall submit a PMA sup- plement for review and approval by FDA before making a change affecting the safety or effectiveness of the device for which the applicant has an ap- proved PMA, unless the change is of a type  for  which  FDA,  under  paragraph

of this section, has advised that an alternate submission is permitted or is of a type which, under section 515(d)(6)(A) of the act and paragraph (f) of this section, does not require a PMA supplement under this  paragraph. While the burden for determining whether a supplement is required is primarily on the PMA holder, changes for which an applicant shall submit a PMA supplement include, but are not limited to, the following types of changes if they affect the safety or ef- fectiveness of the device:

  - New indications for use of the de- vice.

  - Labeling changes.

  - The use of a different facility or establishment to manufacture, process, or package the device.

  - Changes in sterilization proce- dures.

  - Changes in packaging.

  - Changes in the performance or de- sign specifications, circuits, compo- nents, ingredients, principle of oper- ation, or physical layout of the device.

  - Extension of the expiration date of the device based on data obtained under a new or revised stability or ste- rility testing protocol that has  not been approved by FDA. If the protocol has been approved, the change shall be reported to FDA under paragraph (b) of this section.

An applicant may make  a  change in a device after FDA’s approval of a PMA for the device without submitting a PMA supplement if the change does not affect the device’s safety or effec- tiveness and the change is reported to FDA in postapproval periodic reports required as a condition to approval  of the device, e.g., an editorial change in labeling which does not affect the safe- ty or effectiveness of the device.

(c)(1) All procedures and actions that apply to an application under § 814.20 also apply to PMA supplements except that the information required in a sup- plement is limited to that needed to support  the  change.  A  summary  under

§ 814.20(b)(3) is required for only a sup- plement submitted for new indications for use of the device, significant changes in the performance or design specifications, circuits, components, ingredients, principles of operation, or physical layout of the device, or when otherwise required by FDA. The appli- cant shall submit  three  copies  of  a PMA supplement and shall include in- formation relevant to the proposed changes in the device. A PMA supple- ment shall include a separate section that identifies each change for which approval is being requested and ex- plains the reason for each such change. The applicant shall submit additional copies and additional information if re- quested by FDA. The time frames for review of, and FDA action on, a PMA supplement are the same as those pro- vided in § 814.40 for a PMA.

The supplement must include the following information:

  - Information  concerning   pediatric uses as required under § 814.20(b)(13).

  - If information concerning the de- vice that is the subject of the supple- ment was previously submitted under

#### § 814.39

§ 814.20(b)(13) or under this section in a previous supplement, that information may be included by  referencing a  pre- vious application or submission that contains the information. However, if additional  information  required  under

§ 814.20(b)(13) has become readily avail- able to the applicant since the previous submission, the applicant must submit that information as part of the supple- ment.

(d)(1) After  FDA  approves  a  PMA, any change described in  paragraph (d)(2) of this section to reflect newly acquired   information   that    enhances the safety of the device or  the  safety  in the  use  of  the  device  may  be  placed into effect by the applicant prior to the receipt under § 814.17 of a written FDA order approving the PMA supplement provided  that:

    - The PMA supplement and its mail- ing cover are plainly marked ‘‘Special PMA Supplement—Changes Being Ef- fected’’;

    - The PMA supplement provides a full explanation of the basis for the changes;

    - The applicant has received ac- knowledgement from FDA of receipt of the supplement; and

    - The PMA supplement specifically identifies the date that such changes are being effected.

The following changes are per- mitted by paragraph (d)(1) of this sec- tion:

  - Labeling changes that add or strengthen a contraindication, warn- ing, precaution, or information about an adverse reaction for which there is reasonable evidence of a causal asso- ciation.

  - Labeling changes that add or strengthen an instruction that is in- tended to enhance the safe use of the device.

  - Labeling changes that delete misleading, false, or unsupported indi- cations.

  - Changes in quality controls or manufacturing process that add a new specification or test method, or other- wise provide additional assurance of purity, identity, strength, or  reli- ability of the device.

(e)(1) FDA will identify a change to a device for which an applicant has an approved  PMA  and  for  which  a  PMA

#### 21 CFR Ch. I (4–1–17 Edition)

supplement under paragraph (a) is not required. FDA will identify such a change  in  an  advisory  opinion  under

§ 10.85, if the change applies to a ge- neric type of device, or in correspond- ence to the applicant, if the change ap- plies  only  to   the   applicant’s   device. FDA will require that a change  for which a PMA supplement under para- graph (a) is not required be reported to FDA   in:

A periodic report under § 814.84 or

A 30-day PMA supplement under this  paragraph.

(2) FDA will identify, in the advisory opinion or correspondence, the type of information that is to be included  in the report or 30-day PMA supplement. If the change is required to be reported to FDA in a periodic report, the change may be made before it is reported to FDA. If the change is required to be re- ported in a 30-day PMA supplement, the change may be made 30 days after FDA files the 30-day PMA supplement unless FDA requires the PMA holder to provide additional information, in- forms the PMA holder that the supple- ment is not approvable, or disapproves the supplement. The 30-day PMA sup- plement  shall  follow  the   instructions in the correspondence or advisory opin- ion. Any 30-day PMA supplement that does not meet the requirements of the correspondence or advisory opinion will not be filed and, therefore, will not be deemed approved 30 days after re- ceipt.

Under section  515(d)  of the  act, modifications to manufacturing proce- dures or methods of manufacture that affect the safety and effectiveness of a device subject to an approved PMA do not require submission of a PMA sup- plement under paragraph (a) of this section and are eligible to be the  sub- ject of a 30-day notice. A 30-day notice shall describe in detail the change, summarize the data or information supporting  the  change,  and  state  that the change has been made in accord- ance with the requirements  of  part  820 of this chapter. The manufacturer may distribute the device 30 days after the date on which FDA receives the 30-day notice, unless FDA  notifies  the  appli- cant within 30 days from receipt of the notice that the  notice  is  not  adequate. If  the  notice  is  not  adequate,  FDA  shall

#### Food and Drug Administration, HHS	§ 814.42

inform the applicant in writing that a 135-day PMA supplement is needed and shall describe what further information or action is required  for acceptance  of such change. The number of days under review as a 30-day notice shall be de- ducted from the 135-day PMA supple- ment review period if the notice meets appropriate content requirements for a PMA  supplement.

The submission and grant of a written request for an exception or al- ternative under § 801.128 or § 809.11 of this  chapter  satisfies  the   requirement in  paragraph  (a)  of  this  section.

[51 FR 26364, July 22, 1986, as amended at 51

FR 43344, Dec. 2, 1986; 63 FR 54044, Oct. 8, 1998;

67 FR 9587, Mar. 4, 2002; 69 FR 11313, Mar. 10,

2004; 72 FR 73602, Dec. 28, 2007; 73 FR 49610,

Aug. 22, 2008; 79 FR 1740, Jan. 10, 2014]

## Subpart C—FDA Action on a PMA

##### § 814.40   Time  frames  for  reviewing  a PMA.

Within 180 days after receipt of an application that  is  accepted  for  filing and to which the applicant does not submit a major amendment, FDA will review  the  PMA  and,  after  receiving the report and recommendation of the appropriate FDA advisory  committee, send the applicant an approval order under § 814.44(d), an approvable letter under § 814.44(e), a not approvable letter under § 814.44(f), or an order denying ap- proval under § 814.45. The approvable letter and  the  not  approvable  letter will provide an opportunity for the ap- plicant to amend or withdraw the ap- plication,  or  to  consider  the   letter   to be a denial of approval  of  the  PMA under § 814.45 and to request adminis- trative review under section 515 (d)(3) and (g) of the act.

##### § 814.42   Filing a PMA.

The filing of an application means that FDA has made a threshold deter- mination that the application is suffi- ciently complete to permit a sub- stantive review. Within 45 days after a PMA is received by FDA, the agency will notify the applicant whether the application has been filed.

If FDA does not find that  any  of the reasons in paragraph (e) of this sec- tion for refusing to file the PMA ap- plies, the agency will file the PMA and

will notify the applicant in writing of the filing. The notice will include the PMA reference number and the date FDA filed the PMA. The date of filing is the date that a PMA accepted for fil- ing was received by the agency. The 180-day period for review of a PMA starts on the date of filing.

If FDA refuses to file a PMA, the agency will notify the applicant of the reasons for the refusal. This notice will identify the deficiencies in the applica- tion that prevent filing and will in- clude the PMA reference number.

If FDA refuses to file the  PMA, the applicant may:

  - Resubmit the PMA with addi- tional information necessary to comply with the requirements of section 515(c)(1) (A)–(G) of the act and § 814.20. A resubmitted PMA shall include the PMA reference number of the original submission. If the resubmitted PMA is accepted for  filing, the date  of filing is the date FDA receives the resubmis- sion;

  - Request in writing within 10 working days of the date of receipt of the notice refusing to file the PMA, an informal conference with the Director of the Office of Device Evaluation to review FDA’s decision not to file the PMA. FDA will hold the informal con- ference within 10 working days of its receipt of the request and will render its decision on filing within 5 working days after the informal conference. If, after the informal conference, FDA ac- cepts the PMA for filing, the date of filing will be the date of the decision to accept the PMA for filing. If FDA does not reverse its decision not to file the PMA, the applicant may request recon- sideration of the decision from the Di- rector of the Center for Devices and Radiological Health,  the  Director  of the Center for Biologics  Evaluation and Research, or the Director of the Center for Drug Evaluation and Re- search, as applicable. The Director’s decision will constitute final adminis- trative action for the purpose of judi- cial review.

FDA may refuse to file a PMA if any of the following applies:

  - The application is incomplete be- cause it does not on its face contain all the information required under section 515(c)(1) (A)–(G) of the act;

§ 814.44

  - The PMA does not contain each of the items required under § 814.20 and justification for omission  of  any  item is inadequate;

  - The applicant has a pending pre- market notification under section 510(k) of the act with respect to the same device, and FDA has not deter- mined whether the device falls within the scope of § 814.1(c).

  - The PMA contains a false state- ment of material fact.

  - The PMA is not accompanied by a statement of either certification or dis- closure as required by part 54 of this chapter.

[51 FR 26364, July 22, 1986, as amended at 63

FR  5254,  Feb.  2,  1998;  73  FR  49942,  Aug.  25,

2008]

##### § 814.44   Procedures   for   review   of   a PMA.

    - FDA will begin substantive review of a PMA after the PMA is accepted for filing under § 814.42. FDA may refer the PMA to a panel on its own  initiative, and will do so upon request of an appli- cant, unless FDA determines that the application substantially duplicates in- formation previously reviewed by a panel. If FDA refers an application to a panel, FDA will forward the PMA, or relevant portions thereof, to each member of the appropriate FDA panel for review. During the review process, FDA may communicate with the appli- cant as set forth under § 814.37(b), or with a panel to respond to questions that may be posed by panel members or to provide  additional  information  to the panel. FDA will  maintain  a  record of all communications with the appli- cant and with the panel.

    - The advisory committee shall submit a report to FDA which includes the committee’s recommendation and the basis for such recommendation on the PMA. Before submission of this re- port, the committee shall hold a public meeting to review the PMA in accord- ance with part 14. This meeting may be held  by  a  telephone  conference  under

§ 14.22(g). The advisory committee re- port and recommendation may be in the form of a meeting transcript signed by the chairperson of the committee.

    - FDA will complete its review of the PMA and the advisory committee report and recommendation and, with-

#### 21 CFR Ch. I (4–1–17 Edition)

in the later of 180 days from the date of filing of the PMA under § 814.42 or the number of days after  the  date  of  filing as determined under § 814.37(c), issue an approval order under paragraph (d) of this section, an approvable letter under paragraph (e) of this section, a not ap- provable letter under paragraph (f) of this section, or an order denying ap- proval      of      the      application      under

§ 814.45(a).

(d)(1) FDA will issue to  the  applicant an order approving a PMA if  none  of the reasons in § 814.45 for denying ap- proval of  the  application  applies.  FDA will  approve  an  application  on   the basis of draft final labeling if the only deficiencies in the application concern editorial  or  similar  minor   deficiencies in the draft final labeling. Such ap- proval will be conditioned upon the ap- plicant incorporating the specified la- beling changes exactly as directed and upon the   applicant  submitting   to  FDA a copy of the final printed labeling be- fore marketing. FDA will also give the public notice of the order, including notice of and opportunity for any in- terested persons to request  review under section 515(d)(3) of the act. The notice of approval will be  placed  on FDA’s home page on the Internet,  and  it  will  state that a detailed  summary  of  informa- tion respecting the safety and effec- tiveness of the device, which was the basis for the order approving the PMA, including information about  any  ad- verse effects of the device on health, is available on the Internet and has been placed on public display, and  that  cop- ies are  available  upon  request.  FDA will publish in  the  FEDERAL  REGISTER after each quarter a list  of  the  approv- als announced in that quarter. When a notice of  approval  is  published,  data and information in the PMA file will be available for public disclosure in ac- cordance with § 814.9.

(2) A request for copies of the current PMA approvals and denials document and for copies of summaries of safety and effectiveness shall be sent in writ- ing to the Division of Dockets Manage- ment (HFA–305), Food and Drug Ad- ministration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852.

FDA will send the applicant an approvable   letter   if   the   application

#### Food and Drug Administration, HHS	§ 814.45

substantially meets the requirements of this part and the agency believes it can approve the application if specific additional information is submitted or specific conditions are agreed to by the applicant.

  - The approvable letter will de- scribe the information FDA requires to be provided by the applicant or the conditions the applicant is required to meet to obtain approval. For example, FDA may require, as a condition to ap- proval:

    - The submission of certain infor- mation identified in the approvable let- ter, e.g., final labeling;

    - The submission of additional in- formation concerning pediatric uses re- quired by § 814.20(b)(13);

    - An FDA inspection that finds the manufacturing facilities, methods, and controls in compliance with part 820 and, if applicable, that verifies records pertinent to the PMA;

    - Restrictions imposed on the de- vice under section 515(d)(1)(B)(ii) or 520(e) of the act;

    - Postapproval requirements as de- scribed in subpart E of this part.

  - In response to an approvable let- ter the applicant may:

    - Amend the PMA as requested in the approvable letter; or

    - Consider the approvable letter to be a denial of approval  of  the  PMA under § 814.45 and request administra- tive review under section 515(d)(3) of the act by filing  a  petition  in  the  form of  a  petition  for  reconsideration  under

§ 10.33; or

    - Withdraw the PMA.

FDA will send the applicant a not approvable letter if the agency believes that the application may not be ap- proved for one or more of the reasons given in § 814.45(a). The not approvable letter will describe the deficiencies in the application, including each applica- ble ground for denial under section 515(d)(2) (A)–(E) of the act, and, where practical, will identify measures re- quired to place the PMA in approvable form. In response to a not approvable letter,  the  applicant  may:

  - Amend the PMA as requested in the not approvable letter (such an amendment will be considered a major amendment under § 814.37(c)(1)); or

  - Consider the not approvable  letter to be a denial of approval of the PMA under § 814.45 and request administra- tive review under section 515(d)(3) of the act by filing  a  petition  in  the  form of  a  petition  for  reconsideration  under

§ 10.33; or

  - Withdraw the PMA.

FDA will consider a PMA to have been withdrawn voluntarily if:

  - The applicant fails to respond in writing to a written request for an amendment within 180 days after the date FDA issues such request;

  - The applicant fails to respond in writing to an approvable or not approv- able letter within 180 days after the date FDA issues such letter; or

  - The applicant submits  a written notice to FDA that the PMA has been withdrawn.

[51 FR 26364, July 22, 1986, as amended at 57

FR 58403,  Dec. 10,  1992;  63  FR 4572,  Jan. 30,

1998; 79 FR 1740, Jan. 10, 2014]

##### § 814.45   Denial of approval of a PMA.

FDA may issue an order denying approval of a PMA if the applicant fails to follow the requirements of this part or if, upon the basis of the information submitted in the PMA or any other in- formation before the agency, FDA de- termines that any of the grounds for denying approval of a PMA specified in section 515(d)(2) (A)–(E) of the act ap- plies. In addition, FDA may deny ap- proval of a PMA for any of the fol- lowing reasons:

  - The PMA contains a false state- ment of material fact;

  - The device’s proposed  labeling does not comply with  the  requirements in part 801 or part 809;

  - The applicant does not permit an authorized FDA employee an oppor- tunity to inspect at a reasonable time and in a reasonable manner the facili- ties, controls, and to have access to and to copy and verify all records per- tinent to the application;

  - A nonclinical laboratory study that is described  in  the  PMA  and  that is essential to show that the device is safe for use under the conditions pre- scribed, recommended, or suggested in its proposed labeling, was not con- ducted in compliance with the good laboratory practice regulations in part 58 and no reason for the noncompliance

#### § 814.46

is provided or, if it is, the differences between the practices used in con- ducting the study and the good labora- tory practice regulations do not sup- port the validity of the study; or

  - Any clinical investigation involv- ing human subjects described in the PMA, subject to the institutional re- view board regulations in part 56 or in- formed consent regulations in part 50, was not conducted in compliance with those regulations such that the rights or safety of human subjects were not adequately protected.

FDA will issue any order denying approval of the PMA  in  accordance with § 814.17. The order will inform the applicant of the deficiencies in  the PMA, including each applicable ground for denial under section 515(d)(2) of the act and the regulations under this part, and, where practical, will identify measures required to place the PMA in approvable form. The order will include a notice of an opportunity to request review under section 515(d)(4) of the act.

FDA will use  the  criteria  specified in § 860.7 to determine the safety and effectiveness of a device in deciding whether to approve or deny approval of a  PMA.  FDA  may  use  information other than that submitted by the appli- cant  in  making  such  determination.

(d)(1) FDA will give  the  public  notice of an order denying  approval  of  the PMA. The notice will be placed on the FDA’s home page on the Internet,  and  it  will   state that a detailed  summary  of  informa- tion respecting the safety and effec- tiveness of the device, including infor- mation about any adverse effects of the device on health, is available on the Internet and has been placed on public display and that copies are  available upon request. FDA will publish in the FEDERAL REGISTER after each  quarter  a list of the denials announced in that quarter. When a notice of denial of ap- proval is made publicly available, data and information in the PMA file will be available for public disclosure in ac- cordance with § 814.9.

(2) A request for copies of the current PMA approvals and denials document and copies of summaries of safety and effectiveness shall be sent in writing to the  Freedom  of  Information  Staff’s  ad-

#### 21 CFR Ch. I (4–1–17 Edition)

dress listed on the Agency’s Web site

FDA will issue an order denying approval of a PMA after an approvable or not approvable letter has been sent and the applicant:

  - Submits a requested amendment but any ground for denying approval  of the application under section  515(d)(2) of  the  act  still  applies;  or

  - Notifies FDA in writing that the requested amendment will not be sub- mitted; or

  - Petitions for review under section 515(d)(3) of the act  by  filing  a  petition in the form of a petition for reconsider- ation under § 10.33.

[51 FR 26364, July 22, 1986, as amended at 63

FR 4572,  Jan. 30,  1998;  73  FR 34859,  June 19,

2008;  76  FR 31470,  June 1,  2011;  79  FR  68115,

Nov. 14, 2014]

##### § 814.46   Withdrawal  of  approval  of  a PMA.

FDA may issue an order with- drawing approval of a PMA if, from any information available to the agency, FDA determines that:

Any of the grounds under section 515(e)(1) (A)–(G) of the act applies.

Any postapproval requirement im- posed by the PMA approval order or by regulation has not been met.

A nonclinical laboratory study that is described  in  the  PMA  and  that is essential to show that the device is safe for use under the conditions pre- scribed, recommended, or suggested in its proposed labeling, was not con- ducted in compliance with the good laboratory practice regulations in part 58 and no reason for the noncompliance is provided or, if it is, the differences between the practices used in con- ducting the study and the good labora- tory practice regulations do not sup- port the validity of the study.

Any clinical investigation involv- ing human subjects described in the PMA, subject to the institutional re- view board regulations in part 56 or in- formed consent regulations in part 50, was not conducted in compliance with those regulations such that the rights or safety of human subjects were not adequately protected.

(b)(1) FDA may seek advice on sci- entific   matters   from   any   appropriate

#### Food and Drug Administration, HHS	§ 814.47

FDA advisory committee in deciding whether to withdraw approval of a PMA.

(2) FDA may use information other than that submitted by  the  applicant in deciding whether to withdraw ap- proval of a PMA.

  - Before issuing an order with- drawing approval of a PMA, FDA will issue the holder of the approved appli- cation a notice of opportunity for an informal hearing under part 16.

  - If the applicant does not request a hearing or if  after  the  part  16  hearing is held the agency decides to proceed with the withdrawal, FDA will issue  to the holder of the approved  application an order withdrawing approval of the application. The order will be issued under § 814.17, will state each ground for withdrawing approval, and will in- clude a  notice  of  an opportunity  for  ad- ministrative review under section 515(e)(2) of the act.

  - FDA will give the public notice of an order withdrawing approval of  a PMA. The notice will be published  in the FEDERAL  REGISTER  and  will  state that a detailed summary of informa- tion respecting the safety and effec- tiveness of the device, including infor- mation about any adverse effects of the device on health, has been placed on public display and that copies are available upon request. When  a  notice of withdrawal of approval is published, data and information in the PMA file will be available for  public  disclosure in accordance with § 814.9.

##### § 814.47   Temporary  suspension  of  ap- proval of a PMA.

Scope. (1) This section describes the procedures that FDA will follow in exercising its authority under section 515(e)(3) of the act (21 U.S.C. 360e(e)(3)). This authority applies to the original PMA, as well as any PMA supple- ment(s), for a medical device.

(2) FDA will issue an order tempo- rarily suspending approval of a PMA if FDA determines that there is a reason- able probability that continued dis- tribution of the device would cause se- rious, adverse health consequences or death.

Regulatory hearing. (1) If FDA be- lieves that there is a reasonable prob- ability that the continued distribution

of a device subject to an approved PMA would cause serious, adverse health consequences or death, FDA may ini- tiate and conduct a regulatory hearing to determine whether to issue an order temporarily suspending approval of the PMA.

Any regulatory hearing to deter- mine whether to issue an order tempo- rarily suspending approval of a PMA shall be initiated  and  conducted  by FDA pursuant to part 16 of this  chap- ter. If FDA believes that immediate ac- tion to remove a dangerous device from the market is necessary to protect the public health, the agency may, in ac- cordance with § 16.60(h) of this chapter, waive, suspend, or modify any part 16 procedure pursuant to § 10.19 of this chapter.

FDA shall deem the PMA holder’s failure to request a hearing within the timeframe specified by FDA in the no- tice of opportunity for hearing to be a waiver.

Temporary suspension order. If  the PMA holder does not request a regu- latory hearing or if, after the hearing, and after consideration of the adminis- trative record of the hearing, FDA de- termines that there is a reasonable probability that the continued dis- tribution of a device under an approved PMA would cause serious,  adverse health consequences or death, the agency shall, under the authority of section 515(e)(3) of the act, issue an order to the PMA holder temporarily suspending approval of the PMA.

Permanent withdrawal of  approval of the PMA. If FDA issues an order tem- porarily suspending approval of a PMA, the agency shall proceed expeditiously, but within 60 days, to hold a hearing on whether to permanently withdraw ap- proval of the PMA in accordance with section 515(e)(1) of the act and the pro- cedures set out in § 814.46.

[61 FR 15190, Apr. 5, 1996]

## Subpart D—Administrative Review

[Reserved]

#### § 814.80

Subpart E—Postapproval Requirements

##### § 814.80   General.

A device may not be manufactured, packaged, stored, labeled, distributed, or advertised in a manner that is in- consistent with any conditions to ap- proval specified in the PMA approval order for the device.

##### § 814.82   Postapproval requirements.

FDA may impose postapproval re- quirements in a PMA approval order or by regulation at the time of approval of the PMA or by regulation subse- quent to approval. Postapproval re- quirements may include as a condition to approval of the device:

  - Restriction of the sale,  distribu- tion, or use of the device as provided by section 515(d)(1)(B)(ii) or 520(e) of the act.

  - Continuing evaluation and peri- odic reporting on the safety, effective- ness, and reliability of the  device  for its intended use. FDA will state in the PMA approval order the reason or pur- pose for such requirement and the number of patients to be evaluated and the reports required to be submitted.

  - Prominent display in the labeling of a device and in the advertising  of any restricted device of warnings, haz- ards, or precautions important for the device’s safe and effective use, includ- ing patient information, e.g., informa- tion provided to the patient on alter- native modes of therapy and on risks and benefits associated with the use of the device.

  - Inclusion of identification codes on the device or its labeling, or in the case of an implant, on cards given to patients if necessary to protect the public health.

  - Maintenance of records that will enable the applicant to submit to FDA information needed to trace patients if such information is necessary to pro- tect the public health. Under section 519(a)(4) of the act, FDA will require that the identity of any patient be dis- closed in records maintained under this paragraph only to the extent required for the medical welfare of the indi- vidual, to determine the safety or ef- fectiveness of the device, or to verify a

#### 21 CFR Ch. I (4–1–17 Edition)

record,   report,   or   information   sub- mitted to the agency.

  - Maintenance of records for speci- fied periods of time and organization and indexing  of  records  into  identifi- able files to enable FDA to determine whether there is reasonable assurance of the continued safety and effective- ness of the device.

  - Submission to FDA at intervals specified in the approval order of peri- odic reports containing the  informa- tion required by § 814.84(b).

  - Batch testing of the device.

  - Such other requirements as FDA determines are necessary to provide reasonable assurance, or continued rea- sonable assurance, of the safety and ef- fectiveness of the device.

An applicant shall grant to FDA access to any records and reports re- quired under the provisions of this part, and shall permit authorized FDA employees to copy and verify such records and reports and to inspect at a reasonable time and in a reasonable manner all manufacturing facilities to verify that the device is being manu- factured, stored, labeled, and shipped under approved conditions.

Failure to comply with any post- approval requirement constitutes a ground for withdrawal of approval of a PMA.

(Approved  by  the  Office  of  Management  and Budget under control number 0910–0231)

[51 FR 26364, July 22, 1986, as amended at 51

FR 43344, Dec. 2, 1986]

##### § 814.84   Reports.

The holder of an approved PMA shall comply with the requirements of part 803 and with any other require- ments applicable to the device by other regulations in this subchapter or by order approving the device.

Unless FDA specifies otherwise, any periodic report shall:

  - Identify    changes    described    in

§ 814.39(a)   and   changes   required   to   be reported to FDA under § 814.39(b).

  - Contain a summary and bibliog- raphy of the following information not previously submitted as part of the PMA:

    - Unpublished reports of data from any clinical investigations or nonclin- ical laboratory studies involving the device or related devices and known to

#### Food and Drug Administration, HHS	§ 814.102

or that reasonably should be known to the  applicant.

    - Reports in the scientific lit- erature concerning the device and known to or that reasonably should be known to the applicant. If, after re- viewing the summary and bibliog- raphy, FDA concludes that the agency needs a copy of the unpublished or pub- lished reports, FDA will notify the ap- plicant that copies of  such  reports shall be submitted.

  - Identify  changes  made  pursuant to an exception or alternative granted under § 801.128 or § 809.11 of this chapter.

  - Identify each device identifier currently in use  for  the  device,  and each device identifier  for  the  device that has been discontinued since the previous periodic report. It is not nec- essary to identify any device identifier discontinued prior to December 23, 2013.

[51 FR 26364, July 22, 1986, as amended at 51

FR 43344, Dec. 2, 1986; 67 FR 9587, Mar. 4, 2002;

72  FR 73602,  Dec. 28,  2007;  78  FR 58822,  Sept.

24, 2013]

## Subparts F–G [Reserved]

Subpart H—Humanitarian Use Devices

SOURCE:  61  FR  33244,  June  26,  1996,  unless

otherwise noted.

##### § 814.100   Purpose and scope.

This subpart H implements sec- tions 515A and 520(m) of the act.

The purpose of section 520(m) is, to the extent consistent with the pro- tection of the public health and safety and with ethical standards, to encour- age the discovery and use of devices in- tended to benefit patients in the treat- ment or diagnosis of diseases or condi- tions that affect or are manifested in fewer than 4,000 individuals in the United States per year. This subpart provides procedures for obtaining:

  - HUD designation of a medical de- vice; and

  - Marketing approval for the HUD notwithstanding the absence of reason- able assurance of effectiveness that would otherwise be required under sec- tions 514 and 515 of the act.

Section 515A of the act is intended to ensure the submission of readily available information concerning:

Any pediatric subpopulations (neonates, infants, children, adoles- cents) that suffer from the disease or condition that the device is intended to treat, diagnose, or cure; and

The number of affected pediatric patients.

Although a HUD may  also  have uses that differ from the humanitarian use, applicants seeking approval of any non-HUD use shall submit a PMA as re- quired under § 814.20, or a premarket notification as required under  part  807 of  this  chapter.

Obtaining marketing approval for a HUD involves two steps:

  - Obtaining designation of the de- vice as a HUD from FDA’s Office of Or- phan Products Development, and

  - Submitting an HDE to the Office of Device Evaluation (ODE), Center for Devices and Radiological Health (CDRH), the Center for Biologics Eval- uation and Research (CBER), or the Center for Drug Evaluation and Re- search (CDER), as applicable.

A person granted an exemption under section 520(m) of the act shall submit periodic reports as described in

§ 814.126(b).

FDA may suspend or withdraw ap- proval of an HDE after providing notice and an opportunity for an informal hearing.

[61 FR 33244, June 26, 1996, as amended at 63

FR  59220,  Nov.  3,  1998;  73  FR  49942,  Aug.  25,

2008; 79 FR 1740, Jan. 10, 2014]

##### § 814.102   Designation of HUD status.

Request for designation. Prior to submitting an HDE application, the ap- plicant shall submit a request for HUD designation to FDA’s Office of Orphan Products Development. The request shall contain the following:

A statement that the applicant re- quests HUD designation for a rare dis- ease or condition or a valid subset of a disease or condition which shall be identified with specificity;

The name and address of the ap- plicant, the name of the applicant’s primary contact person and/or resident agent, including title, address, and telephone number;

§ 814.104

A description of the  rare  disease or condition for which the device is to be used, the proposed indication or in- dications for use of the device, and the reasons why such therapy is needed. If the device is proposed for an indication that represents a subset of a common disease or condition, a demonstration that the subset is medically plausible should be included;

A description of the device and a discussion of  the  scientific  rationale for the use of the device for the rare disease or condition; and

Documentation, with appended authoritative references, to dem- onstrate that the device is designed to treat or diagnose a disease or condition that affects or is manifested in fewer than 4,000 people in the United States per year. If the device is for diagnostic purposes, the documentation must demonstrate that fewer than 4,000 pa- tients per year would be subjected to diagnosis by the device in the United States. Authoritative references in- clude literature citations in specialized medical journals, textbooks, special- ized medical society proceedings, or governmental statistics publications. When no such studies or literature ci- tations exist, the  applicant  may  be able to demonstrate the prevalence of the disease or condition in the United States by providing credible conclu- sions from appropriate research or sur- veys.

FDA action. Within 45 days of re- ceipt of a request for HUD designation, FDA will take  one of  the following ac- tions:

  - Approve the request and notify the applicant that the device has been designated as a HUD based on the in- formation submitted;

  - Return the request to the appli- cant pending further review upon sub- mission of  additional information.  This action will ensue if the request is in- complete because it does not on its face contain all of the information required under § 814.102(a). Upon receipt of this additional information, the review pe- riod may be extended up to 45 days; or

  - Disapprove the request for HUD designation based on a substantive re- view of the information  submitted. FDA may disapprove a request for HUD designation if:

#### 21 CFR Ch. I (4–1–17 Edition)

There is insufficient evidence to support the estimate that  the  disease or condition for which the device is de- signed to treat or diagnose affects or is manifested in fewer than 4,000 people in the United States per year;

FDA determines that, for a diag- nostic device, 4,000 or more patients in the United States would  be  subjected to diagnosis using the device per year; or

FDA determines that the patient population defined in the request is not a medically plausible subset of a larger population.

Revocation of designation. FDA may revoke a HUD designation if the agency finds that:

The request for designation con- tained an untrue statement of material fact or omitted material  information; or

Based on the evidence available, the device is not eligible for HUD des- ignation.

Submission. The applicant shall submit two  copies  of  a  completed, dated, and signed request for HUD des- ignation to: Office of Orphan Products Development (HF–35), Food and Drug Administration, 5600 Fishers Lane, Rockville, MD 20857.

##### § 814.104   Original applications.

United States applicant or represent- ative. The applicant or an authorized representative shall sign the HDE. If the applicant does not reside or have a place of business within the United States, the HDE shall be countersigned by an authorized representative resid- ing or maintaining a  place  of  business in the United States and shall  identify the representative’s name and address.

Contents. Unless the applicant jus- tifies an omission in accordance with paragraph (d) of this section, an HDE shall  include:

  - A copy of or reference to the de- termination made by FDA’s Office of Orphan Products Development (in ac- cordance with § 814.102) that the device qualifies  as  a  HUD;

  - An explanation of why the device would not be available unless an HDE were granted and a statement that no comparable device (other than another HUD approved under this subpart or a

#### Food and Drug Administration, HHS	§ 814.106

device under an approved IDE) is avail- able to treat or diagnose the disease or condition. The application also shall contain a discussion of the risks and benefits of currently available devices or alternative forms of treatment in the United States;

  - An explanation of why the prob- able benefit to health from the use of the device outweighs the risk of injury or illness from its use, taking into ac- count the probable risks and benefits of currently available devices or alter- native forms of treatment. Such expla- nation shall include a description, ex- planation, or theory of the underlying disease process or condition, and known or postulated mechanism(s) of action of the device in relation to the disease process or condition;

  - All of the information required to be submitted under § 814.20(b), except that:

    - In lieu of the summaries, conclu- sions, and results from clinical inves- tigations	required	under

§§ 814.20(b)(3)(v)(B),         (b)(3)(vi),         and

(b)(6)(ii), the applicant shall include the summaries, conclusions, and re- sults of all clinical experience or inves- tigations (whether adverse or sup- portive) reasonably obtainable by the applicant that are relevant to an as- sessment of the risks and probable ben- efits of the device; and

    - In addition to the proposed label- ing       requirement       set       forth       in

§ 814.20(b)(10), the labeling shall bear the following statement: Humanitarian Device. Authorized by Federal law for use in the [treatment or diagnosis] of [specify disease or condition]. The ef- fectiveness of this device for this use has not been demonstrated;

  - The amount to be charged for the device and, if the amount is more than

$250, a report by an independent cer- tified public accountant, made in ac- cordance with the Statement on Stand- ards for Attestation established by the American Institute of Certified Public Accountants, or in lieu of such a re- port, an attestation by a responsible individual of  the  organization, verifying that the amount charged does not exceed the costs of the device’s re- search, development, fabrication, and distribution. If the amount charged is

$250 or less, the requirement for a re-

port by an independent certified public accountant or an attestation by a re- sponsible individual of the  organiza- tion is waived; and

  - Information concerning pediatric uses   of   the   device,   as   required   by

§ 814.20(b)(13).

Omission of information. If the ap- plicant believes that certain informa- tion required under paragraph (b) of this section is not applicable to the de- vice that is the subject of the HDE, and omits any such information from its HDE, the applicant shall submit a statement that  identifies  and justifies the omission. The statement shall be submitted as a separate section in the HDE and identified in the table of con- tents. If the justification for the omis- sion is not accepted by the agency, FDA will so notify the applicant.

Address for submissions and cor- respondence. Copies of all original HDEs amendments and supplements, as  well as any correspondence relating to an HDE, must be sent or delivered to the following:

For devices regulated by the Cen- ter for  Devices  and  Radiological Health, send to Document Mail Center, 10903 New Hampshire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993–0002.

For devices regulated by the Cen- ter for Biologics Evaluation and Re- search, send this information to  the Food  and  Drug  Administration,  Center for Biologics Evaluation and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002.

For devices regulated by the Cen- ter for Drug Evaluation and Research, send this information to the Central Document Control  Room,  Center  for Drug  Evaluation  and   Research,   Food and Drug Administration, 5901–B Ammendale Rd., Beltsville, MD 20705– 1266.

[61 FR 33244, June 26, 1996, as amended at 63

FR  59220,  Nov.  3,  1998;  73  FR  49942,  Aug.  25,

2008;  75  FR  20915,  Apr.  22,  2010;  79  FR  1740,

Jan. 10, 2014; 80 FR 18094, Apr. 3, 2015]

##### § 814.106   HDE amendments and resub- mitted HDE’s.

An HDE or HDE supplement may be amended or resubmitted upon an appli- cant’s own initiative,  or at the  request of  FDA,  for  the  same  reasons  and  in

#### § 814.108

the same manner as prescribed for PMA’s in § 814.37, except that the time- frames set forth in § 814.37(c)(1) and (d) do not apply. If FDA requests an HDE applicant to submit an HDE amend- ment, and a written response to FDA’s request is not received within 75 days of the date of the request, FDA will consider the pending HDE or HDE sup- plement to be withdrawn voluntarily by the applicant. Furthermore, if the HDE applicant, on its own initiative or at FDA’s request, submits a major amendment	as	described	in

§ 814.37(c)(1), the review period may be extended up to 75 days.

[63 FR 59220, Nov. 3, 1998]

##### § 814.108   Supplemental applications.

After FDA approval of an  original HDE, an applicant shall submit supple- ments in accordance with the require- ments for PMA’s under § 814.39, except that a request for a new indication  for use of a HUD shall comply with re- quirements set forth in § 814.110. The timeframes for review of, and FDA ac- tion on, an HDE  supplement  are  the same as those provided in § 814.114 for an  HDE.

[63 FR 59220, Nov. 3, 1998]

##### § 814.110   New indications for use.

  - An applicant seeking a new indi- cation for use of a HUD approved under this subpart H shall obtain a new des- ignation of HUD status  in  accordance with § 814.102 and shall submit an origi- nal HDE in accordance with § 814.104.

  - An application for a new indica- tion for use made under § 814.104 may incorporate by reference any  informa- tion  or  data  previously  submitted  to the  agency  under  an  HDE.

##### § 814.112   Filing an HDE.

The filing of an HDE means that FDA has made a threshold determina- tion that the application is sufficiently complete to permit substantive review. Within 30 days from the date an HDE is received by FDA, the agency will no- tify the applicant whether the applica- tion has been filed. FDA may refuse to file an HDE if any of the following ap- plies:

  - The application is incomplete be- cause it does not on its face contain all

#### 21 CFR Ch. I (4–1–17 Edition)

the       information       required       under

§ 814.104(b);

  - FDA determines that there is a comparable device available  (other than another HUD approved under this subpart or a device under an approved IDE) to treat or diagnose the disease or condition for which approval  of  the HUD is being sought; or

  - The application contains an un- true statement of material fact or omits material information.

  - The HDE is not accompanied by a statement of either certification or dis- closure, or both, as required by part 54 of this chapter.

The   provisions    contained   in

§ 814.42(b), (c), and (d) regarding notifi- cation of filing decisions, filing dates, the start of the 75-day review period, and applicant’s options in response to FDA refuse to file decisions shall apply to  HDE’s.

[61 FR 33244, June 26, 1996, as amended at 63

FR 5254, Feb. 2, 1998; 63 FR 59221, Nov. 3, 1998]

##### § 814.114   Timeframes for reviewing an HDE.

Within 75 days after receipt of an HDE that is accepted for filing and to which the applicant does not submit a major amendment, FDA shall send the applicant an approval order, an approv- able letter, a not approvable letter (under § 814.116), or an order denying approval (under § 814.118).

[63 FR 59221, Nov. 3, 1998]

##### § 814.116   Procedures  for  review  of an HDE.

Substantive review. FDA will begin substantive review of an HDE after the HDE   is   accepted   for   filing   under

§ 814.112. FDA may refer an original HDE application to a panel on its own initiative, and shall do so upon the re- quest of an applicant, unless FDA de- termines that the application substan- tially duplicates information pre- viously reviewed by a panel. If the HDE is referred to a panel, the agency shall follow the procedures set forth under

§ 814.44, with the exception that FDA will complete its review of the HDE and the advisory committee report and recommendations within 75 days from receipt  of  an  HDE  that  is  accepted  for

#### Food and Drug Administration, HHS	§ 814.118

filing under § 814.112 or the date of fil- ing as determined under § 814.106, whichever is later. Within the later of these  two  timeframes,  FDA  will   issue an approval order under  paragraph  (b) of this section, an approvable letter under paragraph (c) of this section, a not  approvable  letter  under  paragraph

  - of this section, or an order denying approval   of   the   application   under

§ 814.118(a).

    - Approval order. FDA will issue to the applicant an order approving an HDE if none of the reasons in § 814.118 for denying approval of the application applies. FDA will approve an applica- tion on the basis of draft final labeling if the only deficiencies in the applica- tion concern editorial or similar minor deficiencies in the draft final labeling. Such approval will be conditioned upon the applicant incorporating the speci- fied labeling changes exactly as di- rected and upon the applicant submit- ting to FDA a copy of the final printed labeling before marketing. The  notice of approval of an HDE will be published in the  FEDERAL  REGISTER  in  accord- ance with the rules and  policies appli- cable     to     PMA’s     submitted     under

§ 814.20. Following the issuance of an approval  order,  data  and   information in the HDE file will  be  available  for public    disclosure    in    accordance    with

§ 814.9(b)  through  (h),  as  applicable.

    - Approvable letter. FDA will send the applicant an approvable letter if the application substantially meets the requirements of this subpart and the agency believes it can approve the ap- plication if specific additional informa- tion is submitted or specific conditions are agreed to by the applicant. The ap- provable letter will describe the infor- mation FDA requires to be provided by the applicant or the conditions the ap- plicant is required to meet to obtain approval. For example, FDA may re- quire as a condition to approval:

      - The submission of certain infor- mation identified in the approvable let- ter, e.g., final labeling;

      - The submission of additional in- formation concerning pediatric uses of the device, as required by § 814.20(b)(13);

      - Restrictions imposed on the de- vice under section 520(e) of the act;

      - Postapproval requirements as de- scribed in subpart E of this part; and

      - An FDA inspection that finds the manufacturing facilities, methods, and controls in compliance with part 820 of this chapter and, if applicable, that verifies records pertinent to the HDE.

    - Not approvable letter. FDA will send the applicant a not approvable letter if the agency believes that the application may not be approved for one  or  more  of  the  reasons  given  in

§ 814.118. The not approvable letter will describe the deficiencies in the applica- tion and, where practical, will identify measures required to place the HDE in approvable form. The applicant may respond to the not approvable letter in the same manner as permitted for not approvable   letters   for   PMA’s   under

§ 814.44(f), with the exception that if a major HDE amendment is submitted, the review period may be extended up to 75 days.

    - FDA will consider an HDE to have been withdrawn voluntarily if:

      - The applicant fails to respond in writing to a written request for an amendment within 75 days after the date FDA issues such request;

      - The applicant fails to respond in writing to an approvable or not approv- able letter within 75 days after the date FDA issues such letter; or

      - The applicant submits  a written notice to FDA that the HDE has been withdrawn.

[61 FR 33244, June 26, 1996, as amended at 63

FR  59221,  Nov.  3,  1998;  79  FR  1741,  Jan.  10,

2014]

##### § 814.118   Denial  of  approval  or  with- drawal of approval of an HDE.

        - FDA may deny approval or with- draw approval of an application if the applicant fails  to  meet  the  require- ments of section 520(m) of the act or of this part, or of any condition  of  ap- proval imposed by an  IRB  or  by  FDA, or any postapproval requirements im- posed under § 814.126. In addition, FDA may deny approval or withdraw ap- proval of an application  if,  upon  the basis  of  the  information  submitted  in the HDE or any other information be- fore  the  agency,  FDA  determines  that:

          - There is a lack of a showing of reasonable assurance that the device is safe under the conditions of use pre- scribed, recommended, or suggested in the labeling thereof;

#### § 814.120

          - The device is ineffective under the conditions of use prescribed, rec- ommended, or suggested in the labeling thereof;

          - The applicant has not dem- onstrated that there is a reasonable basis from which to conclude that the probable benefit to health from the use of the device outweighs the risk of in- jury or illness, taking into account the probable risks and benefits of currently available devices or alternative forms of treatment;

          - The application or a report sub- mitted by or on behalf of the applicant contains an untrue statement of mate- rial fact, or omits material informa- tion;

          - The device’s labeling does not comply with the requirements in part 801 or part 809 of this chapter;

          - A nonclinical laboratory study that is described  in  the  HDE  and  that is essential to show that the device is safe for use under the conditions pre- scribed, recommended, or suggested in its proposed labeling, was not con- ducted in compliance with the good laboratory practice regulations in part 58 of this chapter and no reason for the noncompliance is provided or, if it  is, the differences between the practices used in conducting the study and the good laboratory practice regulations do not support the validity of the study;

          - Any clinical investigation involv- ing human subjects described in the HDE, subject to the institutional re- view board regulations in part 56 of this chapter or the informed consent regulations in part 50 of this chapter, was not conducted in compliance with those regulations such that the rights or safety of human subjects were not adequately protected;

          - The applicant does not permit an authorized FDA employee an oppor- tunity to inspect at a reasonable time and in a reasonable manner the facili- ties and controls, and to have access to and to copy and verify all records per- tinent to the application; or

          - The device’s HUD designation should  be  revoked  in  accordance  with

§ 814.102(c).

        - If FDA issues an order denying approval of an application, the agency will comply with the same notice and disclosure    provisions    required    for

#### 21 CFR Ch. I (4–1–17 Edition)

PMA’s under § 814.45(b) and (d), as ap- plicable.

        - FDA will issue an order denying approval of an HDE after an approvable or not approvable letter has been sent and the applicant:

          - Submits a requested amendment but any ground for denying approval of the application under § 814.118(a) still applies;

          - Notifies FDA in writing that the requested amendment will not be sub- mitted; or

          - Petitions for review under section 515(d)(3) of the act  by  filing  a  petition in the form of a petition for reconsider- ation under § 10.33 of this chapter.

        - Before issuing an order with- drawing approval of an HDE, FDA will provide the applicant with notice  and an opportunity for a hearing as re- quired for PMA’s under § 814.46(c) and (d), and will provide the public with no- tice in accordance with § 814.46(e), as applicable.

[61 FR 33244, June 26, 1996, as amended at 63

FR 59221, Nov. 3, 1998]

##### § 814.120   Temporary suspension of ap- proval of an HDE.

An HDE or HDE supplement may be temporarily suspended for the same reasons and in the same manner as pre- scribed for PMA’s in § 814.47.

[63 FR 59221, Nov. 3, 1998]

##### § 814.122   Confidentiality  of  data  and information.

Requirement for disclosure. The ‘‘HDE file’’ includes all data and infor- mation submitted with or referenced in the  HDE,  any  IDE  incorporated  into the HDE, any HDE amendment or sup- plement,  any  report  submitted  under

§ 814.126, any master file, or any other related submission. Any record  in  the HDE file will be available for public disclosure in accordance with the pro- visions of this section and part 20  of this   chapter.

Extent of disclosure. Disclosure by FDA of the existence and  contents  of an HDE file  shall  be  subject  to  the same rules  that  pertain  to  PMA’s under § 814.9(b) through (h), as applica- ble.

#### Food and Drug Administration, HHS	§ 814.126

##### § 814.124   Institutional   Review   Board requirements.

  - IRB approval. The HDE holder is responsible for ensuring that a HUD approved under this subpart is adminis- tered only in facilities having an Insti- tutional Review Board (IRB) con- stituted and acting pursuant to part 56 of this chapter, including continuing review of use of the device. In addition, a HUD may be administered only if such use has been approved by the IRB located at the facility or by a similarly constituted IRB that has agreed to oversee such use and to which the local IRB has deferred in a letter to the HDE holder, signed by the IRB chair or an authorized designee. If, however, a phy- sician in an emergency situation deter- mines that approval from an IRB can- not be obtained in time to prevent seri- ous harm or death to a patient, a HUD may be administered without prior ap- proval by the IRB located at the facil- ity or by a similarly constituted IRB that has agreed to oversee such use. In such an emergency situation, the phy- sician shall, within 5 days after the use of the device, provide written notifica- tion to the chairman of the IRB of such use. Such written notification shall in- clude the identification of the patient involved, the date on which the device was used, and the reason for the use.

  - Withdrawal of IRB approval. A holder of an approved HDE shall notify FDA of any withdrawal of approval for the use of a HUD by a reviewing IRB within 5 working days after being noti- fied of the withdrawal of approval.

[61 FR 33244, June 26, 1996, as amended at 63

FR 59221, Nov. 3, 1998]

##### § 814.126   Postapproval	requirements and reports.

An HDE approved under this sub- part H shall be subject to the post- approval requirements  and reports  set forth under subpart E of this part, as applicable,     with     the     exception     of

§ 814.82(a)(7). In addition, medical de- vice reports submitted to FDA in com- pliance with the requirements of part 803 of this chapter shall also be sub- mitted to the IRB of record.

In addition to the reports identi- fied in paragraph (a) of this  section, the  holder  of  an  approved  HDE  shall

prepare and submit the following com- plete, accurate, and timely reports:

Periodic reports. An HDE applicant is required to submit reports in accord- ance with the approval order. Unless FDA specifies otherwise, any periodic report shall include:

  - An update of the information re- quired under § 814.102(a) in a separately bound volume;

  - An update of the information re- quired under § 814.104(b)(2), (b)(3), and (b)(5);

  - The number of devices that have been shipped or sold since initial mar- keting approval under this subpart H and, if the number shipped or sold ex- ceeds 4,000,  an explanation  and esti- mate of the number of devices used per patient. If a single device is used on multiple patients, the applicant shall submit an estimate of the number of patients treated or diagnosed using the device together with an explanation of the basis for the estimate;

  - Information describing the appli- cant’s clinical experience with the de- vice since the HDE was initially ap- proved. This information shall include safety information that is known or reasonably should be known to the ap- plicant, medical device reports made under part 803 of this chapter, any data generated from the postmarketing studies, and information (whether pub- lished or unpublished) that is known or reasonably expected to be known by the applicant that may affect an eval- uation of the safety of the device or that may affect the statement of con- traindications, warnings, precautions, and adverse reactions in the device’s labeling; and

  - A summary of any  changes  made to the device in accordance with sup- plements submitted under § 814.108. If information provided in the periodic reports,  or  any  other  information  in the possession of FDA,  gives  the  agen- cy  reason  to  believe  that   a   device raises  public  health  concerns  or  that the criteria for  exemption  are  no longer met, the agency may require the HDE holder to submit additional infor- mation to demonstrate continued com- pliance  with  the  HDE  requirements.

Other. An HDE holder shall main- tain records of the names and addresses of  the  facilities  to  which  the  HUD  has

#### Pt. 820

been shipped, correspondence with re- viewing IRB’s, as well as any other in- formation requested by a reviewing IRB or FDA. Such records shall be maintained in accordance with the HDE approval order.

[61 FR 33244, June 26, 1996, as amended at 63

FR  59221,  Nov.  3,  1998,  71  FR  16228,  Mar.  31,

2006]

## PART 820—QUALITY SYSTEM REGULATION

##### Subpart A—General Provisions

Sec.

820.1    Scope.

820.3    Definitions.

820.5    Quality system.

##### Subpart B—Quality System Requirements

820.20     Management responsibility.

820.22   Quality audit.

820.25    Personnel.

##### Subpart C—Design Controls

820.30    Design controls.

##### Subpart D—Document Controls

820.40    Document controls.

##### Subpart E—Purchasing Controls

820.50    Purchasing controls.

##### Subpart F—Identification and Traceability

820.60    Identification.

820.65    Traceability.

##### Subpart G—Production and Process Controls

820.70    Production and process controls.

820.72   Inspection,	measuring,	and	test equipment.

820.75    Process validation.

##### Subpart H—Acceptance Activities

820.80    Receiving,    in-process,    and    finished device   acceptance.

820.86    Acceptance status.

##### Subpart I—Nonconforming Product

820.90    Nonconforming product.

##### Subpart J—Corrective and Preventive Action

820.100    Corrective and preventive action.

21 CFR Ch. I (4–1–17 Edition)

##### Subpart K—Labeling and Packaging Control

820.120   Device labeling.

820.130    Device packaging.

##### Subpart L—Handling, Storage, Distribution, and Installation

820.140    Handling.

820.150  Storage.

820.160   Distribution.

820.170    Installation.

##### Subpart M—Records

  - General requirements.

  - Device master record.

820.184   Device history record.

820.186   Quality system record.

820.198   Complaint files.

##### Subpart N—Servicing

820.200   Servicing.

##### Subpart O—Statistical Techniques

  - Statistical techniques.

AUTHORITY: 21 U.S.C. 351, 352, 360, 360c, 360d,

360e,  360h,  360i,  360j,  360l,  371,  374,  381,  383;  42

U.S.C. 216, 262, 263a, 264.

SOURCE:  61  FR  52654,  Oct.  7,  1996,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 820.1   Scope.

    - Applicability. (1) Current good manufacturing practice (CGMP) re- quirements are set forth in this quality system regulation. The requirements in this part govern the methods used in, and the facilities and controls used for, the design, manufacture, pack- aging, labeling, storage, installation, and servicing of all finished devices in- tended for human use. The require- ments in this part are intended to en- sure that finished devices will be safe and effective and otherwise in compli- ance with the Federal Food, Drug, and Cosmetic Act (the act). This part es- tablishes basic requirements applicable to manufacturers of finished medical devices. If a manufacturer engages in only some operations subject to the re- quirements in this part, and not in oth- ers, that manufacturer need only com- ply with those requirements applicable to the operations in which it is en- gaged. With respect to class I devices,

Food and Drug Administration, HHS	§ 820.1

design controls apply only to those de- vices listed in § 820.30(a)(2). This regula- tion  does  not  apply   to   manufacturers of components or parts of finished de- vices, but such manufacturers are en- couraged to use appropriate  provisions of this regulation as guidance. Manu- facturers of blood and blood compo- nents used for transfusion or for fur- ther manufacturing  are  not  subject  to this part, but  are  subject  to  subchapter F of this chapter.  Manufacturers  of human cells, tissues, and cellular and tissue-based products (HCT/Ps), as de- fined in § 1271.3(d) of this chapter, that are medical devices (subject to pre- market review or notification, or  ex- empt from notification, under an appli- cation submitted under the device pro- visions of the act or under a biological product license application under  sec- tion 351  of  the  Public  Health  Service Act)  are  subject  to  this  part  and  are also subject to the donor-eligibility procedures set forth in part 1271 sub- part C of this chapter and applicable current good tissue practice  procedures in part 1271 subpart  D  of  this  chapter. In the event of a conflict between ap- plicable regulations in part 1271 and in other parts of this chapter, the regula- tion specifically  applicable  to  the  de- vice in question shall supersede  the more  general.

The provisions of this part shall be applicable to any finished device as defined in this part, intended for human use, that is manufactured, im- ported, or offered for import in any State or Territory  of  the  United States, the District of Columbia, or the Commonwealth of Puerto Rico.

In    this    regulation    the    term ‘‘where  appropriate’’  is  used  several times. When a requirement is qualified by ‘‘where appropriate,’’ it is deemed to be ‘‘appropriate’’ unless the manu- facturer   can   document   justification otherwise.  A  requirement  is  ‘‘appro- priate’’   if   nonimplementation   could reasonably be expected to result in the product  not  meeting  its  specified  re- quirements  or  the  manufacturer  not being able to carry out any necessary corrective action.

    - The quality system regulation in this part supplements regulations in other parts of this chapter except where  explicitly  stated  otherwise.  In

the event of a conflict between applica- ble regulations in this part and in other parts of this chapter, the regula- tions specifically applicable to the de- vice in question shall supersede any other generally applicable require- ments.

    - Authority. Part 820 is established and issued under authority of sections 501, 502, 510, 513, 514, 515, 518, 519, 520, 522, 701, 704, 801, 803 of the act (21 U.S.C. 351, 352, 360, 360c, 360d, 360e, 360h, 360i, 360j, 360l, 371, 374, 381, 383). The failure to comply with  any  applicable  provi- sion in this part renders a device adul- terated under section 501(h) of the act. Such a device, as well as any person re- sponsible for the failure to comply, is subject  to  regulatory  action.

    - Foreign manufacturers. If a manu- facturer who offers devices for import into the United States refuses to per- mit or allow the completion of a Food and Drug Administration (FDA) inspec- tion of the foreign facility for the pur- pose of determining compliance with this part, it shall  appear  for  purposes of section 801(a) of the act, that the methods used in, and the facilities and controls used for, the design, manufac- ture, packaging, labeling, storage, in- stallation, or servicing of any devices produced at such facility that are of- fered for import into the United States do not conform to the requirements of section 520(f) of the act and this  part and that the devices manufactured at that facility are adulterated under sec- tion 501(h) of the act.

    - Exemptions or variances. (1) Any person who wishes to petition for an exemption or variance from any device quality system requirement is  subject to the requirements of section 520(f)(2) of the act. Petitions  for  an  exemption or variance  shall be  submitted accord- ing to the procedures set forth in § 10.30 of this chapter, the FDA’s administra- tive procedures. Guidance is available from the Food and Drug Administra- tion, Center for Devices and Radio- logical Health, Division of Small Man- ufacturers, International and  Con- sumer Assistance, 10903 New Hampshire Ave., Bldg. 66, rm. 4613, Silver  Spring, MD 20993–0002, 1–800–638–2041 or 301–796– 7100, FAX: 301–847–8149.

#### § 820.3

FDA may initiate and grant a variance from any device quality sys- tem requirement when the agency de- termines that such variance is in the best interest of the public health. Such variance will remain in effect only so long as there remains a public health need for the device and the  device would not likely be made sufficiently available  without  the  variance.

[61 FR 52654, Oct. 7, 1996, as amended at 65 FR

17136, Mar. 31, 2000; 65 FR 66636, Nov. 7, 2000;

69 FR 29829, May 25, 2005; 72 FR 17399, Apr. 9,

2007; 75 FR 20915, Apr. 22, 2010; 80 FR 29906,

May 22, 2015]

##### § 820.3   Definitions.

  - Act means  the  Federal  Food, Drug, and Cosmetic Act, as amended (secs. 201–903, 52 Stat. 1040 et seq., as amended (21 U.S.C. 321–394)). All defini- tions in section 201  of  the  act  shall apply  to  the  regulations  in  this  part.

  - Complaint means any written, electronic, or oral communication that alleges deficiencies related to the iden- tity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for dis- tribution.

  - Component means any raw mate- rial, substance, piece, part, software, firmware, labeling, or assembly which is intended to be included as part of the finished, packaged, and labeled device.

  - Control number means any distinc- tive symbols, such as a distinctive combination of letters or numbers, or both, from which the history of the manufacturing, packaging, labeling, and distribution of a unit, lot, or batch of finished devices can be determined.

  - Design history file (DHF) means a compilation of records which describes the design history of a finished device.

  - Design input means the physical and performance requirements of a de- vice that are used as a basis for device design.

  - Design output means the results of a design effort at each design phase and at the end of the total design  effort. The finished design output is the basis for the device master record. The total finished design output consists of the device, its packaging and labeling, and the device master record.

  - Design review means a docu- mented, comprehensive,  systematic  ex-

#### 21 CFR Ch. I (4–1–17 Edition)

amination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems.

Device history record (DHR) means a compilation of records containing the production history of a finished device.

Device master record (DMR) means a compilation of records containing the procedures and specifications for a fin- ished device.

Establish means define, document (in writing or electronically), and im- plement.

Finished device means any device or accessory to any device that is suit- able for use or capable of functioning, whether or not it is packaged, labeled, or sterilized.

Lot or batch means one or more components or finished devices that consist of a single type, model, class, size, composition, or software version that are manufactured under essen- tially the same conditions and that are intended to have uniform characteris- tics and quality within specified limits.

Management with executive respon- sibility means those senior employees of a manufacturer who have the authority to establish or make changes to the manufacturer’s quality policy and quality system.

Manufacturer means any person who designs, manufactures, fabricates, assembles, or processes a finished de- vice. Manufacturer includes but is not limited to those who perform the func- tions of contract sterilization, installa- tion, relabeling, remanufacturing, re- packing, or specification development, and initial distributors of foreign enti- ties performing these functions.

Manufacturing material means any material or substance used in or used to facilitate the manufacturing proc- ess, a concomitant constituent, or a byproduct constituent produced during the manufacturing process, which is present in or on the finished device as a residue or impurity not by design or intent of the manufacturer.

Nonconformity means the non- fulfillment of a specified requirement.

Product means components, manu- facturing materials,  in- process  de- vices, finished devices, and returned devices.

#### Food and Drug Administration, HHS	§ 820.20

Quality means the totality of fea- tures and characteristics that bear on the ability of a device to satisfy fit- ness-for-use, including safety and per- formance.

Quality audit means a systematic, independent examination of a manufac- turer’s quality system that is per- formed at defined intervals and at suf- ficient frequency to determine whether both quality system activities and the results of such activities comply with quality system procedures, that these procedures are implemented effec- tively, and that these procedures are suitable to achieve quality system ob- jectives.

Quality policy means the overall intentions and direction of an organi- zation with respect to quality, as es- tablished by management with execu- tive responsibility.

  - Quality system means the organi- zational structure, responsibilities, procedures, processes, and resources for implementing quality management.

  - Remanufacturer means any person who processes, conditions, renovates, repackages, restores, or does any other act to a finished device that signifi- cantly changes the finished device’s performance or safety specifications, or intended use.

  - Rework means action taken on a nonconforming product so that it will fulfill the specified DMR requirements before it is released for distribution.

  - Specification means any require- ment with which a product, process, service, or other activity must con- form.

  - Validation means confirmation by examination and provision of objective evidence that the particular require- ments for a specific intended use can be consistently fulfilled.

Process validation means estab- lishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications.

Design validation means estab- lishing by objective evidence that de- vice specifications conform with user needs and intended use(s).

(aa) Verification means confirmation by examination and provision of objec- tive evidence that specified require- ments have been fulfilled.

(bb) Human cell, tissue, or cellular or tissue-based product (HCT/P) regulated as a device means an HCT/P as defined in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

(cc) Unique device identifier  (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A unique  device  identifier  is  composed  of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For  an  HCT/P  regulated  as  a  de- vice,    the    distinct    identification    code required by § 1271.290(c) of this chapter. (dd)    Universal   product   code   (UPC) means   the   product   identifier   used   to identify   an   item   sold   at   retail   in   the

United States.

[61 FR 52654, Oct. 7, 1996, as amended at 78 FR

58822, Sept. 24, 2013]

##### § 820.5   Quality system.

Each manufacturer shall establish and maintain a quality system that is appropriate for the specific medical de- vice(s) designed or manufactured, and that meets the requirements of this part.

## Subpart B—Quality System Requirements

##### § 820.20   Management responsibility.

Quality policy. Management with executive responsibility shall establish its policy and objectives for, and com- mitment to, quality. Management with executive responsibility shall ensure that  the  quality  policy  is  understood,

§ 820.22

implemented,  and  maintained  at  all levels of the organization.

Organization. Each manufacturer shall establish and maintain an ade- quate organizational structure to en- sure that devices are designed and pro- duced in accordance with the require- ments of this part.

  - Responsibility and authority. Each manufacturer shall establish the appro- priate responsibility, authority, and interrelation of all personnel who man- age, perform, and assess work affecting quality, and provide the independence and authority necessary to perform these tasks.

  - Resources. Each manufacturer shall provide adequate resources, in- cluding the assignment of trained per- sonnel, for management, performance of work, and assessment activities, in- cluding internal quality audits, to meet the requirements of this part.

  - Management representative. Man- agement with executive responsibility shall appoint, and document such ap- pointment of, a member of  manage- ment who, irrespective of other respon- sibilities, shall have established au- thority over and responsibility for:

    - Ensuring that quality  system re- quirements are effectively established and effectively maintained in accord- ance with this part; and

    - Reporting on the performance of the quality system to management with executive responsibility for re- view.

Management review. Management with executive responsibility shall re- view the suitability and  effectiveness of the quality system at defined inter- vals and with sufficient frequency ac- cording to established procedures to ensure that the quality system satis- fies the requirements of this part and the manufacturer’s established quality policy and objectives. The dates and re- sults of quality system reviews shall be documented.

Quality planning. Each manufac- turer shall establish a quality plan which defines the quality practices, re- sources, and activities relevant to de- vices that are designed and manufac- tured. The manufacturer shall estab- lish how the requirements for quality will be met.

#### 21 CFR Ch. I (4–1–17 Edition)

Quality system procedures. Each manufacturer shall establish quality system procedures and instructions. An outline of the structure of the docu- mentation used in the quality system shall be established where appropriate.

##### § 820.22   Quality audit.

Each manufacturer shall establish procedures for quality audits and con- duct such audits to assure that the quality system is in compliance with the established quality system require- ments and to determine the effective- ness of the quality system. Quality au- dits shall be conducted by individuals who do not have direct responsibility for the matters being audited. Correc- tive action(s), including a reaudit of deficient matters, shall be taken when necessary. A report of the results of each quality audit, and reaudit(s) where taken, shall be made and such reports shall be reviewed by manage- ment having responsibility for the matters audited. The dates and results of quality audits and reaudits shall be documented.

##### § 820.25   Personnel.

General. Each manufacturer shall have sufficient personnel with the nec- essary education, background, train- ing, and experience to assure that all activities required by this part are cor- rectly performed.

Training. Each manufacturer shall establish procedures for identifying training needs and ensure that all per- sonnel are trained to adequately per- form their assigned responsibilities. Training shall be documented.

  - As part of their training, per- sonnel shall be made aware of device defects which may occur from the im- proper performance of their specific jobs.

  - Personnel	who	perform verification and validation  activities shall be made aware of defects and er- rors that may be  encountered  as  part of their job functions.

## Subpart C—Design Controls

##### § 820.30   Design controls.

General. (1) Each manufacturer of any class III or class II device, and the class I devices listed in paragraph (a)(2)

Food and Drug Administration, HHS	§ 820.30

of this section, shall establish and maintain procedures to control the de- sign of the device in order to ensure that specified design requirements are met.

    - The following class I devices are subject to design controls:

      - Devices automated with computer software; and

      - The devices listed in the fol- lowing chart.

| Section | Device |
| --- | --- |
| 868.6810 ........ | Catheter, Tracheobronchial Suction. |
| 878.4460 ........ | Glove, Surgeon’s. |
| 880.6760 ........ | Restraint, Protective. |
| 892.5650 ........ | System, Applicator, Radionuclide, Manual. |
| 892.5740 ........ | Source, Radionuclide Teletherapy. |

Design and development planning. Each manufacturer shall establish and maintain plans that describe or ref- erence the design and development ac- tivities and define responsibility for implementation. The plans shall iden- tify and describe the interfaces with different groups or activities that pro- vide, or result in, input to the design and development process. The plans shall be reviewed, updated, and ap- proved as design and development evolves.

Design input. Each manufacturer shall establish and maintain proce- dures to ensure that the design require- ments relating to a device are appro- priate and address the intended use of the device, including the needs of the user and patient. The procedures shall include a mechanism for addressing in- complete, ambiguous, or conflicting re- quirements. The design input require- ments shall be documented and shall be reviewed and approved by a designated individual(s). The approval,  including the date and signature of the indi- vidual(s) approving the requirements, shall be documented.

Design output. Each manufacturer shall establish and maintain proce- dures for defining and documenting de- sign output in terms that allow an ade- quate evaluation of conformance to de- sign input requirements. Design output procedures shall contain or make ref- erence to  acceptance  criteria and shall ensure that those design outputs that are essential for the proper functioning of the device are identified. Design out- put shall be documented, reviewed, and

approved before release. The approval, including the date and signature of the individual(s) approving the  output, shall be documented.

Design review. Each manufacturer shall establish and maintain proce- dures to ensure that formal docu- mented reviews of the design results are planned and conducted at appro- priate stages of the device’s design de- velopment. The procedures shall ensure that participants at each design review include representatives of all functions concerned with  the  design stage  being reviewed and an individual(s) who does not have direct responsibility for the design stage being reviewed, as well as any specialists needed. The results of a design review, including identification of the design, the date, and the indi- vidual(s) performing the review, shall be documented in the design history file (the DHF).

Design verification. Each manufac- turer shall establish and maintain pro- cedures for verifying the device design. Design verification shall confirm that the design output meets the  design input requirements. The results of the design verification, including identi- fication of the design, method(s), the date, and the individual(s) performing the verification, shall be  documented in the DHF.

Design validation. Each manufac- turer shall establish and maintain pro- cedures for validating the device de- sign. Design validation shall be per- formed under defined operating condi- tions on initial production units,  lots, or batches, or their equivalents. Design validation shall ensure that devices conform to defined user needs and in- tended uses and shall include testing of production units under actual or simu- lated use conditions. Design validation shall include software validation and risk analysis, where appropriate. The results of the design validation, includ- ing identification of the design, meth- od(s), the date, and the individual(s) performing the validation, shall be doc- umented in the DHF.

Design transfer. Each manufac- turer shall establish and maintain pro- cedures to ensure that the device de- sign is correctly translated into pro- duction specifications.

#### § 820.40

  - Design changes. Each manufac- turer shall establish and maintain pro- cedures for the identification, docu- mentation, validation or where appro- priate verification, review, and ap- proval of design changes before their implementation.

(j) Design history file. Each manufac- turer shall establish and  maintain  a DHF for each type of device. The DHF shall contain or reference the records necessary to demonstrate that the de- sign was developed in accordance with the approved design plan and the re- quirements of this part.

## Subpart D—Document Controls

##### § 820.40   Document controls.

Each manufacturer shall establish and maintain procedures to control all documents that are required by this part. The procedures shall provide for the following:

Document approval and distribu- tion. Each manufacturer shall des- ignate an individual(s) to review for adequacy and approve prior to issuance all documents established to meet the requirements of this part. The ap- proval, including the date and signa- ture of the individual(s) approving the document, shall be documented. Docu- ments established to meet the require- ments of this part shall be available at all locations for which they are des- ignated, used, or otherwise necessary, and all obsolete documents shall be promptly removed from all points of use or otherwise prevented from unin- tended use.

Document changes. Changes to doc- uments shall be reviewed and approved by an individual(s) in the same func- tion or organization that performed the original review and approval, un- less specifically designated otherwise. Approved changes shall be commu- nicated to the appropriate personnel in a timely manner. Each manufacturer shall maintain records of changes to documents. Change records shall in- clude a description of the change, iden- tification of the affected documents, the signature of the approving indi- vidual(s), the approval date, and when the change becomes effective.

21 CFR Ch. I (4–1–17 Edition)

Subpart E—Purchasing Controls

##### § 820.50   Purchasing controls.

Each  manufacturer  shall  establish and maintain procedures to ensure that all purchased or otherwise received product and services conform to speci- fied requirements.

Evaluation of suppliers, contractors, and consultants. Each manufacturer shall establish and maintain the re- quirements, including quality require- ments, that must be met by suppliers, contractors, and consultants. Each manufacturer shall:

  - Evaluate and select potential sup- pliers, contractors, and consultants on the basis of their ability to meet speci- fied requirements, including quality re- quirements. The evaluation shall be documented.

  - Define the type and extent of con- trol to be exercised over the product, services, suppliers, contractors, and consultants, based on the evaluation results.

  - Establish and maintain records of acceptable suppliers, contractors, and consultants.

Purchasing data. Each manufac- turer shall establish and maintain data that clearly describe or reference the specified requirements,  including  qual- ity requirements, for purchased or oth- erwise received product and services. Purchasing documents  shall  include, where possible, an agreement that the suppliers, contractors, and consultants agree to notify the manufacturer of changes in  the  product  or  service  so that manufacturers may determine whether the changes may affect the quality of a finished device. Purchasing data shall be approved  in  accordance with § 820.40.

## Subpart F—Identification and Traceability

##### § 820.60   Identification.

Each manufacturer shall establish and maintain procedures for identi- fying product during all  stages of re- ceipt, production, distribution, and in- stallation to prevent mixups.

Food and Drug Administration, HHS	§ 820.70

##### § 820.65   Traceability.

Each manufacturer of a device that is intended for surgical implant into the body or to support or sustain life and whose failure to perform when properly used in accordance with instructions for use provided in the labeling can be reasonably expected to result in a sig- nificant injury to the user shall estab- lish and maintain procedures for iden- tifying with a control number each unit, lot, or batch of finished devices and where appropriate components. The procedures shall facilitate correc- tive action. Such identification shall be documented in the DHR.

## Subpart G—Production and Process Controls

##### § 820.70   Production  and  process  con- trols.

General. Each manufacturer shall develop, conduct, control, and monitor production processes to ensure that a device conforms to its specifications. Where deviations from device specifica- tions could occur as a result of the manufacturing process, the manufac- turer shall establish and maintain process control procedures that de- scribe any process controls necessary to ensure conformance to specifica- tions. Where process controls are need- ed they shall include:

  - Documented instructions, stand- ard operating procedures (SOP’s), and methods that define and control the manner of production;

  - Monitoring and control of process parameters and component and device characteristics during production;

  - Compliance with specified ref- erence standards or codes;

  - The approval of processes and process equipment; and

  - Criteria for workmanship which shall be expressed in documented standards or by means of identified and approved representative samples.

Production and process changes. Each manufacturer shall establish and maintain procedures for changes to a specification, method, process, or pro- cedure. Such changes shall  be  verified or where appropriate validated accord- ing to § 820.75, before implementation and    these    activities    shall    be    docu-

mented.  Changes  shall  be  approved  in accordance with § 820.40.

Environmental control. Where envi- ronmental conditions could reasonably be expected to have an adverse effect on product quality, the manufacturer shall establish and maintain proce- dures to adequately control these envi- ronmental conditions. Environmental control system(s) shall be periodically inspected to verify that the system, in- cluding necessary equipment, is ade- quate and functioning properly. These activities shall be documented and re- viewed.

Personnel. Each manufacturer shall establish and maintain require- ments for the health, cleanliness, per- sonal practices, and clothing of per- sonnel if contact between such per- sonnel and product or environment could reasonably be expected to have an adverse effect on product quality. The manufacturer shall ensure that maintenance and other personnel who are required to work temporarily under special environmental conditions are appropriately trained or supervised by a trained individual.

Contamination control. Each manu- facturer shall establish and maintain procedures to prevent contamination of equipment or product by  substances that could reasonably be expected to have an adverse effect on product qual- ity.

Buildings. Buildings shall be of suitable design and contain sufficient space to perform necessary operations, prevent mixups, and assure orderly handling.

Equipment. Each manufacturer shall ensure that all equipment used in the manufacturing process meets speci- fied requirements and is appropriately designed, constructed, placed, and in- stalled to facilitate maintenance, ad- justment, cleaning, and use.

Maintenance schedule. Each manu- facturer shall establish and maintain schedules for the adjustment, cleaning, and other maintenance of equipment to ensure that manufacturing specifica- tions are met. Maintenance activities, including the date and individual(s) performing the maintenance activities, shall be documented.

Inspection. Each manufacturer shall  conduct  periodic  inspections  in

§ 820.72

accordance with established procedures to ensure adherence to applicable equipment maintenance schedules. The inspections, including the date and in- dividual(s) conducting the inspections, shall be documented.

Adjustment. Each manufacturer shall ensure that any inherent limita- tions or allowable tolerances are visi- bly posted on or near equipment re- quiring periodic adjustments or are readily available to personnel per- forming these adjustments.

Manufacturing material. Where a manufacturing material could reason- ably be expected to have an adverse ef- fect on product quality, the manufac- turer shall establish and maintain pro- cedures for the use and removal of such manufacturing material to ensure that it is removed or limited to an amount that does not adversely affect the de- vice’s quality. The removal or reduc- tion of such manufacturing material shall be documented.

Automated processes. When com- puters or automated data processing systems are used as part of production or the quality system, the  manufac- turer shall validate computer software for its intended use according to an es- tablished  protocol.  All  software changes shall be validated before ap- proval and issuance. These validation activities and results shall be docu- mented.

##### § 820.72   Inspection,	measuring,	and test equipment.

Control of inspection, measuring, and test equipment. Each manufacturer shall ensure that all inspection, meas- uring, and test equipment, including mechanical, automated, or electronic inspection and test equipment, is suit- able for its intended purposes and is ca- pable of producing valid results. Each manufacturer shall establish and main- tain procedures to ensure that equip- ment is routinely calibrated, inspected, checked, and maintained. The proce- dures shall include provisions for han- dling, preservation, and storage of equipment, so that its accuracy and fitness for use  are maintained.  These activities shall be documented.

Calibration. Calibration proce- dures shall include specific directions and limits for accuracy and precision.

#### 21 CFR Ch. I (4–1–17 Edition)

When accuracy and precision limits are not met, there shall be provisions for remedial action to reestablish the lim- its and to evaluate whether there was any adverse effect on the device’s qual- ity. These activities shall be docu- mented.

Calibration standards. Calibration standards used for inspection, meas- uring, and test equipment shall be traceable to national or international standards. If national or international standards are not practical or avail- able, the manufacturer shall use an independent reproducible standard. If no applicable standard exists, the man- ufacturer shall establish and maintain an in-house standard.

Calibration records. The equipment identification, calibration dates, the individual performing each calibration, and the next calibration date shall be documented. These records shall be dis- played on or near each piece of equip- ment or shall be readily available  to the personnel using  such  equipment and to the individuals responsible for calibrating  the  equipment.

##### § 820.75   Process validation.

Where the results of a process cannot be fully verified by subsequent inspection and test, the process  shall be validated with a high degree of as- surance and approved according to es- tablished procedures. The validation activities and results, including the date and signature of the individual(s) approving the validation and where ap- propriate the major equipment vali- dated, shall be documented.

Each manufacturer shall establish and maintain procedures for moni- toring and control of process param- eters for validated processes to ensure that the specified requirements con- tinue to be met.

  - Each manufacturer shall ensure that validated processes are performed by qualified individual(s).

  - For validated processes, the moni- toring and control methods and data, the date performed, and, where appro- priate, the individual(s) performing the process or the major equipment used shall be documented.

When changes or process devi- ations occur, the manufacturer shall review  and  evaluate  the  process  and

#### Food and Drug Administration, HHS	§ 820.90

perform revalidation where appro- priate. These activities shall be docu- mented.

## Subpart H—Acceptance Activities

##### § 820.80   Receiving, in-process, and fin- ished device acceptance.

General. Each manufacturer shall establish and maintain procedures for acceptance activities. Acceptance ac- tivities include inspections, tests, or other verification activities.

Receiving acceptance activities. Each manufacturer shall establish and maintain procedures for acceptance of incoming product. Incoming product shall be inspected, tested, or otherwise verified as conforming to specified re- quirements. Acceptance or rejection shall be documented.

In-process acceptance activities. Each manufacturer shall establish and maintain acceptance procedures, where appropriate, to ensure that specified requirements for in-process product are met. Such procedures shall ensure that in-process product is controlled until the required inspection and tests or other verification activities have been completed, or necessary approvals are received, and are documented.

Final acceptance activities. Each manufacturer shall establish and main- tain procedures for finished device ac- ceptance to ensure that each produc- tion run, lot, or batch of finished de- vices meets acceptance criteria. Fin- ished devices shall be held in quar- antine or otherwise adequately con- trolled until released. Finished devices shall not be released for distribution until:

  - The activities required in the DMR are completed;

  - the associated data and docu- mentation is reviewed;

  - the release is authorized by the signature of a designated individual(s); and

  - the authorization is dated.

Acceptance records. Each manufac- turer shall document acceptance ac- tivities required by this part. These records shall include:

The acceptance activities per- formed;

the dates acceptance  activities are performed;

the results;

the signature of the individual(s) conducting the acceptance activities; and

where appropriate the equipment used. These records shall be part of the DHR.

##### § 820.86   Acceptance status.

Each manufacturer shall identify by suitable means the acceptance  status of product, to indicate the conformance or nonconformance of product with ac- ceptance criteria. The identification of acceptance status shall be maintained throughout manufacturing, packaging, labeling, installation, and servicing of the product to ensure that only prod- uct which has passed the required ac- ceptance activities is distributed, used, or installed.

## Subpart I—Nonconforming Product

##### § 820.90   Nonconforming product.

  - Control of nonconforming product. Each manufacturer shall establish and maintain procedures to control product that does not conform to specified re- quirements. The procedures shall ad- dress the identification,  documenta- tion, evaluation, segregation, and dis- position of  nonconforming  product. The evaluation of  nonconformance shall include a determination of  the need for an investigation and notifica- tion of the persons or organizations re- sponsible for the nonconformance. The evaluation and any investigation shall be documented.

  - Nonconformity review and disposi- tion. (1) Each manufacturer shall estab- lish and maintain procedures that de- fine the responsibility for review and the authority for the disposition of nonconforming product. The proce- dures shall set forth the review and dis- position process. Disposition of non- conforming product shall be docu- mented. Documentation shall  include the justification for use of noncon- forming product and the signature of the individual(s) authorizing the use.

Each manufacturer shall establish and maintain procedures for rework, to include retesting and reevaluation of the nonconforming product after re- work, to ensure that the product meets

§ 820.100

its current approved specifications. Re- work and reevaluation activities, in- cluding a determination of any adverse effect from the rework upon the prod- uct, shall be documented in the DHR.

## Subpart J—Corrective and Preventive Action

##### § 820.100   Corrective and preventive ac- tion.

  - Each manufacturer shall establish and maintain procedures for imple- menting corrective and preventive ac- tion. The procedures shall include re- quirements for:

    - Analyzing processes, work oper- ations, concessions, quality audit re- ports, quality records, service records, complaints, returned  product,  and other sources of quality data to iden- tify existing and potential causes of nonconforming product, or other qual- ity problems. Appropriate statistical methodology shall be employed where necessary to detect  recurring quality problems;

    - Investigating the cause of nonconformities relating to product, processes, and the quality system;

    - Identifying the action(s) needed to correct and prevent recurrence of non- conforming product and other quality problems;

    - Verifying or validating the correc- tive and preventive action to ensure that such action is effective and does not adversely affect the  finished  de- vice;

    - Implementing and recording changes in methods and procedures needed to correct and prevent identi- fied quality problems;

    - Ensuring that information related to quality problems or nonconforming product is disseminated to those di- rectly responsible for assuring the quality of such product or the preven- tion of such problems; and

    - Submitting relevant information on identified quality problems, as well as corrective and preventive actions, for management review.

  - All activities required under this section, and their results, shall be doc- umented.

21 CFR Ch. I (4–1–17 Edition)

Subpart K—Labeling and Packaging Control

##### § 820.120   Device labeling.

Each manufacturer shall establish and maintain procedures to control la- beling  activities.

Label integrity. Labels shall be printed and applied so as to remain leg- ible and affixed during the customary conditions of processing, storage, han- dling, distribution, and where appro- priate use.

Labeling inspection. Labeling shall not be released for storage or use until a designated individual(s) has exam- ined the labeling for accuracy includ- ing, where applicable, the correct unique device identifier (UDI) or uni- versal product code (UPC), expiration date, control number, storage instruc- tions, handling instructions, and any additional processing instructions. The release, including the date and signa- ture of the individual(s) performing the examination, shall  be  documented  in the  DHR.

Labeling storage. Each manufac- turer shall store labeling in a manner that provides proper identification and is designed to prevent mixups.

Labeling operations. Each manu- facturer shall control labeling and packaging operations to prevent label- ing mixups. The label and labeling used for each production unit, lot, or batch shall be documented in the DHR.

Control number. Where a control number is required by § 820.65, that con- trol number shall be on or shall accom- pany  the  device  through  distribution.

[61 FR 52654, Oct. 7, 1996, as amended at 78 FR

58822, Sept. 24, 2013]

##### § 820.130   Device packaging.

Each manufacturer shall ensure that device packaging and shipping con- tainers are designed and constructed to protect the device from alteration or damage during the customary condi- tions of processing, storage, handling, and distribution.

#### Food and Drug Administration, HHS	§ 820.180

Subpart L—Handling, Storage, Distribution, and Installation

##### § 820.140   Handling.

Each manufacturer shall establish and maintain procedures to ensure that mixups, damage, deterioration, con- tamination, or other adverse effects to product do not occur during handling.

##### § 820.150   Storage.

  - Each manufacturer shall establish and maintain procedures for the con- trol of storage areas and stock rooms for product to prevent mixups, damage, deterioration, contamination, or other adverse effects pending use or distribu- tion and to ensure that no obsolete, re- jected, or deteriorated product is used or distributed. When the quality of product deteriorates over time, it shall be stored in a manner to facilitate proper stock rotation, and its condi- tion shall be assessed as appropriate.

  - Each manufacturer shall establish and maintain procedures that describe the methods for authorizing receipt from and dispatch to storage areas and stock rooms.

##### § 820.160   Distribution.

Each manufacturer shall establish and maintain procedures for control and distribution of finished devices to ensure that only those devices ap- proved for release are distributed and that purchase orders are reviewed to ensure that ambiguities and errors are resolved before devices are released for distribution. Where a device’s fitness for use or quality deteriorates over time, the procedures shall ensure that expired devices or devices deteriorated beyond acceptable fitness for use are not distributed.

Each manufacturer shall main- tain distribution records which include or refer to the location of:

The name and address of the ini- tial  consignee;

The identification and quantity of devices shipped;

The date shipped; and

Any control number(s) used.

##### § 820.170   Installation.

  - Each manufacturer of a device re- quiring installation shall establish and

maintain adequate installation and in- spection instructions, and where appro- priate test procedures. Instructions and procedures shall include directions for ensuring proper installation so that the device will perform as intended after installation. The manufacturer shall distribute the instructions and procedures with the device or other- wise make them available to the per- son(s) installing the device.

  - The person installing the device shall ensure that the installation, in- spection, and any required testing are performed in accordance with the man- ufacturer’s instructions and procedures and shall document the inspection and any test results to demonstrate proper installation.

## Subpart M—Records

##### § 820.180   General requirements.

All records required by  this  part shall be maintained at the manufac- turing establishment or other location that is reasonably accessible to respon- sible officials of the manufacturer and to employees of FDA designated to per- form inspections. Such records, includ- ing those not stored at the inspected establishment, shall be made readily available for review and copying by FDA employee(s). Such records shall be legible and shall be stored to minimize deterioration and to prevent loss. Those records stored in automated data processing systems shall be backed up.

Confidentiality. Records deemed confidential by the manufacturer may be marked to aid FDA in determining whether information may be disclosed under the public information  regula- tion in part 20 of this chapter.

Record retention period. All records required by this part shall be retained for a period of time equivalent to the design and expected life of the device, but in no case less than  2  years  from the date of release for commercial dis- tribution  by  the  manufacturer.

Exceptions. This section does not apply    to    the    reports    required    by

§ 820.20(c) Management review, § 820.22 Quality audits, and supplier audit  re- ports  used  to  meet  the  requirements  of

§ 820.50(a) Evaluation of suppliers, con- tractors,     and     consultants,     but     does

§ 820.181

apply to procedures established under these provisions. Upon request of a des- ignated employee of FDA, an employee in management with executive respon- sibility shall certify in  writing that the management reviews and quality au- dits required under this part, and sup- plier audits where applicable, have been performed and documented, the dates on which they were performed, and that any required corrective action has been undertaken.

##### § 820.181   Device master record.

Each manufacturer shall  maintain device master records (DMR’s). Each manufacturer  shall  ensure  that  each DMR is prepared and approved in ac- cordance with § 820.40. The  DMR  for each type of device  shall  include,  or refer to the location of, the following information:

  - Device specifications including appropriate drawings, composition, for- mulation, component specifications, and software specifications;

  - Production process specifications including the appropriate equipment specifications, production methods, production procedures, and production environment  specifications;

  - Quality assurance procedures and specifications including acceptance cri- teria and the quality assurance equip- ment to be used;

  - Packaging and labeling specifica- tions, including methods and processes used; and

  - Installation, maintenance, and servicing procedures and methods.

##### § 820.184   Device history record.

Each manufacturer shall maintain device history records (DHR’s). Each manufacturer shall establish and main- tain procedures to ensure that DHR’s for each batch,  lot, or  unit are  main- tained to demonstrate that the device is manufactured in accordance with the DMR and the requirements of this part. The DHR shall include, or refer to the location of, the following information:

The dates of manufacture;

The quantity manufactured;

The quantity released for dis- tribution;

The acceptance records which demonstrate the device is manufac- tured in accordance with the DMR;

#### 21 CFR Ch. I (4–1–17 Edition)

The primary identification label and labeling used for each production unit; and

Any unique device identifier (UDI) or universal product code (UPC), and any other device identification(s) and control number(s) used.

[61 FR 52654, Oct. 7, 1996, as amended at 78 FR

58822, Sept. 24, 2013]

##### § 820.186   Quality system record.

Each manufacturer shall maintain a quality system record (QSR). The QSR shall include, or refer to  the  location of, procedures and the  documentation of activities required by this part  that are not specific  to  a  particular type of device(s), including, but not limited to, the records required by § 820.20. Each manufacturer shall ensure that  the QSR is prepared and approved in ac- cordance with § 820.40.

##### § 820.198   Complaint files.

Each manufacturer shall main- tain complaint files. Each manufac- turer shall establish and maintain pro- cedures for receiving, reviewing, and evaluating complaints by a formally designated unit. Such procedures shall ensure that:

All complaints are processed in a uniform and timely manner;

Oral complaints are documented upon receipt; and

Complaints are evaluated to de- termine whether the complaint rep- resents an event which is  required  to be reported  to  FDA  under  part  803  of this chapter, Medical  Device  Report- ing.

Each manufacturer shall review and evaluate all complaints to deter- mine whether an investigation is nec- essary. When no investigation is made, the manufacturer shall maintain a record that includes the reason no in- vestigation was made and the name of the individual responsible for the deci- sion not to investigate.

Any complaint involving the pos- sible failure of a device, labeling, or packaging to meet any of its specifica- tions shall be reviewed, evaluated, and investigated, unless such investigation has already been performed for a simi- lar complaint and another investiga- tion is not necessary.

#### Food and Drug Administration, HHS	§ 820.250

Any complaint that represents an event which must be reported to FDA under part 803 of this chapter shall be promptly reviewed, evaluated, and in- vestigated by a designated indi- vidual(s) and shall be maintained in a separate portion of the  complaint  files or otherwise clearly identified. In addi- tion   to   the   information   required   by

§ 820.198(e), records of investigation under this paragraph shall include a determination  of:

  - Whether the device failed to meet specifications;

  - Whether the device was being used for treatment or diagnosis; and

  - The relationship, if any, of the de- vice to the reported incident or adverse event.

When an investigation is made under this section, a record of the in- vestigation shall be maintained by the formally designated  unit identified  in paragraph (a) of this section. The record of investigation shall include:

The name of the device;

The date the complaint was re- ceived;

Any unique device identifier (UDI) or universal product code (UPC), and any other device identification(s) and control number(s) used;

The name, address, and phone number of the complainant;

The nature and details of the com- plaint;

The dates and results of the inves- tigation;

Any corrective action taken; and

Any reply to the complainant.

When the manufacturer’s formally designated complaint unit is located at a site separate from the manufacturing establishment, the investigated com- plaint(s) and the record(s) of investiga- tion shall be reasonably accessible to the manufacturing establishment.

If a manufacturer’s formally des- ignated complaint unit is located out- side of the United  States, records re- quired by this section shall be reason- ably accessible in the United States at either:

  - A location in the United States where the manufacturer’s records are regularly kept; or

  - The location of the initial dis- tributor.

[61 FR 52654, Oct. 7, 1996, as amended at 69 FR

11313, Mar. 10, 2004; 71 FR 16228, Mar. 31, 2006;

78 FR 58822, Sept. 24, 2013]

## Subpart N—Servicing

##### § 820.200   Servicing.

    - Where servicing is a specified re- quirement, each manufacturer shall es- tablish and maintain instructions and procedures for performing  and verifying that the servicing meets the specified requirements.

    - Each manufacturer shall analyze service reports with appropriate statis- tical methodology in accordance with

§ 820.100.

    - Each manufacturer who receives a service report that represents an event which must be reported to FDA under part 803 of this chapter shall automati- cally consider the report a complaint and shall process it in accordance with the requirements of § 820.198.

    - Service reports shall be docu- mented and shall include:

The name of the device serviced;

Any unique device identifier (UDI) or universal product code (UPC), and any other device identification(s) and control number(s) used;

The date of service;

The individual(s) servicing the de- vice;

The service performed; and

The test and inspection data.

[61 FR 52654, Oct. 7, 1996, as amended at 69 FR

11313, Mar. 10, 2004; 78 FR 58822, Sept. 24, 2013]

## Subpart O—Statistical Techniques

##### § 820.250   Statistical techniques.

  - Where appropriate, each manufac- turer shall establish and maintain pro- cedures for identifying valid statistical techniques required for establishing, controlling, and verifying the accept- ability of process capability and prod- uct  characteristics.

  - Sampling plans, when used, shall be written and based on a valid statis- tical rationale. Each  manufacturer shall establish and maintain proce- dures to ensure that sampling methods are adequate for their intended use and to ensure that when changes occur the

Pt. 821

sampling plans are reviewed. These ac- tivities shall be documented.

## PART 821—MEDICAL DEVICE TRACKING REQUIREMENTS

##### Subpart A—General Provisions

Sec.

  - Scope.

  - Exemptions and variances.

  - Definitions.

  - Imported devices.

##### Subpart B—Tracking Requirements

821.20    Devices subject to tracking.

821.25 Device tracking system and content requirements:  manufacturer  require- ments.

##### Subpart C—Additional Requirements and Responsibilities

821.30 Tracking obligations of persons other than device manufacturers: distributor requirements.

##### Subpart D—Records and Inspections

821.50     Availability.

821.55     Confidentiality.

  - Retention of records.

AUTHORITY: 21 U.S.C. 331, 351, 352, 360, 360e,

360h, 360i, 371, 374.

SOURCE:  58  FR  43447,  Aug.  16,  1993,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 821.1   Scope.

    - The regulations in this part im- plement section 519(e) of the Federal Food, Drug, and  Cosmetic  Act  (the act), which provides that the Food and Drug Administration may require a manufacturer to adopt a method of tracking a class II or class III device, if the device meets one of the following three criteria and FDA issues an order to the manufacturer: the failure of the device would be reasonably likely to have serious adverse health con- sequences; or the device is intended to be implanted in the human body for more than 1 year; or the device is a life-sustaining or life-supporting device used outside a device user facility. A device that meets one of these criteria and is  the  subject  of an FDA  order must  comply  with  this  part  and  is  re-

21 CFR Ch. I (4–1–17 Edition)

ferred to, in this part, as a ‘‘tracked de- vice.’’

    - These regulations are intended to ensure that tracked devices can be traced from the device manufacturing facility to the person for whom the de- vice is indicated, that is, the patient. Effective tracking of devices from the manufacturing facility, through the distributor network (including dis- tributors, retailers, rental firms and other commercial enterprises, device user facilities, and licensed practi- tioners) and, ultimately,  to  the  patient is necessary for the effectiveness of remedies prescribed by the act, such as patient notification (section 518(a) of the act) or device recall (section 518(e) of the act). Although these  regulations do not preclude a manufacturer from involving outside organizations in that manufacturer’s device tracking  effort, the legal responsibility for complying with this part rests with manufactur- ers who are subject to tracking orders, and that responsibility cannot be al- tered, modified, or in any way abro- gated by contracts or other agree- ments.

    - The primary burden for ensuring that the tracking system works rests upon the manufacturer. A manufac- turer or any other person, including a distributor, final distributor, or mul- tiple distributor, who distributes a de- vice subject to tracking, who fails to comply with any applicable require- ment of section 519(e) of the act or of this part, or any person who causes such failure, misbrands the device within the meaning of section 502(t)(2) of the act and commits a prohibited act within the meaning of sections 301(e) and 301(q)(1)(B) of the act.

    - Any person subject to this part who permanently discontinues doing business is required to notify FDA at the time the person notifies any gov- ernment agency, court, or supplier, and provide FDA with a complete set of its tracking records and information. However, if a person ceases distribu- tion of a tracked device but continues to do other business, that person con- tinues to be responsible for compliance with this part unless another person, affirmatively and in writing, assumes responsibility for continuing the track- ing  of  devices  previously  distributed

#### Food and Drug Administration, HHS	§ 821.3

under this part. Further, if a person subject to this part goes out of busi- ness completely, but other persons ac- quire the right to manufacture or dis- tribute tracked devices, those other persons are deemed to be responsible for continuing the tracking responsi- bility of the previous person under this part.

[58 FR 43447, Aug. 16, 1993, as amended at 67

FR  5951,  Feb.  8,  2002;  73  FR  34860,  June  19,

2008]

##### § 821.2   Exemptions and variances.

      - A manufacturer, importer, or dis- tributor may seek an exemption or variance from one or more require- ments of this part.

      - A request for an exemption or variance shall be submitted  in  the  form of a  petition under § 10.30 of this  chap- ter and shall comply with the require- ments set out therein, except that a re- sponse shall be issued in 90 days. The Director   or   Deputy   Directors,    CDRH, or the Director, Office of Compliance, CDRH, shall issue responses to requests under this section.  The  petition  shall also  contain  the  following:

The name of the device and device class and representative labeling show- ing the intended use(s) of the device;

The reasons that compliance with the tracking requirements  of  this  part is unnecessary;

A complete description of alter- native steps that are available, or that the petitioner has already taken, to en- sure that an effective tracking system is in place; and

Other information justifying the exemption or variance.

      - An exemption or variance is not effective until the Director, Office of Compliance, CDRH, approves  the  re- quest under § 10.30(e)(2)(i) of this chap- ter.

[58 FR 43447, Aug. 16, 1993, as amended at 59

FR  31138,  June  17,  1994;  67  FR  5951,  Feb.  8,

2002; 72 FR 17399, Apr. 9, 2007]

##### § 821.3   Definitions.

The following definitions and terms apply to this part:

Act means  the  Federal  Food, Drug, and Cosmetic Act, 21 U.S.C. 321 et seq., as amended.

Importer means the initial dis- tributor  of  an  imported  device  who  is

subject to a tracking order. ‘‘Importer’’ does not include anyone who only fur- thers the marketing, e.g., brokers, job- bers, or warehousers.

Manufacturer means any person, including any importer, repacker and/ or relabeler, who manufactures, pre- pares, propagates, compounds, assem- bles, or processes a device or engages in any   of   the   activities   described   in

§ 807.3(d) of this chapter.

Device failure means the failure of a device to perform or function as in- tended, including any deviations from the device’s performance specifications or intended use.

Serious adverse health consequences means any significant adverse experi- ence related to a device, including de- vice-related events which are life- threatening or which involve perma- nent or long-term injuries or illnesses.

Device intended to be implanted in the human body for more than 1 year means a device that is intended to be placed into a surgically or naturally formed cavity of the human body for more than 1 year to continuously as- sist, restore, or replace the function of an organ system or structure of the human body throughout the useful life of the device. The term does not in- clude a device  that  is  intended  and used only for temporary purposes or that is intended for explantation in 1 year or less.

Life-supporting or life-sustaining device used outside a device user facility means a device which is essential, or yields information that is essential, to the restoration or continuation of a bodily function important to the con- tinuation of human life that is in- tended for use outside a hospital, nurs- ing home, ambulatory surgical facility, or diagnostic or outpatient treatment facility. Physicians’ offices are not de- vice user facilities and, therefore, de- vices used therein are subject to track- ing if they otherwise satisfy the statu- tory and regulatory criteria.

Distributor means any person who furthers the distribution of a device from the original place of manufacture to the person who makes delivery or sale to the ultimate user, i.e., the final or multiple distributor, but who does not repackage or otherwise change the

#### § 821.4

container, wrapper, or labeling of the device or device package.

Final distributor means any person who distributes a tracked device in- tended for use by a single patient over the useful life of the device to the pa- tient. This term includes, but is not limited to, licensed practitioners, re- tail pharmacies, hospitals, and other types of device user facilities.

Distributes means any  distribution of a tracked device, including the char- itable distribution of a tracked device. This term does not include the dis- tribution of a device under an effective investigational device exemption in ac- cordance with section 520(g) of the act and part 812 of this chapter or the dis- tribution of a device for teaching, law enforcement, research, or analysis as specified in § 801.125 of this chapter.

Multiple distributor means any de- vice user facility, rental company, or any other entity that distributes a life- sustaining or life-supporting device in- tended for use by more than one pa- tient over the useful life of the device.

Licensed practitioner means a phy- sician, dentist, or other health care practitioner licensed  by the law  of the State in which he or she practices  to use or order the use of the tracked de- vice.

Any term defined in  section  201 of the act shall have the same defini- tion in this part.

Human cell, tissue, or cellular or tis- sue-based product (HCT/P) regulated as a device  means  an  HCT/P  as  defined  in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A unique  device  identifier  is  composed  of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

#### 21 CFR Ch. I (4–1–17 Edition)

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

[58 FR 43447, Aug. 16, 1993, as amended at 67

FR  5951,  Feb.  8,  2002;  78  FR  58822,  Sept.  24,

2013]

##### § 821.4   Imported devices.

For purposes of this part, the im- porter of a tracked device shall be con- sidered the manufacturer and shall be required to comply with all require- ments of this part applicable to manu- facturers. Importers must keep all in- formation required under this part in the United States.

## Subpart B—Tracking Requirements

##### § 821.20   Devices subject to tracking.

A manufacturer of any class II or class III device that fits within  one  of the three criteria within § 821.1(a) must track that device in accordance with this part, if FDA issues  a  tracking order to that manufacturer.

When responding to premarket notification submissions and remarket approval applications, FDA will notify the sponsor by issuing an order that states that FDA believes the device meets the criteria of section 519(e)(1) of the act and, by virtue of the order, the sponsor must track the device.

[67 FR 5951, Feb. 8, 2002]

##### § 821.25 Device tracking system and content requirements: manufac- turer requirements.

A manufacturer of a tracked de- vice shall adopt a method of tracking for each such type of device that it dis- tributes that enables a  manufacturer to provide FDA with the following in- formation in writing for each tracked device distributed:

  - Except as required by order under section 518(e) of the act, within 3 work- ing days of a  request  from  FDA,  prior to the distribution of  a  tracked  device to a patient, the name, address, and telephone   number   of   the   distributor,

Food and Drug Administration, HHS	§ 821.25

multiple distributor, or final dis- tributor holding the device for dis- tribution and the location of the de- vice;

  - Within 10 working days of a re- quest from FDA  for  tracked  devices that are intended for use by a single patient over the life of the device, after distribution to or implantation in a pa- tient:

    - The unique device identifier (UDI), lot number, batch number, model num- ber, or serial number of the device or other identifier  necessary  to  provide for effective tracking of the devices;

    - The date the device was shipped by the manufacturer;

    - The name, address, telephone number, and social security number (if available) of the patient receiving the device, unless not released by the pa- tient under § 821.55(a);

    - The date the device was provided to the patient;

    - The name, mailing address, and telephone number of the prescribing physician;

    - The name, mailing address, and telephone number of the physician reg- ularly following the patient if different than the prescribing physician; and

    - If applicable, the date the device was explanted and the name, mailing address, and telephone number of the explanting physician; the date of the patient’s death; or the date the device was returned to the manufacturer, per- manently retired from use, or other- wise permanently disposed of.

  - Except as required by order under section 518(e) of the act, within 10 working days of a request from  FDA for tracked devices that are intended for use by more than one patient, after the distribution of the device to the multiple  distributor:

    - The unique device identifier (UDI), lot number, batch number, model num- ber, or serial number of the device or other identifier  necessary  to  provide for effective tracking of the devices;

    - The date the device was shipped by the manufacturer;

    - The name, address, and tele- phone number of the multiple dis- tributor;

    - The name, address, telephone number, and social security number (if available) of the patient using the de-

vice,  unless  not  released  by  the  patient under § 821.55(a);

    - The location of the device;

    - The date the device was provided for use by the patient;

    - The name, address, and tele- phone number of the prescribing physi- cian; and

    - If and when applicable, the date the device was returned to the manu- facturer, permanently retired from use, or otherwise permanently disposed of.

A manufacturer of a tracked de- vice shall keep current records in ac- cordance with its standard operating procedure of the information identified in paragraphs (a)(1), (a)(2) and (a)(3)(i) through (a)(3)(iii) of this section on each tracked device released for dis- tribution for as long as such device is in use or in distribution for use.

A manufacturer of a tracked de- vice shall establish a written standard operating procedure for the collection, maintenance, and auditing of the data specified in paragraphs (a) and (b) of this section. A manufacturer shall make this standard operating proce- dure available to FDA upon request. A manufacturer shall incorporate the fol- lowing into the standard operating pro- cedure:

  - Data collection and recording pro- cedures, which shall include a proce- dure for recording when data which is required under this part is missing and could not be collected and the reason why such required data is missing and could not be collected;

  - A method for recording all modi- fications or changes to the tracking system or to the data collected and maintained under the tracking system, reasons for any modification or change, and dates of any modification or change. Modification and changes in- cluded under this requirement include modifications to the data (including termination of tracking), the data for- mat, the recording system, and the file maintenance procedures system; and

  - A quality assurance program that includes an audit procedure to be  run for each device product subject to tracking, at not less than 6-month in- tervals for the first 3 years of distribu- tion and at least once a year there- after. This audit procedure shall pro- vide  for  statistically  relevant  sampling

#### § 821.30

of the data collected to ensure the ac- curacy of data and performance testing of the functioning of the tracking sys- tem.

When a manufacturer becomes aware that a distributor, final dis- tributor, or multiple distributor has not collected, maintained, or furnished any record or information required by this part, the manufacturer shall no- tify the FDA district office responsible for the area in which the distributor, final distributor, or multiple dis- tributor is located of the failure of such persons to comply with the re- quirements of this part. Manufacturers shall have taken reasonable steps to obtain compliance by the distributor, multiple distributor, or final dis- tributor in question before notifying FDA.

A  manufacturer  may  petition  for an exemption or variance from one or more requirements of this part accord- ing to the procedures in § 821.2 of this chapter.

[58 FR 43447, Aug. 16, 1993, as amended at 67

FR  5951,  Feb.  8,  2002;  78  FR  58822,  Sept.  24,

2013]

## Subpart C—Additional Requirements and Responsibilities

##### § 821.30 Tracking obligations of per- sons other than device manufactur- ers: distributor requirements.

A distributor, final distributor, or multiple distributor of any tracked de- vice shall, upon purchasing or other- wise acquiring any interest in such a device, promptly provide the manufac- turer tracking the device with the fol- lowing information:

  - The name and address of the dis- tributor, final distributor or multiple distributor;

  - The unique device identifier (UDI), lot number, batch number, model number, or serial number of the device or other identifier used by the manufacturer to track the device;

  - The date the device was received;

  - The person from whom the device was received;

  - If and when applicable, the date the device was explanted, the date of the patient’s death, or the date the de- vice  was  returned  to  the  distributor,

21 CFR Ch. I (4–1–17 Edition)

permanently retired from use, or other- wise permanently disposed of.

A final distributor, upon sale or other distribution of a tracked device for use in or by the patient, shall promptly provide the manufacturer tracking the device with the following information:

  - The name and address of the final distributor,

  - The unique device identifier (UDI), lot number, batch number, model number, or serial number of the device or other identifier used by the manufacturer to track the device;

  - The name, address, telephone number, and social security number (if available) of the patient receiving the device, unless not released by the pa- tient under § 821.55(a);

  - The date the device was provided to the patient or for use in the patient;

  - The name, mailing address, and telephone number of the prescribing physician;

  - The name, mailing address, and telephone number of the physician reg- ularly following the patient if different than the prescribing physician; and

  - When applicable, the date the de- vice was explanted and the name, mail- ing address, and telephone number of the explanting physician, the date of the patient’s death, or the date the de- vice was returned to the manufacturer, permanently retired from use, or other- wise permanently disposed of.

(c)(1) A multiple distributor shall keep written records of the following each time such device is distributed for use by a patient:

    - The unique device identifier (UDI), lot number, batch number, model num- ber, or serial number of the device or other identifier used by the manufac- turer to track the device;

    - The name, address, telephone number, and social security number (if available) of the patient using the de- vice;

    - The location of the device, un- less not released by the patient under

§ 821.55(a);

    - The date the device was provided for use by the patient;

    - The name, address, and telephone number of the prescribing physician;

#### Food and Drug Administration, HHS	§ 821.60

    - The name, address, and telephone number of the physician regularly fol- lowing the patient if different than the prescribing physician; and

    - When applicable, the date the device was permanently retired from use or otherwise permanently disposed of.

(2) Except as required by order under section 518(e) of the act, any person who is a multiple distributor subject to the recordkeeping requirement of para- graph (c)(1) of this section shall, within

5 working days of a request from the manufacturer  or  within  10  working days of a request from FDA for the in- formation identified in paragraph  (c)(1) of this section,  provide  such  informa- tion  to  the  manufacturer  or  FDA.

A distributor, final distributor, or multiple distributor shall make any records required to be kept under this part available to the manufacturer of the tracked device for audit upon writ- ten request by an authorized represent- ative of the manufacturer.

A distributor, final distributor, or multiple  distributor  may  petition  for an exemption or variance from one or more requirements of this part accord- ing to the procedures in § 821.2.

[58 FR 43447, Aug. 16, 1993, as amended at 67

FR  5951,  Feb.  8,  2002;  78  FR  58822,  Sept.  24,

2013]

## Subpart D—Records and Inspections

##### § 821.50   Availability.

Manufacturers, distributors, mul- tiple distributors, and final  distribu- tors shall, upon the presentation by an FDA representative of official creden- tials and the issuance of Form FDA 482 at the initiation of an  inspection of an establishment or person under section 704 of the act, make each record and all information required to  be  collected and maintained under this part and all records and information related to the events and persons identified in such records available to FDA personnel.

Records and information ref- erenced in paragraph (a) of this section shall be available to FDA personnel for purposes of reviewing, copying, or any other use related to the enforcement of the act and this part. Records required to be kept by this part shall be kept in

a centralized point for each manufac- turer or distributor within the United States.

[58 FR 43447, Aug. 16, 1993, as amended at 65

FR 43690, July 14, 2000]

##### § 821.55   Confidentiality.

- Any patient receiving a device subject to tracking requirements under this part may refuse to release, or refuse permission to release, the pa- tient’s name, address, telephone num- ber, and social security number, or other identifying information for the purpose of tracking.

- Records and other information submitted to FDA under this part shall be protected from public disclosure to the extent permitted under part 20 of this  chapter,  and  in  accordance  with

§ 20.63 of this chapter, information con- tained in such records that would  iden- tify  patient  or  research  subjects  shall not be available for public disclosure except as provided in those parts.

- Patient names or other identifiers may be disclosed to a manufacturer or other person subject to this part or to a physician when the health or safety of the patient requires that such per- sons have access to the information. Such notification will be pursuant to agreement that the record or informa- tion will not be further disclosed ex- cept as the health aspects of the pa- tient requires. Such notification does not constitute public disclosure and will not trigger the availability of the same information to the public gen- erally.

[58 FR 43447, Aug. 16, 1993, as amended at 67

FR 5951, Feb. 8, 2002]

##### § 821.60   Retention of records.

Persons required to maintain records under this part shall maintain such records for the useful life of each tracked device they manufacture or distribute. The useful life of a device is the time a device is in use or in dis- tribution for use. For example, a record may be retired if the person maintain- ing the record becomes aware of the fact that the device is no longer in use, has been explanted, returned to the manufacturer, or the patient has died.

Pt. 822

PART 822—POSTMARKET SURVEILLANCE

##### Subpart A—General Provisions

Sec.

  - What does this part cover?

  - What is the purpose of this part?

  - How do you  define the  terms used  in this  part?

  - Does this part apply to me?

##### Subpart B—Notification

  - How will I know if I must conduct postmarket surveillance?

  - When will you notify me that I am re- quired to conduct postmarket surveil- lance?

  - What should I do if I do not agree that postmarket surveillance is appropriate?

##### Subpart C—Postmarket Surveillance Plan

  - When, where, and how must I submit my postmarket surveillance plan?

  - What must I include in my submis- sion?

  - What must I include in my surveil- lance plan?

  - What should I consider when design- ing my plan to conduct postmarket sur- veillance?

  - Do you have any information that will help me prepare my submission or design my postmarket surveillance plan?

  - [Reserved]

  - May I reference information pre- viously submitted instead of submitting it again?

  - How long must I conduct postmarket surveillance of my device?

##### Subpart D—FDA Review and Action

  - What will you consider in the review of my submission?

  - How long will your review of my sub- mission  take?

  - How will I be notified of your deci- sion?

  - What kinds of decisions may you make?

  - What are the consequences if I fail to submit a postmarket surveillance plan, my plan is disapproved and I fail to sub- mit a new plan, or I fail to conduct sur- veillance in accordance with my ap- proved plan?

  - What must I do if I want to make changes to my postmarket surveillance plan after you have approved it?

  - What recourse do I have if I do not agree with your decision?

  - Is the information in my submission considered confidential?

#### 21 CFR Ch. I (4–1–17 Edition)

##### Subpart E—Responsibilities of Manufacturers

  - What are my responsibilities once I am notified that  I am required to con- duct postmarket surveillance?

  - What are my responsibilities after my postmarket surveillance plan has been approved?

  - If my company changes ownership, what must I do?

  - If I go out of business, what must I do?

  - If I stop marketing the device subject to postmarket surveillance, what must I do?

##### Subpart F—Waivers and Exemptions

  - May I request a waiver of a specific requirement of this part?

  - May I request exemption from the re- quirement to conduct postmarket sur- veillance?

##### Subpart G—Records and Reports

  - What records am I required to keep?

  - What records are the investigators in my surveillance plan required to keep?

  - How long must we keep the records?

  - What must  I do  with the  records if the sponsor of the plan or an investigator in the plan changes?

  - Can you inspect my manufacturing site or other sites involved in my postmarket surveillance plan?

  - Can you inspect and copy the records related to my postmarket surveillance plan?

  - Under what circumstances would you inspect records identifying subjects?

  - What reports must I submit to you?

AUTHORITY: 21 U.S.C. 331, 352, 360i, 360l, 371,

374.

SOURCE:  67  FR  38887,  June  6,  2002,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 822.1   What does this part cover?

This part implements section 522 of the Federal Food, Drug, and Cosmetic Act (the act) by providing procedures and requirements for postmarket sur- veillance of class II and class III de- vices that meet any of the following criteria:

Failure of the device would be reasonably likely to have serious ad- verse health consequences;

The device is intended to be im- planted in the human body for more than 1 year; or

Food and Drug Administration, HHS	§ 822.3

The device is intended to be used outside a user facility to support or sustain life. If you fail to comply with requirements that we order under sec- tion 522 of the act and this part, your device is considered misbranded under section 502(t)(3) of the act and you are in violation of section 301(q)(1)(C) of the  act.

##### § 822.2   What  is   the   purpose   of   this part?

The purpose of this part is to imple- ment our postmarket surveillance au- thority to maximize the  likelihood that postmarket surveillance plans will result in the collection of useful data. These data can reveal unforeseen ad- verse events, the actual rate of antici- pated adverse events, or other informa- tion necessary to protect the public health.

##### § 822.3   How  do  you  define  the  terms used in this part?

Some of the terms we use in this part are specific to postmarket surveillance and reflect the language used in the statute (law). Other terms are more general and reflect our interpretation of the law. This section of the part de- fines the following terms:

Act means  the  Federal  Food, Drug, and Cosmetic Act, 21 U.S.C. 301 et seq., as amended.

Designated person means the indi- vidual who conducts or supervises the conduct of your postmarket surveil- lance. If your postmarket surveillance plan includes a team of investigators, as defined below, the designated person is the responsible leader of that team.

Device failure means a device does not perform or function as intended, and includes any deviation from the de- vice’s performance specifications or in- tended use.

General plan guidance means agen- cy guidance that provides information about the requirement to conduct postmarket surveillance, the submis- sion of a plan to us for approval, the content of the submission, and the con- duct and reporting requirements of the surveillance.

Human cell, tissue, or cellular or tis- sue-based product (HCT/P) regulated as a device  means  an  HCT/P  as  defined  in

§ 1271.3(d)  of  this  chapter  that  does  not

meet the criteria in § 1271.10(a) and that is also regulated as a device.

Investigator means an individual who collects data or information in support of a postmarket surveillance plan.

Life-supporting or life-sustaining device used outside a device user facility means that a device is essential to, or yields information essential to, the res- toration or continuation of a bodily function important to the continuation of human life and is used outside a hos- pital, nursing home, ambulatory sur- gical facility, or diagnostic or out- patient treatment facility. A physi- cian’s office is not a device user facil- ity.

Manufacturer means any person, including any importer, repacker, and/ or relabeler, who manufactures, pre- pares, propagates, compounds, assem- bles, processes a device, or engages in any   of   the   activities   described   in

§ 807.3(d) of this chapter.

Postmarket surveillance means the active, systematic, scientifically valid collection, analysis, and interpretation of data or other information about a marketed device.

Prospective surveillance means that the subjects are identified at the begin- ning of the surveillance and data or other information will be collected from that time forward (as opposed to retrospective  surveillance).

Serious adverse health consequences means any significant adverse experi- ence related to a device, including de- vice-related events that are life-threat- ening or that involve permanent or long-term injuries or illnesses.

Specific guidance means guidance that provides information regarding postmarket surveillance for specific types or categories of devices or spe- cific postmarket  surveillance  issues. This type of guidance may be used to supplement general guidance and may address such topics as the type of sur- veillance approach that is appropriate for the device and the postmarket sur- veillance question, sample size, or spe- cific  reporting  requirements.

Surveillance question means the issue or issues to be addressed by the postmarket   surveillance.

Unforeseen adverse event means any serious adverse health consequence

#### § 822.4

that either is not addressed in the la- beling of the device or occurs at a rate higher than anticipated.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribu- tion and use by meeting the  require- ments of § 830.20 of this chapter. A UDI is composed of:

A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

  - The lot or batch within which a device was manufactured;

  - The serial number of a specific device;

  - The expiration date of a specific device;

  - The date a specific device was manufactured.

  - For an HCT/P regulated as a de- vice, the distinct identification code required by § 1271.290(c) of this chapter.

[67 FR 38887, June 6, 2002, as amended at 78

FR 55823, Sept. 24, 2013]

##### § 822.4   Does this part apply to me?

If we have ordered you to conduct postmarket surveillance of a medical device under section 522 of the act, this part applies to you. We have the au- thority to order postmarket surveil- lance of any class II or class  III  med- ical device, including a device reviewed under the licensing provisions of sec- tion 351 of the Public Health Service Act, that meets any of the following criteria:

Failure of the device would be reasonably likely to have serious ad- verse health consequences;

The device is intended to be im- planted in the human body for more than 1 year; or

The device is intended to be used to support or sustain life and to be used outside a user facility.

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart B—Notification

##### § 822.5 How will I know if I must con- duct postmarket surveillance?

We will send you a letter (the postmarket surveillance order) noti- fying you of the requirement to con- duct postmarket surveillance.  Before we send the order, or as part of the order, we may require that you submit information about your  device  that will allow us better to define the scope of a surveillance order. We will specify the device(s) subject to the  surveil- lance order and the reason that we are requiring postmarket surveillance  of the device under section 522 of the act. We will also provide you with any gen- eral or specific guidance that is avail- able to help you develop your plan for conducting  postmarket  surveillance.

##### § 822.6 When will you notify me that I am required to conduct postmarket surveillance?

We will notify you as soon as we have determined that postmarket surveil- lance of your device is necessary, based on the identification of a surveillance question. This may occur during the review of a marketing application for your device, as your device goes to market, or after your device has been marketed for a period of time.

##### § 822.7 What should I do if I do not agree that postmarket surveillance is appropriate?

If you do not agree with our deci- sion to order postmarket surveillance for a particular device, you may re- quest review of our decision by:

  - Requesting a meeting with the Di- rector, Office of Surveillance and Bio- metrics, who generally issues the order for postmarket surveillance;

  - Seeking internal review  of  the order under § 10.75 of this chapter;

  - Requesting an informal hearing under part 16 of this chapter; or

  - Requesting review by the Medical Devices Dispute Resolution Panel of the Medical Devices Advisory Com- mittee.

You may obtain guidance docu- ments that discuss these mechanisms from the Center for Devices and Radio- logical Health’s (CDRH’s) Web site

#### Food and Drug Administration, HHS	§ 822.10

.

[67 FR 38887, June 6, 2002, as amended at 72

FR  17399,  Apr.  9,  2007;  78  FR  18233,  Mar.  26,

2013]

## Subpart C—Postmarket Surveillance Plan

##### § 822.8 When, where, and how  must  I submit my postmarket surveillance plan?

You must submit your plan to con- duct postmarket surveillance within 30 days of the date you receive the postmarket surveillance order. For de- vices regulated by the Center for Bio- logics Evaluation and Research, send three copies of your submission to the Food  and  Drug  Administration,  Center for Biologics Evaluation and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002. For de- vices regulated by the Center for Drug Evaluation and Research, send  three copies of your submission to the Cen- tral Document Room, Center for Drug Evaluation and  Research,  Food  and Drug	Administration,	5901–B, Ammendale Rd., Beltsville, MD 20705– 1266. For devices regulated by the  Cen- ter  for   Devices   and   Radiological Health, send three copies of your sub- mission to the Document Mail Center, 10903 New Hampshire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993–0002. When we receive your original submis- sion, we will send you an acknowledg- ment letter identifying the unique doc- ument number assigned to your sub- mission. You must use this number in any correspondence related to this sub- mission.

[75 FR 20915, Apr. 22, 2010, as amended at 80

FR 18094, Apr. 3, 2015]

##### § 822.9   What must I include in my sub- mission?

Your submission must include the following:

Organizational/administrative in- formation:

  - Your name and address;

  - Generic and trade names of your device;

  - Name and address of the contact person for the submission;

  - Premarket application/submission number and device identifiers for your device;

  - Table of contents identifying the page numbers for each section of the submission;

  - Description of the device  (this may be incorporated by reference to the appropriate premarket application/ submission);

  - Product codes and a list of all rel- evant model numbers; and

  - Indications for use and claims for the device;

Postmarket surveillance plan;

Designated person information;

Name, address, and telephone number; and

Experience and qualifications.

[67 FR 38887, June 6, 2002, as amended at 78

FR 55823, Sept. 24, 2013]

##### § 822.10   What  must  I  include  in  my surveillance plan?

Your surveillance plan must include a discussion of:

The plan objective(s) addressing the surveillance question(s) identified in our order;

The subject of the study, e.g., pa- tients, the device, animals;

The variables and endpoints that will be used to answer the surveillance question, e.g., clinical parameters or outcomes;

The surveillance approach or methodology to be used;

Sample size and units of observa- tion;

The investigator agreement, if ap- plicable;

Sources of data, e.g., hospital records;

The data collection plan and forms;

The consent document, if applica- ble;

Institutional Review Board infor- mation, if applicable;

The patient followup plan, if ap- plicable;

The procedures for monitoring conduct and progress of the surveil- lance;

An estimate of the duration of surveillance;

§ 822.11

All data analyses and statistical tests planned;

The content and timing of re- ports.

##### § 822.11 What should I consider when designing my plan to conduct postmarket surveillance?

You must design your surveillance to address the postmarket surveillance question identified in the order you re- ceived. You should consider what, if any,  patient  protection  measures should be incorporated into your plan. You should also consider the function, operating characteristics, and intended use of your device when designing a surveillance  approach.

##### § 822.12 Do you have any information that will help me prepare my sub- mission or design my postmarket surveillance plan?

Guidance documents that discuss our current thinking on preparing a postmarket    surveillance    submission and designing  a  postmarket  surveil- lance plan are available on the  Center for Devices and Radiological Health’s Web site and from the Food and Drug Administration, Center for Devices and Radiological Health, Office of Surveil- lance and Biometrics, 10903 New Hamp- shire Ave., Bldg. 66, rm. 3219, Silver Spring, MD 20993–0002. Guidance docu- ments represent our current interpre- tation of, or policy on,  a  regulatory issue. They do not establish legally en- forceable rights or responsibilities and do  not  legally  bind  you  or  FDA.  You may choose to use an approach other than the one set forth in a guidance document, as long as your alternative approach complies with the relevant statutes (laws) and regulations. If you wish, we will meet with you to discuss whether an  alternative  approach  you are considering will satisfy the require- ments  of  the  act  and  regulations.

[75 FR 20915, Apr. 22, 2010]

##### § 822.13   [Reserved]

§ 822.14 May I reference information previously submitted instead of submitting it again?

Yes, you may reference information that you have submitted in premarket submissions      as      well      as      other

#### 21 CFR Ch. I (4–1–17 Edition)

postmarket surveillance submissions. You must specify the information to be incorporated and the document number and pages where the information is lo- cated.

##### § 822.15 How long must I conduct postmarket surveillance of my de- vice?

The   length   of   postmarket   surveil- lance  will  depend  on  the  postmarket surveillance  question  identified  in  our order.  We  may  order  prospective  sur- veillance for a period up to 36 months; longer periods require your agreement. If we believe that a prospective period of greater than 36 months is necessary to   address   the   surveillance   question, and you do not agree, we will use the Medical    Devices    Dispute    Resolution Panel  to  resolve  the  matter.  You  may obtain guidance regarding dispute reso- lution  procedures  from  the  Center  for Devices	and	Radiological	Health’s (CDRH’)   Web   site

The 36-month period refers to the sur- veillance period, not the length of time from the issuance of the order.

[72  FR  17400,  Apr.  9,  2007,  as  amended  at  78

FR 18233, Mar. 26, 2013]

## Subpart D—FDA Review and Action

##### § 822.16 What will you consider in the review of my submission?

First, we will determine that the sub- mission is administratively complete. Then, in accordance with the law, we must determine whether the des- ignated person has appropriate quali- fications and experience to conduct the surveillance and whether the surveil- lance plan will result in the collection of useful data that will answer the sur- veillance question.

##### § 822.17 How long will your review of my submission take?

We will review your submission with- in 60 days of receipt.

Food and Drug Administration, HHS	§ 822.21

##### § 822.18   How will I be notified of your decision?

We will send you a letter notifying you of our decision and identifying any action you must take.

##### § 822.19   What kinds of decisions may you make?

| If your plan: | Then we will send you: | And you must: |
| --- | --- | --- |
| Should result in the collection of useful data that will address the postmarket surveillance question Should result in the collection of useful data that will address the postmarket surveillance question after specific revisions are made or specific information is provided Does not meet the requirements specified in this part Is not likely to result in the collection of useful data that will address the postmarket surveil- lance question | An approval order, identifying any specific requirements related to your postmarket surveillance An approvable letter identifying the specific revisions or information that must be submitted before your plan can be approved A letter disapproving your plan and identifying the reasons for dis- approval A letter disapproving your plan and identifying the reasons for dis- approval | Conduct postmarket surveillance of your device in accordance with the approved plan Revise your postmarket surveillance submission to address the con- cerns in the approvable letter and submit it to us within the specified timeframe. We will determine the timeframe case-by-case, based on the types of revisions or informa- tion that you must submit Revise your postmarket surveillance submission and submit it to us within the specified timeframe. We will determine the timeframe case- by-case, based on the types of re- visions or information that you must submit Revise your postmarket surveillance submission and submit it to us within the specified timeframe. We will determine the timeframe case- by-case, based on the types of re- visions or information that you must submit |

##### § 822.20 What are the consequences if I fail to submit a postmarket surveil- lance plan, my plan is disapproved and I fail to submit a new plan, or I fail to conduct surveillance in ac- cordance with my approved plan?

The failure to have an approved postmarket surveillance plan or failure to conduct postmarket surveillance in accordance with the approved plan con- stitutes failure to comply with section 522 of the act. Your failure would be a prohibited	act	under	section 301(q)(1)(C) of the act, and your device would be misbranded under section 502(t)(3) of the act. We have the author- ity to initiate actions against products that are adulterated  or  misbranded, and against persons who commit pro- hibited acts. Adulterated or mis- branded devices can be seized. Persons who commit prohibited acts can be en- joined from committing such acts, re- quired to pay civil money penalties, or prosecuted.

##### § 822.21 What must I do if I want to make changes to my postmarket surveillance plan after you have ap- proved it?

You must receive our approval in writing before making changes in your plan that will affect the nature or va- lidity of the data collected in accord- ance with the plan. To obtain our ap- proval, you must submit  three  copies of the request to make the proposed change and revised postmarket surveil- lance plan  to  the applicable address listed in § 822.8. You may reference in- formation already submitted in accord- ance with § 822.14. In your cover letter, you must identify your submission as a supplement and cite the unique docu- ment number that we assigned in our acknowledgment letter for your origi- nal  submission,  specifically   identify the changes to the plan, and  identify the reasons and justification for mak- ing the changes. You must report changes in your plan that will not  af- fect  the  nature  or  validity  of  the  data

#### § 822.22

collected in accordance with  the  plan in the  next  interim  report required  by your approval order.

##### § 822.22   What  recourse  do  I  have  if  I do not agree with your decision?

If you disagree with us about the content of your plan or if we dis- approve your plan, or if you believe there is a less burdensome approach that will answer the surveillance ques- tion, you may request review of our de- cision by:

  - Requesting a meeting with the Di- rector, Office of Surveillance and Bio- metrics, Center for Devices and Radio- logical Health (CDRH), who generally issues the order for postmarket surveil- lance;

  - Seeking internal review  of  the order under § 10.75 of this chapter;

  - Requesting an informal hearing under part 16 of this chapter; or

  - Requesting review by the Medical Devices Dispute Resolution Panel of the Medical Devices Advisory Com- mittee.

You may obtain guidance docu- ments that discuss these mechanisms from the Center for Devices and Radio- logical Health’s (CDRH’s) Web site.

[67 FR 38887, June 6, 2002, as amended at 72

FR 17400, Apr. 9, 2007]

##### § 822.23   Is the information in my sub- mission considered confidential?

We consider the content of your sub- mission confidential until we have ap- proved your postmarket surveillance plan. After we have approved  your plan, the contents of the original sub- mission and any amendments, supple- ments, or reports may be disclosed in accordance with the Freedom of Infor- mation Act. We will continue to pro- tect trade secret and confidential com- mercial information after your plan is approved. We will not disclose informa- tion  identifying  individual  patients. You may wish to indicate in your sub- mission which information you con- sider trade  secret or  confidential com- mercial.

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart E—Responsibilities of Manufacturers

##### § 822.24 What are my responsibilities once I am notified that I am re- quired to conduct postmarket sur- veillance?

You must submit your plan to con- duct postmarket surveillance to us within 30 days from receipt of the order (letter) notifying you that you are re- quired to conduct postmarket surveil- lance of a device.

##### § 822.25 What are my responsibilities after my postmarket surveillance plan has been approved?

After we have approved your plan, you must conduct the postmarket sur- veillance of your device in accordance with your approved plan. This means that you must ensure that:

Postmarket surveillance is initi- ated in a timely manner;

The surveillance is conducted with due diligence;

The data identified in the plan is collected;

Any reports required as part  of your approved plan are submitted to us in a timely manner; and

Any information that we request prior to your submission of a report or in response to our review of a report is provided in a timely manner.

##### § 822.26 If my company changes own- ership, what must I do?

You must notify us within 30 days of any change in ownership of your com- pany. Your notification should identify any changes to the name or address of the  company,  the  contact  person,  or the  designated  person  (as  defined  in

§ 822.3(b)). Your obligation to conduct postmarket surveillance will generally transfer to the new owner, unless you and the new owner have both agreed that you will continue to conduct the surveillance. If you will continue to conduct the postmarket surveillance, you still must notify  us  of  the  change in ownership.

##### § 822.27 If I go out of business, what must I do?

You must notify us within 30 days of the date of your decision to close your

#### Food and Drug Administration, HHS	§ 822.33

business. You should provide the ex- pected date of closure and discuss your plans to complete or terminate postmarket surveillance of your de- vice. You must also identify who will retain the records related to the sur- veillance (described in subpart G of this part) and where the records will be kept.

##### § 822.28 If I stop marketing the device subject to postmarket surveillance, what must I do?

You must continue to conduct postmarket surveillance in accordance with your approved plan even if you no longer market the device. You may re- quest that we allow you to terminate postmarket surveillance or  modify your postmarket surveillance because you no longer market the device. We will make these decisions on a case-by- case basis, and you must continue to conduct the postmarket surveillance unless we notify you that you may stop your surveillance study.

## Subpart F—Waivers and Exemptions

##### § 822.29 May I request a waiver of a specific requirement of this part?

You may request that we waive any specific requirement of this part. You may submit your request, with sup- porting documentation, separately or as a part of your postmarket surveil- lance  submission  to  the  address  in

§ 822.8.

##### § 822.30 May I request exemption from the requirement to conduct postmarket surveillance?

You may request exemption from the requirement to conduct postmarket surveillance for your device or any spe- cific model of that device at any  time. You must comply with the require- ments of this part unless and until we grant an exemption for your  device. Your request for exemption must ex- plain why you believe we should ex- empt the device or model from postmarket surveillance. You should demonstrate why the surveillance question does not apply to your device or does not need to be answered for the device for which you are requesting ex- emption.  Alternatively,  you  may  pro-

vide information that answers the sur- veillance question for your device, with supporting documentation, to the ad- dress in § 822.8.

## Subpart G—Records and Reports

##### § 822.31 What records am I required to keep?

You must keep copies of:

All correspondence with your in- vestigators or FDA, including required reports;

Signed agreements from each of your investigators, if your surveillance plan uses investigators, stating the commitment to conduct the surveil- lance in accordance with the approved plan, any applicable FDA regulations, and any conditions of approval for your plan, such as reporting requirements;

Your approved postmarket sur- veillance plan, with documentation of the date and reason for any deviation from the plan;

All data collected and analyses conducted in support of your postmarket surveillance plan; and

Any other records that we require to be maintained by regulation or by order, such as copies of signed consent documents, evidence of Institutional Review Board review and approval, etc.

##### § 822.32 What records are the inves- tigators in my surveillance plan re- quired to keep?

Your investigator must keep copies of:

All correspondence between inves- tigators, FDA, the manufacturer, and the designated person, including re- quired reports.

The approved postmarket surveil- lance plan, with documentation of the date and reason for any deviation from the plan.

All data collected and analyses conducted at that site for postmarket surveillance.

Any other records that we require to be maintained by regulation or by order.

##### § 822.33 How long must we keep the records?

You, the designated person, and your investigators must keep all records for

§ 822.34

a period of 2 years after we have ac- cepted your final report, unless we specify otherwise.

##### § 822.34 What must I do with the records if the sponsor of the plan or an investigator in the plan changes?

If the sponsor of the plan or an inves- tigator in the plan changes, you must ensure that all records related to the postmarket surveillance have been transferred to the new sponsor or in- vestigator and notify us within 10 working days of the effective date  of the change. You must provide the name, address, and telephone  number of the new sponsor or investigator, cer- tify that all records have been trans- ferred, and provide the date of transfer.

##### § 822.35 Can you inspect my manufac- turing site or other sites involved in my postmarket surveillance plan?

We can review your postmarket sur- veillance programs during regularly scheduled inspections, inspections ini- tiated to investigate recalls or other similar actions, and inspections initi- ated specifically to review your postmarket surveillance plan. We may also inspect any other person or site involved in your postmarket surveil- lance, such as investigators or contrac- tors. Any person authorized to grant access to a facility must permit au- thorized FDA employees to enter and inspect any facility where the device is held or where records regarding postmarket surveillance are held.

##### § 822.36 Can you inspect and copy the records related to my postmarket surveillance plan?

We may, at a reasonable time and in a reasonable manner, inspect and copy any records pertaining to  the  conduct of postmarket surveillance that are re- quired to be kept by this regulation. You must be able to produce records and information required by this regu- lation that are in the possession of oth- ers under contract with you to conduct the postmarket surveillance.  Those who have signed agreements or are under contract with you must also produce the records and information upon our request. This information must be produced within 72 hours  of the   initiation   of   the   inspection.   We

#### 21 CFR Ch. I (4–1–17 Edition)

generally will redact information per- taining to individual subjects prior to copying those records, unless there are extenuating circumstances.

##### § 822.37 Under what circumstances would you inspect records identi- fying subjects?

We can inspect and copy records identifying subjects under the same circumstances that we can inspect any records relating to postmarket surveil- lance. We are likely to be interested in such records if we have reason to be- lieve that required reports have not been submitted, or are incomplete, in- accurate, false, or misleading.

##### § 822.38   What reports must I submit to you?

You must submit interim and final reports as specified in your approved postmarket surveillance plan. In addi- tion, we may ask you to submit addi- tional information when we believe that the information is necessary for the protection of the public health and implementation of the act. We will also state the reason or purpose for the re- quest and how we will use the informa- tion.

## PART 830—UNIQUE DEVICE IDENTIFICATION

##### Subpart A—General Provisions

830.3    Definitions.

##### Subpart B—Requirements for a Unique Device Identifier

Sec.

830.10    Incorporation by reference.

830.20    Requirements   for   a   unique   device identifier.

830.40   Use  and  discontinuation  of  a  device identifier.

830.50    Changes that require use of a new de- vice   identifier.

830.60    Relabeling of a device that is required to  bear  a  unique  device  identifier.

##### Subpart C—FDA Accreditation of an Issuing Agency

830.100    FDA    accreditation    of    an    issuing agency.

830.110    Application  for  accreditation  as  an issuing   agency.

  - Responsibilities  of  an  FDA-accred- ited issuing agency.

Food and Drug Administration, HHS	§ 830.3

830.130 Suspension or revocation of the ac- creditation  of  an  issuing  agency.

##### Subpart D—FDA as an Issuing Agency

830.200 When FDA will act as an issuing agency.

830.210 Eligibility for use of FDA as an issuing agency.

830.220 Termination of  FDA service  as  an issuing   agency.

##### Subpart E—Global Unique Device Identification Database

830.300 Devices subject to device identifica- tion  data  submission  requirements.

830.310 Information required for unique de- vice    identification.

830.320 Submission of unique device identi- fication     information.

830.330 Times for submission of unique de- vice    identification    information.

830.340 Voluntary submission of ancillary device    identification    information.

830.350 Correction  of information submitted to the  Global  Unique  Device  Identifica- tion   Database.

830.360 Records to be maintained by the la- beler.

AUTHORITY: 21 U.S.C. 321, 331, 352, 353, 360, 360d, 360i, 360j, 371.

SOURCE: 78 FR 55823, Sept. 24, 2013, unless

otherwise noted.

## Subpart A—General Provisions

SOURCE: 78 FR 55825, Sept. 24, 2013, unless

otherwise noted.

##### § 830.3   Definitions.

As used in this part:

Automatic identification and data cap- ture (AIDC) means any technology that conveys the unique device identifier or the device identifier of a device in a form that can be entered into an elec- tronic patient record or other com- puter system via an automated proc- ess.

Center Director means the Director of the Center for Devices and Radio- logical Health or the Director of the Center for Biologics Evaluation and Research, depending on which Center has been assigned lead responsibility for the device.

Device package means a package that contains a fixed quantity of a par- ticular version or model of a device.

Expiration date means the date by which the label of a device states the device must or should be used.

FDA, we, or us means the Food and Drug Administration.

Federal Food, Drug, and Cosmetic Act means 21 U.S.C. 321 et seq., as amended. Finished device means  any  device  or accessory to any device that is suitable

for use or capable of functioning.

Global Unique Device Identification Database (GUDID) means the database that serves as a repository of informa- tion to facilitate the identification of medical devices through their distribu- tion and use.

Human cell, tissue, or cellular or tissue- based product (HCT/P) regulated as a de- vice  means   an   HCT/P   as   defined   in

§ 1271.3(d) of this chapter that does not meet the criteria in § 1271.10(a) and that is also  regulated  as a  device.

Issuing agency means an organization accredited by FDA to operate a system for the issuance of unique device iden- tifiers.

Label has the meaning set forth in section 201(k) of the Federal Food, Drug, and Cosmetic Act.

Labeler means:

    - Any person who causes a label to be applied to a device with the intent that the device will be commercially distributed without any subsequent re- placement or modification of the label; and

    - Any person who causes the label of a device to be replaced or modified with the intent that the device will be commercially distributed without any subsequent replacement or modifica- tion of the label, except that the addi- tion of the name of, and contact infor- mation for, a person who distributes the device, without making any other changes to the label, is not a modifica- tion for the purposes of determining whether a person is a labeler.

Lot or batch means one finished device or more that consist of a single type, model, class, size, composition, or soft- ware version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits.

Shipping container means a container used during the shipment or transpor- tation  of  devices,  and  whose  contents

§ 830.10

may  vary  from  one  shipment  to  an- other.

Small business means a medical device manufacturer with 500 or fewer em- ployees, or a medical  device  relabeler or repackager with 100 or fewer em- ployees.

Specification means any requirement with which a device must conform.

Unique device identifier (UDI) means an identifier that adequately identifies a device through its distribution and use  by  meeting  the  requirements  of

§ 830.20. A UDI is composed of:

      - A device identifier—a mandatory, fixed portion of a UDI that  identifies the specific version or model of a de- vice and the labeler of that device; and

      - A production identifier—a condi- tional, variable portion of a UDI that identifies one or more of the following when included on the label of the de- vice:

        - The lot or batch within which a device was manufactured;

        - The serial number of a specific device;

        - The expiration date of a specific device;

        - The date a specific device was manufactured.

        - For  an  HCT/P  regulated  as  a  de- vice,   the   distinct   identification   code required by § 1271.290(c) of this chapter. Universal  product  code  (UPC)  means the  product  identifier  used  to  identify an  item  sold  at  retail  in  the  United

States.

Version or model means all devices that have specifications, performance, size, and composition, within limits set by the labeler.

## Subpart B—Requirements for a Unique Device Identifier

##### § 830.10   Incorporation by reference.

Certain material is incorporated by reference into this part with the ap- proval of the Director of the Federal Register in accordance with 5 U.S.C. 552(a) and 1 CFR part 51. To  enforce any edition other than that specified in this section, the Food and Drug Admin- istration must publish notice of change in the FEDERAL REGISTER and the ma- terial must be available to  the  public. All approved material is available for inspection  at  the  Division  of  Dockets

21 CFR Ch. I (4–1–17 Edition)

Management (HFA–305), Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852, 301–827–6860,

and is available from the source listed in paragraph (b) of this section. Copies are  also  available  for  purchase  from the  American  National  Standards  In- stitute  (ANSI),  mailing  address:  ANSI, Attn: Customer Service Department, 25 West 43rd St., 4th floor, New York, NY 10036, phone: 212–642–4980, and may be ordered	online		at	http:// webstore.ansi.org/. The material is also available for inspection at the National Archives  and  Records  Administration (NARA). For information on the avail- ability  of  this  material  at  NARA,  call 202–741–6030.

International Organization for Standardization (ISO), mailing address: ISO,  Attn:  ISO  Central  Secretariat,   1, ch. de la Voie-Creuse, Case postale 56, CH–1211 Geneva 20, Switzerland, phone (dialing from the United States): 011– 41–22–749–0111, and may be ordered on- line

  - ISO/IEC 646:1991(E), Information technology—ISO 7-bit  coded  character set for information interchange (third edition;     December     15,     1991),     into

§§ 830.20(c) and 830.100(b);

(2)  ISO/IEC  15459–2:2006(E),  Informa-

tion technology—Unique  identifiers— Part 2: Registration procedures (second edition; March 1, 2006), into §§ 830.20(b) and 830.100(b);

(3)  ISO/IEC  15459–4:2008(E),  Informa-

tion technology—Unique  identifiers— Part 4: Individual items  (second  edi- tion; July 15, 2008), into §§ 830.20(b) and 830.100(b);

(4)  ISO/IEC  15459–6:2007(E),  Informa-

tion technology—Unique  identifiers— Part 6: Unique identifier for product groupings (first edition; June 15, 2007), into §§ 830.20(b) and 830.100(b).

##### § 830.20   Requirements for a unique de- vice identifier.

A unique device identifier (UDI) must:

Be issued under a system operated by FDA or an FDA-accredited issuing agency;

Conform to each of the following international standards:

#### Food and Drug Administration, HHS	§ 830.110

  - ISO/IEC 15459–2, which is incor- porated by reference at § 830.10;

  - ISO/IEC 15459–4, which is incor- porated by reference at § 830.10; and

  - ISO/IEC 15459–6, which is incor- porated by reference at § 830.10.

Use only characters and numbers from the invariant character set of ISO/IEC 646, which is incorporated by reference at § 830.10.

[78 FR 55825, Sept. 24, 2013]

##### § 830.40   Use  and  discontinuation  of  a device identifier.

Only one device identifier from any particular system for the issuance of unique device identifiers (UDIs) may be used to identify a particular version or model of a device. A particular version or model may be identified by UDIs from two or more systems for the issuance of UDIs.

A device identifier shall be used to identify only one version or model.

In the event that a version or model of a device is discontinued, its device identifier may not be reassigned to another device. If a discontinued version or model is re-introduced and no changes have been made that would require the use of a new device identi- fier, the device identifier that was pre- viously in use may be used to identify the device.

In the event that  an  issuing  agen- cy relinquishes or does not renew its accreditation, you  may  continue  to  use a previously issued UDI  until  such  time as § 830.50 requires you to assign a new device   identifier.

[78 FR 55825, Sept. 24, 2013]

##### § 830.50   Changes that require use of a new device identifier.

Whenever you make a change to a device that is required to bear a unique device identifier (UDI) on its label, and the change results in a new version or model, you must assign a new device identifier to the new version or model.

Whenever you create a new device package, you must assign a new device identifier to the new device package.

[78 FR 55825, Sept. 24, 2013]

##### § 830.60 Relabeling of a device that is required to bear a unique device identifier.

If you relabel a device that is re- quired to bear a unique device identi- fier (UDI), you must:

Assign a new device identifier to the device, and

Keep a record showing the rela- tionship of the prior device identifier to your new device identifier.

[78 FR 55825, Sept. 24, 2013]

## Subpart C—FDA Accreditation of an Issuing Agency

##### § 830.100 FDA accreditation of an issuing agency.

Eligibility. A private organization may apply for accreditation as an issuing agency.

Accreditation criteria. FDA may ac- credit an organization as an issuing agency, if the system it will operate:

  - Will employ unique device identi- fiers (UDIs) that meet the require- ments of this part to adequately iden- tify a device through its distribution and use;

  - Conforms to each of the following international standards:

    - ISO/IEC 15459–2, which is incor- porated by reference at § 830.10;

    - ISO/IEC 15459–4, which is incor- porated by reference at § 830.10;

    - ISO/IEC 15459–6, which is incor- porated by reference at § 830.10.

  - Uses only characters and numbers from the invariant character set of ISO/IEC 646, which is incorporated by reference at § 830.10.

  - Will be available to all users ac- cording to a single set of consistent, fair, and reasonable terms and condi- tions.

  - Will protect against conflicts of interest between the issuing agency (and its officers, employees, and other agents) and labelers (and their officers, employees, and other agents)  seeking to use UDIs that may impede the appli- cant’s ability  to  independently  operate a fair and neutral identifier system.

##### § 830.110 Application for accreditation as an issuing agency.

Application for initial accreditation.

An  applicant  seeking  initial  FDA

§ 830.110

accreditation   as   an   issuing   agency shall notify FDA of its desire to be ac- credited by sending a notification by email   to:    GUDIDSupport@fda.hhs.gov, or by correspondence to: UDI Regu- latory Policy Support, Center for  De- vices  and  Radiological   Health,   Food and Drug Administration, 10903 New Hampshire Ave., Bldg. 66, Rm. 3303, Sil- ver Spring, MD 20993–0002.

FDA will provide the applicant with additional information to aid in submission of an application for ap- proval as an issuing agency, together with an email address for submission of an application.

The applicant shall furnish to FDA, via email to the email address provided in paragraph (a)(1) of this sec- tion, an application containing the fol- lowing information, materials, and supporting  documentation:

Name, address, and phone number of the applicant;

Detailed descriptions of any standards or criteria the applicant will apply to participating labelers;

A detailed description of the guidelines that govern assignment of a unique device identifier (UDI) to a de- vice;

A detailed description of the re- view and decisionmaking process the applicant will apply when determining whether a particular labeler may use the applicant’s UDI system, including:

  - Copies of the application forms, guidelines, instructions, and other ma- terials the applicant will send to med- ical device labelers who wish to use the applicant’s unique device identification system;

  - Policies and procedures for noti- fying a labeler of deficiencies in its use of UDIs;

  - Procedures for monitoring a la- beler’s correction of deficiencies in its use of UDIs;

  - Policies and procedures for sus- pending or revoking a labeler’s use of the applicant’s UDI system, including any appeals process.

Description of the applicant’s electronic data management  system with respect to its review and decision processes and the applicant’s ability to provide electronic data in a format compatible with FDA data systems;

#### 21 CFR Ch. I (4–1–17 Edition)

Fee schedules, if any, together with an explanation of any fee waivers or reductions that are available;

Detailed information regarding any financial or other relationship be- tween the applicant and any labeler(s) or governmental entity(ies); and

Other information required by FDA to clarify the application for ac- creditation.

Application for renewal of accredita- tion. An accredited issuing agency that intends to continue to serve as an issuing agency beyond its current term shall apply to FDA for renewal or no- tify FDA of its plans not to apply for renewal in accordance with the fol- lowing procedures and schedule:

  - At least 9 months  before  the  date of expiration of its accreditation, an issuing  agency  shall  inform  FDA,  at the address given in paragraph (a)(1) of this section, of its intent to  seek  re- newal.

  - FDA will notify the issuing agen- cy of the relevant information, mate- rials, and supporting documentation that we will require the issuing agency to submit as part of the renewal proce- dure. We will tailor these requirements to reflect our experience with the issuing agency during the current and any prior period of accreditation. We will limit our request to the types of the information required by paragraph (a)(3) of this section, and we will re- quire less information if experience shows that we need only a subset of that information.

  - At least 6 months before the date of expiration of its accreditation, an issuing agency shall furnish to FDA, at the email address we provide, a copy of a renewal application containing the information, materials, and supporting documentation requested by FDA in accordance with paragraph  (b)(2)  of this  section.

  - Any issuing agency that does not plan to renew its accreditation shall so notify FDA at the address given in paragraph (a)(1) of this section at least

9 months before the expiration of the issuing agency’s term of accreditation and shall include a description of its plans for allowing continued use  of UDIs issued prior to the expiration  of the current term of accreditation.

#### Food and Drug Administration, HHS	§ 830.130

FDA action on an application for initial or renewal accreditation. (1) FDA will conduct a review and evaluation to determine whether the applicant meets the requirements of this subpart and whether the UDI system proposed by the applicant will meet the require- ments of this subpart.

Within 60 days of receipt of an ap- plication for accreditation, FDA will notify the applicant of any deficiencies in its application and will request cor- rection of those deficiencies within 60 days. The applicant may request an ex- tension if it needs additional time to correct deficiencies in  its  application. If the deficiencies are not resolved to FDA’s satisfaction within the specified time period, the application for accred- itation as an issuing agency may be de- nied.

FDA shall  notify the  applicant whether the application for accredita- tion has been granted or denied. That notification shall list any conditions of approval or state the reasons for de- nial.

If FDA denies an application, we will advise the applicant of the cir- cumstances under which a denied appli- cation may be resubmitted.

If FDA does not reach a final deci- sion on a renewal application before the expiration of an issuing agency’s current accreditation, the approval will be deemed extended until FDA reaches a final decision on the applica- tion.

Relinquishment of accreditation. If an issuing agency decides to relinquish its accreditation  before  expiration  of the current term of  accreditation,  it shall submit a letter of such intent to FDA, at the address provided in para- graph (a)(1) of this section, at least 9 months before relinquishing its accred- itation.

Notice of termination of accredita- tion. An issuing agency that does not apply for renewal of its accreditation, is denied renewal of accreditation by FDA, or relinquishes its accreditation and duties before expiration of the cur- rent term of accreditation, shall notify all labelers that are using the issuing agency’s UDI system, in a manner and time period approved  by FDA,  of the date that the issuing agency will cease

to serve as an FDA-accredited issuing agency.

Term of accreditation. The initial term of accreditation for an issuing agency shall be for a period of 3 years. An issuing agency’s term of accredita- tion may be periodically renewed for a period of 7 years.

[78 FR 55825, Sept. 24, 2013, as amended at 81

FR 11429, Mar. 4, 2016]

##### § 830.120 Responsibilities of an FDA- accredited issuing agency.

To maintain its accreditation, an issuing agency must:

Operate a system for  assignment of unique device identifiers (UDIs) that meets the requirements of § 830.20;

Make available information con- cerning its system for the assignment of UDIs;

Maintain a  list  of  labelers  that use its system for the assignment of UDIs and provide FDA a copy  of  such list in electronic form by December 31 of each year;

Upon request, provide FDA with information concerning a labeler that is employing the issuing agency’s sys- tem for assignment of UDIs; and

Remain in compliance with the eligibility  and  accreditation   criteria set forth in § 830.100.

##### § 830.130 Suspension or revocation of the accreditation of an issuing agency.

FDA may suspend or revoke the ac- creditation of an issuing agency if FDA finds, after providing the issuing agen- cy with notice and opportunity for an informal hearing in accordance  with part 16 of this chapter, that the issuing agency or any officer, employee, or other agent of the issuing agency:

Has been guilty of misrepresenta- tion or failure to disclose required in- formation  in  obtaining  accreditation;

Has failed to fulfill the respon- sibilities outlined in § 830.120;

Has failed to protect against con- flicts of interest that may impede the issuing agency’s ability to independ- ently operate a fair and neutral identi- fier system;

In the operation of the issuing agency, has engaged in any anti- competitive activity to restrain trade; or

#### § 830.200

Has violated or aided and abetted in the violation  of  any  regulation issued under section 510(e) or section 519(f) of the Federal Food, Drug, and Cosmetic  Act.

## Subpart D—FDA as an Issuing Agency

SOURCE: 78 FR 55826, Sept. 24, 2013, unless

otherwise noted.

##### § 830.200   When   FDA   will   act   as   an issuing agency.

During any period where there is no accredited issuing agency, FDA will act as an issuing agency.

If FDA determines that a signifi- cant number of small businesses would be substantially and adversely affected by the fees required by all accredited issuing agencies, FDA will act as an issuing agency.

FDA may, in its discretion, act as an issuing agency if we determine it is necessary for us to do so to ensure the continuity or the effectiveness of the system for the identification of med- ical devices.

FDA may, in its discretion,  act  as an issuing agency if we determine it is appropriate for us to do so in order to facilitate or implement an alternative granted under § 801.55 of this chapter.

##### § 830.210   Eligibility for use of FDA as an issuing agency.

When FDA acts as an issuing agency, any labeler will be permitted to use FDA’s unique device identification sys- tem, regardless of whether the labeler is considered a small business.

##### § 830.220   Termination  of  FDA  service as an issuing agency.

FDA may end our services as an issuing agency if we determine that the conditions that prompted us to act no longer exist and that ending our serv- ices would not be likely to lead to a re- turn of the conditions that prompted us to act.

If FDA has ended our services as an issuing agency, a labeler may con- tinue to use a device identifier as- signed under FDA’s unique device iden- tification  system  until  such  time  as

§ 830.50 requires the use of a new device identifier.

21 CFR Ch. I (4–1–17 Edition)

Subpart E—Global Unique Device Identification Database

SOURCE: 78 FR 55826, Sept. 24, 2013, unless

otherwise noted.

##### § 830.300 Devices subject to device identification data submission re- quirements.

- In general. The labeler of a device must provide the information required by this subpart for each version or model required to bear a unique device identifier (UDI).

- Voluntary submission of informa- tion. If a labeler voluntarily includes a UDI  on  the  label  of  a  device  under

§ 801.40, the labeler may also volun- tarily submit information concerning that device under this part.

- Exclusions. FDA may reject or re- move any device identification data where:

  - The device identifier  submitted does not conform to § 830.20;

  - The information concerns a device that is neither manufactured in the United States nor in interstate com- merce in the United States,

  - The information concerns a prod- uct that FDA determines is not a de- vice or a combination product that in- cludes a device constituent part,

  - The information concerns a device or a combination product that re- quires, but does not have, FDA pre- market approval, licensure, or clear- ance;

  - A device that FDA has banned under section 516 of the Federal Food,

Drug, and Cosmetic Act; or

  - FDA has suspended the accredita- tion of the issuing agency that oper- ates the system used by the labeler.

##### § 830.310 Information required for unique device identification.

The contact for device identification designated under § 830.320(a) shall pro- vide FDA with  the  following  informa- tion concerning each version  or  model of a device required to bear a unique device  identifier  (UDI)  on  its  label:

    - Concerning the labeler:

      - The name of the labeler;

#### Food and Drug Administration, HHS	§ 830.320

      - A telephone number or email ad- dress that will allow FDA to commu- nicate with the contact for device iden- tification designated under § 830.320(a); and

      - The name of each issuing agency whose system is used by the labeler to assign UDIs used by the labeler.

    - Concerning each version or model of a device with a UDI on its label:

      - The device identifier portion of the UDI assigned to the version or model;

      - When reporting a substitution of a new device identifier that will be used in lieu of a previously reported identi- fier, the device identifier that was pre- viously assigned to the version or model;

      - If § 801.45 of this chapter requires the device to bear a UDI as a perma- nent marking on the device itself, ei- ther:

        - A statement that the device iden- tifier that appears as a permanent marking on the device is identical to that reported under paragraph (b)(1) of this section, or

        - The device identifier portion of the UDI that appears as a permanent marking on the device;

      - The proprietary, trade, or brand name of the device as it appears on the label of the device;

      - Any version or model number or similar reference that appears on the label of the device;

      - If the device is labeled as sterile, a statement to that effect;

      - If the device is labeled as con- taining natural rubber latex  that  con- tacts humans, or  is labeled as having packaging containing natural  rubber latex that contacts humans,  as  de- scribed by §§ 801.437(b)(1), 801.437(b)(3), and 801.437(f) of this chapter, a state- ment  to  that  effect;

      - Whether a patient may be safely exposed to magnetic resonance imag- ing, nuclear magnetic resonance imag- ing, or magnetic resonance tomog- raphy while using the device, or while the device is implanted in patient.

      - If the device is available in more than one size, the size of the particular version or model, together with the unit of measure, as it appears on the label of the device;

      - The type of production identi- fiers that appear on the label of the de- vice;

      - The FDA premarket submission number of a cleared or approved device, or a statement that FDA has by regula- tion exempted the device from pre- market notification;

      - The FDA listing number assigned to the device;

      - The Global Medical Device No- menclature (GMDN) term or code for the device;

      - The total number of individual devices contained in the device pack- age.

##### § 830.320   Submission of unique device identification information.

Designation of contact for device identification. Each labeler must des- ignate an individual to serve as the point of contact with FDA on matters relating to the identification of med- ical devices marketed by the labeler. The contact for device information is responsible for ensuring FDA is pro- vided with all information required by this part. The contact for device infor- mation may authorize an issuing agen- cy or any other person to provide infor- mation to FDA on behalf of the labeler.

Information shall be submitted via electronic means. All information re- quired by this subpart shall be sub- mitted electronically to FDA’s Global Unique Device Identification Database (GUDID) in a format that we can proc- ess, review, and archive, unless the la- beler has obtained a waiver from elec- tronic submission of unique device identifier (UDI) data.

Waiver from  electronic  submission.

A labeler may request a waiver from electronic submission of UDI data by submitting a letter addressed to the appropriate Center Director explaining why electronic submission is not tech- nologically  feasible;  send  the  request by email to: udi@fda.hhs.gov, or by cor- respondence to: UDI Regulatory Policy Support, Center for Devices and Radio- logical Health, Food and Drug Admin- istration, Bldg. 66, Rm. 3303, 10903 New Hampshire Ave., Silver Spring, MD 20993–0002.

If the establishment where the la- beler is located has obtained a waiver

#### § 830.330

from electronic submission of registra- tion and listing information under sec- tion 510(p) of the Federal Food, Drug, and Cosmetic Act, the labeler  is deemed to have a waiver from elec- tronic submission of UDI data.

A labeler that has a waiver from electronic submission of UDI data must send a letter containing all of the in- formation required by § 830.310, as well as any ancillary  information  permitted to be submitted under § 830.340 that the labeler wishes to submit,  within  the time permitted by § 830.330, addressed to:  UDI  Regulatory  Policy  Support, Center for Devices and Radiological Health, Food and Drug Administration, Bldg. 66, Rm. 3303, 10903 New Hampshire Ave., Silver Spring, MD 20993–0002.

##### § 830.330 Times for submission of unique device identification infor- mation.

  - The  labeler  shall  submit  to  FDA the information required by § 830.310 no later than the date the label of the  de- vice must bear a unique device identi- fier under § 801.20 of this chapter.

  - The labeler of a device shall sub- mit to FDA an update to the informa- tion required by § 830.310 whenever the information changes. The updated in- formation must be submitted no later than the date a device is first labeled with the changed information. If the information does not appear  on  the label of a device, the updated informa- tion must be submitted within 10 busi- ness days of the change.

##### § 830.340 Voluntary submission of an- cillary device identification infor- mation.

You may not submit any informa- tion to the Global Unique Device Iden- tification  Database  (GUDID)   other than that specified by § 830.310, except where FDA acts to permit the submis- sion of specified additional types of in- formation, termed ancillary informa- tion.

FDA will provide information through the FDA Web site concerning the types of ancillary information that may be submitted to the GUDID.

FDA may periodically change the types of ancillary information  that may  be  submitted  to  the  GUDID.  We

#### 21 CFR Ch. I (4–1–17 Edition)

will announce any change on the FDA Web site at least 60 days before making the change.

##### § 830.350 Correction of information submitted to the Global Unique De- vice Identification Database.

If FDA becomes aware that any information submitted to the Global Unique Device Identification Database (GUDID) appears to be incorrect or po- tentially misleading,  we  may  notify the labeler of the specific information that appears to be incorrect, and re- quest that the labeler provide cor- rected information or explain why the information is correct. The labeler must provide corrected information or provide a satisfactory explanation of why the information is correct within

30 days of receipt of FDA’s notifica- tion.

If the labeler does not respond to FDA’s notification within 30 days of re- ceipt, or if FDA  determines,  at  any time, that any  information  in  the GUDID is incorrect or could be mis- leading, we may delete or correct the information. Any action taken by FDA under this paragraph does not relieve the labeler of its responsibility under paragraph (a) of this section to provide corrected information or an expla- nation of why the information pre- viously submitted is correct.

##### § 830.360 Records to be maintained by the labeler.

Each labeler shall retain, and sub- mit to FDA upon specific  request, records showing all unique device iden- tifiers (UDIs) used to identify devices that must bear a UDI on their  label, and the particular version or model as- sociated with each device identifier. These records must be retained for 3 years from  the  date the  labeler  ceases to market the version or model.

Compliance with this section does not relieve the labeler of the need to comply with recordkeeping require- ments of any other FDA regulation.

## PART 860—MEDICAL DEVICE CLASSIFICATION PROCEDURES

##### Subpart A—General

Sec.

  - Scope.

Food and Drug Administration, HHS	§ 860.3

860.3    Definitions.

860.5 Confidentiality and use of data and in- formation submitted in connection with classification  and  reclassification.

860.7 Determination of safety and effective- ness.

##### Subpart B—Classification

860.84 Classification procedures for ‘‘old de- vices.’’

860.93 Classification of implants, life-sup- porting  or  life-sustaining  devices.

860.95 Exemptions from  sections  510,  519, and 520(f) of the act.

##### Subpart C—Reclassification

860.120  General.

860.123  Reclassification   petition:   Content and  form.

860.125    Consultation with panels.

860.130 General procedures under section 513(e) of the act.

860.132 Procedures when the Commissioner initiates a performance standard or pre- market approval proceeding under  sec- tion 514(b) or 515(b) of the act.

860.134 Procedures for ‘‘new devices’’ under section 513(f) of the act and reclassifica- tion  of  certain  devices.

  - Procedures for transitional products under section 520(l) of the act.

AUTHORITY: 21 U.S.C. 360c, 360d, 360e, 360i, 360j, 371, 374.

SOURCE: 43 FR 32993, July 28, 1978, unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 860 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General

##### § 860.1   Scope.

    - This part implements sections 513, 514(b), 515(b), and 520(l) of the act with respect to the classification and reclas- sification of  devices  intended  for human use.

    - This part prescribes the criteria and procedures to be used by classifica- tion panels in making their rec- ommendations and by the Commis- sioner in making the Commissioner’s determinations regarding the class of regulatory control (class I, class II, or class III) appropriate for particular de- vices. Supplementing the general Food and Drug Administration procedures governing advisory committees (part 14 of this chapter), this part also provides procedures for manufacturers, import- ers,   and   other   interested   persons   to

participate in proceedings  to  classify and reclassify devices. This part also describes the kind of data required for determination of the safety and effec- tiveness of a device, and the cir- cumstances under which information submitted to classification panels or to the Commissioner in connection with classification and reclassification pro- ceedings will be available to the public.

##### § 860.3   Definitions.

For the purposes of this part:

    - Act means the Federal Food, Drug, and Cosmetic Act.

    - Commissioner means the Commis- sioner of Food and Drugs, Food and Drug Administration, United States Department of Health and Human Services, or the Commissioner’s des- ignee.

    - Class means one of the three cat- egories of regulatory control for med- ical devices, defined below:

      - Class I means the class of devices that are subject to only the general controls authorized by or under sec- tions 501 (adulteration), 502 (mis- branding), 510 (registration), 516 (banned devices), 518 (notification and other remedies), 519 (records and re- ports), and 520 (general provisions) of the act. A device is in class I if (i) gen- eral controls are sufficient to provide reasonable assurance of the safety and effectiveness of the device, or (ii) there is insufficient information from  which to determine that general controls are sufficient to provide reasonable assur- ance of the safety and effectiveness  of the device or to establish special con- trols to provide such assurance, but the device is not life-supporting or life-sus- taining or  for a  use which  is  of substanial importance in preventing impairment of human health,  and which does not present a potential un- reasonable risk of illness of injury.

      - Class II means the class of devices that is or eventually will be subject to special controls. A device is in class II if general controls alone are insuffi- cient to provide reasonable assurance of its safety and effectiveness and there is sufficient information to establish special controls, including the promul- gation of performance standards, postmarket surveillance, patient reg- istries, development and dissemination

§ 860.3

of guidance documents (including guid- ance on the submission of clinical data in premarket notification submissions in accordance with section 510(k) of the act), recommendations, and other ap- propriate actions as the Commissioner deems necessary to provide such assur- ance. For a device that is purported or represented to be for use in supporting or sustaining human life, the Commis- sioner shall examine and identify the special controls, if any, that are nec- essary to provide adequate assurance of safety and effectiveness and describe how such controls provide such assur- ance.

      - Class III means the class of devices for which premarket approval is or will be required in accordance with section 515 of the act. A device is in class III if insufficient information exists to de- termine that general controls are suffi- cient to provide reasonable  assurance of its safety and effectiveness or that application of special controls de- scribed in paragraph (c)(2) of this sec- tion would provide such assurance and if, in addition, the device is life-sup- porting or life-sustaining, or for a use which is of substantial importance in preventing impairment of human health, or if the device presents a po- tential unreasonable risk of illness or injury.

    - Implant means a device that is placed into a surgically or naturally formed cavity of the human body. A de- vice is regarded as an implant for the purpose of this part only if it is in- tended to remain implanted continu- ously for a period of 30 days or more, unless the Commissioner determines otherwise in order to protect human health.

    - Life-supporting or life-sustaining de- vice means a device that is essential to, or that yields information that is es- sential to, the restoration or continu- ation of a bodily function important to the continuation of human life.

    - Classification questionnaire means a specific series of questions prepared by the Commissioner for use as guidelines by classification panels preparing rec- ommendations to the Commissioner re- garding classification and by peti- tioners submitting petitions for reclas- sification. The questions relate to the safety  and  effectiveness  characteristics

#### 21 CFR Ch. I (4–1–17 Edition)

of a device and the answers are de- signed to help the Commissioner deter- mine the proper classification of the device.

    - Supplemental data sheet means in- formation compiled by a classification panel or submitted in a petition for re- classification,  including:

      - A summary of the reasons for the recommendation (or petition);

      - A summary of the data upon which the recommendation (or peti- tion) is based;

      - An identification of the risks to health (if any) presented by the device;

      - To the extent practicable in the case of a class II or class III device, a recommendation for the  assignment of a priority for the application of the re- quirements of performance standards or premarket approval;

      - In the case of a class I device, a recommendation whether the device should be exempted from any of the re- quirements of registration, record- keeping and reporting, or good manu- facturing practice requirements of the quality  system  regulation;

      - In the case of an implant or a life- supporting or life-sustaining device for which classification in class III is not recommended, a statement of the rea- sons for not recommending that the de- vice be classified in class III;

      - Identification of any needed re- strictions on the use of the device, e.g., whether the device requires special la- beling, should be banned, or should be used only upon authorization of a prac- titioner licensed by law  to  administer or use such device; and

      - Any known existing standards ap- plicable to the device, device compo- nents, or device materials.

    - Classification panel means one of the several advisory committees estab- lished by the Commissioner under sec- tion 513 of the act and part 14 of this chapter for the purpose of making rec- ommendations to the Commissioner on the  classification  and  reclassification of devices and for other purposes pre- scribed by the act or by the Commis- sioner.

    - Generic type of device means a grouping of devices that do not differ significantly in purpose, design, mate- rials,  energy  source,  function,  or  any

#### Food and Drug Administration, HHS	§ 860.5

other feature related to safety and ef- fectiveness, and for which similar regu- latory controls are sufficient to pro- vide reasonable assurance of safety and effectiveness.

    - Petition means a submission seek- ing reclassification of a device in ac- cordance with § 860.123.

[43 FR 32993, July 28, 1978, as amended at 57

FR 58403, Dec. 10, 1992; 65 FR 56480, Sept. 19,

2000]

##### § 860.5 Confidentiality and use of data and information  submitted  in  con- nection with classification and re- classification.

This section governs the avail- ability for public disclosure and the use by the Commissioner of data and infor- mation submitted to classification panels or to the Commissioner in con- nection with the classification or re- classification  of  devices  under  this part.

In general, data and information submitted to classification panels in connection with the classification of devices under § 860.84 will be available immediately for public disclosure upon request.  However,  except  as  provided by the special rules in paragraph (c)  of this section, this provision does not apply to data and information exempt from public disclosure in accordance with part 20 of this chapter: Such  data and information will be  available  only in accordance with part 20.

(c)(1) Safety and effectiveness data submitted to classification panels or to the Commissioner in connection with the   classification   of   a   device   under

§ 860.84, which have not been disclosed previously  to  the  public,   as   described in § 20.81 of this chapter, shall be re- garded as confidential if the device is classified in to class III. Because the classification of a device under § 860.84 may be ascertained only upon publica- tion of a final regulation, all safety and effectiveness data that have not been disclosed previously  are  not  available for  public  disclosure  unless  and  until the device is classified  into  class  I  or II, in which case the procedure in para- graph (c)(2) of this section applies.

Thirty days after publication of a final regulation under  § 860.84 classifying a device into  class  I  or  class II,   safety   and   effectiveness   data   sub-

mitted for that device that had been regarded as confidential under para- graph (c)(1) of this section will be available  for  public  disclosure  and placed on public display in the  office  of the Division of Dockets Management, Food and Drug Administration unless, within that 30-day  period,  the  person who submitted the data  demonstrates that the data still fall within the ex- emption for trade secrets and confiden- tial  commercial  information  described in § 20.61 of this chapter. Safety and ef- fectiveness data submitted for a device that is classified into class III by regu- lation in accordance with § 860.84 will remain   confidential    and    unavailable for public disclosure  so  long  as  such data have not been disclosed to  the public as described in § 20.81 of this chapter.

Because device classification  af- fects generic types of devices, in  mak- ing determinations under § 860.84 con- cerning the initial classification of a device,  the   classification   panels   and the Commissioner may consider safety and effectiveness data developed for another  device  in  the  same   generic type, regardless  of  whether  such  data are regarded currently as confidential under paragraph (c)(1) of this section.

(d)(1) The fact of its existence  and the contents of a petition for reclassi- fication     filed     in     accordance     with

§ 860.130 or § 860.132 are available for public disclosure at the time the peti- tion is received by the Food and Drug Administration.

(2) The fact of the existence of a peti- tion for reclassification filed in accord- ance with § 860.134 or § 860.136 is avail- able for public disclosure at the  time the  petition  is  received  by  the  Food and Drug Administration. The  contents of such a petition are not available for public disclosure for the period of time following its receipt (not longer than 30 days) during which the petition is re- viewed for any deficiencies preventing the Commissioner from making a deci- sion on it. Once it is  determined  that the petition contains no deficiencies preventing the Commissioner from making a  decision  on  it,  the  petition will be filed with the  Division  of  Dock- ets Management and its entire  con- tents will be available for public disclo- sure   and   subject   to   consideration   by

#### § 860.7

classification panels and by the Com- missioner in making a decision on the petition. If, during this 30-day period of time, the petition is found to contain deficiencies that prevent the Commis- sioner from making a decision on it, the petitioner will be so notified and afforded an opportunity to correct the deficiencies.

Thirty days after notice to the peti- tioner of deficiencies in the petition, the contents of the petition will be available for public disclosure unless, within that 30 days, the petitioner sub- mits  supplemental  material  intended to correct the deficiencies in the peti- tion. The Commissioner, in the Com- missioner’s discretion, may allow with- drawal of a deficient petition during the 30-day period provided for cor- recting deficiencies. Any supplemental material submitted by the petitioner, together with the material in the origi- nal petition, is considered as a new pe- tition. The new petition is reviewed for deficiencies in the same manner as the original petition, and the same proce- dures for notification and correction of deficiencies are followed. Once the pe- titioner has corrected the deficiencies, the entire contents of the petition will be available for public disclosure and subject to consideration by classifica- tion panels and by the Commissioner in making a decision on the petition. De- ficient petitions which have not been corrected within 180 days after notifi- cation of deficiency will be returned to the petitioner and will not be consid- ered further unless resubmitted.

The Commissioner may not dis- close, or use as the basis for reclassi- fication of a  device  from  class  III  to class II, any information reported to or otherwise obtained by the  Commis- sioner under section 513, 514, 515, 516, 518, 519, 520(f), 520(g), or 704 of the act that falls within the  exemption  de- scribed in § 20.61 of  this  chapter  for trade secrets and confidential commer- cial information. The exemption de- scribed in § 20.61 does not apply to data or information contained in a  petition for reclassification submitted in ac- cordance with § 860.130 or § 860.132, or in a  petition  submitted  in  accordance with § 860.134 or § 860.136 that has been determined to contain no  deficiencies that   prevent    the   Commissioner   from

#### 21 CFR Ch. I (4–1–17 Edition)

making a decision on it.  Accordingly, all data and information contained in such petitions may be disclosed by the Commissioner and used as the basis for reclassification of a device from class III to class II.

For purposes of this section, safe- ty and effectiveness data include data and results derived from all studies and tests of a device on animals and hu- mans and from all studies and tests of the device itself intended  to  establish or determine its safety and effective- ness.

##### § 860.7   Determination of safety and ef- fectiveness.

The classification panels, in re- viewing evidence concerning the safety and effectiveness of a device and in pre- paring advice to the Commissioner, and the Commissioner, in making deter- minations concerning the safety and effectiveness of a device, will apply the rules in this section.

In determining the safety and ef- fectiveness of a device for purposes of classification, establishment of per- formance standards for class II devices, and premarket approval of class III de- vices, the Commissioner and the classi- fication panels will consider the fol- lowing, among other relevant factors:

  - The persons for whose use the de- vice is represented or intended;

  - The conditions of use for the de- vice, including conditions of use pre- scribed, recommended, or suggested in the labeling or advertising of the de- vice, and other intended conditions of use;

  - The probable benefit to health from the use of the device weighed against any probable injury or illness from such use; and

  - The reliability of the device.

(c)(1) Although the manufacturer may submit any form of evidence  to the Food and Drug Administration in an attempt to substantiate the safety and effectiveness of a device, the agen- cy relies upon only valid scientific evi- dence to determine whether there is reasonable assurance that the device is safe and effective. After considering the nature of the device and the rules in this section, the Commissioner will determine whether the evidence sub- mitted  or  otherwise  available  to  the

#### Food and Drug Administration, HHS	§ 860.7

Commissioner is valid scientific evi- dence for the purpose of determining the safety or effectiveness of a par- ticular device and whether the avail- able evidence, when taken as a whole, is adequate to support a determination that there is reasonable assurance that the device is safe and effective for its conditions of use.

(2) Valid scientific evidence is evi- dence from well-controlled investiga- tions, partially controlled studies, studies and objective trials without matched controls, well-documented case histories conducted by qualified experts, and reports of significant human experience with a marketed de- vice, from which it can fairly and re- sponsibly be concluded by qualified ex- perts that there is reasonable assur- ance of the safety and effectiveness of a device under its conditions of use. The evidence required may vary ac- cording to the characteristics of the device, its conditions of use, the exist- ence and adequacy of warnings and other restrictions, and the extent of ex- perience with its use. Isolated case re- ports, random experience, reports lack- ing sufficient details to permit sci- entific evaluation, and unsubstantiated opinions are not regarded as valid sci- entific evidence to show safety or effec- tiveness. Such information may be con- sidered, however, in identifying a de- vice the safety and effectiveness of which is questionable.

(d)(1) There is reasonable assurance that a device is safe when it can be de- termined, based upon valid scientific evidence, that  the probable benefits  to health from use of the device for its in- tended uses and conditions of use, when accompanied by adequate directions and warnings against unsafe use, out- weigh any probable risks. The valid sci- entific evidence used to determine the safety of a device shall adequately demonstrate the absence of unreason- able risk of illness or injury associated with the use of the device for its in- tended uses and conditions of use.

(2) Among the types of evidence that may be required, when appropriate, to determine that there is reasonable as- surance that a device is safe are inves- tigations using laboratory animals, in- vestigations  involving  human  subjects,

and  nonclinical  investigations  includ- ing in vitro studies.

(e)(1) There is reasonable assurance that a device is effective when it can be determined, based upon valid scientific evidence, that in a  significant  portion of the target population, the use of the device for its intended uses and condi- tions of use, when accompanied by ade- quate directions for use and warnings against unsafe use, will provide clini- cally  significant  results.

(2) The valid scientific evidence used to determine the effectiveness of a de- vice shall consist principally of well- controlled investigations, as defined in paragraph (f) of this section, unless the Commissioner authorizes reliance upon other valid  scientific  evidence  which the Commissioner has determined is sufficient evidence from which to de- termine the effectiveness of a device, even in the absence of well-controlled investigations. The Commissioner may make such a determination where the requirement of well-controlled inves- tigations in paragraph (f) of this sec- tion is not reasonably applicable to the device.

The following principles have been developed over a period of years and are recognized by the scientific com- munity as the essentials of a well-con- trolled clinical investigation. They provide the basis for the Commis- sioner’s determination whether there is reasonable assurance that a device is effective based upon well-controlled in- vestigations and are also useful in as- sessing the weight to be given to other valid scientific evidence  permitted under this section.

  - The plan or protocol for the study and the report of the results of a well- controlled investigation shall include the following:

    - A clear statement of the objec- tives of the study;

    - A method of selection of the sub- jects that:

Provides adequate assurance that the subjects are suitable for the pur- poses of the study, provides diagnostic criteria of the condition to be treated or diagnosed, provides confirmatory laboratory tests where appropriate and, in the case of a device to prevent a disease or condition, provides evi- dence of susceptibility and exposure to

#### § 860.7

the  condition  against  which  prophy- laxis is desired;

Assigns the subjects to test groups, if used, in such a way as to minimize any possible bias;

Assures comparability between test groups and any control groups of pertinent variables such as sex, sever- ity or duration of the disease, and use of therapy other than the test device;

    - An explanation of the methods of observation and recording of results utilized, including the variables meas- ured, quantitation, assessment of any subject’s response, and steps taken to minimize any possible bias of subjects and observers;

    - A  comparison  of  the  results  of treatment  or  diagnosis  with  a  control in  such  a  fashion  as  to  permit  quan- titative  evaluation.  The  precise  nature of the control must be specified and an explanation  provided  of  the  methods employed   to   minimize   any   possible bias  of  the  observers  and  analysts  of the data. Level and methods of ‘‘blind- ing,’’ if appropriate and used, are to be documented.   Generally,   four   types   of comparisons are recognized:

No treatments. Where objective measurements of effectiveness are available and placebo effect is neg- ligible, comparison  of the  objective re- sults in comparable groups of treated and untreated patients;

Placebo control. Where there may be a placebo effect with the use of a de- vice, comparison of the results of use of the device with an ineffective device used under conditions designed to re- semble the conditions of use under in- vestigation as far as possible;

Active treatment control. Where an effective regimen of therapy may be used for comparison, e.g., the condition being treated is such that the use of a placebo or the withholding of  treat- ment would be inappropriate or con- trary to the interest of the patient;

Historical control. In certain cir- cumstances, such as those involving diseases with high and predictable mor- tality or signs and symptoms of pre- dictable duration or severity, or in the case of prophylaxis where morbidity is predictable, the results of use of the de- vice may be compared quantitatively with prior experience historically de- rived from the adequately documented

#### 21 CFR Ch. I (4–1–17 Edition)

natural history of the disease or condi- tion in comparable patients or popu- lations who received no treatment or who followed an established effective regimen (therapeutic, diagnostic, pro- phylactic).

    - A summary of the methods of analysis and an evaluation of the data derived from the study, including any appropriate statistical methods uti- lized.

  - To insure the reliability of the re- sults of an investigation, a well-con- trolled investigation shall involve the use of a test device that is standardized in its composition or design and per- formance.

(g)(1) It is the responsibility of each manufacturer and importer of a device to assure that adequate, valid sci- entific evidence exists, and to furnish such evidence to the Food and Drug Administration to provide reasonable assurance that the device is safe and effective for its intended uses and con- ditions of use. The failure of a manu- facturer or importer of a device to present to the Food and Drug Adminis- tration adequate, valid scientific evi- dence showing that there is reasonable assurance of the safety and effective- ness of the device, if regulated by gen- eral controls alone, or by general con- trols and performance standards, may support a determination that the de- vice be classified into class III.

The Commissioner may require that a manufacturer, importer, or dis- tributor make reports or provide other information bearing on the classifica- tion of a device and indicating whether there is reasonable assurance of the safety and effectiveness of the device or whether it is adulterated or mis- branded under the act.

A requirement for a report or other information under this paragraph will comply with section 519 of the act. Accordingly,  the  requirement  will state the reason or purpose for such re- quest; will describe the required report or information as clearly as possible; will not be imposed on a manufacturer, importer, or distributor of a classified device that has been exempted from such a requirement in accordance with

§ 860.95; will prescribe the time for com- pliance with the requirement; and will prescribe    the    form    and    manner    in

#### Food and Drug Administration, HHS	§ 860.84

which the report or information is to be provided.

Required information that has been submitted previously to the Cen- ter for Devices and Radiological Health, the  Center  for  Biologics Eval- uation and Research, or the Center for Drug Evaluation and Research, as ap- plicable, need not be resubmitted, but may be incorporated by reference.

[43 FR 32993, July 28, 1978, as amended at 53

FR  11253,  Apr.  6,  1988;  73  FR  49942,  Aug.  25,

2008]

## Subpart B—Classification

##### § 860.84   Classification  procedures  for ‘‘old devices.’’

  - This subpart sets forth the proce- dures for  the  original  classification  of a device that either was in commercial distribution before May 28, 1976, or is substantially equivalent to a  device that was in commercial distribution before that date. Such a device will be classified  by  regulation  into  either class I (general controls), class II (spe- cial controls) or class III (premarket approval), depending upon the level of regulatory control required to provide reasonable assurance of the safety and effectiveness of the device (§ 860.3(c)). This subpart does not apply to a device that is classified into class III  by  stat- ute under section 513(f) of the act be- cause the Food and Drug Administra- tion has determined that the device  is not ‘‘substantially equivalent’’ to any device subject to this subpart or under section 520(l) (1) through (3) of the act because the device was regarded pre- viously as a new  drug. In classifying a device under this section, the Food and Drug Administration will follow the procedures described in paragraphs (b) through (g) of this section.

  - The Commissioner refers the de- vice to the appropriate classification panel organized and operated in accord- ance with section 513 (b) and (c) of the act and part 14 of this chapter.

  - In order to make recommenda- tions to the Commissioner on the class of regulatory control (class I, class II, or class III) appropriate for the device, the panel reviews the device for safety and effectiveness. In so doing, the panel:

Considers the factors set forth in

§ 860.7  relating  to  the  determination  of safety  and  effectiveness;

Determines the safety and effec- tiveness of the device on the  basis  of the  types  of  scientific  evidence   set forth in § 860.7;

Answers the questions in the clas- sification questionnaire applicable to the device being classified;

Completes a supplemental data sheet for the device;

Provides, to the maximum extent practicable, an opportunity for inter- ested persons to submit data and views on the classification of the device in accordance with part 14 of this chapter.

  - Based upon its review of evidence of the safety and effectiveness of the device, and applying the definition of each class in § 860.3(c), the panel sub- mits to the Commissioner a rec- ommendation regarding the classifica- tion of the device. The recommenda- tion will include:

    - A summary of the reasons for the recommendation;

    - A summary of the data upon which the recommendation is based, accompanied by references to the sources containing such data;

    - An identification of the risks to health (if any) presented by the device;

    - In the case of a recommendation for classification into class I, a rec- ommendation as to whether the device should be exempted from the require- ments of one or more of the following sections of the act: section 510 (reg- istration, product listing, and pre- market notification) section 519 (records and reports) and section 520(f) (good manufacturing practice require- ments of the quality  system  regula- tion) in accordance with § 860.95;

    - In the case of a recommendation for classification into class II  or  class III, to the extent practicable, a rec- ommendation for  the  assignment  to the device of a priority for the applica- tion of a performance standard or a premarket approval requirement;

    - In the case of a recommendation for classification of an implant or a life-supporting or life-sustaining device into class I or class II, a statement  of why premarket approval is not nec- essary to provide reasonable assurance of  the  safety  and  effectiveness  of  the

#### § 860.93

device, accompanied by references to supporting documentation and data satisfying the requirements  of  § 860.7, and an identification of the risks to health,  if  any,  presented  by  the  device.

  - A panel recommendation is re- garded as preliminary until the Com- missioner has reviewed it, discussed it with the panel if appropriate, and pub- lished a proposed regulation classifying the device. Preliminary panel rec- ommendations are filed in the Division of Dockets Management’s office upon receipt and are available to the public upon request.

  - The Commissioner publishes the panel’s recommendation in the FED- ERAL REGISTER, together with a pro- posed regulation classifying the device, and other devices of that generic type, and provides interested persons an op- portunity to submit comments on the recommendation and proposed regula- tion.

  - The Commissioner reviews the comments and issues a final regulation classifying the device and other devices of that generic type.  The  regulation will:

If classifying  the  device  into  class I, prescribe which, if any, of the re- quirements of sections 510, 519, and 520(f) of the act will not apply to the device and state the reasons for  mak- ing the requirements inapplicable, in accordance with § 860.95;

If classifying the device into class II or class III, at the discretion of the Commissioner, establish priorities for the application to the device of a per- formance standard  or a premarket  ap- proval requirement;

If classifying an implant, or life- supporting or life-sustaining device, comply with § 860.93(b).

[43 FR 32993, July 28, 1978, as amended at 57

FR 58404, Dec. 10, 1992; 64 FR 404, Jan. 5, 1999]

##### § 860.93 Classification of implants, life- supporting or life-sustaining de- vices.

  - The classification panel will rec- ommend classification into class III of any implant or life-supporting or life- sustaining device unless the panel de- termines that  such  classification is not necessary to provide reasonable assur- ance of the safety and effectiveness  of the  device.  If  the  panel  recommends

#### 21 CFR Ch. I (4–1–17 Edition)

classification  or  reclassification   of such a device into a class other than class III, it shall set forth in its rec- ommendation the reasons for so doing together with  references  to supporting documentation and data satisfying the requirements of § 860.7, and an identi- fication of the risks to health, if any, presented by the device.

  - The Commissioner will classify an implant or life-supporting or life-sus- taining device into class III unless the Commissioner determines that such classification is not necessary to pro- vide reasonable assurance of the safety and effectiveness of the device. If the Commissioner proposes to classify or reclassify such a device into a class other than class III, the regulation or order effecting such classification or reclassification will be accompanied by a full statement of the reasons for so doing. A statement of the reasons for not classifying or retaining  the  device in class III may be in the form of con- currence with the reasons for the rec- ommendation of the  classification panel, together with supporting docu- mentation and data satisfying the re- quirements of § 860.7 and an identifica- tion of the risks to health, if any, pre- sented by the device.

##### § 860.95   Exemptions from sections 510, 519, and 520(f) of the act.

A panel recommendation to the Commissioner that a device be classi- fied or reclassified into class I will in- clude a recommendation as to whether the device should be exempted from some or all of the requirements of one or more of the following sections of the act: Section 510 (registration, product listing and premarket notification), section 519 (records and reports), and section 520(f) (good manufacturing practice requirements of the quality system  regulation).

A regulation or an order classifying or reclassifying a  device into class  I will  specify which require- ments, if any, of sections 510, 519, and 520(f) of the act the device is to be ex- empted from, together with the reasons for such exemption.

The Commissioner will grant ex- emptions under this section only if the Commissioner determines that the re- quirements  from  which  the  device  is

#### Food and Drug Administration, HHS	§ 860.123

exempted are not necessary to provide reasonable assurance of the safety and effectiveness of the device.

## Subpart C—Reclassification

##### § 860.120   General.

  - Sections 513(e) and (f), 514(b), 515(b), and 520(l) of the act provide for reclassification of a device and pre- scribe the procedures to be followed to effect reclassification. The purposes of subpart C are to:

Set forth the requirements as to form and content of petitions for re- classification;

Describe the circumstances in which each of the five statutory reclas- sification provisions applies; and

Explain the procedure for reclassi- fication prescribed in the five statu- tory reclassification provisions.

  - The criteria for determining the proper class for a  device  are  set  forth in § 860.3(c). The reclassification of any device within a generic type of device causes the reclassification of all sub- stantially equivalent  devices  within that generic type. Accordingly, a peti- tion for the reclassification of a spe- cific device will be considered a peti- tion for reclassification of all substan- tially equivalent devices within the same generic type.

  - Any interested person may submit a petition for reclassification under section 513(e), 514(b), or 515(b). A manu- facturer or importer may submit a pe- tition for reclassification under section 513(f) or 520(l). The Commissioner may initiate the reclassification of a device classified into class III under sections 513(f) and 520(l) of the act.

[43 FR 32993, July 28, 1978, as amended at 57

FR 58404, Dec. 10, 1992]

##### § 860.123   Reclassification	petition: Content and form.

Unless otherwise provided in writ- ing by the Commissioner, any petition for reclassification of a device, regard- less of the section of the act under which it is filed, shall include the fol- lowing:

  - A specification of the type of de- vice for which reclassification is re- quested;

  - A statement  of the  action  re- quested  by  the  petitioner,  e.g.,  ‘‘It  is

requested that l device(s) be reclassi- fied from class III to a class II’’;

  - A completed supplemental data sheet applicable to the device for which reclassification is requested;

  - A completed classification ques- tionnaire applicable to  the device  for which reclassification is requested;

  - A statement of the basis for dis- agreement with the present classifica- tion status of the device;

  - A full statement of the reasons, together with  supporting  data  satis- fying the requirements of § 860.7, why the device should not be  classified  into its present classification and how the proposed classification will provide reasonable assurance of the safety and effectiveness  of  the  device;

  - Representative data and informa- tion known by the petitioner that are unfavorable to the petitioner’s posi- tion;

  - If the petition is based upon new information under section 513(e), 514(b), or 515(b) of the act, a summary of the new information;

  - Copies of source documents from which new information used to support the petition has been obtained (at- tached as appendices to the petition).

  - A financial certification or dis- closure statement or both as  required by part 54 of this chapter.

Each petition submitted pursuant to this section shall be:

  - For devices regulated by the Cen- ter for  Devices  and  Radiological Health, addressed to the Food and Drug Administration, Center for Devices and Radiological Health, Regulations Staff, 10903 New Hampshire Ave., Bldg. 66, Rm. 4438, Silver Spring, MD 20993–0002; for devices regulated by the Center for Biologics Evaluation and Research, ad- dressed to the Food and Drug Adminis- tration, Center for Biologics  Evalua- tion and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002; for devices regulated by the Center for Drug Evaluation and Re- search, addressed to the Food and Drug Administration, Center for Drug Eval- uation and Research, Central Docu- ment	Control	Room,	5901–B Ammendale Rd., Beltsville, MD 20705– 1266, as applicable.

§ 860.125

  - Marked clearly with the section of the act under which the petition is being submitted, i.e., ‘‘513(e),’’ ‘‘513(f),’’ ‘‘514(b),’’ ‘‘515(b),’’ or ‘‘520(l) Petition’’;

  - Bound in a volume or volumes, where necessary; and

  - Submitted in an original and two copies.

[43 FR 32993, July 28, 1978, as amended at 49

FR  14505,  Apr.  12,  1984;  53  FR  11253,  Apr.  6,

1988;  55  FR  11169,  Mar. 27,  1990;  63  FR  5254,

Feb. 2, 1998; 65 FR 17137, Mar. 31, 2000; 73 FR

49942, Aug. 25, 2008; 75 FR 20916, Apr. 22, 2010;

79 FR 77388, Dec. 24, 2014]

##### § 860.125   Consultation with panels.

When the Commissioner is re- quired to refer  a  reclassification  peti- tion to a classification panel for its recommendation under § 860.134, or is required, or chooses, to consult with a panel concerning a reclassification pe- tition, such as under § 860.130, § 860.132, or § 860.136, the Commissioner will dis- tribute a copy of the petition, or its relevant portions, to each panel mem- ber and will consult with the panel  in one of the following ways:

Consultation by telephone with at least a majority of current voting panel members and, when possible, nonvoting panel members;

Consultation by mail with at least a majority of current voting panel members and, when possible, nonvoting panel members; and

Discussion at a panel meeting.

The method of consultation cho- sen by the Commissioner will depend upon the importance and complexity of the subject matter involved and the time available for action. When time and circumstances permit, the Com- missioner will consult with a panel through discussion at a panel meeting.

When a petition is  submitted under § 860.134 for  a  post-enactment, not  substantially  equivalent   device (‘‘new device’’), in consulting with the panel the Commissioner will obtain a recommendation that includes the in- formation described in § 860.84(d). In consulting with  a  panel  about  a  peti- tion submitted under § 860.130, § 860.132, or § 860.136, the Commissioner may or may not obtain a formal recommenda- tion.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 860.130   General    procedures    under section 513(e) of the act.

Section 513(e) of the act applies to reclassification proceedings under the act based upon new information.

A proceeding to reclassify a de- vice under section 513(e) may be initi- ated:

  - On the initiative of the Commis- sioner alone;

  - On the initiative of the Commis- sioner in response to a request for change in classification based  upon new information, under section 514(b) or 515(b) of the act (see § 860.132); or

  - In response to the petition of an interested person, based upon new in- formation,   filed   in   accordance   with

§ 860.123.

By regulation promulgated under this section, the Commissioner may change the classification from class III into:

  - Class II if the Commissioner de- termines that special controls in addi- tion to general controls would provide reasonable assurance of the safety and effectiveness of the device and there is sufficient information to establish spe- cial controls to provide assurance; or

  - Class I if the Commissioner deter- mines that general controls would pro- vide reasonable assurance of the safety and effectiveness of the device.

The  rulemaking  procedures  in

§ 10.40 of this chapter apply to pro- ceedings to reclassify a device under section 513(e), except that the Commis- sioner may secure a recommendation with respect to a proposed reclassifica- tion from the classification panel to which the device was last referred. The panel will consider a proposed reclassi- fication submitted to it by the Com- missioner in accordance with the con- sultation procedures of § 860.125. Any recommendation submitted to the Commissioner by the panel will be pub- lished in the FEDERAL REGISTER  when the Commissioner promulgates a regu- lation under this section.

Within 180 days after the filing of a petition for  reclassification  under this section, the  Commissioner,  by order published in the FEDERAL REG- ISTER, will either deny the petition or give notice of his intent to initiate a change in the classification of the de- vice.

#### Food and Drug Administration, HHS	§ 860.134

If a device is reclassified under this section, the regulation effecting the reclassification may revoke any special control or premarket approval requirement that previously applied to the device but that is no longer appli- cable because of the change in classi- fication.

A regulation under this section changing the classification of a device from class III to class II may provide that such classification will not take effect until the effective date of a spe- cial control for the device established under section 514 of the act.

[43 FR 32993, July 28, 1978, as amended at 57

FR 58404, Dec. 10, 1992]

##### § 860.132 Procedures when the Com- missioner initiates a performance standard or premarket approval proceeding under section 514(b) or 515(b) of the act.

Sections 514(b) and 515(b) of the act require the Commissioner  to  pro- vide, by notice in  the  FEDERAL  REG- ISTER, an opportunity for interested parties to request a change in the clas- sification of a device based upon new information relevant to  its  classifica- tion when the Commissioner initiates a proceeding either to develop a perform- ance standard for the  device  if  in  class II, or to promulgate a regulation re- quiring premarket approval for the de- vice if in class III. In either case, if the Commissioner agrees that the new in- formation warrants a change in classi- fication, the Commissioner will  publish in the FEDERAL REGISTER notice of the Commissioner’s intent  to initiate  a proceeding under section 513(e) of the act and § 860.130 to effect such a change.

The procedures for effecting a change in classification under sections 514(b) and 515(b) of the act are as fol- lows:

Within 15 days after publication of the Commissioner’s notice referred  to in paragraph (a) of this section, an in- terested person files a petition for re- classification      in      accordance      with

§ 860.123.

The Commissioner consults  with the  appropriate  classification   panel with regard to the petition in accord- ance with § 860.125.

Within 60 days after publication of the notice referred to in paragraph (a)

of this section, the Commissioner, by order published in the FEDERAL REG- ISTER, either denies the petition  or gives notice of his intent to initiate a change in classification in accordance with § 860.130.

##### § 860.134 Procedures for ‘‘new devices’’ under section 513(f) of the act and reclassification of certain devices.

Section 513(f)(3) of the act applies to proceedings for reclassification of a device currently in class III by oper- ation of section 513(f)(1) of the act. This category includes any device that is to be first introduced or delivered for in- troduction into interstate  commerce for commercial distribution after May 28, 1976, unless:

  - It is substantially equivalent to another device that was in commercial distribution before that date and had not been regulated before that date as a new drug; or

  - It is substantially equivalent to another device that was not in com- mercial distribution before such date but which has been classified into class I or class II; or

  - The Commissioner has classified the device into class I or class II in re- sponse to a petition for reclassification under this section.

The Commissioner determines whether a  device  is  ‘‘substantially  equivalent’’ for purposes of the application of this section. If a manufacturer or importer believes that a device is not ‘‘substan- tially  equivalent’’  but  that  it  should not be in class III under the criteria in

§ 860.3(c), the manufacturer or importer may petition for reclassification under this section. A manufacturer or im- porter who believes that a device is ‘‘substantially equivalent’’ and wishes to proceed to market the device shall submit a premarket notification in ac- cordance with part 807 of this chapter. After considering a premarket notifica- tion, the Commissioner will determine whether the device is ‘‘substantially equivalent’’ and will notify the manu- facturer or importer of such determina- tion in accordance with part 807 of this chapter.

The procedures for effecting re- classification under section  513(f)  of the act are as follows:

#### § 860.136

  - The manufacturer or importer of the device petitions for reclassification of   the   device   in   accordance   with

§ 860.123.

  - Within 30 days  after  the  petition is filed, the Commissioner notifies the petitioner of any deficiencies in the pe- tition that prevent the Commissioner from making a decision on it and al- lows the petitioner to supplement a de- ficient petition. Within 30 days after any supplemental material is received, the Commissioner notifies the peti- tioner whether the petition, as supple- mented, is adequate for review.

  - After determining that the peti- tion contains no deficiencies pre- cluding a decision on it, the Commis- sioner may for good cause shown refer the petition to the appropriate classi- fication panel for its review and rec- ommendation whether to approve or deny the petition.

  - Within 90 days after the date the petition is referred to the panel, fol- lowing the review procedures  set  forth in § 860.84(c) for the original classifica- tion of an ‘‘old’’ device, the panel sub- mits to the Commissioner its rec- ommendation containing the informa- tion set forth in § 860.84(d). A panel rec- ommendation is regarded as prelimi- nary until the Commissioner has re- viewed it, discussed  it  with  the  panel, if appropriate, and developed a pro- posed reclassification order.  Prelimi- nary  panel  recommendations  are   filed in the Division of Dockets Management upon receipt and are available to the public  upon  request.

  - The panel recommendation is pub- lished  in  the   FEDERAL   REGISTER   as soon as practicable and interested per- sons are provided an opportunity to comment on the recommendation.

  - Within 90 days after the panel’s recommendation is received (and no more than 210 days after the date the petition was filed), the Commissioner denies or approves the petition by order in the form of a letter to the pe- titioner. If the Commissioner approves the petition, the order will classify the device into class I or class II in accord- ance   with   the   criteria   set   forth   in

§ 860.3(c) and subject to the applicable requirements of § 860.93, relating to the classification of implants,  life-sup- porting   or   life-sustaining   devices,   and

#### 21 CFR Ch. I (4–1–17 Edition)

§ 860.95,  relating  to  exemptions  from certain requirements of the act.

  - Within a reasonable time after issuance of an order under this section, the Commissioner announces the order by notice published in the FEDERAL REGISTER.

[43 FR 32993, July 28, 1978, as amended at 57

FR 58404,  Dec. 10,  1992;  73  FR 34860,  June 19,

2008]

##### § 860.136 Procedures for transitional products under section 520(l) of the act.

    - Section 520(l)(2) of the act applies to reclassification  proceedings  initi- ated by a manufacturer or importer for reclassification of a device currently in class III by operation of section 520(l)(1) of the act. This section applies only to devices that the Food and Drug Administration regarded as ‘‘new drugs’’ before May 28, 1976.

    - The procedures for effecting re- classification under section 520(l)  are as follows:

      - The manufacturer or importer of the device files a petition for reclassi- fication of the device  in  accordance with § 860.123.

      - Within 30 days  after  the  petition is filed, the Commissioner notifies the petitioner of any deficiencies in the pe- tition that prevent the Commissioner from making a decision on it, allowing the petitioner to supplement a  defi- cient petition. Within 30 days after any supplemental material is received, the Commissioner notifies the petitioner whether the petition, as supplemented, is adequate for review.

      - The Commissioner provides the petitioner an opportunity for a regu- latory hearing conducted in accordance with part 16 of this chapter.

      - The Commissioner consults  with the  appropriate  classification   panel with regard to the petition in accord- ance with § 860.125.

      - Within 180 days after the petition is filed (where the Commissioner has determined it to be adequate for re- view), the Commissioner, by order in the form of a letter to the petitioner, either denies the petition or classifies the device into class I or class II in ac- cordance  with  the  criteria  set  forth  in

§ 860.3(c).

#### Food and Drug Administration, HHS	§ 861.7

      - Within a reasonable time after issuance of an order under this section, the Commissioner announces the order by notice published in the FEDERAL REGISTER.

## PART 861—PROCEDURES FOR PER- FORMANCE STANDARDS DEVEL- OPMENT

##### Subpart A—General

Sec.

861.1   Purpose and scope.

861.5   Statement of policy.

861.7    Contents of standards.

##### Subpart B—Procedures for Performance Standards Development and Publication

861.20    Summary   of   standards   development process.

861.24   Existing    standard    as    a    proposed standard.

861.30    Development of standards.

861.34    Amendment or revocation of a stand- ard.

861.36    Effective dates.

  - Standards advisory committees.

AUTHORITY:  21  U.S.C.  351,  352,  360c,  360d,

360gg–360ss, 371, 374; 42 U.S.C. 262, 264.

SOURCE: 45 FR 7484, Feb. 1, 1980, unless oth-

erwise noted.

## Subpart A—General

##### § 861.1   Purpose and scope.

    - This part implements section 514 of the Federal Food, Drug, and Cos- metic Act (the act) with respect to the establishment, amendment, and rev- ocation of performance standards ap- plicable to devices intended for human use.

    - The Food and Drug  Administra- tion may determine that a performance standard, as described under special controls for class II devices in § 860.7(b) of this chapter, is necessary to provide reasonable assurance of the safety and effectiveness  of  the  device.  Perform- ance  standards  may  be  established  for:

A class II device;

A class III device which, upon the effective date of the standard, is reclas- sified into class II; and

A class III device,  as  a  condition to premarket approval under section 515 of the act, to reduce or eliminate a

risk  or  risks  associated  with  such  de- vice.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

[45 FR 7484, Feb. 1, 1980, as amended at 45 FR

23686, Apr. 8, 1980; 57 FR 58404, Dec. 10, 1992]

##### § 861.5   Statement of policy.

In carrying out its duties under this section, the Food and Drug Adminis- tration will, to the maximum extent practical:

Use personnel, facilities, and other technical support available in other Federal agencies;

Consult with other Federal agen- cies concerned with standard setting and other nationally or internationally recognized standard-setting entities; and

Invite participation, through con- ferences, workshops, or other means, by representatives of scientific, profes- sional, industry, or consumer organiza- tions who can make a significant con- tribution.

##### § 861.7   Contents of standards.

Any performance standard estab- lished under this part will include such provisions as the Food and Drug Ad- ministration determines are necessary to provide reasonable assurance of the safety and effectiveness of  the  device or devices for which it is established. Where necessary to provide such assur- ance, a standard will address (but need not be limited to):

Performance characteristics of the device;

The design, construction, compo- nents, ingredients, and properties of the device, and its compatibility with power systems and connections to such systems;

The manufacturing processes and quality control procedures applicable to the device;

Testing of the device on either a sample or a 100-percent basis by the manufacturer, or, if it is determined that no other more practical means are available to the Food and Drug Admin- istration to assure the conformity of the device to the standard, providing

§ 861.20

for testing by the Food and Drug Ad- ministration or a third person to en- sure that the device conforms to the standard;

The publication of the results of each test or of certain tests of the de- vice to show that the device conforms to the portions of the standard for which the test or tests were required;

Manufacturers’ certification to purchasers or to the Food and Drug Ad- ministration that the device conforms to the applicable performance stand- ard;

Restrictions on the sale and dis- tribution of the device, but only to the extent authorized under section 520(e) of the act;

The use, and the form and con- tent, of labeling for the proper installa- tion, maintenance, operation, and use of the device. Among the provisions that may be required in the labeling are warnings; storage and transpor- tation information; expiration dates; the date and place of manufacture; the results that may be expected if the de- vice is used properly; the ranges of ac- curacy of diagnostic information; in- structions regarding the proper care of, and the proper components, acces- sories, or other equipment to be used with the device; and statements con- cerning the appropriate patient popu- lation, for example, a statement that the device is considered safe and effec- tive only when used by, or in the treat- ment of, a patient who has been tested by particular designated procedures and found to have an illness or condi- tion for which use of the device is indi- cated by a person skilled in the use of the device.

## Subpart B—Procedures for Per- formance Standards Develop- ment and Publication

##### § 861.20 Summary of standards devel- opment process.

The procedure by which a perform- ance standard for a device may be es- tablished, amended, or revoked is as follows:

The Food and Drug Administra- tion (FDA) will publish in the FEDERAL REGISTER a notice of proposed rule- making for the establishment, amend-

21 CFR Ch. I (4–1–17 Edition)

ment,  or  revocation  of  any  perform- ance standard for a device.

A notice of proposed rulemaking for the establishment or amendment of a performance standard for a device will:

  - Set forth a finding, with sup- porting justification, that the perform- ance standard is appropriate and nec- essary to provide reasonable assurance of the safety and effectiveness of the device;

  - Set forth proposed findings with respect to the risk of illness or injury that the performance standard is in- tended to reduce or eliminate;

  - Invite interested persons to sub- mit to the Food and Drug Administra- tion, within 30 days  of  the  publication of the notice, requests for changes  in the  classification  of  the  device  pursu- ant to § 860.132 of this chapter, based on new information relevant to the classi- fication;   and

  - Invite interested persons to sub- mit an existing performance standard for the device, including a draft or pro- posed performance standard, for con- sideration by  the  Commissioner  of Food and Drugs.

A notice of proposed rulemaking for the revocation of a performance standard will set forth a finding, with supporting justification, that the per- formance standard is no longer nec- essary to provide reasonable assurance of the safety and effectiveness of a de- vice.

A notice under this section will provide for a comment period of not less than 60 days.

If, after publication of a notice under paragraph (a) of this section, FDA receives a request to change the classification of the device, FDA will, within 60 days of the publication of the notice and after consultation with the appropriate panel under § 860.125 of this chapter, either deny the request or give notice of its intent to initiate a change in the classification under § 860.130.

If FDA initiates a rulemaking proceeding under paragraph (a) of this section, FDA will:

  - Complete the proceeding and es- tablish the  performance  standard  for the device in accordance with this part and § 10.40 of this chapter; or

#### Food and Drug Administration, HHS	§ 861.34

  - Terminate the proceeding by pub- lishing in the FEDERAL REGISTER a  no- tice announcing such termination and the reasons therefor and, unless the proceeding is terminated because the device is a banned device, initiate a proceeding in accordance with section 513(e) of the act to reclassify the de- vice; or

  - Take other appropriate action.

[57 FR 58404, Dec. 10, 1992]

##### § 861.24   Existing  standard  as  a  pro- posed standard.

The Food and Drug Administra- tion may accept an  existing  standard or a proposed or draft standard if it in- cludes:

  - A description of the procedures used to develop the standard and a list of the persons and organizations that participated in its development, to the extent that such information is avail- able or reasonably obtainable;

  - An identification of the specific portions of the existing standard that the person submitting the standard be- lieves are appropriate for adoption as, or inclusion in, the proposed standard; and

  - A summary of the test data, or, if requested by the Food  and Drug Ad- ministration, all such data or other in- formation supporting the specific por- tions of the standard identified by the person submitting the standard.

The Food and Drug Administra- tion will publish a notice in the FED- ERAL REGISTER stating either that  it has accepted, or accepted with modi- fication, as a proposed standard, an ex- isting standard or one that has been developed, or that an existing standard is not acceptable, together with the reasons therefor.

[45 FR 7484, Feb. 1, 1980, as amended at 57 FR

58405, Dec. 10, 1992]

##### § 861.30   Development of standards.

The Food and Drug Administration (FDA), while engaged in the develop- ment of a proposed standard under this section will:

Support its proposed performance standard by such test data or other documents or materials as may reason- ably be required;

Provide interested persons an op- portunity to participate in the develop- ment of the standard by accepting comments and, where appropriate, holding public meetings on issues re- lating to development of the standard. Notice of the opportunity  to  partici- pate in the development of the stand- ard will be furnished in a manner rea- sonably calculated to reach the  major- ity of persons interested in the devel- opment of the standard. This require- ment shall be satisfied by publishing such a notice in  the  FEDERAL  REG- ISTER. Whenever it is appropriate, FDA will  use  the  FEDERAL   REGISTER   to make announcements  about  the stand- ard development process of standard developers other than Federal agen- cies.

Maintain records disclosing the course of development of the proposed standard, the comments and other in- formation submitted by a person in connection with such development (in- cluding comments and information re- garding the need for a standard), and such other information as may be re- quired to evaluate the standard.

[45 FR 7484, Feb. 1, 1980, as amended at 57 FR

58405, Dec. 10, 1992]

##### § 861.34   Amendment or revocation of a standard.

The Food and Drug Administra- tion will  provide  for  periodic  evalua- tion of performance standards to deter- mine whether such standards should be changed to reflect new medical, sci- entific, or other technological data.

The Food and Drug Administra- tion may, on its own initiative or upon petition of an interested party, amend or revoke by regulation a standard es- tablished under this part.

Any petition to amend or revoke a standard shall:

  - Identify the specific device and standard for which the amendment or revocation is sought; and

  - Be submitted  in  accordance  with the requirements of § 10.30.

Proceedings to amend or revoke a performance standard shall be con- ducted in accordance with the rule- making procedures of § 10.40. In addi- tion,  a  notice  of   proposed   rulemaking to amend or revoke a standard shall set forth  proposed  findings  with  respect  to

#### § 861.36

the degree of risk or illness to be elimi- nated or reduced and the benefit the public will derive from the proposed amendment or revocation.

##### § 861.36   Effective dates.

A regulation establishing, amend- ing, or revoking a performance stand- ard will set forth the date upon which it will take effect. To the extent prac- tical, consistent with the public health and safety, such effective date will be established so as to minimize economic loss to, and disruption or  dislocation of, domestic and international trade.

Except  as  provided  in  paragraph

of this section, no regulation estab- lishing, amending, or revoking a stand- ard may take effect before 1 year after the date of its publication unless:

  - The Food and Drug Administra- tion determines  that  an  earlier  effec- tive date is necessary to protect the public health and safety; or

  - The standard has been established for a device that, by the effective date of the standard, has been reclassified from class III to class II.

(c) The Food and Drug Administra- tion may declare a proposed regulation amending a standard effective on publi- cation in the FEDERAL REGISTER if it determines that making the regulation so effective is in the public interest. A proposed amendment of a performance standard made effective upon publica- tion may not prohibit the introduction or delivery for introduction into inter- state commerce of a device that con- forms to the standard without the change or changes provided in the pro- posed amendment until the effective date of any final action on the pro- posal.

[45 FR 7484, Feb. 1, 1980, as amended at 57 FR

58405, Dec. 10, 1992]

##### § 861.38   Standards   advisory   commit- tees.

The Food and Drug Administra- tion will establish advisory  commit- tees to which proposed regulations may be referred, and these committees shall consider such referrals in accordance with this section and part 14 of this chapter. Such advisory committees, which may not be classification panels, shall be considered ad hoc advisory committees.   Their   members   shall   be

#### 21 CFR Ch. I (4–1–17 Edition)

selected in accordance with §§ 14.82 and 14.84, except that no member may be a regular full-time FDA employee. Each advisory committee established under this section shall include as nonvoting members a representative of consumer interests and a representative of inter- ests of the device manufacturing indus- try.

A proposed regulation to estab- lish, amend, or revoke a performance standard shall be referred to an advi- sory committee for a report and rec- ommendation with respect to any mat- ter involved in the proposed regulation which requires the exercise of sci- entific  judgment  if:

  - The Food and Drug Administra- tion determines that such referral is necessary or appropriate under the cir- cumstances; or

  - Requested by an interested  per- son, in the form of a citizen petition in accordance with § 10.30 of this chapter, which is made within the period pro- vided for comment on the proposed reg- ulation and which demonstrates good cause  for  referral.

When a proposed regulation is re- ferred to an advisory committee, the Food and Drug Administration will fur- nish the committee with the data and information upon which the proposed regulation is based. After independ- ently reviewing the materials fur- nished by the Food and Drug Adminis- tration and  any  other  available  data and information, the advisory com- mittee shall, within 60 days of the re- ferral, submit a report and rec- ommendation on the proposed regula- tion, together with all underlying data and information and a statement of the reason or basis for the recommenda- tion. A copy of the report and rec- ommendation will be publicly dis- played in the office of the Division of Dockets Management, Food and Drug Administration.

Where appropriate, each proposed regulation establishing a standard pub- lished in the FEDERAL REGISTER will include a call for nominations to the advisory committee for that particular standard.

[45 FR 7484, Feb. 1, 1980, as amended at 57 FR

58405, Dec. 10, 1992]

#### Food and Drug Administration, HHS	Pt. 862

PART 862—CLINICAL CHEMISTRY AND CLINICAL TOXICOLOGY DEVICES

##### Subpart A—General Provisions

Sec.

  - Scope.

  - Regulation of calibrators.

  - Effective dates of requirement for pre- market  approval.

862.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Clinical Chemistry Test Systems

862.1020 Acid phosphatase (total or  pros- tatic)   test   system.

862.1025   Adrenocorticotropic	hormone (ACTH)  test  system.

862.1030 Alanine amino transferase (ALT/ SGPT)   test   system.

862.1035    Albumin test system. 862.1040   Aldolase test system. 862.1045   Aldosterone test system.

862.1050 Alkaline phosphatase or isoenzymes test  system.

862.1055 Newborn screening test system for amino acids, free carnitine, and acylcarnitines using tandem mass spec- trometry.

862.1060 Delta-aminolevulinic acid test sys- tem.

862.1065   Ammonia test system. 862.1070   Amylase test system.

862.1075    Androstenedione test system. 862.1080   Androsterone test system.

862.1085 Angiotensin I and renin test sys- tem.

862.1090 Angiotensin converting enzyme (A.C.E.)  test  system.

862.1095    Ascorbic acid test system.

862.1100 Aspartate amino transferase (AST/ SGOT)   test   system.

862.1110 Bilirubin (total or direct) test sys- tem.

862.1113 Bilirubin (total and unbound) in the neonate test system.

862.1115 Urinary bilirubin and its conjugates (nonquantitative)   test   system.

862.1117 B-type natriuretic peptide test sys- tem.

862.1118    Biotinidase test system.

862.1120 Blood gases (PCO2PO2) and blood pH test   system.

862.1130    Blood volume test system.

862.1135 C-peptides of  proinsulin  test  sys- tem.

862.1140    Calcitonin test system. 862.1145    Calcium test system. 862.1150  Calibrator.

862.1155 Human chorionic  gonadotropin (HCG)  test  system.

862.1160 Bicarbonate/carbon dioxide test system.

862.1163    Cardiac   allograft   gene   expression profiling   test   system.

862.1165    Catecholamines  (total)  test  system. 862.1170   Chloride test system.

862.1175    Cholesterol (total) test system. 862.1177    Cholylglycine test system.

862.1180    Chymotrypsin test system.

862.1185   Compound S (11-deoxycortisol) test system.

862.1187    Conjugated     sulfolithocholic     acid (SLCG)   test   system.

862.1190   Copper test system. 862.1195    Corticoids test system.

862.1200    Corticosterone test system.

862.1205 Cortisol	(hydrocortisone	and hydroxycorticosterone)   test   system.

862.1210   Creatine test system.

862.1215 Creatine phosphokinase/creatine ki- nase or isoenzymes test system.

862.1225    Creatinine test system. 862.1230    Cyclic AMP test system. 862.1235    Cyclosporine test system. 862.1240    Cystine test system.

862.1245   Dehydroepiandrosterone   (free   and sulfate)   test   system.

862.1250    Desoxycorticosterone test system. 862.1255   2,3-Diphosphoglyceric acid test sys-

tem.

862.1260   Estradiol test system. 862.1265   Estriol test system.

862.1270    Estrogens (total, in pregnancy) test system.

862.1275    Estrogens     (total,     nonpregnancy) test   system.

862.1280    Estrone test system.

862.1285    Etiocholanolone test system. 862.1290    Fatty acids test system.

862.1295    Folic acid test system.

862.1300   Follicle-stimulating   hormone   test system.

862.1305    Formiminoglutamic    acid    (FIGLU) test   system.

862.1310    Galactose test system.

862.1315    Galactose-1-phosphate uridyl trans- ferase  test  system.

862.1320    Gastric acidity test system. 862.1325   Gastrin test system.

862.1330   Globulin test system. 862.1335   Glucagon test system.

862.1340    Urinary   glucose   (nonquantitative) test   system.

862.1345   Glucose test system.

862.1350    Continuous   glucose   monitor   sec- ondary   display.

862.1360   Gamma-glutamyl	transpeptidase and  isoenzymes  test  system.

862.1365    Glutathione test system.

862.1370   Human  growth  hormone  test  sys- tem.

862.1373    Hemoglobin A1c test system. 862.1375    Histidine test system.

862.1377    Urinary  homocystine  (nonquantita- tive)  test  system.

862.1380    Hydroxybutyric dehydrogenase test system.

862.1385  17-Hydroxycorticosteroids	(17- ketogenic  steroids)  test  system.

#### Pt. 862

| acid/sero- | 862.1635 | Total protein test system. |
| --- | --- | --- |
|  | 862.1640 | Protein-bound iodine test system. |

862.1390   5-Hydroxyindole   acetic tonin  test  system.

862.1395   17-Hydroxyprogesterone    test    sys- tem.

862.1400    Hydroxyproline test system.  862.1405    Immunoreactive   insulin   test   sys-

tem.

862.1410   Iron (non-heme) test system. 862.1415    Iron-binding capacity test system. 862.1420    Isocitric   dehydrogenase   test   sys-

tem.

862.1430   17-Ketosteroids test system.

862.1435    Ketones  (nonquantitative)  test  sys- tem.

862.1440    Lactate dehydrogenase test system. 862.1445    Lactate  dehydrogenase  isoenzymes

test system.

862.1450    Lactic acid test system.

862.1455   Lecithin/sphingomyelin	ratio	in amniotic  fluid  test  system.

862.1460    Leucine   aminopeptidase   test   sys- tem.

862.1465   Lipase test system. 862.1470    Lipid (total) test system. 862.1475    Lipoprotein test system.

862.1485    Luteinizing hormone test system. 862.1490   Lysozyme   (muramidase)   test   sys-

tem.

862.1495    Magnesium test system.

862.1500    Malic dehydrogenase test system. 862.1505    Mucopolysaccharides  (nonquantita-

tive) test system.

862.1509    Methylmalonic   acid   (nonquantita- tive)  test  system.

862.1510    Nitrite   (nonquantitative)   test   sys- tem.

862.1515    Nitrogen (amino-nitrogen) test sys- tem.

862.1520   5-Nucleotidase test system. 862.1530    Plasma oncometry test system.

862.1535    Ornithine carbamyl transferase test system.

862.1540    Osmolality test system. 862.1542    Oxalate test system.

862.1545    Parathyroid hormone test system. 862.1550    Urinary   pH   (nonquantitative)   test

system.

862.1555    Phenylalanine test system.

862.1560 Urinary	phenylketones	(non- quantitative)   test   system.

862.1565   6-Phosphogluconate  dehydrogenase test   system.

862.1570   Phosphohexose  isomerase  test  sys- tem.

862.1575   Phospholipid test system.

862.1580    Phosphorus (inorganic) test system. 862.1585    Human  placental  lactogen  test  sys-

tem.

862.1590    Porphobilinogen test system. 862.1595   Porphyrins test system.

862.1600    Potassium test system. 862.1605    Pregnanediol test system. 862.1610    Pregnanetriol test system. 862.1615    Pregnenolone test system. 862.1620    Progesterone test system.

862.1625    Prolactin (lactogen) test system. 862.1630     Protein (fractionation) test system.

#### 21 CFR Ch. I (4–1–17 Edition)

862.1645 Urinary protein or albumin (non- quantitative)   test   system.

862.1650    Pyruvate kinase test system. 862.1655    Pyruvic acid test system.

862.1660 Quality  control  material  (assayed and  unassayed).

862.1665   Sodium test system.

862.1670 Sorbitol dehydrogenase test  sys- tem.

862.1675    Blood specimen collection device. 862.1678    Tacrolimus test system.

862.1680    Testosterone test system.

862.1685 Thyroxine-binding globulin test system.

862.1690 Thyroid-stimulating hormone test system.

862.1695    Free thyroxine test system. 862.1700    Total thyroxine test system. 862.1705    Triglyceride test system.

862.1710     Total  triiodothyronine  test  system. 862.1715    Triiodothyronine   uptake   test   sys-

tem.

862.1720 Triose phosphate isomerase test system.

862.1725   Trypsin test system. 862.1730   Free tyrosine test system. 862.1770   Urea nitrogen test system. 862.1775   Uric acid test system.

862.1780 Urinary calculi (stones) test sys- tem.

862.1785 Urinary urobilinogen (nonquantita- tive)  test  system.

862.1790   Uroporphyrin test system. 862.1795    Vanilmandelic acid test system. 862.1805    Vitamin A test system.

862.1810    Vitamin B12 test system. 862.1815    Vitamin E test system. 862.1820   Xylose test system.

862.1825    Vitamin D test system.

##### Subpart C—Clinical Laboratory Instruments

862.2050 General purpose laboratory equip- ment labeled or promoted for a specific medical   use.

862.2100 Calculator/data processing module for  clinical  use.

862.2140 Centrifugal  chemistry  analyzer  for clinical  use.

862.2150    Continuous    flow    sequential    mul- tiple  chemistry  analyzer  for  clinical  use. 862.2160    Discrete photometric chemistry an-

alyzer for clinical use.

862.2170 Micro chemistry analyzer for  clin- ical  use.

862.2230 Chromatographic separation mate- rial  for  clinical  use.

862.2250 Gas liquid chromatography  system for  clinical  use.

862.2260 High pressure liquid chroma- tography system for clinical use.

862.2265 High throughput genomic sequence analyzer for clinical use.

862.2270 Thin-layer chromatography system for  clinical  use.

#### Food and Drug Administration, HHS	§ 862.1

862.2300    Colorimeter,  photometer,  or  spec- trophotometer  for  clinical  use.

862.2310    Clinical sample concentrator. 862.2320    Beta or gamma counter for clinical

use.

862.2400 Densitometer/scanner (integrating, reflectance, TLC, or radiochromatogram) for clinical use.

862.2485    Electrophoresis  apparatus  for  clin- ical  use.

862.2500    Enzyme analyzer for clinical use. 862.2540   Flame    emission    photometer    for

clinical use.

862.2560    Fluorometer for clinical use. 862.2570     Instrumentation  for  clinical  multi-

plex test systems.

862.2680     Microtitrator for clinical use. 862.2700    Nephelometer for clinical use. 862.2720    Plasma oncometer for clinical use. 862.2730    Osmometer for clinical use.

862.2750   Pipetting  and  diluting  system  for clinical  use.

862.2800    Refractometer for clinical use. 862.2850    Atomic  absorption  spectrophotom-

eter for clinical use.

862.2860    Mass spectrometer for clinical use. 862.2900    Automated urinalysis system.

862.2920   Plasma viscometer for clinical use.

##### Subpart D—Clinical Toxicology Test Systems

862.3030    Acetaminophen test system. 862.3035    Amikacin test system.

862.3040    Alcohol test system. 862.3050    Breath-alcohol test system.

862.3080    Breath nitric oxide test system. 862.3100    Amphetamine test system.

862.3110   Antimony test system. 862.3120   Arsenic test system.

862.3150    Barbiturate test system. 862.3170   Benzodiazepine test system. 862.3200     Clinical toxicology calibrator. 862.3220   Carbon monoxide test system. 862.3240    Cholinesterase test system.

862.3250    Cocaine and cocaine metabolite test system.

862.3270    Codeine test system.

862.3280   Clinical  toxicology  control  mate- rial.

862.3300   Digitoxin test system. 862.3320   Digoxin test system.

862.3350    Diphenylhydantoin test system. 862.3360  Drug	metabolizing	enzyme

genotyping system.

862.3380    Ethosuximide test system. 862.3450    Gentamicin test system. 862.3520    Kanamycin test system.

862.3550  Lead test system. 862.3555   Lidocaine test system. 862.3560   Lithium test system.

862.3580    Lysergic   acid   diethylamide   (LSD) test   system.

862.3600   Mercury test system.

862.3610    Methamphetamine test system. 862.3620   Methadone test system.

862.3630    Methaqualone test system.

862.3640    Morphine test system.

862.3645 Neuroleptic	drugs	radioreceptor assay  test  system.

862.3650   Opiate test system.  862.3660    Phenobarbital test system. 862.3670    Phenothiazine test system. 862.3680    Primidone test system.

862.3700   Propoxyphene test system. 862.3750   Quinine test system.

862.3830   Salicylate test system. 862.3840   Sirolimus test system. 862.3850   Sulfonamide test system. 862.3870  Cannabinoid test system. 862.3880   Theophylline test system. 862.3900  Tobramycin test system.

862.3910    Tricyclic  antidepressant  drugs  test system.

862.3950    Vancomycin test system.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE:  52  FR  16122,  May  1,  1987,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 862 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 862.1   Scope.

This part sets forth the classifica- tion of clinical chemistry and clinical toxicology devices intended for human use that are in commercial distribu- tion.

The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required in § 807.87.

References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 16122, May 1, 1987, as amended at 67

FR 58329, Sept. 16, 2002; 78 FR 18233, Mar. 26,

2013; 79 FR 50552, Aug. 25, 2014]

§ 862.2

##### § 862.2   Regulation of calibrators.

Many devices classified in this part are intended to be used with a cali- brator. A calibrator has a reference value assigned to it which serves as the basis by which test results of  patients are derived or calculated. The cali- brator for a device may be (a) manufac- tured and distributed separately from the device with which it is intended to be used, (b) manufactured and distrib- uted as one of several device compo- nents,  such  as  in  a  kit  of  reagents,  or

  - built-in as an integral part of the device. Because of the central role that a calibrator plays in the measurement process and the critical effect cali- brators have on accuracy of test re- sults, elsewhere in this  part,  all  three of these types of calibrators (§§ 862.1150 and 862.3200 of this part) are classified into class II, notwithstanding the clas- sification of the device with which it is intended to be used. Thus, a device and its calibrator may have different clas- sifications, even if the  calibrator  is built into the device.

##### § 862.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

    - Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a device, has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except as provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after  its  promulgation  or  on  the  last

#### 21 CFR Ch. I (4–1–17 Edition)

day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

    - Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

##### § 862.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly,    manufacturers    of    any

#### Food and Drug Administration, HHS	§ 862.1030

commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

      - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

      - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

      - The device is an in vitro device that is intended:

        - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

        - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

        - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

        - For assessing the risk of cardio- vascular diseases;

        - For use in diabetes management;

        - For identifying or inferring the identity of a microorganism directly from clinical material;

        - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the

assay  is  intended  for  use  in  matrices other than serum or plasma;

        - For noninvasive testing as defined in § 812.3(k) of this chapter; and

        - For near patient testing (point of care).

[65 FR 2304, Jan. 14, 2000]

## Subpart B—Clinical Chemistry Test Systems

##### § 862.1020   Acid  phosphatase  (total  or prostatic) test system.

          - Identification. An acid phosphatase (total or prostatic) test system is a de- vice intended to measure the activity of the acid phosphatase enzyme in plas- ma and serum.

          - Classification. Class II.

##### § 862.1025   Adrenocorticotropic	hor- mone (ACTH) test system.

Identification.	An adrenocorticotropic hormone (ACTH) test system is a device intended to measure adrenocorticotropic hormone in plasma and serum. ACTH measure- ments are used in the differential diag- nosis and treatment of certain dis- orders of the adrenal glands such as Cushing’s syndrome, adrenocortical in- sufficiency, and the ectopic ACTH syn- drome.

Classification. Class II.

##### § 862.1030   Alanine   amino   transferase (ALT/SGPT) test system.

Identification. An alanine amino transferase (ALT/SGPT) test system is a device intended to measure the activ- ity of the enzyme alanine amino trans- ferase (ALT) (also known as a serum glutamic pyruvic transaminase or SGPT) in serum and plasma. Alanine amino transferase measurements are used in the diagnosis and treatment of certain liver diseases (e.g., viral hepa- titis and cirrhosis) and heart diseases.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

§ 862.1035

##### § 862.1035   Albumin test system.

  - Identification. An albumin test system is a device intended to measure the albumin concentration in serum and plasma. Albumin measurements are used in the diagnosis and treat- ment of numerous diseases involving primarily the liver or kidneys.

  - Classification. Class II.

##### § 862.1040   Aldolase test system.

Identification. An aldolase test system is a device intended to measure the activity of the enzyme aldolase in serum or plasma. Aldolase measure- ments are used in the diagnosis and treatment of the early stages of acute hepatitis and for certain muscle dis- eases such as progressive Duchenne- type muscular dystrophy.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1045   Aldosterone test system.

Identification. An aldosterone test system is a device intended to measure the hormone aldosterone in serum and urine. Aldosterone measurements are used in the diagnosis and treatment of primary aldosteronism (a disorder caused by the excessive secretion of aldosterone by the adrenal gland), hy- pertension caused by primary aldosteronism,	selective hypoaldosteronism, edematous states, and other conditions of electrolyte im- balance.

Classification. Class II.

##### § 862.1050   Alkaline	phosphatase	or isoenzymes test system.

Identification. An alkaline phos- phatase or isoenzymes test system is a device intended to measure alkaline phosphatase or its isoenzymes (a group of enzymes with similar biological ac- tivity) in serum or plasma. Measure- ments of alkaline phosphatase or its isoenzymes are used in the diagnosis and treatment of liver, bone, parathy- roid, and intestinal diseases.

Classification. Class II.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 862.1055 Newborn screening test sys- tem for amino acids, free carnitine, and acylcarnitines using tandem mass spectrometry.

Identification. A newborn screen- ing test system for amino acids, free carnitine, and acylcarnitines using tandem mass spectrometry is a device that consists of stable isotope internal standards, control materials, extrac- tion solutions, flow solvents, instru- mentation, software packages, and other reagents and materials. The de- vice is intended for the measurement and evaluation of amino acids, free car- nitine, and acylcarnitine concentra- tions from newborn whole blood filter paper samples. The quantitative anal- ysis of amino acids, free carnitine, and acylcarnitines and their relationship with each other provides analyte con- centration profiles that may aid in screening newborns for one or more in- born errors of amino acid, free carni- tine, and acyl-carnitine metabolism.

Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Newborn Screening Test Systems for Amino Acids, Free Carnitine, and Acylcarnitines Using Tandem Mass Spectrometry.’’ See  § 862.1(d) for the availability  of  this  guidance  document.

[69 FR 68255, Nov. 24, 2004]

##### § 862.1060 Delta-aminolevulinic  acid test system.

Identification.        A	delta- aminolevulinic acid test system is a device intended to measure the level of delta-aminolevulinic acid (a precursor of porphyrin) in urine. Delta- aminolevulinic acid measurements are used in the diagnosis and treatment of lead poisoning and certain porphyrias (diseases affecting the liver, gastro- intestinal, and nervous systems that are accompanied by increased urinary excretion of various heme compounds including delta-aminolevulinic acid).

Classification. Class I (general con- trols). The device is exempt from pre- market notification  procedures in sub- part E of part 807 of this  chapter  sub- ject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 862.1100

##### § 862.1065   Ammonia test system.

  - Identification. An ammonia test system is a device intended to measure ammonia levels in blood, serum, and plasma, Ammonia measurements are used in the diagnosis and treatment of severe liver disorders, such as cir- rhosis, hepatitis, and Reye’s syndrome.

  - Classification. Class I.

##### § 862.1070   Amylase test system.

Identification. An amylase test system is a device intended to measure the activity of the enzyme amylase in serum and urine. Amylase measure- ments are used primarily for the diag- nosis and treatment of pancreatitis (in- flammation of the pancreas).

Classification. Class II.

##### § 862.1075   Androstenedione   test   sys- tem.

Identification. An androstenedione test system is a device intended to measure androstenedione (a substance secreted by the testes, ovary, and adre- nal glands) in serum. Adrostenedione measurements are used  in the diag- nosis and treatment of females with ex- cessive levels of androgen (male sex hormone) production.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1080   Androsterone test system.

Identification. An androsterone test system is a device intended to measure the hormone adrosterone in serum, plasma, and urine. Andros- terone measurements are used in the diagnosis and treatment of gonadal and adrenal diseases.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1085   Angiotensin I and renin test system.

Identification. An angiotensin I and renin test system is a device in- tended to measure the level of angiotensin I generated by renin in plasma. Angiotensin I measurements are used in the diagnosis and treat- ment of certain types of hypertension.

Classification. Class II.

##### § 862.1090   Angiotensin  converting  en- zyme (A.C.E.) test system.

Identification. An angiotensin con- verting enzyme (A.C.E.) test system is a device intended to measure the activ- ity of angiotensin converting  enzyme in serum and plasma. Measurements obtained by this device are used in the diagnosis and treatment of diseases such as sarcoidosis, a disease charac- terized by the formation of nodules in the lungs, bones, and skin, and Gaucher’s disease, a hereditary dis- order affecting the spleen.

Classification. Class II.

##### § 862.1095   Ascorbic acid test system.

Identification. An  ascorbic acid test system is a device intended to measure the level of ascorbic acid (vi- tamin C) in plasma, serum, and urine. Ascorbic acid measurements are used in the diagnosis and treatment of ascorbic acid dietary deficiencies.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1100   Aspartate amino transferase (AST/SGOT) test system.

Identification. An aspartate amino transferase (AST/SGOT) test system is a device intended to measure the activ- ity of the enzyme aspartate amino transferase (AST) (also known as a serum glutamic oxaloacetic transferase or SGOT) in serum and plasma. Aspartate amino transferase measure- ments are used in the diagnosis and treatment of certain types of liver and heart disease.

Classification. Class II.

#### § 862.1110

##### § 862.1110   Bilirubin   (total   or   direct) test system.

  - Identification. A bilirubin (total or direct) test system is a device intended to measure the levels of bilirubin (total or direct) in plasma or serum. Measure- ments of the levels of bilirubin, an or- ganic compound formed during the nor- mal and abnormal distruction of red blood cells, if used in the diagnosis and treatment of liver, hemolytic hematological, and metabolic dis- orders, including hepatitis and gall bladder block.

  - Classification. Class II.

##### § 862.1113   Bilirubin	(total	and	un- bound) in the neonate test system.

Identification. A bilirubin (total and unbound) in the neonate test sys- tem is a device intended to measure the levels of bilirubin (total and un- bound) in the blood (serum) of newborn infants to aid in indicating the risk of bilirubin encephalopathy (kernicterus).

Classification. Class I.

[54 FR 30206, July 19, 1989]

##### § 862.1115 Urinary bilirubin and its conjugates (nonquantitative) test system.

Identification. A urinary bilirubin and its conjugates (nonquantitative) test system is a device intended to measure the levels of bilirubin con- jugates in urine. Measurements of uri- nary bilirubin and its conjugates (non- quantitative) are used in the diagnosis and treatment of certain liver diseases.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1117   B-type   natriuretic   peptide test system.

Identification. The B-type natriuretic peptide (BNP) test system is an in vitro diagnostic device in- tended to measure BNP in whole blood and plasma. Measurements of BNP are used as an aid in the diagnosis of pa- tients with congestive heart failure.

Classification. Class II (special controls).  The  special  control  is  ‘‘Class

#### 21 CFR Ch. I (4–1–17 Edition)

II Special Control Guidance Document for B-Type Natriuretic Peptide Pre- market Notifications; Final Guidance for Industry and FDA Reviewers.’’

[66 FR 12734, Feb. 28, 2001]

##### § 862.1118   Biotinidase test system.

- Identification. The biotinidase test system is an in vitro diagnostic device intended to measure the activity of the enzyme biotinidase in blood. Measure- ments of biotinidase are used in the treatment and diagnosis of biotinidase deficiency, an inborn error of metabo- lism in infants, characterized by the inability to utilize dietary protein bound vitamin or to recycle endoge- nous biotin. The deficiency may result in irreversible neurological impair- ment.

- Classification. Class II (special controls). The special control is sale, distribution,  and  use  in  accordance with the prescription device require- ments in § 801.109 of this chapter.

[65 FR 16521, Mar. 29, 2000]

##### § 862.1120    Blood  gases  (PCO2,  PO2)  and blood pH test system.

Identification. A blood gases (PCO2, PO2) and blood pH test system is a de- vice intended to measure certain gases in blood, serum, plasma or pH of blood, serum, and plasma. Measurements of blood gases (PCO2, PO2) and blood pH are used in the diagnosis and treatment of life-threatening  acid-base  disturbances.

Classification. Class II.

##### § 862.1130   Blood volume test system.

Identification. A blood volume test system is a device intended to measure the circulating blood volume. Blood volume measurements are used in the diagnosis and treatment of shock, hem- orrhage, and polycythemia vera (a dis- ease characterized by an absolute in- crease in erythrocyte mass and total blood volume).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 862.1163

##### § 862.1135   C-peptides of proinsulin test system.

  - Identification. A C-peptides of proinsulin test system is a device in- tended to measure C-peptides of proinsulin levels in serum, plasma, and urine. Measurements of C-peptides of proinsulin are used in the diagnosis and treatment of patients with abnor- mal insulin secretion, including diabe- tes mellitus.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1140   Calcitonin test system.

Identification. A calcitonin test system is a device intended to measure the thyroid hormone calcitonin (thyrocalcitonin) levels in plasma and serum. Calcitonin measurements are used in the diagnosis and treatment of diseases involving the thyroid and parathyroid glands, including car- cinoma and hyperparathyroidism (ex- cessive activity of the parathyroid gland).

Classification. Class II.

##### § 862.1145   Calcium test system.

Identification. A calcium test sys- tem is a device intended to measure the total calcium level in serum. Cal- cium measurements are used in the di- agnosis and treatment of parathyroid disease, a variety of bone diseases, chronic renal disease and tetany (inter- mittent muscular contractions or spasms).

Classification. Class II.

##### § 862.1150   Calibrator.

Identification. A calibrator is a de- vice intended for medical purposes for use in a test  system  to  establish  points of reference that are used in the deter- mination of values in  the  measurement of substances in human specimens. (See also § 862.2 in this part.)

Classification. Class II.

##### § 862.1155   Human	chorionic gonadotropin (HCG) test system.

Human chorionic gonadotropin (HCG) test system intended for the early detection of pregnancy—(1) Identification. A human chorionic gonadotropin (HCG) test system is a device intended for the early detection of pregnancy is in- tended to measure HCG, a placental hormone, in plasma or urine.

(2) Classification. Class II.

Human chorionic gonadotropin (HCG) test system intended for any uses other than early detection of pregnancy—

Identification. A human chorionic goadotropin (HCG) test system is a de- vice intended for any uses other than early detection of pregnancy (such as an aid in the diagnosis, prognosis, and management of treatment of persons with certain tumors or carcinomas) is intended to measure HCG, a placental hormone, in plasma or urine.

Classification. Class III.

Date PMA or notice of completion of a PDP is required. As of the enactment date of the amendments, May  28,  1976, an approval under section 515 of the act is required before the  device  described in paragraph (b)(1) may be  commer- cially distributed. See § 862.3.

##### § 862.1160   Bicarbonate/carbon   dioxide test system.

Identification. A bicarbonate/car- bon dioxide test system is a device in- tended to measure bicarbonate/carbon dioxide in plasma, serum, and whole blood. Bicarbonate/carbon dioxide measurements are used  in the diag- nosis and treatment of numerous po- tentially serious disorders associated with changes in body acid-base  bal- ance.

Classification. Class II.

##### § 862.1163   Cardiac  allograft  gene  ex- pression profiling test system.

Identification. A cardiac allograft gene expression profiling test system is a device that measures the ribonucleic acid (RNA) expression level of multiple genes and combines this information to yield a signature (pattern, classifier, index, score) to aid in the identifica- tion of a low probability of acute cel- lular rejection (ACR) in heart trans- plant recipients with stable allograft function.

#### § 862.1165

Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Cardiac Allograft Gene Expression Profiling Test Systems.’’ See § 862.1(d) for the availability of this guidance document.

[74 FR 53885, Oct. 21, 2009]

##### § 862.1165   Catecholamines  (total)  test system.

Identification. A catecholamines (total) test system is a device intended to determine whether a group of simi- lar	compounds	(epinephrine, norepinephrine, and dopamine) are present in urine and plasma. Catecholamine determinations are used in the diagnosis and treatment of adrenal medulla and hypertensive dis- orders, and for catecholamine-secret- ing tumors (pheochromo-cytoma, neu- roblastoma, ganglioneuroma, and retinoblastoma).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1170   Chloride test system.

Identification. A chloride test sys- tem is a device intended to measure the level of chloride in plasma, serum, sweat, and urine. Chloride measure- ments are used in the diagnosis and treatment of electrolyte and metabolic disorders such as cystic fibrosis and di- abetic acidosis.

Classification. Class II.

##### § 862.1175   Cholesterol  (total)  test  sys- tem.

Identification. A cholesterol (total) test system is a device intended to measure cholesterol in plasma and serum. Cholesterol measurements are used in the diagnosis and treatment of disorders involving excess cholesterol in the blood and lipid and lipoprotein metabolism disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1177   Cholylglycine test system.

Identification. A cholylglycine test system is a device intended to measure the bile acid cholylglycine in serum. Measurements obtained by this device are used in the diagnosis and treat- ment of liver disorders, such as cir- rhosis or obstructive liver disease.

Classification. Class II.

##### § 862.1180   Chymotrypsin test system.

Identification. A chymotrypsin test system is a device intended to measure the activity of the enzyme chymotrypsin in blood and other body fluids and in feces. Chymotrypsin measurements are used in the diag- nosis and treatment of pancreatic exo- crine insufficiency.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1185   Compound	S	(11- deoxycortisol) test system.

Identification. A compound S (11- dioxycortisol) test system is a device intended to measure the level of com- pound S (11-dioxycortisol) in plasma. Compound S is a steroid intermediate in the biosynthesis of the adrenal hor- mone cortisol. Measurements of com- pound S are used in the diagnosis and treatment of certain adrenal and pitui- tary gland disorders resulting in clin- ical symptoms of masculinization and hypertension.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 862.1215

##### § 862.1187   Conjugated   sulfolithocholic acid (SLCG) test system.

  - Identification. A conjugated sulfolithocholic acid (SLCG) test sys- tem is a device intended to measure the bile acid SLCG in serum. Measure- ments obtained by this device are used in the diagnosis and treatment of liver disorders, such as cirrhosis or obstruc- tive liver disease.

  - Classification. Class II.

##### § 862.1190   Copper test system.

Identification. A copper test sys- tem is a device intended to measure copper levels in plasma, serum, and urine. Measurements of copper are used in the diagnosis and treatment of ane- mia, infections, inflammations, and Wilson’s disease (a hereditary disease primarily of the liver and nervous sys- tem). Test results are also used in mon- itoring patients with Hodgkin’s disease (a disease primarily of the lymph sys- tem).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1195   Corticoids test system.

Identification. A corticoids test system is a device intended to measure the levels of corticoids (hormones of the adrenal cortex) in serum and p lasma. Measurements of corticoids are used in the diagnosis and treatment of disorders of the cortex of the adrenal glands, especially those  associated with hypertension and electrolyte dis- turbances.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1200   Corticosterone test system.

Identification. A corticosterone test system is a device intended to measure corticosterone (a steroid se- creted by the adrenal gland) levels in

plasma. Measurements of  corticos- terone are used in the diagnosis and treatment of adrenal disorders such as adrenal cortex  disorders  and blocks in cortisol  synthesis.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1205  Cortisol   (hydrocortisone and hydroxycorticosterone) test system.

Identification. A cortisol (hydro- cortisone and hydroxycorticosterone) test system is a device intended to measure the cortisol hormones  se- creted by the adrenal gland in plasma and urine. Measurements  of  cortisol are used in the diagnosis and treat- ment of disorders of the adrenal gland.

Classification. Class II.

##### § 862.1210   Creatine test system.

Identification. A creatine test sys- tem is a device intended to measure creatine (a substance synthesized in the liver and pancreas and found in bio- logical fluids) in plasma, serum, and urine. Measurements of creatine are used in the diagnosis and treatment of muscle diseases and endocrine dis- orders including hyperthyroidism.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1215 Creatine phosphokinase/cre- atine kinase or isoenzymes test sys- tem.

Identification. A creatine phosphokinase/creatine kinase or isoenzymes test system is a device in- tended to measure the activity of the enzyme creatine phosphokinase or its isoenzymes (a group of enzymes with similar biological activity)  in  plasma and serum. Measurements of creatine phosphokinase and its isoenzymes are used in the diagnosis and treatment of

#### § 862.1225

myocardial infarction and muscle dis- eases such as progressive, Duchenne- type muscular dystrophy.

Classification. Class II.

##### § 862.1225   Creatinine test system.

Identification. A creatinine test system is a device intended to measure creatinine levels in plasma and urine. Creatinine measurements are used in the diagnosis and treatment of renal diseases, in monitoring renal dialysis, and as a calculation basis for meas- uring other urine analytes.

Classification. Class II.

##### § 862.1230   Cyclic AMP test system.

Identification. A cyclic AMP test system is a device intended to measure the level of adenosine 3, 5- monophosphate (cyclic AMP) in plas- ma, urine, and other body fluids. Cyclic AMP measurements are used in the di- agnosis and treatment of endocrine dis- orders, including hyperparathyroidism (overactivity of the parathyroid gland). Cyclic AMP measurements may also be used in the diagnosis and treatment of Graves’ disease  (a  disorder of  the  thy- roid) and in  the  differentiation  of causes of hypercalcemia (elevated lev- els of serum calcium.)

Classification. Class II.

##### § 862.1235   Cyclosporine test system.

Identification. A cyclosporine test system is a device intended to quan- titatively determine cyclosporine con- centrations as an aid in the manage- ment of transplant patients receiving therapy with this drug. This generic type of device includes immunoassays and chromatographic assays for cyclosporine.

Classification. Class II (special controls). The special  control  is  ‘‘Class II Special Controls  Guidance  Docu- ment: Cyclosporine and Tacrolimus As- says; Guidance for Industry and FDA.’’ See § 862.1(d) for the availability of this guidance document.

[67 FR 58329, Sept. 16, 2002]

##### § 862.1240   Cystine test system.

Identification. A cystine test sys- tem is a device intended to measure the amino acid cystine in urine. Cys- tine measurements are used in the di-

#### 21 CFR Ch. I (4–1–17 Edition)

agnosis of cystinuria (occurrence of cystine in urine). Patients with cystinuria frequently develop kidney calculi (stones).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2305, Jan. 14, 2000]

##### § 862.1245   Dehydroepiandrosterone (free and sulfate) test system.

Identification.		A dehydroepiandrosterone (free and sul- fate) test system is  a  device  intended to measure dehydroepiandrosterone (DHEA) and its sulfate in urine, serum, plasma,	and	amniotic	fluid. Dehydroepiandrosterone measurements are used in the diagnosis and treat- ment of DHEA-secreting adrenal car- cinomas.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1250   Desoxycorticosterone	test system.

Identification.		A desoxycorticosterone test system is a device	intended	to	measure desoxycorticosterone (DOC) in plasma and urine. DOC measurements are used in the diagnosis and treatment of pa- tients with hypermineralocorticoidism (excess retention of sodium and loss of potassium) and other disorders of the adrenal gland.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1255   2,3-Diphosphoglyceric	acid test system.

Identification.	A	2,3- diphosphoglyceric acid test system is a device intended to measure 2,3- diphosphoglyceric    acid    (2,3-DPG)    in

#### Food and Drug Administration, HHS	§ 862.1280

erythrocytes (red blood cells). Meas- urements of 2,3-diphosphoglyceric acid are used in the diagnosis and treat- ment of blood disorders that affect the delivery of oxygen by erythrocytes to tissues and in monitoring  the  quality of stored blood.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1260   Estradiol test system.

Identification. An estradiol test system is a device intended to measure estradiol, an estrogenic steroid, in plasma. Estradiol measurements are used in the diagnosis and treatment of various hormonal sexual disorders and in assessing placental function in com- plicated pregnancy.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1265   Estriol test system.

Identification. An estriol test sys- tem is a device intended to measure es- triol, an estrogenic steroid, in plasma, serum, and urine of pregnant females. Estriol measurements are used in the diagnosis	and	treatment	of fetoplacental distress in certain cases of high-risk pregnancy.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1270   Estrogens   (total,   in   preg- nancy) test system.

Identification. As estrogens (total, in pregnancy) test system is a device intended to measure total estrogens in plasma, serum, and urine during preg- nancy. The device primarily measures estrone  plus  estradiol.  Measurements

of total estrogens are used to aid in the diagnosis	and	treatment	of fetoplacental distress in certain cases of high-risk pregnancy.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1275   Estrogens   (total,   nonpreg- nancy) test system.

Identification. As estrogens (total, nonpregnancy) test system is a device intended to measure the level of estro- gens (total estrone, estradiol, and es- triol) in plasma, serum, and urine of males and nonpregnant females. Meas- urement of estrogens (total, nonpreg- nancy) is used in the diagnosis and treatment of numerous disorders, in- cluding infertility, amenorrhea (ab- sence of menses) differentiation of pri- mary and secondary ovarian malfunc- tion, estrogen secreting testicular and ovarian tumors, and precocious pu- berty in females.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1280   Estrone test system.

Identification. An estrone test sys- tem is a device intended to measure estrone, an estrogenic steroid, in plas- ma. Estrone measurements are used in the diagnosis and treatment of numer- ous disorders, including infertility, amenorrhea, differentiation of primary and secondary ovarian malfunction, es- trogen secreting testicular and ovarian tumors, and precocious puberty in fe- males.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

#### § 862.1285

##### § 862.1285   Etiocholanolone test system.

  - Identification. An etiocholanolone test system is a device intended to measure etiocholanolone in serum and urine. Etiocholanolone is a metabolic product of the hormone testosterone and is excreted in the urine. Etiocholanolone measurements are used in the diagnosis and treatment of disorders of the testes and ovaries.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1290   Fatty acids test system.

Identification. A fatty acids test system is a device intended to measure fatty acids in plasma and serum. Meas- urements of fatty acids are used in the diagnosis and treatment of various dis- orders of lipid metabolism.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1295   Folic acid test system.

Identification. A folic acid test sys- tem is a device intended to measure the vitamin folic acid in plasma and serum. Folic acid measurements are used in the diagnosis and treatment of megaloblastic anemia, which is charac- terized by the presence of megaloblasts (an abnormal red blood cell series) in the bone marrow.

Classification. Class II.

[52 FR 16122, May 1, 1987; 53 FR 11645, Apr. 8,

1988]

##### § 862.1300   Follicle-stimulating	hor- mone test system.

Identification. A follicle-stimu- lating hormone test system is a device intended to measure follicle-stimu- lating hormone (FSH) in plasma, serum, and urine. FSH measurements are used in the diagnosis and treat- ment of pituitary gland and gonadal disorders.

#### 21 CFR Ch. I (4–1–17 Edition)

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1305   Formiminoglutamic	acid (FIGLU) test system.

Identification.	A formiminoglutamic acid (FIGLU) test system is a device intended to measure formiminolutamic acid in  urine. FIGLU measurements obtained by this device are used in the diagnosis of anemias, such as pernicious anemia and congenital hemolytic anemia.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1310   Galactose test system.

Identification. A galactose test system is a device intended to measure galactose in blood and urine. Galactose measurements are used in the diag- nosis and treatment of the hereditary disease galactosemia (a disorder of ga- lactose metabolism) in infants.

Classification. Class I.

##### § 862.1315   Galactose-1-phosphate uridyl transferase test system.

Identification. A galactose-1-phos- phate uridyl  transferase  test  system  is a device intended to measure the activ- ity of the enzyme galactose-1-phos- phate	uridyl	transferase	in erythrocytes (red blood cells). Meas- urements of galactose-1-phosphate uridyl transferase are used in the diag- nosis and treatment of the hereditary disease galactosemia (disorder of galac- tose metabolism) in infants.

Classification. Class II.

##### § 862.1320   Gastric acidity test system.

Identification. A gastric acidity test system is a device intended to measure the acidity of gastric fluid. Measurements of gastric acidity are used in the diagnosis and treatment of

#### Food and Drug Administration, HHS	§ 862.1350

patients with peptic ulcer, Zollinger- Ellison syndrome (peptic ulcer due to gastrin-secreting tumor of the pan- creas), and related gastric disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1325   Gastrin test system.

Identification. A gastrin test sys- tem is a device intended to measure the hormone gastrin in plasma and serum. Measurements of gastrin are used in the diagnosis and treatment of patients with ulcers, pernicious ane- mia, and the Zollinger-Ellison syn- drome (peptic ulcer due to a gastrin-se- creting tumor of the pancreas).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1330   Globulin test system.

Identification. A globulin test sys- tem is a device intended to measure globulins (proteins) in plasma and serum. Measurements of globulin are used in the diagnosis and treatment of patients with numerous illnesses in- cluding severe liver and renal disease, multiple myeloma, and other disorders of blood globulins.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1335   Glucagon test system.

Identification. A glucagon test sys- tem is a device intended to measure the pancreatic hormone glucagon in plasma and serum. Glucagon measure- ments are used in the diagnosis and treatment of patients with various dis- orders of carbohydrate metabolism, in-

cluding	diabetes	mellitus,	hypo- glycemia, and hyperglycemia.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1340   Urinary	glucose	(non- quantitative) test system.

Identification. A urinary glucose (nonquantitative) test system is a de- vice intended to measure glucosuria (glucose in urine). Urinary glucose (nonquantitative) measurements are used in the diagnosis and treatment of carbohydrate metabolism disorders in- cluding diabetes mellitus, hypo- glycemia, and hyperglycemia.

Classification. Class II.

##### § 862.1345   Glucose test system.

Identification. A glucose test sys- tem is a device intended to measure glucose quantitatively in blood and other body fluids. Glucose measure- ments are used in the diagnosis and treatment of carbohydrate metabolism disorders including diabetes mellitus, neonatal hypoglycemia, and idiopathic hypoglycemia, and of pancreatic islet cell carcinoma.

Classification. Class II.

##### § 862.1350   Continuous glucose monitor secondary display.

Identification. A continuous glu- cose monitor secondary display is iden- tified as a device intended to be used for passive real-time monitoring of continuous glucose monitoring data. It must not be capable of serving as a stand-alone primary display device. The primary display device, which is not a part of the continuous glucose monitor secondary display, directly re- ceives the glucose data (for example, it communicates directly with trans- mitter) from the continuous glucose meter, which is not a part of the con- tinuous glucose monitor secondary dis- play, and is the primary means of view- ing the continuous glucose monitor data and alerting the patient to a low or high glucose value. A continuous glucose monitor secondary display can

#### § 862.1360

be used by caregivers of people with di- abetes to monitor a person’s contin- uous glucose monitoring data. A device is not a continuous glucose monitor secondary display if the data from the primary display device is modified (for example, predicting future glucose val- ues) or the patient can use the sec- ondary display in lieu of a primary dis- play device (for example, the primary display device is blinded or the pri- mary display does not have to be near the person wearing the sensor and transmitter).

Classification. Class II (special controls). The special controls for this device are:

Devices being marketed must in- clude appropriate measures to protect against unauthorized access to  data and  unauthorized  modification  of  data.

The labeling must prominently and conspicuously display a warning that states ‘‘Dosing  decisions  should not be made based on this device. The user should follow instructions on the continuous glucose monitoring sys- tem.’’

The labeling for the device must include a statement that reads ‘‘This device is not intended to replace self- monitoring practices as advised by a physician.’’

[82 FR 13550, Mar. 14, 2017]

##### § 862.1360   Gamma-glutamyl transpeptidase and isoenzymes test system.

  - Identification. A gamma-glutamyl transpeptidase and isoenzymes test system is a device intended to measure the activity of the enzyme gamma- glutamyl transpeptidase (GGTP) in plasma and serum. Gamma-glutamyl transpeptidase and isoenzymes meas- urements are used in the diagnosis and treatment of liver diseases such as al- coholic cirrhosis and primary and sec- ondary liver tumors.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 862.1365   Glutathione test system.

Identification. A glutathione test system is a device intended to measure glutathione (the tripeptide of glycine, cysteine, and glutamic acid) in erythrocytes (red blood cells). Gluta- thione measurements are used in the diagnosis and treatment of certain drug-induced hemolytic (erythrocyte destroying) anemias due to an inher- ited enzyme deficiency.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1370   Human	growth	hormone test system.

Identification. A human growth hormone test system is a device in- tended to measure the levels of human growth hormone in plasma. Human growth hormone measurements are used in the diagnosis and treatment of disorders involving the anterior lobe of the pituitary gland.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1373   Hemoglobin A1c test system.

Identification. A hemoglobin A1c test system is a device used to measure the percentage concentration of hemo- globin A1c in blood. Measurement of hemoglobin A1c is used as an aid in the diagnosis of diabetes mellitus and as an aid in the identification of patients at risk for developing diabetes mellitus.

Classification. Class II (special controls). The special controls for this device are:

The device must have initial and annual standardization verification by a certifying glycohemoglobin standard- ization organization deemed acceptable by FDA.

The premarket notification sub- mission must include performance testing  to  evaluate  precision,  accuracy,

#### Food and Drug Administration, HHS	§ 862.1385

linearity,  and  interference,  including the following:

  - Performance testing of device pre- cision must, at a minimum, use blood samples with concentrations near 5.0 percent, 6.5 percent, 8.0 percent, and 12 percent hemoglobin A1c. This testing must evaluate precision over  a min- imum of 20 days using at least  three lots of the device and three instru- ments, as applicable.

  - Performance testing of device ac- curacy must include a minimum of 120 blood samples that span the measuring interval of the device and compare re- sults of  the  new  device to  results  of  a standardized  test  method.  Results must demonstrate little or no bias versus the standardized method.

  - Total error of the new device must be evaluated using single meas- urements by the new device compared to results of the standardized test method, and this evaluation must dem- onstrate a total error less  than  or equal to 6 percent.

  - Performance testing must dem- onstrate that there is little to no inter- ference from common hemoglobin variants, including Hemoglobin C, He- moglobin D, Hemoglobin E, Hemo- globin A2, and Hemoglobin S.

When assay interference from He- moglobin F or interference with other hemoglobin variants with low fre- quency in the population is observed, a warning statement must be placed in a black box and must appear in all label- ing material for these devices describ- ing the interference and any affected populations.

[79 FR 50551, Aug. 25, 2014]

##### § 862.1375   Histidine test system.

Identification. A histidine test sys- tem is a device intended to measure free histidine (an amino acid) in plas- ma and urine. Histidine measurements are used in the diagnosis and treat- ment of hereditary histidinemia char- acterized by excess histidine in the blood and urine often resulting in men- tal retardation and disordered speech development.

Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

subpart  E  of  part  807  of  this  chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1377   Urinary  homocystine  (non- quantitative) test system.

Identification. A urinary homocystine (nonquantitative) test system is a device intended to identify homocystine (an analogue of the amino acid cystine) in urine. The identifica- tion of urinary homocystine is used in the diagnosis and treatment of homocystinuria (homosystine in urine), a heritable metabolic disorder which may cause mental retardation.

Classification. Class II.

##### § 862.1380   Hydroxybutyric	dehydro- genase test system.

Identification. A hydroxybutyric dehydrogenase test system is a device intended to measure the activity of the enzyme alpha-hydroxybutric dehydro- genase (HBD) in plasma or serum. HBD measurements are used in the diag- nosis and treatment of myocardial in- farction, renal damage (such as rejec- tion	of	transplants),	certain hematological diseases (such as acute leukemias and megaloblastic anemias) and, to a lesser degree, liver disease.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38787,  July 25,

2001]

##### § 862.1385   17-Hydroxycorticosteroids (17-ketogenic steroids) test system.

Identification. A 17-hydroxy- corticosteroids (17-ketogenic steroids) test system is a device intended to measure corticosteroids that possess a dihydroxyacetone

moiety on the steroid nucleus in urine. Corticosteroids with this chemical con- figuration   include   cortisol,   cortisone 11-desoxycortisol, desoxycorticosterone,	and	their

#### § 862.1390

tetrahydroderivatives. This group of hormones is synthesized by the adrenal gland. Measurements of 17-hydroxy- corticosteroids (17-ketogenic steroids) are used in the diagnosis and treat- ment of various diseases of the adrenal or pituitary glands and gonadal dis- orders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987; 52 FR 29468, Aug. 7,

1987, as amended at 65 FR 2306, Jan. 14, 2000]

##### § 862.1390   5-Hydroxyindole acetic acid/ serotonin test system.

Identification. A 5-hydroxyindole acetic acid/serotonin test system is a device intended to measure 5- hydroxyindole acetic acid/serotonin in urine.	Measurements	of	5- hydroxyindole acetic acid/serotonin are used in the diagnosis and treatment of carcinoid tumors of endocrine tissue.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1395   17-Hydroxyprogesterone test system.

Identification.	A	17- hydroxyprogesterone test system is a device intended to measure 17- hydroxyprogesterone (a steroid) in plasma and serum. Measurements of 17- hydroxyprogesterone are used in the diagnosis and treatment of various dis- orders of the adrenal glands or the ova- ries.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1400   Hydroxyproline test system.

Identification. A hydroxyproline test system is a device intended to measure	the	amino		acid hydroxyproline	in	urine.

#### 21 CFR Ch. I (4–1–17 Edition)

Hydroxyproline measurements are used in the diagnosis and treatment of var- ious collagen (connective tissue) dis- eases, bone disease such as Paget’s dis- ease, and endocrine disorders such as hyperparathyroidism and hyper- thyroidism.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1405   Immunoreactive insulin test system.

Identification. An immunoreactive insulin test system is a device intended to measure immunoreactive insulin in serum and plasma. Immunoreactive in- sulin measurements are used in the di- agnosis and treatment of various car- bohydrate metabolism disorders, in- cluding diabetes mellitus, and hypo- glycemia.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2306, Jan. 14, 2000]

##### § 862.1410   Iron (non-heme) test system.

Identification. An  iron  (non-heme) test system is a device intended to measure iron (non-heme) in serum and plasma. Iron (non-heme) measurements are used in the diagnosis and treat- ment of diseases such as iron  defi- ciency anemia, hemochromatosis (a disease associated with widespread de- posit in the tissues of two iron-con- taining pigments, hemosiderin and hemofuscin, and characterized by pig- mentation of the skin), and  chronic renal disease.

Classification. Class I.

##### § 862.1415   Iron-binding   capacity   test system.

Identification. An iron-binding ca- pacity test system is a device intended to measure iron-binding capacity in serum. Iron-binding capacity measure- ments are used in the diagnosis and treatment of anemia.

#### Food and Drug Administration, HHS	§ 862.1450

Classification. Class I.

##### § 862.1420   Isocitric dehydrogenase test system.

Identification. An isocitric dehy- drogenase test system is a device in- tended to measure the activity of the enzyme isocitric dehydrogenase in serum and plasma. Isocitric dehydro- genase measurements are used in the diagnosis and treatment of liver dis- ease such as viral hepatitis, cirrhosis, or acute inflammation of the biliary tract; pulmonary disease such as pul- monary infarction (local arrest or sud- den insufficiency of the blood supply to the lungs), and diseases associated with pregnancy.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1430   17-Ketosteroids test system.

Identification. A 17-ketosteroids test system is a device intended to measure 17-ketosteroids in urine. Meas- urements of 17-ketosteroids are used in the diagnosis and treatment of dis- orders of the adrenal cortex and gonads and of other endocrine disorders, in- cluding hypertension, diabetes, and hypothyroidism.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1435   Ketones	(nonquantitative) test system.

Identification. A ketones (non- quantitative) test system is a device intended to identify ketones in urine and other body fluids. Identification of ketones is used in the diagnosis and treatment of acidosis (a condition characterized by abnormally high acid- ity of body fluids) or ketosis (a condi- tion characterized by increased produc- tion of ketone bodies such as acetone) and     for     monitoring     patients     on

ketogenic diets and patients with dia- betes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1440   Lactate  dehydrogenase  test system.

Identification. A lactate dehydro- genase test system is a device intended to measure the activity of the enzyme lactate dehydrogenase in serum. Lac- tate dehydrogenase measurements are used in the diagnosis and treatment of liver diseases such as acute viral hepa- titis, cirrhosis, and metastatic car- cinoma of the liver, cardiac diseases such as myocardial infarction, and tu- mors of the lung or kidneys.

Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 63

FR 59225, Nov. 3, 1998]

##### § 862.1445   Lactate	dehydrogenase isoenzymes test system.

Identification. A lactate dehydro- genase isoenzymes test system is a de- vice intended to measure the activity of lactate dehydrogenase isoenzymes (a group of enzymes with similar biologi- cal activity) in serum.  Measurements of lactate dehydrogenase isoenzymes are used in the diagnosis and treat- ment of liver diseases, such as viral hepatitis, and myocardial infarction.

Classification. Class II.

##### § 862.1450   Lactic acid test system.

Identification. A lactic acid test system is a device intended to measure lactic acid in whole blood and plasma. Lactic acid measurements that evalu- ate the acid-base status are used in the diagnosis and treatment of lactic aci- dosis (abnormally high acidity of the blood).

Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 862.1455

subpart  E  of  part  807  of  this  chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1455   Lecithin/sphingomyelin ratio in amniotic fluid test system.

  - Identification. A lecithin/ sphingomyelin ratio in amniotic fluid test system is a device intended to measure the lecithin/sphingomyelin ratio in amniotic fluid. Lecithin and sphingomyelin are phospholipids (fats or fat-like substances containing phos- phorus). Measurements of the lecithin/ sphingomyelin ratio in amniotic fluid are used in evaluating fetal maturity.

  - Classification. Class II.

##### § 862.1460   Leucine	aminopeptidase test system.

Identification. A leucine aminopeptidase test system is a device intended to measure the activity of the enzyme leucine amino-peptidase in serum, plasma, and urine. Leucine aminopeptidase measurements are used in the diagnosis and treatment of liver diseases such as viral hepatitis and ob- structive  jaundice.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1465   Lipase test system.

Identification. A lipase test system is a device intended to measure the ac- tivity of the enzymes lipase in serum. Lipase measurements are used in diag- nosis and treatment of diseases of the pancreas such as acute pancreatitis and obstruction of the pancreatic duct.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1470   Lipid (total) test system.

Identification. A lipid (total) test system is a device intended to measure total lipids (fats or fat-like substances)

#### 21 CFR Ch. I (4–1–17 Edition)

in serum and plasma. Lipid (total) measurements are used  in the diag- nosis and treatment of various diseases involving lipid metabolism and athero- sclerosis.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1475   Lipoprotein test system.

Identification. A lipoprotein test system is a device intended to measure lipoprotein in serum and plasma. Lipoprotein measurements are used in the diagnosis and treatment of lipid disorders (such as diabetes mellitus), atherosclerosis, and various liver and renal diseases.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1485   Luteinizing    hormone    test system.

Identification. A luteinizing hor- mone test system is a device intended to measure luteinizing hormone in serum and urine. Luteinizing hormone measurements are used in the diag- nosis and treatment of gonadal dys- function.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1490   Lysozyme (muramidase) test system.

Identification. A lysozyme (muramidase) test system is a device intended to measure the activity of the bacteriolytic	enzyme	lysozyme (muramidase) in serum, plasma, leu- kocytes, and urine. Lysozyme measure- ments  are  used  in  the  diagnosis  and

#### Food and Drug Administration, HHS	§ 862.1515

treatment of monocytic leukemia and kidney disease.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1495   Magnesium test system.

Identification. A magnesium test system is a device intended to measure magnesium levels in serum and plas- ma. Magnesium measurements are used in the diagnosis and treatment of hypomagnesemia (abnormally low plas- ma levels of magnesium) and hypermagnesemia (abnormally high plasma levels of magnesium).

Classification. Class I.

##### § 862.1500   Malic    dehydrogenase    test system.

Identification. A malic dehydro- genase test system is a device that is intended to measure the activity of the enzyme malic dehydrogenase in serum and plasma. Malic dehydrogenase measurements are used in the diag- nosis and treatment of muscle and liver diseases, myocardial infarctions, cancer, and blood disorders such as myelogenous (produced in the bone marrow) leukemia.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1505   Mucopolysaccharides   (non- quantitative) test system.

| measure the | levels | of |
| --- | --- | --- |
| mucopolysaccharides | in | urine. |

Identification.	A mucopolysaccharides (nonquantitative) test  system  is  a  device  intended  to

Mucopolysaccharide measurements in urine are used in the diagnosis and treatment of various inheritable dis- orders that affect bone and connective tissues, such as Hurler’s, Hunter’s, Sanfilippo’s, Scheie’s Morquio’s and Maroteaux-Lamy  syndromes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1509   Methylmalonic    acid    (non- quantitative) test system.

Identification. A methylmalonic acid (nonquantitative) test system is a device	intended	to	identify methylmalonic acid in urine. The iden- tification of methylmalonic acid in urine is used in the diagnosis and treatment of  methylmalonic  aciduria, a heritable metabolic disorder which, if untreated, may cause mental retarda- tion.

Classification. Class II.

##### § 862.1510   Nitrite	(nonquantitative) test system.

Identification. A nitrite (non- quantitative) test system is a device intended to identify nitrite in urine. Nitrite identification is used in the di- agnosis and treatment of uninary tract infection of bacterial origin.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1515   Nitrogen	(amino-nitrogen) test system.

Identification. A nitrogen (amino- nitrogen) test system is a device in- tended to measure amino acid nitrogen levels in serum, plasma, and urine. Ni- trogen (amino-nitrogen) measurements are used in the diagnosis and treat- ment of certain forms of severe liver disease and renal disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

#### § 862.1520

##### § 862.1520   5-Nucleotidase test system.

  - Identification. A 5-nucleotidase test system is a device intended to measure the activity of the enzyme 5- nucleotidase in serum and plasma. Measurements of 5-nucleotidase are used in the diagnosis and treatment of liver diseases and in the differentia- tions between liver and  bone  diseases in the presence of elevated serum alka- line  phosphatase  activity.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1530   Plasma  oncometry  test  sys- tem.

Identification. A plasma oncometry test system is a device in- tended to measure plasma oncotic pres- sure. Plasma oncotic pressure is that portion of the total fluid pressure con- tributed by proteins and other mol- ecules too large to pass through a spec- ified membrane. Measurements of plas- ma oncotic pressure are used in the di- agnosis and treatment of dehydration and circulatory disorders related to low serum protein levels and increased capillary permeability, such as edema and shock.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1535   Ornithine   carbamyl   trans- ferase test system.

Identification. An ornithine carbamyl transferase test system is a device intended to measure the activ- ity of the enzyme ornithine carbamyl transferase (OCT) in serum. Ornithine carbamyl transferase measurements are used in the diagnosis and treat- ment of liver diseases, such as infec- tious hepatitis, acute cholecystitis (in- flammation of the gall bladder), cir- rhosis, and liver metastases.

Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### 21 CFR Ch. I (4–1–17 Edition)

premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1540   Osmolality test system.

Identification. An osmolality test system is a device intended to measure ionic and nonionic solute concentra- tion in body fluids, such as serum and urine. Osmolality measurement is used as an adjunct to other tests in the eval- uation of a variety of diseases, includ- ing kidney diseases (e.g., chronic pro- gressive renal failure), diabetes insipidus, other endocrine and meta- bolic disorders, and fluid imbalances.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1542   Oxalate test system.

Identification. An oxalate test sys- tem is a device intended to measure the concentration of oxalate in urine. Measurements of oxalate are used to aid in the diagnosis or treatment of urinary stones or certain other meta- bolic disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1545   Parathyroid   hormone   test system.

Identification. A parathyroid hor- mone test system is a device intended to measure the levels of parathyroid hormone in serum and plasma. Meas- urements of parathyroid hormone lev- els are used in the differential diag- nosis of hypercalcemia (abnormally high levels of calcium in the blood) and hypocalcemia (abnormally low levels of calcium in the blood) resulting from disorders of calcium metabolism.

Classification. Class II.

#### Food and Drug Administration, HHS	§ 862.1580

##### § 862.1550   Urinary   pH   (nonquantita- tive) test system.

  - Identification. A urinary pH (non- quantitative) test system is a device intended to estimate the pH of urine. Estimations of pH are used to evaluate the acidity or alkalinity of urine as it relates to numerous renal and meta- bolic disorders and in the monitoring of patients with certain diets.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1555   Phenylalanine test system.

Identification. A phenylalanine test system is a device intended to measure free phenylalanine (an amino acid) in serum, plasma, and urine. Measurements of phenylalanine are used in the diagnosis and treatment of congenital phenylketonuria which, if untreated, may cause mental retarda- tion.

Classification. Class II.

##### § 862.1560   Urinary	phenylketones (nonquantitative) test system.

Identification. A urinary phenylketones (nonquantitative) test system is a device intended to identify phenylketones (such as phenylpyruvic acid) in urine. The identification of urinary phenylketones is used in the diagnosis and treatment of congenital phenylketonuria which, if untreated, may cause mental retardation.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1565   6-Phosphogluconate	dehy- drogenase test system.

Identification.	A	6- phosphogluconate dehydrogenase test system is a device intended to measure the activity of the enzyme 6- phosphogluconate dehydrogenase (6 PGD) in serum and erythrocytes. Meas- urements of 6-phosphogluconate dehy-

drogenase are used in the diagnosis and treatment of certain  liver  diseases (such as hepatitis) and anemias.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1570   Phosphohexose	isomerase test system.

Identification. A phosphohexose isomerase test system  is a  device in- tended to measure the activity of the enzyme phosphohexose isomerase in serum.	Measurements	of phosphohexose isomerase are used in the diagnosis and treatment of muscle diseases such as muscular dystrophy, liver diseases such as hepatitis or cir- rhosis, and metastatic carcinoma.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1575   Phospholipid test system.

Identification. A phospholipid test system is a device intended to measure phospholipids in serum and plasma. Measurements of phospholipids are used in the diagnosis and treatment of disorders involving lipid (fat) metabo- lism.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1580   Phosphorus (inorganic) test system.

Identification. A phosphorus (inor- ganic) test system is a device intended to measure inorganic phosphorus in serum, plasma, and urine. Measure- ments of phosphorus (inorganic) are used in the diagnosis and treatment of

#### § 862.1585

various disorders, including parathy- roid gland and kidney diseases, and vi- tamin D imbalance.

Classification. Class I.

##### § 862.1585   Human   placental   lactogen test system.

Identification. A human placental lactogen test system is a device in- tended to measure the hormone human placental lactogen (HPL), (also  known as	human	chorionic somatomammotrophin (HCS)), in ma- ternal serum and maternal plasma. Measurements of human placental lactogen are used in the diagnosis and clinical management of high-risk preg- nancies involving fetal distress associ- ated with placental insufficiency. Measurements of HPL are also used in pregnancies complicated by hyper- tension, proteinuria, edema, post-ma- turity, placental insufficiency, or pos- sible  miscarriage.

Classification. Class II.

##### § 862.1590   Porphobilinogen    test    sys- tem.

Identification. A porphobilinogen test system is a device intended to measure porphobilinogen (one of the derivatives of hemoglobin which can make the urine a red color) in urine. Measurements obtained by this device are used in the diagnosis and treat- ment of porphyrias (primarily inher- ited diseases associated with disturbed porphyrine metabolism), lead poi- soning, and other diseases character- ized by alterations in the heme path- way.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2307, Jan. 14, 2000]

##### § 862.1595   Porphyrins test system.

Identification. A porphyrins test system is a device intended to measure porphyrins (compounds formed during the biosynthesis of heme, a constituent of hemoglobin, and related compounds) in urine and feces. Measurements ob- tained by this device are used in the di- agnosis and treatment of lead poi- soning, porphyrias (primarily  inherited

#### 21 CFR Ch. I (4–1–17 Edition)

diseases associated with disturbed por- phyrin metabolism), and other diseases characterized by alterations  in  the heme pathway.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1600   Potassium test system.

Identification. A potassium test system is a device intended to measure potassium in serum, plasma, and urine. Measurements obtained by this device are used to monitor electrolyte balance in the diagnosis and treatment of dis- eases conditions characterized by  low or high blood potassium levels.

Classification. Class II.

##### § 862.1605   Pregnanediol test system.

Identification. A pregnanediol test system is a device intended to measure pregnanediol (a major urinary meta- bolic product of progesterone) in urine. Measurements obtained by this device are used in the diagnosis and treat- ment of disorders of the ovaries or pla- centa.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1610   Pregnanetriol test system.

Identification. A pregnanetriol test system is a device intended to measure pregnanetriol (a precursor in the bio- synthesis of the adrenal hormone cortisol) in urine. Measurements ob- tained by this device are used in the di- agnosis and treatment of congenital adrenal hyperplasia (congenital en- largement of the adrenal gland).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 862.1645

##### § 862.1615   Pregnenolone test system.

  - Identification. A pregnenolone test system is a device intended to measure pregnenolone (a precursor in the bio- synthesis of the adrenal hormone cortisol and adrenal androgen) in serum and plasma. Measurements ob- tained by this device are used in the di- agnosis and treatment of diseases of the adrenal cortex or the gonads.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1620   Progesterone test system.

Identification. A progesterone test system is a device intended to measure progesterone (a female hormone) in serum and plasma. Measurements ob- tained by this device are used in the di- agnosis and treatment of disorders of the ovaries or placenta.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1625   Prolactin (lactogen) test sys- tem.

Identification. A prolactin (lactogen) test system is a device in- tended to measure the anterior pitui- tary polypeptide hormone prolactin in serum and plasma. Measurements ob- tained by this device are used in the di- agnosis and treatment of disorders of the anterior pituitary gland or of the hypothalamus portion of the brain.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1630   Protein  (fractionation)  test system.

Identification. A protein (fraction- ation) test system is a device intended to measure protein fractions in blood,

urine, cerebrospinal fluid, and other body fluids. Protein fractionations are used as an aid in recognizing abnormal proteins in body fluids and genetic variants of proteins produced in dis- eases with tissue destruction.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1635   Total protein test system.

Identification. A total protein test system is a device intended to measure total protein(s) in serum or plasma. Measurements obtained by this device are used in the diagnosis and treat- ment of a variety of diseases involving the liver,  kidney,  or  bone  marrow  as well as other metabolic or nutritional disorders.

Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 63

FR 59225, Nov. 3, 1998]

##### § 862.1640   Protein-bound   iodine   test system.

Identification. A protein-bound io- dine test system is a device intended to measure protein-bound iodine  in serum. Measurements of protein-bound iodine obtained by this device are used in the diagnosis and treatment of thy- roid disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1645   Urinary protein or albumin (nonquantitative) test system.

Identification. A urinary protein or albumin (nonquantitative) test system is a device intended to identify pro- teins or albumin in urine. Identifica- tion   of   urinary   protein   or   albumin

#### § 862.1650

(nonquantitative) is used in the diag- nosis and treatment of disease condi- tions such as renal or heart diseases or thyroid disorders, which are character- ized by proteinuria or albuminuria.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1650   Pyruvate kinase test system.

Identification. A pyruvate kinase test system is a device intended to measure the activity of the enzyme pyruvate kinase in erythrocytes (red blood cells). Measurements obtained by this device are used in the diagnosis and treatment of various inherited anemias due to pyruvate kinase defi- ciency or of acute leukemias.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1655   Pyruvic acid test system.

Identification. A pyruvic acid test system is a device intended to measure pyruvic acid (an intermediate com- pound in the metabolism of carbo- hydrate) in plasma. Measurements ob- tained by this device are used in the evaluation of electrolyte metabolism and in the diagnosis and treatment of acid-base and electrolyte disturbances or anoxia (the reduction of oxygen in body tissues).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1660   Quality control material (as- sayed and unassayed).

Identification. A quality control material (assayed and unassayed) for clinical chemistry is a device intended for medical purposes for use in a test system to estimate test precision and

#### 21 CFR Ch. I (4–1–17 Edition)

to detect systematic analytical devi- ations that may arise from reagent or analytical instrument variation. A quality control material (assayed and unassayed) may be used for proficiency testing in interlaboratory  surveys. This generic type of device includes controls (assayed and unassayed) for blood gases, electrolytes, enzymes, multianalytes (all kinds), single (speci- fied) analytes, or urinalysis controls.

Classification. Class I (general con- trols). Except when used in donor screening tests, unassayed material is exempt from the premarket  notifica- tion procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1665   Sodium test system.

Identification. A sodium test sys- tem is a device intended to measure so- dium in serum, plasma, and urine. Measurements obtained by this device are used in the diagnosis and treat- ment of aldosteronism (excessive secre- tion of the hormone aldosterone), dia- betes insipidus (chronic excretion of large amounts of dilute urine, accom- panied by extreme thirst), adrenal hy- pertension, Addison’s disease (caused by destruction of the adrenal glands), dehydration, inappropriate antidiuretic hormone secretion, or other diseases involving  electrolyte  imbalance.

Classification. Class II.

##### § 862.1670   Sorbitol dehydrogenase test system.

Identification. A sorbitol dehydro- genase test system is a device intended to measure the activity of the enzyme sorbitol dehydrogenase in serum. Meas- urements obtained by this device are used in the diagnosis and treatment of liver disorders such as cirrhosis or acute hepatitis.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 862.1705

##### § 862.1675   Blood   specimen   collection device.

  - Identification. A blood specimen collection device is a device intended for medical purposes to collect and to handle blood specimens and to separate serum from nonserum (cellular) compo- nents prior to further testing. This ge- neric type device may include blood collection tubes, vials, systems, serum separators, blood collection trays, or vacuum sample tubes.

  - Classification. Class II.

##### § 862.1678   Tacrolimus test system.

Identification. A tacrolimus test system is a device intended to quan- titatively determine tacrolimus con- centrations as an aid in the manage- ment of transplant patients receiving therapy with this drug. This generic type of device includes immunoassays and chromatographic assays for tacrolimus.

Classification. Class II (special controls). The special  control  is  ‘‘Class II Special Controls  Guidance  Docu- ment: Cyclosporine and Tacrolimus As- says; Guidance for Industry and FDA.’’ See § 862.1(d) for the availability of this guidance document.

[67 FR 58329, Sept. 16, 2002]

##### § 862.1680   Testosterone test system.

Identification. A testosterone test system is a device intended to measure testosterone (a male sex hormone) in serum, plasma, and urine. Measure- ment of testosterone are used in the di- agnosis and treatment of disorders in- volving the male sex hormones (androgens), including primary and sec- ondary hypogonadism, delayed or pre- cocious puberty, impotence in males and, in females hirsutism (excessive hair) and virilization (masculinization) due to tumors, polycystic ovaries, and adrenogenital syndromes.

Classification. Class I.

[52 FR 16122, May 1, 1987; 53 FR 11645, Apr. 8,

1988]

##### § 862.1685   Thyroxine-binding   globulin test system.

Identification. A thyroxine-binding globulin test system is a device in- tended to measure thyroxine (thyroid)- binding globulin (TBG), a plasma pro-

tein which binds thyroxine, in serum and plasma. Measurements obtained by this device are used in the diagnosis and treatment of thyroid diseases.

Classification. Class II.

##### § 862.1690   Thyroid	stimulating	hor- mone test system.

Identification. A thyroid stimu- lating hormone test system is a device intended to measure thyroid stimu- lating hormone, also known as thyrot- rophin and thyrotrophic hormone, in serum and plasma. Measurements of thyroid stimulating hormone produced by the anterior pituitary are used  in the diagnosis of thyroid or pituitary disorders.

Classification. Class II.

##### § 862.1695   Free thyroxine test system.

Identification. A free thyroxine test system is a device intended to measure free (not protein bound) thyroxine (thyroid hormone) in serum or plasma. Levels of free thyroxine in plasma are thought to reflect the amount of thyroxine hormone avail- able to the cells and may therefore de- termine the clinical metabolic  status of thyroxine. Measurements obtained by this device are used in the diagnosis and treatment of thyroid diseases.

Classification. Class II.

##### § 862.1700   Total thyroxine test system.

Identification. A total thyroxine test system is a device intended to measure total (free and protein bound) thyroxine (thyroid hormone) in serum and plasma. Measurements obtained by this device are used in the diagnosis and treatment of thyroid diseases.

Classification. Class II.

##### § 862.1705   Triglyceride test system.

Identification. A triglyceride test system is a device intended to measure triglyceride (neutral fat) in serum and plasma. Measurements obtained by this device are used in the diagnosis and treatment of patients with diabetes mellitus, nephrosis, liver obstruction, other diseases involving lipid metabo- lism, or various endocrine disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 862.1710

subpart  E  of  part  807  of  this  chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1710   Total  triiodothyronine  test system.

  - Identification.	A	total triiodothyronine test system is a de- vice intended to measure the hormone triiodothyronine in serum and plasma. Measurements obtained by this device are used in the diagnosis and treat- ment of thyroid diseases such as hyper- thyroidism.

  - Classification. Class II. This device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter subject to the limita- tions in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 62286, Oct. 18, 2000]

##### § 862.1715   Triiodothyronine	uptake test system.

Identification. A triiodothyronine uptake test system is a device intended to measure the total amount of binding sites available for binding thyroid hor- mone on the thyroxine-binding pro- teins,	thyroid-binding	globulin, thyroxine-binding prealbumin, and al- bumin of serum and plasma. The device provides an indirect measurement of thyrkoxine levels in serum and plasma. Measurements of triiodothyronine up- take are used in the diagnosis and treatment of thyroid disorders.

Classification. Class II. The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter subject to the limita- tions in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 64

FR 1124, Jan. 8, 1999]

##### § 862.1720   Triose phosphate isomerase test system.

Identification. A triose phosphate isomerase test system  is a  device in- tended to measure the activity of the enzyme triose phosphate isomerase in erythrocytes (red blood cells). Triose phosphate isomerase is an enzyme im- portant in glycolysis (the energy-yield- ing conversion of glucose to lactic acid in various tissues). Measurements ob- tained by this device are used in the di-

#### 21 CFR Ch. I (4–1–17 Edition)

agnosis and treatment of congenital triose phosphate isomerase enzyme de- ficiency, which causes a type of hemo- lytic anemia.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1725   Trypsin test system.

Identification. A trypsin test sys- tem is a device intended to measure the activity of trypsin (a pancreatic enzyme important in digestion for the breakdown of  proteins)  in  blood  and other body fluids and in feces. Measure- ments obtained by this device are used in the diagnosis and treatment of pan- creatic disease.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1730   Free tyrosine test system.

Identification. A free tyrosine test system is a device intended to measure free tyrosine (an amono acid) in serum and urine. Measurements obtained by this device are used in the diagnosis and treatment of diseases such as con- genital tyrosinemia (a disease that can cause liver/kidney disorders) and as an adjunct to the measurement of phenylalanine in detecting congenital phenylketonuria (a disease that can cause brain damage).

Classification. Class I.

##### § 862.1770   Urea nitrogen test system.

Identification. A urea nitrogen test system is a device intended to measure urea nitrogen (an end-product of nitro- gen metabolism) in whole blood, serum, plasma, and urine. Measure- ments obtained by this device are used in the diagnosis and treatment of cer- tain renal and metabolic diseases.

Classification. Class II.

#### Food and Drug Administration, HHS	§ 862.1815

##### § 862.1775   Uric acid test system.

  - Identification. A uric acid test sys- tem is a device intended to measure uric acid in serum, plasma, and urine. Measurements obtained by this device are used in the diagnosis and treat- ment of numerous renal and metabolic disorders, including renal failure, gout, leukemia, psoriasis, starvation or other wasting conditions, and of pa- tients receiving cytotoxic drugs.

  - Classification. Class I.

##### § 862.1780   Urinary calculi (stones) test system.

Identification. A urinary calculi (stones) test system is a device in- tended for the analysis of urinary calculi. Analysis of urinary calculi is used in the diagnosis and treatment of calculi of the urinary tract.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1785   Urinary  urobilinogen  (non- quantitative) test system.

Identification. A urinary urobi- linogen (nonquantitative) test  system is a device intended to detect and esti- mate urobilinogen (a bile pigment deg- radation product of red cell hemo- globin) in urine. Estimations obtained by this device are used in the diagnosis and treatment of liver diseases and he- molytic (red cells) disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1790   Uroporphyrin test system.

Identification. A uroporphyrin test system is a device intended to measure uroporphyrin in urine. Measurements obtained by this device are used in the diagnosis and treatment of porphyrias (primarily inherited diseases associ- ated with disturbed porphyrin metabo- lism),  lead  poisoning,  and  other  dis-

eases   characterized   by   alterations   in the heme pathway.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1795   Vanilmandelic acid test sys- tem.

Identification. A vanilmandelic acid test system is a device intended to measure vanilmandelic acid in urine. Measurements of vanilmandelic acid obtained by this device are used in the diagnosis and treatment of neuro- blastoma, pheochromocytoma, and cer- tain  hypertensive  conditions.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1805   Vitamin A test system.

Identification. A vitamin A test system is a device intended to measure vitamin A in serum or plasma. Meas- urements obtained by this device are used in the diagnosis and treatment of vitamin A  deficiency conditions,  in- cluding night blindness, or skin, eye, or intestinal disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1810   Vitamin B 12  test system.

Identification. A vitamin B12 test system is a device intended to measure vitamin B12 in serum, plasma, and urine. Measurements obtained by this device are used in the diagnosis and treatment of anemias of gastro- intestinal malabsorption.

Classification. Class II.

##### § 862.1815   Vitamin E test system.

Identification. A vitamin E test system is a device intended to measure

#### § 862.1820

vitamin E (tocopherol) in serum. Meas- urements obtained by this device are used in the diagnosis and treatment of infants with vitamin E deficiency syn- drome.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.1820   Xylose test system.

Identification. A xylose test sys- tem is a device intended to measure xylose (a sugar) in serum, plasma, and urine. Measurements obtained by this device are used in the diagnosis and treatment of gastrointestinal mal- absorption syndrome (a group of dis- orders in which there is subnormal ab- sorption of dietary constituents and thus excessive loss from the body  of the nonabsorbed substances).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.1825   Vitamin D test system.

Identification. A vitamin D test system is a device intended for use in clinical laboratories for the quan- titative determination of 25- hydroxyvitamin D (25-OH-D) and other hydroxylated metabolites of vitamin D in serum or plasma to be used in the assessment of vitamin D sufficiency.

Classification. Class II (special controls). Vitamin D test systems must comply with the following special con- trols:

Labeling in conformance with 21 CFR 809.10 and

Compliance with existing stand- ards of the National Committee on Clinical Laboratory Standards.

[63 FR 40366, July 29, 1998]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart C—Clinical Laboratory Instruments

##### § 862.2050 General purpose laboratory equipment labeled or promoted for a specific medical use.

Identification. General purpose laboratory equipment labeled or pro- moted for  a specific  medical use is  a device that is intended to prepare or examine specimens from the human body and that is labeled or promoted for a specific medical use.

Classification. Class I (general con- trols). The device is identified in para- graph (a) of this section and is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter   subject   to   the   limitations   in

§ 862.9. The device is also exempt  from the current good  manufacturing  prac- tice requirements  of  the  quality  sys- tem regulation in part 820 of this chap- ter, with the exception of § 820.180, with respect to general requirements con- cerning records, and § 820.198, with re- spect  to  complaint  files.

[52  FR  16122,  May  1,  1987,  as  amended  at  66

FR 38788, July 25, 2001]

##### § 862.2100   Calculator/data	processing module for clinical use.

Identification. A calculator/data processing module for clinical use is an electronic device intended to store, re- trieve, and process laboratory data.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 53

FR 21449,  June 8,  1988;  66  FR 38788,  July 25,

2001]

##### § 862.2140   Centrifugal  chemistry  ana- lyzer for clinical use.

Identification. A centrifugal chem- istry analyzer for clinical use is an automatic device intended to cen- trifugally mix a sample and a reagent and spectrophotometrically measure concentrations of the sample constitu- ents. This device is intended for use in conjunction with certain materials to measure a variety of analytes.

Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	§ 862.2250

premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.2150 Continuous flow sequential multiple chemistry analyzer for clinical use.

  - Identification. A continuous flow sequential multiple chemistry analyzer for clinical use is a modular analytical instrument intended to simultaneously perform multiple chemical procedures using the principles of automated con- tinuous flow systems. This device is in- tended for use in conjunction with cer- tain materials to measure a variety of analytes.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2308, Jan. 14, 2000]

##### § 862.2160   Discrete  photometric  chem- istry analyzer for clinical use.

Identification. A discrete photo- metric chemistry analyzer for clinical use is a device intended to duplicate manual analytical procedures by per- forming automatically various steps such as pipetting, preparing filtrates, heating, and measuring color intensity. This device is intended for use in con- junction with certain materials to measure a variety of analytes. Dif- ferent models of the device incorporate various instrumentation such as micro analysis apparatus, double beam, sin- gle, or dual channel photometers, and bichromatic	2-wavelength photometers. Some models of the de- vice may include reagent-containing components that may also serve as re- action units.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2170   Micro    chemistry    analyzer for clinical use.

Identification. A micro chemistry analyzer for clinical use is a device in- tended to duplicate manual analytical procedures by performing automati- cally various steps such as pipetting, preparing filtrates, heating, and meas- uring color intensity. The distin- guishing characteristic of the device is that it requires only micro volume samples obtainable from pediatric pa- tients. This device is intended for use in conjunction with certain  materials to measure a variety of analytes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2230   Chromatographic	separa- tion material for clinical use.

Identification. A chromatographic separation material for clinical use is a device accessory (e.g., ion exchange absorbents, ion exchagne resins, and ion papers) intended for use in ion ex- change chromatography, a procedure in which a compound is separated from a solution.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38788,  July 25,

2001]

##### § 862.2250   Gas  liquid  chromatography system for clinical use.

Identification. A gas liquid chro- matography system for clinical use is a device intended to separate one or more drugs or compounds from a mix- ture. Each of the constituents in a va- porized mixture of compounds is sepa- rated according to its vapor pressure. The device may include accessories such as columns, gases, column sup- ports, and liquid coating.

Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 862.2260

subpart  E  of  part  807  of  this  chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2260   High  pressure  liquid  chro- matography system for clinical use.

  - Identification. A high pressure liq- uid chromatography system for clinical use is a device intended to separate one or more drugs or compounds from a so- lution by processing the mixture of compounds (solutes) through a column packed with materials of uniform size (stationary phase) under the  influence of a high pressure  liquid  (mobile phase). Separation of the solutes oc- curs either by absorption, sieving, par- tition,  or  selective  affinity.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2265   High   throughput   genomic sequence analyzer for clinical use.

Identification. A high throughput genomic sequence analyzer for clinical use is an analytical instrument system intended to generate, measure and sort signals in order to analyze nucleic acid sequences in a clinical sample. The de- vice may include a signal reader unit; reagent handling, dedicated instrument control, and other hardware compo- nents; raw data storage mechanisms; data acquisition software; and software to process detected signals.

Classification. Class II (special controls). The special controls for this device are:

The labeling for the instrument system must reference legally mar- keted pre-analytical and analytical re- agents to be used with the instrument system and include or reference legally marketed analytical software that in- cludes sequence alignment and variant calling functions, to be used with the instrument system.

The labeling for the instrument system must include a description of the following information:

  - The specimen type(s) validated as an appropriate source of nucleic acid for  this  instrument.

#### 21 CFR Ch. I (4–1–17 Edition)

  - The type(s) of nucleic acids (e.g., germline DNA, tumor DNA) validated with this instrument.

  - The type(s) of sequence vari- ations (e.g. single nucleotide variants, insertions, deletions) validated with this instrument.

  - The type(s) of sequencing (e.g., targeted sequencing)  validated  with this  instrument.

  - The appropriate read depth for the sensitivity claimed and validation information supporting those claims.

  - The nucleic acid extraction method(s) validated for use with the in- strument system.

  - Limitations must specify the types of sequence variations that the instrument cannot detect with the claimed accuracy and precision (e.g., insertions or deletions larger than a certain size, translocations).

  - Performance characteristics of the instrument system must include:

Reproducibility data generated using multiple instruments and mul- tiple operators, and  at multiple sites. Samples tested must include all claimed specimen types, nucleic acid types, sequence variation types, and types of sequencing. Variants queried shall be located in varying sequence context (e.g., different chromosomes, GC-rich regions). Device results shall be compared to reference sequence data with high confidence.

Accuracy data for all claimed specimen types and nucleic acid types generated by testing a panel of well characterized samples to query all claimed sequence  variation  types, types of sequencing, and sequences lo- cated in varying sequence context (e.g., different chromosomes, GC-rich re- gions). The well-characterized sample panel shall include samples  from  at least two sources that have highly con- fident sequence based on well-validated sequencing methods. At least one ref- erence source shall have sequence gen- erated independently of the manufac- turer with respect to technology and analysis. Percent agreement and per- cent disagreement with the reference sequences must be described for all re- gions queried by the instrument.

If applicable, data describing en- dogenous or exogenous substances that

#### Food and Drug Administration, HHS	§ 862.2400

may interfere with the instrument sys- tem.

If applicable, data demonstrating the ability of the system to consist- ently generate an accurate result for a given sample across different indexing primer combinations.

  - The upper and lower limit of input nucleic acid that will achieve the claimed accuracy and reproducibility. Data supporting such claims must also be summarized.

[82 FR 13552, Mar. 14, 2017]

##### § 862.2270   Thin-layer   chromatography system for clinical use.

Identification. A thin-layer chro- matography (TLC) system for clinical use is a device intended to separate one or more drugs or compounds from a mixture. The mixture of compounds is absorbed onto a stationary phase or thin layer of inert material (e.g., cel- lulose, alumina, etc.) and eluted off by a moving solvent (moving phase) until equilibrium occurs between the two phases.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9. Particular compo- nents of TLC systems,  i.e.,  the  thin- layer chromatography apparatus, TLC atomizer, TLC developing tanks,  and TLC ultraviolet light, are exempt from the current good manufacturing prac- tice requirements of the quality sys- tem regulation in part 820 of this chap- ter, with the exception of § 820.180 of this chapter, with respect to general requirements  concerning  records,  and

§ 820.198 of this chapter, with respect to complaint    files.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2300   Colorimeter, photometer, or spectrophotometer for clinical use.

Identification. A colorimeter, a photometer, or a  spectrophotometer for clinical  use is an  instrument in- tended to measure radiant energy emitted, transmitted, absorbed, or re- flected under  controlled  conditions. The device may include a monochromator to produce light of a specific  wavelength.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2310   Clinical	sample	concen- trator.

Identification. A clinical sample concentrator is a device intended to concentrate (by dialysis, evaporation, etc.) serum, urine, cerebrospinal fluid, and other body fluids before the fluids are analyzed.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38899,  July 28,  1995;  66 FR 38788, July 25,

2001]

##### § 862.2320   Beta or gamma counter for clinical use.

Identification. A beta or gamma counter for clinical use is a device in- tended to detect and count beta or gamma radiation emitted by clinical samples. Clinical samples are prepared by addition of a radioactive reagent to the sample. These measurements are useful in the diagnosis and treatment of various disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38788, July 25,

2001]

##### § 862.2400 Densitometer/scanner (inte- grating, reflectance, TLC,  or radiochromatogram) for clinical use.

Identification. A densitometer/ scanner (integrating, reflectance, thin- layer	chromatography,	or radiochromatogram) for clinical use is device intended to measure the con- centration of a substance on the sur- face of a film or  other  support  media by either a photocell measurement of the light transmission through a given area  of  the  medium  or,  in  the  case  of

#### § 862.2485

the radiochromatogram scanner, by measurement of the distribution of a specific radio-active element on a radiochromatogram.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2485   Electrophoresis	apparatus for clinical use.

Identification. An electrophoresis apparatus for clinical use is a device intended to separate molecules or par- ticles, including plasma proteins, lipoproteins,	enzymes,	and hemoglobulins on the basis of their net charge in specified buffered media. This device is used in conjunction with certain materials to measure  a  variety of analytes as an aid in the  diagnosis and treatment of certain disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38788, July 25,

2001]

##### § 862.2500   Enzyme analyzer for clinical use.

Identification. An enzyme analyzer for clinical use is a device intended to measure enzymes in plasma or  serum by nonkinetic or kinetic  measurement of enzyme-catalyzed reactions. This de- vice is used in conjunction with certain materials to measure a variety of en- zymes as an aid in the diagnosis and treatment of certain enzyme-related disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2540   Flame  emission  photometer for clinical use.

Identification. A flame emission photometer for clinical use is a device

#### 21 CFR Ch. I (4–1–17 Edition)

intended to measure the concentration of sodium, potassium, lithium, and other metal ions in body fluids. Abnor- mal variations in the concentration of these substances in the body are indic- ative of certain disorders (e.g., electro- lyte imbalance and heavy metal intoxi- cation) and are, therefore, useful in di- agnosis and treatment of those dis- orders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2560   Fluorometer for clinical use.

Identification. A fluorometer for clinical use is a device intended to measure by fluorescence certain analytes. Fluorescence is the property of certain substances of radiating, when illuminated, a light of a different wavelength. This device is used in con- junction with certain materials to measure a variety of analytes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2570   Instrumentation for clinical multiplex test systems.

Identification. Instrumentation for clinical multiplex test systems is a de- vice intended to measure and sort mul- tiple signals generated by an assay from a clinical sample. This instru- mentation is used with a specific assay to measure multiple similar analytes that establish a single indicator to aid in diagnosis.  Such instrumentation may be compatible with more than one specific assay. The device includes a signal reader unit, and may also inte- grate reagent handling, hybridization, washing, dedicated instrument control, and other hardware components, as well as raw data storage mechanisms, data acquisition software, and software to process detected signals.

Classification. Class II (special controls).  The  special  control  is  FDA’s

#### Food and Drug Administration, HHS	§ 862.2750

guidance document entitled ‘‘Class II Special Controls Guidance Document: Instrumentation for Clinical Multiplex Test Systems.’’ See § 862.1(d) for the availability  of  this  guidance  document.

[70 FR 11868, Mar. 10, 2005]

##### § 862.2680   Microtitrator	for	clinical use.

  - Identification. A microtitrator for clinical use is a device intended for use in micronanalysis to measure the con- centration of a substance by reacting it with a measure ‘‘micro’’ volume of a known standardized solution.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2700   Nephelometer    for    clinical use.

Identification. A nephelometer for clinical use is a device intended to esti- mate the concentration of particles in a suspension by measuring their light scattering properties (the deflection of light rays by opaque particles in their path). The device is used in conjunc- tion with certain materials to measure the concentration of a variety of analytes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2720   Plasma  oncometer  for  clin- ical use.

Identification. A plasma oncometer for clinical use is a device intended to measure plasma oncotic pressure, which is that portion of the total plas- ma osmotic pressure contributed by protein and other molecules too large to pass through a specified semipermeable membrane. Because variations in plasma oncotic pressure are indications of certain disorders, measurements of the variations are useful in the diagnosis and treatment of these disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38788, July 25,

2001]

##### § 862.2730   Osmometer for clinical use.

Identification. An osmometer for clinical use is a device intended to measure the osmotic pressure of body fluids. Osmotic pressure is the pressure required to prevent the passage of a so- lution with a lesser solute concentra- tion into a solution with greater solute concentration when the two solutions are separated by a semipermeable membrane. The concentration of a so- lution affects its osmotic pressure, freezing	point,	and		other physiochemical		properties. Osmometers determine osmotic pres- sure by methods such as the measure- ment of the freezing point. Measure- ments obtained by this device are used in the diagnosis and treatment of body fluid disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2750   Pipetting  and  diluting  sys- tem for clinical use.

Identification. A pipetting and di- luting system for clinical use is a de- vice intended to provide an accurately measured volume of liquid at a speci- fied temperature for use in certain test procedures. This generic type of device system includes serial, manual, auto- mated, and semi-automated dilutors, pipettors, dispensers, and pipetting sta- tions.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

#### § 862.2800

##### § 862.2800   Refractometer   for   clinical use.

  - Identification. A refractometer for clinical use is a device intended to de- termine the amount of solute in a solu- tion by measuring the index of refrac- tion (the ratio of the velocity of light in a vacuum to the velocity of light in the solution). The index of refraction is used to measure the concentration of certain analytes (solutes), such a plas- ma total proteins and urinary total solids. Measurements obtained by this device are used in the diagnosis and treatment of certain conditions.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38788, July 25,

2001]

##### § 862.2850   Atomic  absorption  spectro- photometer for clinical use.

Identification. An atomic absorp- tion spectrophotometer for clinical use is a device intended to identify and measure elements and metals  (e.g., lead and mercury) in human specimens. The metal elements are identified ac- cording to the wavelength and inten- sity of the light that is absorbed when the specimen is converted to the atom- ic vapor phase. Measurements obtained by this device are used in the diagnosis and treatment of certain conditions.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2860   Mass spectrometer  for  clin- ical use.

Identification. A mass spectrom- eter for clinical use is a device in- tended to identify inorganic or organic compounds (e.g., lead, mercury, and drugs) in human specimens by ionizing the compound under investigation and separating the resulting ions by means of an electrical and magnetic field ac- cording to their mass.

#### 21 CFR Ch. I (4–1–17 Edition)

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2900   Automated   urinalysis   sys- tem.

Identification. An automated uri- nalysis system is a device intended to measure certain of the physical prop- erties and chemical constituents of urine by procedures that duplicate manual urinalysis systems. This device is used in conjunction with certain ma- terials to measure a variety of urinary analytes.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.2920   Plasma  viscometer  for clin- ical use.

Identification. A plasma viscom- eter for clinical use is a device in- tended to measure the viscosity of plasma by determining the time period required for the plasma to flow a meas- ured distance through a calibrated glass tube. Measurements obtained by this device are used to monitor changes in the amount of solids present in plas- ma in various disorders.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 862.9.

[52 FR 16122, May 1, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38788, July 25,

2001]

## Subpart D—Clinical Toxicology Test Systems

##### § 862.3030   Acetaminophen test system.

Identification. An acetaminophen test system is a device intended to measure acetaminophen, an analgestic and fever reducing drug, in serum. Measurements obtained by this device

Food and Drug Administration, HHS	§ 862.3170

are  used  in  the  diagnosis  and  treat- ment of acetaminophen overdose.

Classification. Class II.

##### § 862.3035   Amikacin test system.

Identification. An amikacin test system is a device intended to measure amikacin, an aminoglycoside anti- biotic drug, in serum and plasma. Measurements obtained by this device are used in the diagnosis and treat- ment of amikacin overdose and in mon- itoring levels of amikacin to ensure ap- propriate therapy.

Classification. Class II.

##### § 862.3040   Alcohol test system.

Identification. An alcohol test sys- tem is a device intented to measure al- cohol (e.g., ethanol, methanol, isopropanol, etc.) in human body fluids (e.g., serum, whole blood, and urine). Measurements obtained by this device are used in the diagnosis and treat- ment of alcohol intoxication and poi- soning.

Classification. Class II.

##### § 862.3050   Breath-alcohol test system.

Identification. A breath-alcohol test system is a device intened to measure alcohol in the human breath. Measurements obtained by this device are used in the diagnosis of alcohol in- toxication.

Classification. Class I.

##### § 862.3080   Breath nitric oxide test sys- tem.

Identification. A breath nitric oxide test system is  a  device  intended to measure fractional  nitric oxide in human  breath.  Measurement  of changes in fractional nitric oxide con- centration in expired breath aids in evaluating an asthma patient’s re- sponse  to  anti-inflammatory  therapy, as an  adjunct  to  established  clinical and laboratory assessments  of  asthma. A breath nitric oxide test system com- bines chemiluminescence detection of nitric oxide with a pneumotachograph, display, and dedicated software.

Classification. Class II (special controls). The special control is FDA’s guidance entitled ‘‘Class II Special Controls Guidance Document: Breath Nitric     Oxide     Test     System.’’     See

§ 862.1(d)   for   the   availability   of   this guidance document.

[68 FR 40127, July 7, 2003]

##### § 862.3100   Amphetamine test system.

Identification. An amphetamine test system is a device intended to measure amphetamine, a central nerv- ous system stimulating drug, in plasma and urine. Measurements obtained by this device are used in the diagnosis and treatment of amphetamine use or overdose and in monitoring levels of amphetamine to  ensure  appropriate therapy.

Classification. Class II.

##### § 862.3110   Antimony test system.

Identification. An antimony test system is a device intended to measure antimony, a heavy metal, in urine, blood, vomitus, and stomach contents. Measurements obtained by this device are used in the diagnosis and treat- ment of antimony poisoning.

Classification. Class I.

##### § 862.3120   Arsenic test system.

Identification. An arsenic test sys- tem is a device intended to measure ar- senic, a poisonous heavy metal, in urine, vomitus, stomach contents, nails, hair, and blood. Measurements obtained by this device are used in the diagnosis and treatment of arsenic poi- soning.

Classification. Class I.

##### § 862.3150   Barbiturate test system.

Identification. A barbiturate test system is a device intended to measure barbiturates, a class of hypnotic and sedative drugs, in serum, urine, and gastric contents. Measurements ob- tained by this device are used in the di- agnosis and treatment of barbiturate use or overdose and in monitoring lev- els of barbiturate to ensure appropriate therapy.

Classification. Class II.

##### § 862.3170   Benzodiazepine test system.

Identification. A benzodiazepine test system is a device intended to measure any of the benzodiazepine compounds, sedative  and hypnotic drugs, in blood, plasma, and urine. The benzodiazepine    compounds    include

#### § 862.3200

chlordiazepoxide, diazepam, oxazepam, chlorzepate,	flurazepam,	and nitrazepam. Measurements obtained by this device are used in the diagnosis and treatment of benzodiazepine use or overdose and in monitoring levels of benzodiazepines to ensure appropriate therapy.

Classification. Class II.

##### § 862.3200   Clinical	toxicology	cali- brator.

Identification. A clinical  toxi- cology  calibrator  is  a  device  intended for medical purposes for use in a test system to establish points of reference that are used in the determination of values in the measurement of  sub- stances in human specimens. A clinical toxicology calibrator can be  a  mixture of drugs or a specific material for a particular  drug  (e.g.,   ethanol,   lido- caine, etc.). (See also § 862.2  in  this part.)

Classification. Class II.

##### § 862.3220   Carbon  monoxide  test  sys- tem.

Identification. A carbon monoxide test system is a device intended to measure carbon monoxide or carboxyhemoglobin (carbon monoxide bound to the hemoglobin in the blood) in blood. Measurements  obtained  by this device are used in the diagnosis and treatment of or confirmation of carbon monoxide poisoning.

Classification. Class I.

##### § 862.3240   Cholinesterase test system.

Identification. A cholinesterase test system is a device intended to measure cholinesterase (an enzyme that catalyzes the hydrolysis of acetyl- choline to choline) in human speci- mens. There are two principal types of cholinesterase in human tissues. True cholinesterase is present at nerve endings and in erythrocytes (red blood cells) but is not present in plasma. Pseudo cholinesterase is present in plasma and liver but is not present in erythrocytes. Measurements obtained by this device are used in the diagnosis and treatment of cholinesterase inhibi- tion disorders (e.g., insecticide poi- soning and succinylcholine poisoning).

Classification. Class I.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 862.3250   Cocaine   and   cocaine   me- tabolite test system.

Identification. A cocaine and co- caine metabolite test system is a de- vice intended to measure cocaine and a cocaine metabolite  (benzoylecgonine) in serum, plasma, and urine. Measure- ments obtained by this device are used in the diagnosis and treatment of co- caine use or overdose.

Classification. Class II.

##### § 862.3270   Codeine test system.

Identification. A codeine test sys- tem is a device intended to measure co- deine (a narcotic pain-relieving  drug) in serum and urine. Measurements ob- tained by this device are used in the di- agnosis and treatment of  codeine  use or overdose and in monitoring levels of codeine to ensure appropriate therapy.

Classification. Class II.

##### § 862.3280   Clinical   toxicology   control material.

Identification. A clinical toxi- cology control material is a device in- tended to provide an estimation of the precision of a device test system and to detect and monitor systematic devi- ations from accuracy resulting from re- agent or instrument defects. This ge- neric type of device includes various single, and multi-analyte control ma- terials.

Classification. Class I (general con- trols). Except when used in donor screening, unassayed  material  is  ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 862.9.

[52 FR 16122, May 1, 1987, as amended at 65

FR 2309, Jan. 14, 2000]

##### § 862.3300   Digitoxin test system.

Identification. A digitoxin test sys- tem is a device intended to measure digitoxin, a cardiovascular drug, in serum and plasma. Measurements ob- tained by this device are used in the di- agnosis and treatment of digitoxin overdose and in monitoring levels of digitoxin to ensure appropriate ther- apy.

Classification. Class II.

#### Food and Drug Administration, HHS	§ 862.3560

##### § 862.3320   Digoxin test system.

  - Identification. A digoxin test sys- tem is a device intended to measure digoxin, a cardiovascular drug,  in serum and plasma. Measurements ob- tained by this device are used in the di- agnosis and treatment of digoxin over- dose and in monitoring levels of dig- oxin to ensure appropriate therapy.

  - Classification. Class II.

##### § 862.3350   Diphenylhydantoin test sys- tem.

Identification. A diphenylhydan- toin test system is a device intended to measure	diphenylhydantoin,	an antiepileptic drug, in human speci- mens. Measurements obtained by this device are used in the diagnosis and treatment of diphenylhydantoin over- dose and in monitoring levels of di- phenylhydantoin to ensure appropriate therapy.

Classification. Class II.

##### § 862.3360   Drug  metabolizing  enzyme genotyping system.

Identification. A drug metabolizing enzyme genotyping system is a device intended for use in testing deoxyribonucleic acid (DNA) extracted from clinical  samples to identify  the presence or absence  of  human genotypic markers encoding a drug me- tabolizing enzyme. This device  is  used as an aid in determining treatment choice and individualizing treatment dose for therapeutics that are metabo- lized primarily by the specific enzyme about which the system provides genotypic   information.

Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Drug Metabolizing Enzyme Genotyping Test System.’’ See § 862.1(d) for the availability  of  this  guidance  document.

[70 FR 11867, Mar. 10, 2005]

##### § 862.3380   Ethosuximide test system.

Identification. An ethosuximide test system is a device intended to measure ethosuximide, an antiepileptic drug, in human specimens. Measure- ments obtained by this device are used in the diagnosis and treatment of ethosuximide  overdose  and  in  moni-

toring levels of ethosuximide to ensure appropriate therapy.

Classification. Class II.

##### § 862.3450   Gentamicin test system.

Identification. A gentamicin test system is a device intended to measure gentamicin, an antibiotic drug, in human specimens. Measurements ob- tained by this device are used in the di- agnosis and treatment of gentamicin overdose and in monitoring levels of gentamicin to ensure appropriate ther- apy.

Classification. Class II.

##### § 862.3520   Kanamycin test system.

Identification. A kanamycin test system is a device intended to measure kanamycin, an antibiotic drug, in plas- ma and serum. Measurements obtained by this device are used in the diagnosis and treatment of kanamycin overdose and in monitoring levels of kanamycin to ensure appropriate therapy.

Classification. Class II.

##### § 862.3550   Lead test system.

Identification. A lead test system is a device intended to measure lead, a heavy metal, in blood and urine. Meas- urements obtained by this device are used in the diagnosis and treatment of lead poisoning.

Classification. Class II.

##### § 862.3555   Lidocaine test system.

Identification. A lidocaine test sys- tem is a device intended to measure lidocaine, an antiarrythmic and anticonvulsant drug, in serum and plasma. Measurements obtained by this device are used in the diagnosis and treatment of lidocaine overdose or in monitoring levels of lidocaine to en- sure appropriate therapy.

Classification. Class II.

##### § 862.3560   Lithium test system.

Identification. A lithium test sys- tem is a device intended to measure lithium (from the drug lithium car- bonate) in serum or plasma. Measure- ments of lithium are used to assure that the proper drug dosage is adminis- tered in the treatment of patients with mental disturbances, such as manic-de- pressive illness (bipolar disorder).

Classification. Class II.

#### § 862.3580

##### § 862.3580   Lysergic  acid  diethylamide (LSD) test system.

  - Identification. A lysergic acid diethylamide (LSD) test system is a device intended to measure lysergic acid diethylamide, a hallucinogenic drug, in serum, urine, and gastric con- tents. Measurements obtained by this device are used in the diagnosis and treatment of LSD use or overdose.

  - Classification. Class II.

##### § 862.3600   Mercury test system.

Identification. A mercury test sys- tem is a device intended to measure mercury, a heavy metal, in human specimens. Measurements obtained by this device are used in the diagnosis and treatment of mercury poisoning.

Classification. Class I.

##### § 862.3610   Methamphetamine  test  sys- tem.

Identification. A methamphet- amine test system is a device intended to measure methamphetamine, a cen- tral nervous system stimulating drug, in serum, plasma, and urine. Measure- ments obtained by this device are used in the diagnosis and treatment of methamphetamine use or overdose.

Classification. Class II.

##### § 862.3620   Methadone test system.

Identification. A methadone test system is a device intended to measure methadone, an addictive narcotic pain- relieving drug, in serum and urine. Measurements obtained by this device are used in the diagnosis and treat- ment of methadone use or overdose and to determine compliance with regula- tions in methadone maintenance treat- ment.

Classification. Class II.

##### § 862.3630   Methaqualone test system.

Identification. A methaqualone test system is a device intended to measure methaqualone, a hypnotic and sedative drug, in urine. Measurements obtained by this device are used in the diagnosis and treatment of methaqua- lone use or overdose.

Classification. Class II.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 862.3640   Morphine test system.

Identification. A morphine test system is a device intended to measure morphine, an addictive narcotic pain- relieving drug, and its analogs in serum, urine, and gastric contents. Measurements obtained by this device are used in the diagnosis and treat- ment of morphine use or overdose and in monitoring levels of morphine and its analogs to ensure appropriate ther- apy.

Classification. Class II.

##### § 862.3645   Neuroleptic	drugs radioreceptor assay test system.

Identification. A neuroleptic drugs radioceptor assay test system is a de- vice intended to measure in serum or plasma the dopamine receptor blocking activity of neuroleptic drugs and their active metabolites. A neuroleptic drug has anti-psychotic action affecting principally psychomotor activity, is generally without hypnotic effects, and is a tranquilizer. Measurements ob- tained by this device are used to aid in determining whether a patient is tak- ing the prescribed dosage level of such drugs.

Classification. Class II.

##### § 862.3650   Opiate test system.

Identification. An opiate test sys- tem is a device intended to measure any of the addictive narcotic pain-re- lieving opiate drugs in blood, serum, urine, gastric contents, and saliva. An opiate is any natural or synthetic drug that	has	morphine-like pharmocological actions. The opiates include drugs such as morphine, mor- phine glucoronide, heroin, codeine, nalorphine, and meperedine. Measure- ments obtained by this device are used in the diagnosis and treatment of opi- ate use or overdose and in monitoring the levels of opiate administration to ensure appropriate therapy.

Classification. Class II.

##### § 862.3660   Phenobarbital test system.

Identification. A phenobarbitol test system is a device intended to measure	phenobarbital,	an antiepileptic and sedative-hypnotic drug, in human specimens. Measure- ments obtained by this device are used

#### Food and Drug Administration, HHS	§ 862.3870

in the diagnosis and treatment of phe- nobarbital use or overdose and in moni- toring levels of phenobarbital to ensure appropriate therapy.

Classification. Class II.

##### § 862.3670   Phenothiazine test system.

Identification. A phenothiazine test system is a device intended to measure any of the drugs of the phenothiazine class in human speci- mens. Measurements obtained by this device are used in the diagnosis and treatment of phenothiazine use or overdose.

Classification. Class II.

##### § 862.3680   Primidone test system.

Identification. A primidone test system is a device intended to measure primidone, an antiepileptic drug, in human specimens. Measurements ob- tained by this device are used in the di- agnosis and treatment of primidone overdose and in monitoring levels of primidone to ensure appropriate ther- apy.

Classification. Class II.

##### § 862.3700   Propoxyphene test system.

Identification. A propoxyphene test system is a device intended to measure propoxyphene, a pain-reliev- ing drug, in serum, plasma, and urine. Measurements obtained by this device are used in the diagnosis and treat- ment of propoxyphene use or overdose or in monitoring levels of propoxyphene to ensure appropriate therapy.

Classification. Class II.

##### § 862.3750   Quinine test system.

Identification. A quinine test sys- tem is a device intended to measure quinine, a fever-reducing and pain-re- lieving drug intended in the treatment of malaria, in serum and urine. Meas- urements obtained by this device are used in the diagnosis and treatment of quinine overdose and malaria.

Classification. Class I.

[52 FR 16122, May 1, 1987, as amended at 53

FR  21450,  June  8,  1988;  65  FR  2310,  Jan.  14,

2000]

##### § 862.3830   Salicylate test system.

Identification. A salicylate test system is a device intended to measure salicylates, a class of analgesic, anti- pyretic and anti-inflammatory drugs that includes aspirin, in human speci- mens. Measurements obtained by this device are used in diagnosis and treat- ment of salicylate overdose and in monitoring salicylate levels to ensure appropriate therapy.

Classification. Class II.

##### § 862.3840   Sirolimus test system.

Identification. A sirolimus test system is a device intended to quan- titatively determine sirolimus con- centrations in whole blood. Measure- ments are used as an aid in manage- ment of transplant patients receiving therapy with sirolimus.

Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Sirolimus Test Systems.’’ See § 862.1(d) for the availability of this guidance document.

[69 FR 58259, Sept. 30, 2004]

##### § 862.3850   Sulfonamide test system.

Identification. A sulfonamide test system is a device intended to measure sulfonamides, any of the antibacterial drugs derived from sulfanilamide, in human specimens. Measurements ob- tained by this device are used in the di- agnosis and treatment of sulfonamide overdose and in monitoring sul- fonamide levels to ensure appropriate therapy.

Classification. Class I.

[52 FR 16122, May 1, 1987, as amended at 53

FR  21450,  June  8,  1988;  65  FR  2310,  Jan.  14,

2000]

##### § 862.3870   Cannabinoid test system.

Identification.  A  cannabinoid  test system is a device intended to measure any   of   the   cannabinoids,   hallucino- genic  compounds  endogenous  to  mari- huana,  in  serum,   plasma,  saliva,   and urine.  Cannabinoid  compounds  include delta-9-tetrahydrocannabinol, cannabidiol,	cannabinol,	and cannabichromene.    Measurements    ob- tained by this device are used in the di- agnosis  and  treatment  of  cannabinoid

#### § 862.3880

use or abuse and  in  monitoring  levels of cannabinoids during clinical inves- tigational  use.

Classification. Class II.

##### § 862.3880   Theophylline test system.

Identification. A theophylline test system is a device intended to measure theophylline (a drug used for stimula- tion of the muscles in the cardio- vascular, respiratory, and central nerv- ous systems) in serum and plasma. Measurements obtained by this device are used in the diagnosis and treat- ment of theophylline overdose or in monitoring levels of theophylline to ensure appropriate therapy.

Classification. Class II.

##### § 862.3900   Tobramycin test system.

Identification. A tobramycin test system is a device intended to measure tobramycin, an aminoglycoside anti- biotic drug, in plasma and serum. Measurements obtained by this device are used in the diagnosis and treat- ment of tobramycin overdose and in monitoring levels of tobramycin to en- sure appropriate therapy.

Classification. Class II.

##### § 862.3910   Tricyclic	antidepressant drugs test system.

Identification. A tricyclic antidepressant drugs test system is a device intended to measure any of the tricyclic antidepressant drugs in serum. The tricyclic antidepressant drugs	include		imipramine, desipramine,	amitriptyline, nortriptyline, protriptyline, and doxepin. Measurements obtained by this device are used in the diagnosis and treatment of chronic depression to ensure appropriate therapy.

Classification. Class II.

##### § 862.3950   Vancomycin test system.

Identification. A vancomycin test system is a device intended to measure vancomycin, an antibiotic drug, in serum. Measurements obtained by this device are used in the diagnosis and treatment of vancomycin overdose and in monitoring the level of vancomycin to ensure appropriate therapy.

Classification. Class II.

#### 21 CFR Ch. I (4–1–17 Edition)

PART 864—HEMATOLOGY AND PATHOLOGY DEVICES

##### Subpart A—General Provisions

Sec.

864.1    Scope.

864.3 Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Biological Stains

864.1850  Dye and chemical solution stains. 864.1860   Immunohistochemistry	reagents

and kits.

864.1870 Early growth response  1  (EGR1) gene fluorescence in-situ hybridization (FISH) test system for specimen charac- terization.

##### Subpart C—Cell and Tissue Culture Products

864.2220 Synthetic cell and tissue culture media and components.

864.2240 Cell and tissue culture supplies and equipment.

864.2260   Chromosome culture kit.

864.2280    Cultured animal and human cells. 864.2360    Mycoplasma   detection   media   and

components.

864.2800    Animal and human sera.

864.2875 Balanced salt solutions or formula- tions.

##### Subpart D—Pathology Instrumentation and Accessories

864.3010   Tissue processing equipment. 864.3250   Specimen    transport    and    storage

container.

864.3260 OTC test sample collection systems for  drugs  of  abuse  testing.

864.3300   Cytocentrifuge.

864.3400    Device for sealing microsections. 864.3600   Microscopes and accessories.

864.3800    Automated slide stainer. 864.3875   Automated tissue processor.

##### Subpart E—Specimen Preparation Reagents

864.4010   General purpose reagent. 864.4020    Analyte specific reagents. 864.4400   Enzyme preparations.

##### Subpart F—Automated and Semi- Automated Hematology Devices

864.5200    Automated cell counter.

864.5220  Automated	differential	cell counter.

864.5240 Automated blood cell diluting appa- ratus.

#### Food and Drug Administration, HHS	Pt. 864

864.5260     Automated cell-locating device. 864.5300   Red cell indices device.

864.5350    Microsedimentation centrifuge.

864.5400    Coagulation instrument.

864.5425    Multipurpose   system   for   in   vitro coagulation    studies.

864.5600     Automated hematocrit instrument. 864.5620    Automated hemoglobin system.

864.5680    Automated heparin analyzer. 864.5700    Automated      platelet      aggregation

system.

864.5800    Automated  sedimentation  rate  de- vice.

864.5850   Automated slide spinner. 864.5950   Blood volume measuring device.

##### Subpart G—Manual Hematology Devices

864.6100    Bleeding time device.

864.6150    Capillary blood collection tube. 864.6160    Manual blood cell counting device. 864.6400    Hematocrit measuring device.

864.6550   Occult blood test. 864.6600    Osmotic fragility test. 864.6650    Platelet adhesion test. 864.6675    Platelet aggregometer.

864.6700  Erythrocyte	sedimentation	rate test.

##### Subpart H—Hematology Kits and Packages

864.7040   Adenosine	triphosphate	release assay.

864.7060    Antithrombin III assay. 864.7100   Red blood cell enzyme assay.

864.7140  Activated whole blood clotting time tests.

864.7250    Erythropoietin assay. 864.7275    Euglobulin lysis time tests.

864.7280    Factor V Leiden DNA mutation de- tection  systems.

864.7290   Factor deficiency test.

864.7300    Fibrin    monomer    paracoagulation test.

864.7320   Fibrinogen/fibrin degradation prod- ucts   assay.

864.7340    Fibrinogen determination system. 864.7360   Erythrocytic	glucose-6-phosphate

dehydrogenase assay.

864.7375    Glutathione reductase assay. 864.7400   Hemoglobin A2 assay.

864.7415    Abnormal hemoglobin assay. 864.7425   Carboxyhemoglobin assay.

864.7440    Electrophoretic    hemoglobin    anal- ysis   system.

864.7455    Fetal hemoglobin assay.  864.7470    Glycosylated hemoglobin assay. 864.7490   Sulfhemoglobin assay.

864.7500   Whole blood hemoglobin assays. 864.7525  Heparin assay.

864.7660   Leukocyte    alkaline    phosphatase test.

864.7675    Leukocyte peroxidase test.

864.7695   Platelet	factor	4

radioimmunoassay.

864.7720    Prothrombin consumption test.

864.7735    Prothrombin-proconvertin  test  and thrombotest.

864.7750    Prothrombin time test. 864.7825    Sickle cell test.

864.7875   Thrombin time test.

864.7900    Thromboplastin generation test. 864.7925     Partial thromboplastin time tests.

##### Subpart I—Hematology Reagents

864.8100   Bothrops atrox reagent. 864.8150    Calibrator for cell indices.

864.8165    Calibrator  for  hemoglobin  or  hem- atocrit    measurement.

864.8175     Calibrator for platelet counting. 864.8185    Calibrator  for   red  cell  and  white

cell counting.

864.8200    Blood cell diluent.

864.8500    Lymphocyte separation medium. 864.8540    Red cell lysing reagent.

864.8625    Hematology   quality   control   mix- ture.

864.8950    Russell viper venom reagent.

##### Subpart J—Products Used In Establishments That  Manufacture  Blood  and  Blood Products

864.9050    Blood bank supplies.

864.9100 Empty container for the  collection and processing of blood and blood compo- nents.

864.9125    Vacuum-assisted    blood    collection system.

864.9145   Processing system for frozen blood. 864.9160  Blood	group	substances	of nonhuman   origin   for   in   vitro   diagnostic

use.

864.9175   Automated    blood    grouping    and antibody   test   system.

864.9185   Blood grouping view box.

864.9195   Blood   mixing   devices   and   blood weighing devices.

864.9205   Blood and plasma warming device. 864.9225   Cell-freezing     apparatus     and     re-

agents for in vitro diagnostic use.

864.9245    Automated blood cell separator. 864.9275    Blood  bank  centrifuge  for  in  vitro

diagnostic use.

864.9285    Automated  cell-washing  centrifuge for    immuno-hematology.

864.9300    Automated Coombs test systems. 864.9320    Copper  sulfate  solution  for  specific

gravity determinations.

864.9400    Stabilized enzyme solution. 864.9550    Lectins and protectins.

864.9575    Environmental chamber for storage of   platelet   concentrate.

864.9600   Potentiating media for in vitro di- agnostic  use.

864.9650   Quality control kit for blood bank- ing reagents.

864.9700   Blood     storage     refrigerator     and blood  storage  freezer.

864.9750   Heat-sealing device.

#### § 864.1

864.9875   Transfer set.

##### Subpart K—Products Used In Establish- ments That Manufacture Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/Ps)

864.9900   Cord  blood  processing  system  and storage    container.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

EDITORIAL NOTE: Nomenclature changes  to part 864 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 864.1   Scope.

    - This part sets forth the classifica- tion of hematology and pathology de- vices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 17732, May 11, 1987, as amended at 69

FR 12273, Mar. 16, 2004; 78 FR 18233, Mar. 26,

2013; 79 FR 50552, Aug. 25, 2014]

##### § 864.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of

21 CFR Ch. I (4–1–17 Edition)

FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows that a device being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device  into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of

#### Food and Drug Administration, HHS	§ 864.1850

the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17732, May 11, 1987]

##### § 864.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2310, Jan. 14, 2000]

## Subpart B—Biological Stains

##### § 864.1850   Dye  and  chemical  solution stains.

      - Identification. Dye and chemical solution stains for medical purposes are mixtures of synthetic or natural dyes or nondye chemicals in solutions used in staining cells and tissues for di- agnostic	histopathology, cytopathology, or hematology.

      - Classification. Class I (general con- trols). These devices are exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9. These devices are also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with re- spect to general requirements con- cerning records, and § 820.198, with re- spect to complaint files.

[45 FR 60583, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

§ 864.1860

##### § 864.1860   Immunohistochemistry	re- agents and kits.

        - Identification.

Immunohistochemistry test systems (IHC’s) are in vitro diagnostic devices consisting of polyclonal or monoclonal antibodies labeled with directions for use and performance claims, which may be packaged with ancillary re- agents in kits. Their intended use is to identify, by immunological techniques, antigens in tissues or cytologic speci- mens. Similar devices intended for use with flow cytometry devices are not considered IHC’s.

        - Classification		of immunohistochemistry devices. (1) Class I (general controls). Except as described in paragraphs (b)(2) and (b)(3) of this section, these devices are exempt from the premarket notification require- ments in part 807, subpart E of this chapter. This  exemption  applies  to IHC’s that provide the pathologist with adjunctive diagnostic information that may be incorporated into the patholo- gist’s report, but that is not ordinarily reported to the clinician as an inde- pendent finding. These IHC’s are used after the primary diagnosis of tumor (neoplasm) has been made by conven- tional	histopathology	using nonimmunologic histochemical stains, such as hematoxylin and eosin. Exam- ples of class I IHC’s are differentiation markers that are used  as  adjunctive tests to subclassify tumors, such as keratin.

Class II (special control, guidance document: ‘‘FDA Guidance for Submis- sion of Immunohistochemistry Appli- cations to the FDA,’’ Center for De- vices and Radiologic Health, 1998). These IHC’s are intended for the detec- tion and/or measurement of certain target analytes in order to provide prognostic or predictive data that are not directly confirmed by routine histopathologic internal and external control specimens. These IHC’s provide the pathologist  with information  that is ordinarily reported as independent diagnostic information to the ordering clinician, and the claims associated with these data are widely accepted and supported by valid scientific evi- dence. Examples of class II IHC’s are those   intended   for   semiquantitative

#### 21 CFR Ch. I (4–1–17 Edition)

measurement   of   an   analyte,   such   as hormone receptors in breast cancer.

Class III (premarket approval). IHC’s intended for any use not de- scribed in paragraphs (b)(1) or (b)(2) of this  section.

        - Date of PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the Federal Food,  Drug,  and   Cosmetic  Act is required for any device described in paragraph (b)(3) of this section before this device may be commercially dis- tributed. See § 864.3.

[63 FR 30142, June 3, 1998]

##### § 864.1870 Early growth response 1 (EGR1) gene fluorescence in-situ hybridization (FISH) test system for specimen characterization.

Identification. An early growth re- sponse 1 (EGR1) gene fluorescence in- situ hybridization (FISH) test system for specimen characterization is a de- vice intended to detect the EGR1 probe target on chromosome 5q in bone mar- row specimens from  patients  with acute myeloid leukemia (AML) or myelodysplastic syndrome (MDS). The assay results are intended to be inter- preted only by a  qualified  pathologist or cytogeneticist. These devices do not include automated systems that di- rectly report results  without  review and interpretation by a qualified pa- thologist or cytogeneticist. These de- vices also do not include any device in- tended for use to select patient ther- apy, predict patient response to ther- apy, or to screen for disease as well as any device with a claim for  a  par- ticular diagnosis, prognosis, moni- toring, or risk assessment.

Classification. Class II (special controls). The special controls for this device are:

  - Premarket notification submis- sions must also include the following information:

    - A detailed description of all probes included in the kit;

    - Purpose of each probe;

    - Probe molecular specificity;

    - Probe specificity;

    - Probe limits;

    - Probe sensitivity;

    - Specification of required ancil- lary reagents, instrumentation, and equipment;

#### Food and Drug Administration, HHS	§ 864.2240

    - Specification of the specimen collection,  processing,  storage   and slide preparation methods;

    - Specification of the assay proce- dure;

    - Specification of control elements that are incorporated into the rec- ommended testing procedures;

    - Specification of risk mitigation elements: Description of all additional procedures, methods, and practices in- corporated into the directions for use that mitigate risks associated with testing;

    - Specification of the criteria for test result interpretation and report- ing;

    - Device analytical sensitivity data;

    - Device analytical specificity data;

    - Device reference limit data;

    - Device precision/reproducibility data;

    - Device stability data to in- clude:

Real-time stability,

Freeze-thaw stability,

Transport and temperature sta- bility,

Post-hybridization signal sta- bility,

Photostability of probe, and

    - Documentation that dem- onstrates the clinical validity of the device. The documentation must in- clude data from clinical studies, a min- imum of two peer-reviewed published literature references using the specific device seeking marketing clearance, or both. Documentation for the clinical studies and peer-reviewed published lit- erature references cited must include the following elements:

Documentation that the spon- sor’s probe was used in the literature reference,

Number and type of specimens,

Target population studied,

Upper reference limit, and

Range of positive probe results.

  - Your § 809.10(b)(12) of this chapter compliant labeling must include a statement summarizing the data iden- tified in paragraphs (b)(1)(xiii) through

(xviii) of this section and a description of the studies supporting the informa- tion, including the pre-specified ac- ceptance criteria for these performance

studies, justification for the pre-speci- fied acceptance criteria, and whether the pre-specified acceptance criteria were met.

  - Your § 809.10 of this chapter com- pliant  labeling  must  include:

    - A warning that reads ‘‘The assay results are intended to be interpreted only by a qualified pathologist or cytogeneticist.’’

    - A warning that reads ‘‘This de- vice is not for high-risk uses such as selecting therapy, predicting thera- peutic response or disease screening.’’

    - A warning that reads ‘‘The use of this device for diagnosis, monitoring or risk assessment has not been estab- lished.’’

[79 FR 52196, Sept. 3, 2014]

## Subpart C—Cell And Tissue Culture Products

##### § 864.2220   Synthetic   cell   and   tissue culture media and components.

Identification. Synthetic cell and tissue culture media and components are substances that are composed en- tirely of defined components (e.g., amino acids, vitamins, inorganic salts) that are essential for the survival and development of cell lines of humans and other animals. This does not in- clude tissue culture media for human ex vivo tissue and cell culture proc- essing   applications   as   described   in

§ 876.5885 of this chapter.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60583, Sept. 12, 1980, as amended at 54

FR 25044,  June 12,  1989;  66  FR 27024,  May 16,

2001; 66 FR 38789, July 25, 2001]

##### § 864.2240   Cell and tissue culture sup- plies and equipment.

Identification. Cell and tissue cul- ture supplies and equipment are de- vices that are used to examine, propa- gate, nourish, or grow cells and tissue cultures. These include such articles as slide culture chambers, perfusion and roller apparatus, cell culture suspen- sion systems, and tissue culture flasks, disks, tubes, and roller bottles.

§ 864.2260

Classification. Class I (general con- trols). These devices are exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.  If the devices are not labeled or otherwise represented as sterile, they are exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 60584, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.2260   Chromosome culture kit.

Identification. A chromosome cul- ture kit is a device containing the nec- essary ingredients (e.g., Minimum Es- sential Media (MEM) of McCoy’s 5A culture media, phytohemagglutinin, fetal calf  serum, antibiotics,  and hep- arin) used to culture tissues for diag- nosis of congenital chromosome abnor- malities.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60585, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.2280   Cultured animal and human cells.

Identification. Cultured animal and human cells are in vitro cultivated cell lines from the tissue of humans or other animals which are used in var- ious diagnostic procedures, particu- larly diagnostic virology and cyto- genetic studies.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60585, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.2360   Mycoplasma	detection media and components.

Identification. Mycoplasma detec- tion media and components are used to

#### 21 CFR Ch. I (4–1–17 Edition)

detect and isolate mycoplasma pleuropneumonia-like	organisms (PPLO), a common microbial contami- nant in cell cultures.

Classification. Class I (general con- trols). These devices are exempt from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60586, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.2800   Animal and human sera.

Identification. Animal and human sera are biological products, obtained from the blood of humans or other ani- mals, that provide the necessary growth-promoting nutrients in a cell culture  system.

Classification. Class I (general con- trols). These devices are exempt from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60586, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.2875   Balanced  salt  solutions  or formulations.

Identification. A balanced salt so- lution or formulation is a defined mix- ture of salts and glucose in a simple medium. This device is included as a necessary component of most cell cul- ture systems. This media component controls for pH, osmotic pressure, en- ergy source, and inorganic ions.

Classification. Class I (general con- trols). These devices are exempt from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60586, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

## Subpart D—Pathology Instrumentation and Accessories

##### § 864.3010   Tissue	processing	equip- ment.

Identification. Tissue processing equipment consists of devices used to prepare human tissue specimens for di- agnostic  histological  examination  by

Food and Drug Administration, HHS	§ 864.3300

processing specimens through the var- ious stages of decalcifying, infiltrating, sectioning, and mounting on micro- scope slides.

Classification. Class I (general con- trols). These devices are exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9. The devices are also exempt from the cur- rent good manufacturing practice re- quirements of the quality system regu- lation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint  files.

[45 FR 60587, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.3250   Specimen	transport	and storage container.

Identification. A specimen trans- port and storage container, which may be empty or prefilled, is a device in- tended to contain biological specimens, body waste, or body exudate during storage and transport in order that the matter contained therein can be de- stroyed or used effectively for diag- nostic examination. If prefilled, the de- vice contains  a  fixative  solution or other general purpose reagent to pre- serve the condition of a biological spec- imen added to the container. This sec- tion does not apply to specimen trans- port and storage containers that are intended for use as part of an over-the- counter test sample collection system for drugs of abuse testing.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9. If the device is not la- beled or otherwise represented as ster- ile, it is exempt from the current good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,  with  respect  to  complaint files.

[54 FR 47206, Nov. 13, 1989, as amended at 65

FR  2310,  Jan.  14,  2000;  65  FR  18234,  Apr.  7,

2000]

##### § 864.3260   OTC  test  sample  collection systems for drugs of abuse testing.

Identification. An over-the-counter (OTC) test  sample  collection  system for drugs of abuse testing is a device intended to: Collect biological speci- mens (such as hair, urine, sweat, or sa- liva), outside of a medical setting and not on order of a health care profes- sional (e.g., in the home, insurance, sports, or workplace setting); maintain the integrity of  such  specimens  during storage and transport in order that the matter contained therein can be tested in a laboratory for the  presence  of drugs of  abuse  or  their  metabolites; and provide access to test results and counseling. This section does not apply to collection, transport, or laboratory testing of biological specimens for the presence of drugs of abuse or their me- tabolites that is performed to develop evidence for law enforcement purposes.

Classification. Class I (general con- trols). The device is exempt from the premarket  notification   requirements in part 807, subpart E of this chapter subject to the limitations in § 864.9 if it is sold, distributed, and used in accord- ance  with  the  restrictions  set  forth  in

§ 809.40 of this chapter. If the device is not labeled or otherwise represented as sterile, it is exempt  from  the  current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.198 of this chapter with respect  to  complaint  files.

[65 FR 18234, Apr. 7, 2000]

##### § 864.3300   Cytocentrifuge.

Identification. A cytocentrifuge is a centrifuge used to concentrate cells from biological cell suspensions (e.g., cerebrospinal fluid) and to deposit these cells on a glass microscope slide for cytological examination.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60588, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

#### § 864.3400

##### § 864.3400   Device for sealing microsec- tions.

  - Identification. A device for sealing microsections is an automated instru- ment used to seal stained cells and microsections for histological and cytological  examination.

  - Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60589, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.3600   Microscopes	and	acces- sories.

Identification. Microscopes and ac- cessories are optical instruments used to enlarge images of specimens, prep- arations, and cultures for medical pur- poses. Variations of microscopes and accessories (through a change in the light source) used for medical purposes include the following:

Phase contrast microscopes, which permit visualization of unstained preparations by altering the phase relationship of light that passes around the object and through the ob- ject.

Fluorescense microscopes, which permit examination of specimens stained with fluorochromes that fluo- resce under ultraviolet light.

Inverted stage microscopes, which permit examination of tissue  cultures or other biological specimens  con- tained in bottles  or  tubes  with  the light source mounted above the speci- men.

Classification. Class I (general con- trols). These devices are exempt from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.  If the device is not labeled or otherwise represented  as  sterile,  it  is   exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with   respect   to   general   requirements

#### 21 CFR Ch. I (4–1–17 Edition)

concerning  records,  and  § 820.198,  with respect  to  complaint  files.

[45 FR 60590, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.3800   Automated slide stainer.

Identification. An automated slide stainer is a device used to stain his- tology, cytology,  and  hematology slides for diagnosis.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60591, Sept. 12, 1980, as amended at 54

FR 25044, June 12, 1989; 66 FR 38789, July 25,

2001]

##### § 864.3875   Automated tissue processor.

Identification. An automated tis- sue processor is an automated system used to process tissue specimens for ex- amination through fixation, dehydra- tion, and infiltration.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60591, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38789, July 25,

2001]

## Subpart E—Specimen Preparation Reagents

##### § 864.4010   General purpose reagent.

A general purpose reagent is a chemical reagent that has general lab- oratory application, that is used to col- lect, prepare, and examine specimens from the human body for diagnostic purposes, and that is not labeled or otherwise intended for a specific diag- nostic application. It may be either an individual substance, or multiple sub- stances reformulated, which, when combined with or used in conjunction with an appropriate analyte specific re- agent (ASR) and other general purpose reagents, is part of a diagnostic test procedure or system constituting a fin- ished in vitro diagnostic (IVD) test. General purpose reagents are appro- priate for combining with one or more

Food and Drug Administration, HHS	§ 864.4020

than one ASR in producing such sys- tems and include labware or disposable constituents of tests; but they do not include laboratory machinery, auto- mated or powered systems. General purpose reagents include cytological preservatives, decalcifying reagents, fixative and adhesives, tissue proc- essing reagents, isotonic solutions and pH buffers. Reagents used in tests for more than one individual chemical sub- stance or ligand are general purpose re- agents (e.g., Thermus aquaticus (TAQ) polymerase, substrates for enzyme immunoassay  (EIA)).

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 60592, Sept. 12, 1980, as amended at 54

FR 25045,  June 12,  1989;  62  FR 62260,  Nov.  21,

1997; 66 FR 38789, July 25, 2001]

##### § 864.4020   Analyte specific reagents.

Identification. Analyte specific re- agents (ASR’s) are antibodies, both polyclonal and monoclonal, specific re- ceptor proteins, ligands, nucleic acid sequences, and similar reagents which, through specific binding or chemical reaction with substances in a speci- men, are intended for use in a diag- nostic application for identification and quantification of an individual chemical substance or ligand in bio- logical specimens. ASR’s that other- wise fall within this definition are not within the scope of subpart E of this part when they are sold to:

  - In vitro diagnostic manufacturers; or

  - Organizations that use the re- agents to make tests for purposes other than providing diagnostic information to patients and practitioners, e.g., fo- rensic, academic, research, and other nonclinical laboratories.

Classification. (1) Class I (general controls).  Except  as  described  in  para-

graphs (b)(2) and (b)(3) of this section, these devices are exempt from the pre- market notification  requirements  in part 807, subpart E of this chapter.

Class II (special controls/guidance documents), when the analyte is used in blood banking tests that have been classified as class II devices (e.g., cer- tain cytomegalovirus serological and treponema pallidum nontreponemal test reagents). Guidance Documents:

‘‘Specifications for Immunological Test- ing for Infectious Disease; Approved Guide- line,’’ NCCLS Document I/LA18–A, December 1994.

‘‘Assessment of the Clinical Accuracy of Laboratory Tests Using Receiver Operating Characteristic (ROC) Plots; Tentative Guide- line,’’ NCCLS Document KGP10–T, December 1993.

‘‘Review Criteria for Assessment of In Vitro Diagnostic Devices for Direct Detec- tion of Mycobacterium spp,’’ FDA, July 6, 1993, and its ‘‘Attachment 1,’’ February 28, 1994.

‘‘Draft Review Criteria for Nucleic Acid Amplification-Based In Vitro Diagnostic De- vices for Direct Detection of Infectious Microorganisms,’’ FDA, July 6, 1993.

The Center for Biologics Evaluation and Research, FDA, ‘‘Points to Consider in the Manufacture and Clinical Evaluation of  In Vitro Tests  to  Detect  Antibodies  to  the Human Immunodeficiency Virus, Type I’’  (54 FR 48943, November 28, 1989).

Class III (premarket approval), when:

The analyte is intended as a com- ponent in a test intended for use in the diagnosis of a contagious  condition that is highly likely to result in a fatal outcome and prompt, accurate diag- nosis offers the opportunity to miti- gate the public health impact of the condition (e.g., human immuno- deficiency virus (HIV/AIDS)or tuber- culosis (TB)); or

The analyte is intended as a com- ponent in a test intended for use in donor screening for conditions  for which FDA has recommended or re- quired testing in order to safeguard the blood supply or establish  the  safe  use of blood and blood products (e.g., tests for hepatitis or tests for identifying blood groups).

Date of 510(k), or date of PMA or notice of completion of a product develop- ment protocol is required. (1) Preamendments    ASR’s;    No    effective

#### § 864.4400

date has been established for the re- quirement for  premarket  approval  for the device described in paragraph (b)(3) of this section. See § 864.3.

(2) For postamendments ASR’s; No- vember 23, 1998.

Restrictions. Restrictions on  the sale, distribution and use of  ASR’s  are set forth in § 809.30 of this chapter.

[62 FR 62260, Nov. 21, 1997]

##### § 864.4400   Enzyme preparations.

Identification. Enzyme prepara- tions are products that are used in the histopathology laboratory for the fol- lowing purposes:

To disaggregate tissues and cells already in established cultures for preparation into subsequent cultures (e.g., trypsin);

To disaggregate fluid specimens for cytological examination (e.g., papain for gastric lavage or trypsin for sputum liquefaction);

To aid in the selective staining of tissue specimens (e.g., diastase for gly- cogen  determination).

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60592, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38789, July 25,

2001]

## Subpart F—Automated and Semi- Automated Hematology Devices

##### § 864.5200   Automated cell counter.

Identification. An automated cell counter is a fully-automated or semi- automated device used to count red blood cells, white blood cells, or blood platelets using a sample of the pa- tient’s peripheral blood (blood circu- lating in one of the body’s extremities, such as the arm). These devices may also measure  hemoglobin  or hemato- crit and may also calculate or measure one or more of the red cell indices (the erythrocyte mean corpuscular volume, the mean corpuscular hemoglobin, or the mean corpuscular hemoglobin con- centration). These devices may use ei- ther an electronic particle counting method or an optical counting method.

21 CFR Ch. I (4–1–17 Edition)

Classification. Class II (perform- ance standards).

[45 FR 60593, Sept. 12, 1980]

##### § 864.5220   Automated  differential  cell counter.

Identification. An automated dif- ferential cell counter is a device used to identify one or more of the formed elements of the blood. The device may also have the capability to flag, count, or classify immature or abnormal hematopoietic cells of the blood, bone marrow, or other body fluids. These de- vices may combine an electronic par- ticle counting method, optical method, or a flow cytometric method utilizing monoclonal CD (cluster designation) markers. The device includes accessory CD markers.

Classification. Class II (special controls). The special control for this device is the FDA document entitled ‘‘Class II Special Controls Guidance Document: Premarket Notifications for Automated Differential Cell Counters for Immature or Abnormal Blood Cells; Final Guidance for Industry and FDA.’’

[67 FR 1607, Jan. 14, 2002]

##### § 864.5240   Automated  blood  cell  dilut- ing apparatus.

Identification. An automated blood cell diluting apparatus is a fully auto- mated or semi-automated device used to make appropriate dilutions of a blood sample for further testing.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60596, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.5260   Automated  cell-locating  de- vice.

Identification. An automated cell- locating device is a device used to lo- cate blood cells on a peripheral blood smear, allowing the operator to iden- tify and classify each cell according to type. (Peripheral blood is blood circu- lating in one of the body’s extremities, such as the arm.)

Classification. Class II (perform- ance standards).

[45 FR 60597, Sept. 12, 1980]

#### Food and Drug Administration, HHS	§ 864.5700

##### § 864.5300   Red cell indices device.

  - Identification. A red cell indices device, usually part of a larger system, calculates or directly measures the erythrocyte mean corpuscular volume (MCV), the mean corpuscular hemo- globin (MCH), and the mean corpus- cular	hemoglobin	concentration (MCHC). The red cell indices are used for the differential diagnosis of anemias.

  - Classification. Class II (perform- ance standards).

[45 FR 60597, Sept. 12, 1980]

##### § 864.5350   Microsedimentation	cen- trifuge.

Identification. A microsedimenta- tion centrifuge is a device used to sedi- ment red cells for the microsedimenta- tion rate test.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60598, Sept. 12, 1980, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38789,  July 25,

2001]

##### § 864.5400   Coagulation instrument.

Identification. A coagulation in- strument is an automated or semiauto- mated device used to determine the onset of clot formation for in vitro co- agulation studies.

Classification. Class II (perform- ance standards).

[45 FR 60598, Sept. 12, 1980]

##### § 864.5425   Multipurpose  system  for  in vitro coagulation studies.

Identification. A multipurpose sys- tem for in vitro coagulation studies is a device consisting of one automated or semiautomated instrument and its as- sociated reagents and controls. The system is used to perform a series of coagulation studies and coagulation factor assays.

Classification. Class II (perform- ance standards).

[45 FR 60599, Sept. 12, 1980]

##### § 864.5600   Automated   hematocrit   in- strument.

Identification. An automated hem- atocrit instrument  is  a  fully  auto- mated or semi-automated device which may or may not be part of a larger sys- tem. This device measures the packed red cell volume of a blood sample to distinguish normal from  abnormal states, such as anemia  and erythrocytosis (an increase in the num- ber of red cells).

Classification. Class II (perform- ance standards).

[45 FR 60600, Sept. 12, 1980]

##### § 864.5620   Automated  hemoglobin  sys- tem.

Identification. An automated he- moglobin system is a  fully  automated or semi-automated device which  may or may not be part of a larger system. The generic type of device consists  of the reagents, calibrators, controls, and instrumentation used to determine the hemoglobin content of human blood.

Classification. Class II (perform- ance standards).

[45 FR 60601, Sept. 12, 1980]

##### § 864.5680   Automated	heparin	ana- lyzer.

Identification. An automated hep- arin analyzer is a device used to deter- mine the heparin level in a blood sam- ple by mixing the sample with prot- amine (a heparin-neutralizing sub- stance)	and	determining photometrically the onset of air-acti- vated clotting. The analyzer also deter- mines the amount of protamine nec- essary to neutralize the heparin in the patient’s  circulation.

Classification. Class II (special controls).

[45 FR 60601, Sept. 12, 1980, as amended at 52

FR  17733,  May  11,  1987;  58  FR  51571,  Oct.  4,

1993]

##### § 864.5700   Automated platelet aggrega- tion system.

Identification. An automated platelet aggregation system is a device used to determine changes in platelet shape and platelet aggregation fol- lowing the addition of an aggregating reagent to a platelet-rich plasma.

#### § 864.5800

Classification. Class II (perform- ance standards).

[45 FR 60602, Sept. 12, 1980]

##### § 864.5800   Automated	sedimentation rate device.

Identification. An automated sedi- mentation rate device is an instrument that measures automatically the erythrocyte sedimentation rate in whole blood. Because an increased sedi- mentation rate indicates tissue dam- age or inflammation, the erythrocyte sedimentation rate device is useful in monitoring treatment of a disease.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60602, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.5850   Automated slide spinner.

Identification. An automated slide spinner is a device that prepares auto- matically a blood film on a microscope slide using a small amount of periph- eral blood (blood circulating in one of the body’s extremities, such as the arm).

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60603, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.5950   Blood volume measuring de- vice.

Identification. A blood volume measuring device is a manual, semi- automated, or automated system that is used to calculate the red cell mass, plasma volume, and total blood vol- ume.

Classification. Class II (perform- ance standards).

[45 FR 60603, Sept. 12, 1980]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart G—Manual Hematology Devices

##### § 864.6100   Bleeding time device.

Identification. A bleeding time de- vice is a device, usually employing two spring-loaded blades, that produces two small incisions in the patient’s skin. The length of time required for the bleeding to stop is a measure of the ef- fectiveness of the coagulation system, primarily the platelets.

Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60604, Sept. 12, 1980, as amended at 63

FR 59225, Nov. 3, 1998]

##### § 864.6150   Capillary   blood   collection tube.

Identification. A capillary blood collection tube is a plain or heparinized glass tube of very small di- ameter used to collect blood by cap- illary  action.

Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60604, Sept. 12, 1980, as amended at 54

FR 25045,  June 12,  1989;  65  FR 2310,  Jan. 14,

2000]

##### § 864.6160   Manual  blood  cell  counting device.

Identification. A manual blood cell counting device is a device used to count red blood cells, white blood cells, or blood platelets.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60605, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.6400   Hematocrit   measuring   de- vice.

Identification. A hematocrit meas- uring device is a system consisting of instruments, tubes, racks, and a sealer and a holder. The device is used to measure the packed red cell volume in

#### Food and Drug Administration, HHS	§ 864.7060

blood to determine whether the pa- tient’s total red cell volume is normal or abnormal. Abnormal states include anemia (an abnormally low total red cell volume) and erythrocytosis (an ab- normally high total red cell mass). The packed red cell volume is produced by centrifuging a given volume of blood.

Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60606, Sept. 12, 1980, as amended at 63

FR 59225, Nov. 3, 1998]

##### § 864.6550   Occult blood test.

Identification. An occult blood test is a device used to detect occult blood in urine or feces. (Occult blood is blood present in such small quantities that it can be detected only by chemical tests of suspected material, or by micro- scopic or spectroscopic examination.)

Classification. Class II (perform- ance standards).

[45 FR 60606, Sept. 12, 1980]

##### § 864.6600   Osmotic fragility test.

Identification. An osmotic fragility test is a device used to determine the resistance of red blood cells to hemol- ysis (destruction) in varying con- centrations of hypotonic saline solu- tions.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60607, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.6650   Platelet adhesion test.

Identification. A platelet adhesion test is a device used to determine in vitro platelet function.

Classification. Class II (perform- ance standards).

[45 FR 60608, Sept. 12, 1980]

##### § 864.6675   Platelet aggregometer.

Identification. A platelet aggregometer is a device, used to de- termine changes in platelet shape and platelet aggregation following the ad-

dition  of  an  aggregating  reagent  to  a platelet rich plasma.

Classification. Class II (perform- ance standards).

[45 FR 60608, Sept. 12, 1980]

##### § 864.6700   Erythrocyte	sedimentation rate test.

Identification. An erythrocyte sedimentation rate test is a device that measures the length of time required for the red cells in a blood sample to fall a specified distance or a device that measures the degree of sedimenta- tion taking place in a given length of time. An increased rate indicates tis- sue damage or inflammation.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60608, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

## Subpart H—Hematology Kits and Packages

##### § 864.7040   Adenosine  triphosphate  re- lease assay.

Identification. An adenosine triphosphate release assay is a device that measures the release of adenosine triphosphate (ATP) from platelets fol- lowing aggregation. This measurement is made on platelet-rich plasma using a photometer and a luminescent firefly extract. Simultaneous  measurements of platelet aggregation and ATP re- lease are used to evaluate platelet function disorders.

Classification. Class I (general controls).

[45 FR 60609, Sept. 12, 1980]

##### § 864.7060   Antithrombin III assay.

Identification. An antithrombin III assay is a device that is used to deter- mine the plasma level of antithrombin III (a substance which acts with the anticoagulant heparin to prevent co- agulation). This determination is used to monitor the administration of hep- arin in the treatment of thrombosis. The determination may also be used in the diagnosis of thrombophilia (a con- genital deficiency of antithrombin III).

§ 864.7100

Classification. Class II (perform- ance standards).

[45 FR 60609, Sept. 12, 1980]

##### § 864.7100   Red	blood	cell	enzyme assay.

Identification. Red blood cell en- zyme assay is a device used to measure the activity in red blood cells of clini- cally important  enzymatic  reactions and their products, such as pyruvate kinase or 2,3-diphosphoglycerate. A red blood cell enzyme assay is used to de- termine the enzyme defects responsible for a patient’s hereditary hemolytic anemia.

Classification. Class II (perform- ance standards).

[45 FR 60610, Sept. 12, 1980]

##### § 864.7140   Activated  whole  blood  clot- ting time tests.

Identification. An activated whole blood clotting time tests is a device, used to monitor heparin therapy for the treatment of venous thrombosis or pulmonary embolism by measuring the coagulation time of whole blood.

Classification. Class II (perform- ance standards).

[45 FR 60611, Sept. 12, 1980]

##### § 864.7250   Erythropoietin assay.

Identification. A erythropoietin assay is a device that measures the concentration of erythropoietin (an en- zyme that regulates the production of red blood cells) in serum or urine. This assay provides diagnostic information for the evaluation of erythrocytosis (increased total red cell mass) and ane- mia.

Classification. Class II. The special control for this device is FDA’s ‘‘Docu- ment  for  Special  Controls  for  Erythro- poietin   Assay   Premarket   Notification (510(k)s).’’

[45 FR 60612, Sept. 12, 1980, as amended at 52

FR 17733,  May 11,  1987;  65  FR 17144,  Mar. 31,

2000]

##### § 864.7275   Euglobulin lysis time tests.

Identification. A euglobulin lysis time test is a device that measures the length of time required for the lysis (dissolution) of a clot formed from fibrinogen  in  the  euglobulin  fraction

#### 21 CFR Ch. I (4–1–17 Edition)

(that fraction of the plasma responsible for the formation of plasmin, a clot lysing enzyme). This test evaluates natural fibrinolysis (destruction of a blood clot after bleeding has been ar- rested). The test also will detect accel- erated fibrinolysis.

Classification. Class II (perform- ance standards).

[45 FR 60612, Sept. 12, 1980]

##### § 864.7280   Factor V Leiden DNA muta- tion detection systems.

Identification. Factor V Leiden deoxyribonucleic acid (DNA) mutation detection systems are devices that con- sist of different reagents and instru- ments which include polymerase chain reaction (PCR) primers, hybridization matrices,  thermal  cyclers,  imagers, and software packages. The  detection of the Factor V  Leiden  mutation  aids in the diagnosis of patients with sus- pected  thrombophilia.

Classification. Class II (special controls). The special control is FDA’s guidance entitled ‘‘Class II Special Controls Guidance Document: Factor V Leiden DNA Mutation Detection Sys- tems.’’ (See § 864.1(d) for the avail- ability of this guidance document.)

[69 FR 12273, Mar. 16, 2004]

##### § 864.7290   Factor deficiency test.

Identification. A factor deficiency test is a device used to diagnose spe- cific coagulation defects, to monitor certain types of therapy, to detect co- agulation inhibitors, and to detect a carrier state (a person carrying both a recessive gene for a coagulation factor deficiency such as hemophilia and the corresponding normal gene).

Classification. Class II (perform- ance standards).

[45 FR 60613, Sept. 12, 1980]

##### § 864.7300   Fibrin	monomer paracoagulation test.

Identification. A fibrin monomer paracoagulation test is a device used to detect fibrin monomer in the diagnosis of disseminated intravascular coagula- tion (nonlocalized clotting within a blood vessel) or in the differential diag- nosis	between	disseminated intravascular coagulation and primary

#### Food and Drug Administration, HHS	§ 864.7425

fibrinolysis (dissolution of the fibrin in a blood clot).

Classification. Class II. The special control for this device is FDA’s ‘‘In Vitro Diagnostic Fibrin  Monomer Paracoagulation  Test.’’

[45 FR 60614, Sept. 12, 1980, as amended at 52

FR 17733,  May 11,  1987;  65  FR 17144,  Mar. 31,

2000]

##### § 864.7320   Fibrinogen/fibrin	degrada- tion products assay.

Identification. A fibrinogen/fibrin degradation products assay is a device used to detect and measure fibrinogen degradation products and fibrin deg- radation products (protein fragments produced by the enzymatic action of plasmin on fibrinogen and fibrin) as an aid in detecting the presence and de- gree of intravascular coagulation and fibrinolysis (the dissolution  of  the fibrin in a blood clot) and in moni- toring therapy for disseminated intravascular coagulation  (nonlocal- ized clotting in the blood vessels).

Classification. Class II (perform- ance standards).

[45 FR 60615, Sept. 12, 1980]

##### § 864.7340   Fibrinogen	determination system.

Identification. A fibrinogen deter- mination system is a device that con- sists of the instruments, reagents, standards, and controls used to deter- mine the fibrinogen levels in dissemi- nated intravascular coagulation (non- localized clotting within the blood ves- sels) and primary fibrinolysis (the dis- solution of fibrin in a blood clot).

Classification. Class II (perform- ance standards).

[45 FR 60615, Sept. 12, 1980]

##### § 864.7360   Erythrocytic glucose-6-phos- phate dehydrogenase assay.

Identification. An erythrocytic glucose-6-phosphate	dehydrogenase assay is a device used to measure the activity of the enzyme glucose-6-phos- phate dehydrogenase or of glucose-6- phosphate dehydrogenase isoenzymes. The results of this assay are  used  in the diagnosis and treatment of nonspherocytic congenital hemolytic anemia or drug-induced hemolytic ane- mia  associated  with  a  glucose-6-phos-

phate dehydrogenase deficiency. This generic device includes assays based on fluorescence,	electrophoresis, methemoglobin reduction, catalase in- hibition,  and  ultraviolet  kinetics.

Classification. Class II (perform- ance standards).

[45 FR 60616, Sept. 12, 1980]

##### § 864.7375   Glutathione	reductase assay.

Identification. A glutathione re- ductase assay is a device used to deter- mine the activity of the enzyme gluta- thione reductase in serum, plasma, or erythrocytes by such techniques as flu- orescence and photometry. The results of this assay are used in the diagnosis of liver disease, glutathione reductase deficiency, or riboflavin deficiency.

Classification. Class II (perform- ance standards).

[45 FR 60616, Sept. 12, 1980]

##### § 864.7400  Hemoglobin A 2 assay.

Identification. A hemoglobin A2 assay is a device used to determine the hemoglobin A2 content of human blood. The measurement of hemoglobin A2 is used in the diagnosis of  the thalassemias (hereditary hemolytic anemias characterized  by decreased synthesis of one or more types of he- moglobin polypeptide chains).

Classification. Class II (perform- ance standards).

[45 FR 60617, Sept. 12, 1980]

##### § 864.7415   Abnormal hemoglobin assay.

Identification. An abnormal hemo- globin assay is a device consisting of the reagents, apparatus, instrumenta- tion, and controls necessary to isolate and identify abnormal genetically de- termined hemoglobin types.

Classification. Class II (perform- ance standards).

[45 FR 60618, Sept. 12, 1980]

##### § 864.7425   Carboxyhemoglobin assay.

Identification.		A carboxyhemoglobin assay is a device used	to		determine	the carboxyhemoglobin (the compound formed when hemoglobin is exposed to carbon   monoxide)   content   of   human

#### § 864.7440

blood as an aid in the diagnosis of car- bon monoxide poisoning. This measure- ment may be made using methods such as	spectroscopy,	colorimetry, spectrophotometry,  and  gasometry.

Classification. Class II (perform- ance standards).

[45 FR 60619, Sept. 12, 1980]

##### § 864.7440   Electrophoretic  hemoglobin analysis system.

Identification. An electrophoretic hemoglobin analysis system is a device that electrophoretically separates and identifies normal and abnormal hemo- globin types as an aid in the diagnosis of anemia or erythrocytosis (increased total red cell mass) due to a hemo- globin abnormality.

Classification. Class II (perform- ance standards).

[45 FR 60620, Sept. 12, 1980]

##### § 864.7455   Fetal hemoglobin assay.

Identification. A fetal hemoglobin assay is a device that is used to deter- mine the presence and distribution of fetal hemoglobin (hemoglobin F) in red cells or to measure the amount of fetal hemoglobin present. The assay may be used to detect fetal red cells in the ma- ternal circulation or to detect the ele- vated levels of fetal hemoglobin exhib- ited in cases of hemoglobin abnormali- ties such as thalassemia (a hereditary hemolytic anemia characterized by a decreased synthesis of one or more types of hemoglobin polypeptide chains). The hemoglobin determination may be made by methods such as elec- trophoresis, alkali denaturation, col- umn chromatography, or radial immunodiffusion.

Classification. Class II (perform- ance standards).

[45 FR 60620, Sept. 12, 1980]

##### § 864.7470   Glycosylated	hemoglobin assay.

Identification. A glycosylated he- moglobin assay is a device used to measure the glycosylated hemoglobins (A1a, A1b, and A1c) in a patient’s blood by a column chromatographic proce- dure. Measurement of glycosylated he- moglobin is used to assess the level of control of a patient’s diabetes and to determine the proper insulin dosage for

#### 21 CFR Ch. I (4–1–17 Edition)

a patient. Elevated levels of glycosylated hemoglobin indicate un- controlled diabetes in a patient.

Classification. Class II (perform- ance standards).

[45 FR 60621, Sept. 12, 1980]

##### § 864.7490   Sulfhemoglobin assay.

Identification. A sulfhemoglobin assay is a device consisting of the re- agents, calibrators, controls, and in- strumentation used to determine the sulfhemoglobin (a compound of sulfur and hemoglobin) content of human blood as  an  aid  in  the  diagnosis of sulfhemoglobinemia (presence of sulfhemoglobin in the blood  due  to drug administration or exposure to a poison). This measurement may be made using methods such  as spectros- copy, colorimetry,  spectrophotometry, or  gasometry.

Classification. Class II (perform- ance standards).

[45 FR 60621, Sept. 12, 1980]

##### § 864.7500   Whole blood hemoglobin as- says.

Identification. A whole blood he- moglobin assay is a  device  consisting or reagents, calibrators, controls, or photometric or spectrophotometric in- strumentation used to measure the he- moglobin content of whole blood for the detection of anemia. This generic device category does not include auto- mated hemoglobin systems.

Classification. Class II (perform- ance standards).

[45 FR 60622, Sept. 12, 1980]

##### § 864.7525   Heparin assay.

Identification. A heparin assay is a device used to determine the level of the anticoagulant heparin in the pa- tient’s circulation. These assays are quantitative clotting time procedures using the effect of heparin on activated coagulation factor X (Stuart factor) or procedures based on the neutralization of heparin by protamine sulfate (a pro- tein that neutralizes heparin).

Classification. Class II (perform- ance standards).

[45 FR 60623, Sept. 12, 1980]

#### Food and Drug Administration, HHS	§ 864.7750

##### § 864.7660   Leukocyte	alkaline	phos- phatase test.

  - Identification. A leukocyte alka- line phosphatase test is a device used to identify the enzyme leukocyte alka- line phosphatase in neutrophilic granulocytes (granular leukocytes stainable by neutral dyes). The cytochemical identification of alkaline phosphatase depends on the formation of blue granules in cells containing al- kaline phosphatase. The results of this test are used to differentiate chronic granulocytic leukemia (a malignant disease characterized by excessive overgrowth of granulocytes in the bone marrow) and reactions that resemble true leukemia, such as those occuring in severe infections and polycythemia (increased total red cell mass).

  - Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60623, Sept. 12, 1980, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38790,  July 25,

2001]

##### § 864.7675   Leukocyte peroxidase test.

Identification. A leukocyte peroxi- dase test is a device used to distinguish certain myeloid cells derived from the bone marrow, i.e., neutrophils, eosinophils, and monocytes, from lymphoid cells of the lymphatic system and erythroid cells of the red blood cell series on the basis of their peroxidase activity as evidenced by staining. The results of this test are used in the dif- ferential diagnosis of the leukemias.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60624, Sept. 12, 1980, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38790,  July 25,

2001]

##### § 864.7695   Platelet	factor	4

radioimmunoassay.

Identification. A platelet factor 4 radioimmunoassay is a device used to measure the level of platelet factor 4, a protein released during platelet activa- tion by radioimmunoassay. This device measures    platelet    activiation,    which

may indicate a coagulation disorder, such as myocardial infarction or coro- nary artery disease.

Classification. Class II (perform- ance standards).

[45 FR 60625, Sept. 12, 1980; 46 FR 14890, Mar.

3, 1981]

##### § 864.7720   Prothrombin	consumption test.

Identification. A prothrombin con- sumption tests is a device that meas- ures the patient’s capacity to generate thromboplastin in the coagulation process. The test also is an indirect in- dicator of qualitative or quantitative platelet abnormalities. It is a screen- ing test for thrombocytopenia (de- creased number of blood platelets) and hemophilia A and B.

Classification. Class II (perform- ance standards).

[45 FR 60625, Sept. 12, 1980]

##### § 864.7735   Prothrombin-proconvertin test and thrombotest.

Identification. The prothrombin- proconvertin test and thrombotest are devices used in the regulation of cou- marin therapy (administration of a coumarin anticoagulant such as so- dium warfarin in the treatment of ve- nous thrombosis and pulmonary embo- lism) and as a diagnostic test in con- junction with, or in place of, the Quick prothrombin time test to detect coagu- lation disorders.

Classification. Class II (perform- ance standards).

[45 FR 60626, Sept. 12, 1980]

##### § 864.7750   Prothrombin time test.

Identification. A prothrombin time test is a device used as a general screening procedure for the detection of possible clotting factor deficiencies in the extrinsic coagulation pathway, which involves the reaction between coagulation factors III and VII, and to monitor patients receiving coumarin therapy (the administration of one of the coumarin anticoagulants in the treatment of venous thrombosis or pul- monary embolism).

Classification. Class II (perform- ance standards).

[45 FR 60626, Sept. 12, 1980]

#### § 864.7825

##### § 864.7825   Sickle cell test.

  - Identification. A sickle cell test is a device used to determine the sickle cell hemoglobin content of human blood to detect sickle cell trait or sick- le cell diseases.

  - Classification. Class II (perform- ance standards).

[45 FR 60627, Sept. 12, 1980]

##### § 864.7875   Thrombin time test.

Identification. A thrombin time test is a device used to measure fibrinogen concentration and detect fibrin or fibrinogen split products for the evaluation of bleeding disorders.

Classification. Class II (perform- ance standards).

[45 FR 60628, Sept. 12, 1980]

##### § 864.7900   Thromboplastin   generation test.

Identification. A thromboplastin generation test is a device used to de- tect and identify coagulation factor de- ficiencies and coagulation inhibitors.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60628, Sept. 12, 1980, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38790,  July 25,

2001]

##### § 864.7925   Partial thromboplastin time tests.

Identification. A partial thrombo- plastin time test is a device used for primary screening for coagulation ab- normalities, for evaluation of the ef- fect of therapy on procoagulant dis- orders, and as an assay for coagulation factor deficiencies of the intrinsic co- agulation pathway.

Classification. Class II (perform- ance standards).

[45 FR 60629, Sept. 12, 1980]

## Subpart I—Hematology Reagents

##### § 864.8100   Bothrops atrox reagent.

Identification. A Bothrops atrox reagent is a device made from snake venom and used to determine blood fibrinogen levels to aid in the evalua- tion of disseminated intravascular co-

21 CFR Ch. I (4–1–17 Edition)

agulation (nonlocalized clotting in the blood vessels) in patients receiving heparin therapy (the administration of the anticoagulant heparin in the treat- ment of thrombosis) or as an aid in the classification of dysfibrinogenemia (presence in the plasma of functionally defective fibrinogen).

Classification. Class II (perform- ance standards).

[45 FR 60629, Sept. 12, 1980]

##### § 864.8150   Calibrator for cell indices.

Identification. A calibrator for cell indices is a device that approximates whole blood or certain blood cells and that is used to set an instrument in- tended to  measure mean cell  volume (MCV), mean corpuscular hemoglobin (MCH), and mean corpuscular hemo- globin concentration (MCHC), or other cell indices. It is a suspension of par- ticles or cells whose size, shape, con- centration, and other characteristics have been precisely and accurately de- termined.

Classification. Class II (perform- ance standards).

[45 FR 60631, Sept. 12, 1980]

##### § 864.8165   Calibrator   for   hemoglobin or hematocrit measurement.

Identification. A calibrator for he- moglobin or hematocrit measurement is a device that approximates whole blood, red blood cells, or a hemoglobin derivative and that  is used  to set  in- struments intended to measure hemo- globin, the hematocrit, or both. It is a material whose characteristics have been precisely and accurately deter- mined.

Classification. Class II (perform- ance standards).

[45 FR 60632, Sept. 12, 1980]

##### § 864.8175   Calibrator	for	platelet counting.

Identification. A calibrator for platelet counting is a device that re- sembles platelets in plasma or whole blood and that is used to set a platelet counting instrument. It is a suspension of particles or cells whose size, shape concentration, and other characteris- tics have been precisely and accurately determined.

#### Food and Drug Administration, HHS	§ 864.9050

Classification. Class II (perform- ance standards).

[45 FR 60633, Sept. 12, 1980]

##### § 864.8185   Calibrator  for  red  cell  and white cell counting.

Identification. A calibrator for red cell and white cell counting is a device that resembles red or white blood cells and that is used to set instruments in- tended to count red cells, white cells, or both. It is a suspension of particles or cells whose size, shape, concentra- tion, and other characteristics have been precisely and accurately deter- mined.

Classification. Class II (perform- ance standards).

[45 FR 60634, Sept. 12, 1980]

##### § 864.8200   Blood cell diluent.

Identification. A blood cell diluent is a device used to dilute blood for fur- ther testing, such as blood cell count- ing.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60635, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.8500   Lymphocyte  separation  me- dium.

Identification. A lymphocyte sepa- ration medium is a device used to iso- late lymphocytes from whole blood.

Classification. Class I (general con- trols). This device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 864.9.

[45 FR 60636, Sept. 12, 1980, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38790,  July 25,

2001]

##### § 864.8540   Red cell lysing reagent.

Identification. A red cell lysing re- agent is a device used to lyse (destroy) red blood cells for hemoglobin deter- minations or aid in the counting of white blood cells.

Classification. Class I (general con- trols). This device is exempt from the premarket   notification   procedures   in

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 864.9.

[45 FR 60636, Sept. 12, 1980, as amended at 54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 864.8625   Hematology  quality  control mixture.

- Identification. A hematology qual- ity control mixture is a device used to ascertain the accuracy and precision of manual, semiautomated, and auto- mated determinations of cell param- eters such as white cell count (WBC), red cell count (RBC), platelet count (PLT), hemoglobin, hematocrit (HCT), mean corpuscular volume (MCV), mean corpuscular hemoglobin (MCH), and mean corpuscular hemoglobin con- centration (MCHC).

- Classification. Class II (perform- ance standards).

[45 FR 60637, Sept. 12, 1980]

##### § 864.8950   Russell   viper   venom   rea- gent.

- Identification. Russell viper venom reagent is a device used to determine the cause of an increase in the pro- thrombin time.

- Classification. Class I (general con- trols).

[45 FR 60637, Sept. 12, 1980]

## Subpart J—Products Used In Es- tablishments That Manufac- ture Blood and Blood Prod- ucts

##### § 864.9050   Blood bank supplies.

- Identification. Blood bank supplies are general purpose devices intended for in vitro use in blood banking. This generic type of device includes prod- ucts such as blood bank pipettes, blood grouping slides, blood typing tubes, blood typing racks, and cold packs for antisera reagents. The device does not include articles that are licensed by the Center for Biologics  Evaluation and Research of the Food and Drug Ad- ministration.

- Classification. Class I (general con- trols).

[45 FR 60638, Sept. 12, 1980, as amended at 53

FR 11253, Apr. 6, 1988]

§ 864.9100

##### § 864.9100 Empty container for the col- lection and processing of blood and blood components.

  - Identification. An empty container for the collection and processing of blood and blood components is a device intended for medical purposes that is an empty plastic bag or plastic or glass bottle used to collect, store, or transfer blood and blood components for further processing.

  - Classification. Class II (perform- ance standards).

[45 FR 60638, Sept. 12, 1980]

##### § 864.9125 Vacuum-assisted blood col- lection system.

- Identification. A vacuum-assisted blood collection system is a device in- tended for medical purposes that uses a vacuum to draw blood for subsequent reinfusion.

- Classification. Class I (general con- trols). The manual device is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60639, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.9145 Processing system for fro- zen blood.

- Identification. A processing system for frozen blood is a device used to glycerolize red blood cells prior to freezing to minimize hemolysis (disrup- tion of the red cell membrane accom- panied by the release of hemoglobin) due to freezing and thawing  of  red blood cells and to deglycerolize  and wash thawed cells for subsequent re- infusion.

- Classification. Class II (perform- ance standards).

[45 FR 60639, Sept. 12, 1980]

##### § 864.9160 Blood group substances of nonhuman origin for in vitro diag- nostic use.

- Identification. Blood group sub- stances of nonhuman origin for in vitro diagnostic use are materials, such as blood group specific substances pre- pared from nonhuman sources (e.g., pigs, cows, and horses) used to detect, identify, or neutralize antibodies to various human blood group antigens. This generic type of device does not in-

#### 21 CFR Ch. I (4–1–17 Edition)

clude materials that are licensed by the Center for Biologics  Evaluation and Research of the Food and Drug Ad- ministration.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60640, Sept. 12, 1980, as amended at 53

FR  11253,  Apr.  6,  1988;  63  FR  59225,  Nov.  3,

1998]

##### § 864.9175   Automated  blood  grouping and antibody test system.

- Identification. An automated blood grouping and antibody test system is a device used to group erythrocytes (red blood cells) and to detect antibodies to blood group antigens.

- Classification. Class II (perform- ance standards).

[45 FR 60641, Sept. 12, 1980]

##### § 864.9185   Blood grouping view box.

- Identification. A blood grouping view box is a device with a glass or plastic viewing surface, which may be illuminated and heated, that is used to view cell reactions in antigen-antibody testing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60641, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.9195   Blood   mixing   devices   and blood weighing devices.

- Identification. A blood mixing de- vice is a device intended for medical purposes that is used to mix blood or blood components by agitation. A blood weighing device is a device in- tended for medical purposes that is used to weigh blood or blood compo- nents as they are collected.

- Classification. Class I (general con- trols). The manual device is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60642, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 864.9300

##### § 864.9205   Blood and plasma warming device.

  - Nonelectromagnetic blood or plasma warming device—(1) Identification. A nonelectromagnetic blood and plasma warming device is a device that warms blood or plasma, by means other than electromagnetic radiation, prior to ad- ministration.

(2) Classification. Class II (perform- ance standards).

  - Electromagnetic blood and plasma warming device—(1) Identification. An electromagnetic blood and plasma warming device is a device that em- ploys	electromagnetic	radiation (radiowaves or microwaves) to warm a bag or bottle of blood or plasma prior to  administration.

(2) Classfication. Class III (premarket approval).

  - Date PMA or notice of completion of a PDP is required. No effective date has been established of the requirement for premarket approval for the device de- scribed in paragraph (b)(1). See § 864.3.

[45 FR 60642, Sept. 12, 1980, as amended at 52

FR 17733, May 11, 1987]

##### § 864.9225   Cell-freezing apparatus and reagents for in vitro diagnostic use.

- Identification. Cell-freezing appa- ratus and reagents for in vitro diag- nostic use are devices used to freeze human red blood cells for in vitro diag- nostic use.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60643, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.9245   Automated  blood  cell  sepa- rator.

- Identification. An automated blood cell separator is a device that uses a centrifugal or filtration separation principle to automatically withdraw whole blood from a donor, separate the whole blood into blood components, collect one or more of the blood compo- nents, and return to the donor the re- mainder of the whole blood and blood components. The automated blood cell separator device is intended for routine collection of blood and blood compo-

nents for transfusion or further manu- facturing use.

- Classification. Class II (special controls). The special control for this device is a guidance for industry and FDA staff entitled ‘‘Class II Special Controls Guidance Document: Auto- mated Blood Cell Separator Device Op- erating by Centrifugal or Filtration Separation   Principle.’’

[72 FR 67644, Nov. 30, 2007]

##### § 864.9275   Blood bank centrifuge for in vitro diagnostic use.

- Identification. A blood bank cen- trifuge for in vitro diagnostic use is a device used only to separate blood cells for further diagnostic testing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60645, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.9285   Automated	cell-washing centrifuge for immuno-hematology.

- Identification. An automated cell- washing centrifuge for immuno-hema- tology is a device used to separate and prepare cells and sera for further in vitro diagnostic testing.

- Classification. Class II (perform- ance standards).

[45 FR 60646, Sept. 12, 1980]

##### § 864.9300   Automated Coombs test sys- tems.

- Identification. An automated Coombs test system is a device used to detect and identify antibodies in pa- tient sera or antibodies bound to red cells. The Coombs test is used for the diagnosis of hemolytic disease of the newborn, and autoimmune hemolytic anemia. The test is also used in crossmatching and in investigating transfusion reactions and drug-induced red cell sensitization.

- Classification. Class II (perform- ance standards).

[45 FR 60646, Sept. 12, 1980]

#### § 864.9320

##### § 864.9320   Copper  sulfate  solution  for specific gravity determinations.

  - Identification. A copper sulfate so- lution for specific gravity determina- tions is a device used to determine whether the hemoglobin content of a potential donor’s blood meets the re- quired level (12.5 grams per 100 milli- liters of blood for women and  13.5 grams per 100 milliliters of blood for men).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60647, Sept. 12, 1980, as amended at 65

FR 2310, Jan. 14, 2000]

##### § 864.9400   Stabilized enzyme solution.

- Identification. A stabilized enzyme solution is a reagent intended for med- ical purposes that is used to enhance the reactivity of red blood cells with certain antibodies, including anti- bodies that are not detectable by other techniques. These enzyme solutions in- clude papain, bromelin, ficin, and trypsin.

- Classification. Class II (perform- ance standards).

[45 FR 60647, Sept. 12, 1980]

##### § 864.9550   Lectins and protectins.

- Identification. Lectins and protectins are proteins derived from plants and lower animals that cause cell agglutination in the presence of certain antigens. These substances are used to detect blood group antigens for in vitro diagnostic purposes.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60648, Sept. 12, 1980, as amended at 63

FR 59226, Nov. 3, 1998]

##### § 864.9575   Environmental chamber for storage of platelet concentrate.

- Identification. An environmental chamber for storage of platelet con- centrate is a device used to hold plate- let-rich plasma within a preselected temperature range.

- Classification. Class II (special controls).  The  device  is  exempt  from

#### 21 CFR Ch. I (4–1–17 Edition)

the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60648, Sept. 12, 1980, as amended at 63

FR 59226, Nov. 3, 1998]

##### § 864.9600   Potentiating   media   for   in vitro diagnostic use.

- Identification. Potentiating media for in vitro diagnostic use are media, such as bovine albumin, that are used to suspend red cells and to enhance cell reactions for antigen-antibody testing.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60649, Sept. 12, 1980, as amended at 63

FR 59226, Nov. 3, 1998]

##### § 864.9650   Quality control kit for blood banking reagents.

- Identification. A quality control kit for blood banking reagents is a de- vice that consists of sera, cells, buffers, and antibodies used to determine the specificity, potency, and reactivity of the cells and reagents used for blood banking.

- Classification. Class II (perform- ance standards).

[45 FR 60649, Sept. 12, 1980]

##### § 864.9700   Blood   storage   refrigerator and blood storage freezer.

- Identification. A blood storage re- frigerator and a blood storage freezer are devices intended for medical pur- poses that are used to preserve blood and blood products by storing them at cold or freezing temperatures.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60650, Sept. 12, 1980, as amended at 63

FR 59226, Nov. 3, 1998]

##### § 864.9750   Heat-sealing device.

- Identification. A heat-sealing de- vice is a device intended for medical purposes that uses heat to seal plastic bags containing blood or blood compo- nents.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	Pt. 866

premarket notification procedures in subpart E of part 807 of this chapter subject to § 864.9.

[45 FR 60650, Sept. 12, 1980, as amended at 65

FR 2311, Jan. 14, 2000]

##### § 864.9875   Transfer set.

  - Identification. A transfer set is a device intended for medical purposes that consists of a piece of tubing with suitable adaptors used to transfer blood or plasma from one container to another.

##### Subpart B—Diagnostic Devices

866.1620  Antimicrobial  susceptibility   test disc.

866.1640 Antimicrobial susceptibility test powder.

866.1645 Fully automated short-term incu- bation cycle antimicrobial susceptibility system.

866.1700 Culture medium for antimicrobial susceptibility     tests.

##### Subpart C—Microbiology Devices

| (b) Classification. Class II | (perform- 866.2120 | Anaerobic chamber. |
| --- | --- | --- |
| ance standards). | 866.2160 | Coagulase plasma. |
| [45 FR 60651, Sept. 12, 1980] | 866.2170 | Automated colony counter. |
|  | 866.2180 | Manual colony counter. |

866.2050   Staphylococcal	typing bacteriophage.

## Subpart K—Products Used In Es- tablishments That Manufac- ture Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/Ps)

##### § 864.9900 Cord blood processing sys- tem and storage container.

- Identification. A cord blood proc- essing system and storage container is a device intended for use in the proc- essing and the storage of cord blood. This device is a functionally closed processing system that includes con- tainers, other soft goods, and a cen- trifugation system for cord blood con- centration,  and  a  final  container  for the cryopreservation and the storage of a cord blood product.

- Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document: Cord Blood Processing System and  Storage  Container.’’  For the availability of this guidance docu- ment, see § 864.1(d).

[72 FR 4638, Feb. 1, 2007]

## PART 866—IMMUNOLOGY AND MICROBIOLOGY DEVICES

##### Subpart A—General Provisions

Sec.

866.1    Scope.

866.3 Effective dates of requirement for pre- market  approval.

866.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

866.2300    Multipurpose culture medium. 866.2320    Differential culture medium. 866.2330   Enriched culture medium.

866.2350    Microbiological   assay   culture   me- dium.

866.2360    Selective culture medium. 866.2390   Transport culture medium.

866.2410    Culture    medium    for    pathogenic

Neisseria spp.

866.2420   Oxidase   screening   test   for   gonor- rhea.

866.2440   Automated  medium  dispensing  and stacking    device.

866.2450    Supplement for culture media. 866.2480    Quality    control    kit    for    culture

media.

866.2500    Microtiter  diluting  and  dispensing device.

866.2540 Microbiological incubator. 866.2560 Microbial growth monitor. 866.2580   Gas-generating device.

866.2600   Wood’s fluorescent lamp.

866.2660    Microorganism   differentiation   and identification    device.

866.2850   Automated zone reader.

866.2900    Microbiological specimen collection and  transport  device.

##### Subpart D—Serological Reagents

866.3010 Acinetobacter	calcoaceticus	sero- logical    reagents.

866.3020    Adenovirus serological reagents. 866.3035    Arizona spp. serological reagents. 866.3040    Aspergillus spp. serological reagents. 866.3050    Beta-glucan serological assays.

866.3060    Blastomyces  dermatitidis  serological reagents.

866.3065    Bordetella spp. serological reagents. 866.3085    Brucella spp. serological reagents.

866.3110   Campylobacter  fetus  serological  re- agents.

866.3120    Chlamydia serological reagents. 866.3125   Citrobacter spp. serological reagents. 866.3130    Clostridium difficile toxin gene am-

plification assay.

866.3135   Coccidioides  immitis  serological  re- agents.

Pt. 866

866.3140 Corynebacterium spp. serological re- agents.

866.3145 Coxsackievirus  serological  re- agents.

866.3165 Cryptococcus neoformans serological reagents.

866.3175 Cytomegalovirus serological re- agents.

866.3200 Echinococcus spp. serological re- agents.

866.3205    Echovirus serological reagents. 866.3210   Endotoxin assay.

866.3220 Entamoeba histolytica serological re- agents.

866.3225    Enterovirus nucleic acid assay. 866.3235    Epstein-Barr   virus   serological   re-

agents.

866.3240 Equine encephalomyelitis virus se- rological    reagents.

866.3250 Erysipelothrix rhusiopathiae sero- logical    reagents.

866.3255    Escherichia coli serological reagents. 866.3270   Flavobacterium  spp.  serological  re-

agents.

866.3280 Francisella tularensis serological re- agents.

866.3290    Gonococcal antibody test (GAT). 866.3300   Haemophilus   spp.    serological    re-

agents.

866.3305 Herpes simplex virus serological as- says.

866.3310 Hepatitis A virus (HAV) serological assays.

866.3320 Histoplasma capsulatum serological reagents.

866.3328 Influenza virus antigen  detection test  system.

866.3330 Influenza virus serological  re- agents.

866.3332 Reagents for detection of specific novel influenza A viruses.

866.3336 John  Cunningham Virus serological reagents.

866.3340    Klebsiella spp. serological reagents. 866.3350    Leptospira spp. serological reagents. 866.3355    Listeria spp. serological reagents.

866.3360   Lymphocytic	choriomeningitis virus   serological   reagents.

866.3365 Multiplex nucleic acid assay for identification of microorganisms and re- sistance markers from positive blood cul- tures.

866.3370 Mycobacterium	tuberculosis

immunofluorescent  reagents.

866.3372 Nucleic acid-based in vitro diag- nostic devices for the detection of Mycobacterium tuberculosis complex in respiratory   specimens.

866.3373 Nucleic acid-based in vitro diag- nostic devices for the detection of Mycobacterium tuberculosis complex (MTB-complex)  and   the   genetic mutations associated with MTB-complex antibiotic resistance in respiratory specimens.

866.3375 Mycoplasma spp. serological re- agents.

#### 21 CFR Ch. I (4–1–17 Edition)

866.3380    Mumps virus serological reagents. 866.3390    Neisseria spp. direct serological test

reagents.

866.3395    Norovirus serological reagents. 866.3400    Parainfluenza  virus  serological  re-

agents.

866.3402 Plasmodium species antigen detec- tion   assays.

866.3405    Poliovirus serological reagents. 866.3410    Proteus spp. (Weil-Felix) serological

reagents.

866.3415 Pseudomonas spp. serological re- agents.

866.3460 Rabiesvirus immunofluorescent re- agents.

866.3470    Reovirus serological reagents. 866.3480    Respiratory   syncytial   virus   sero-

logical reagents.

866.3490   Rhinovirus serological reagents. 866.3500   Rickettsia serological reagents. 866.3510   Rubella virus serological reagents. 866.3520  Rubeola (measles) virus serological

reagents.

866.3550    Salmonella spp. serological reagents. 866.3600   Schistosoma    spp.    serological    re-

agents.

866.3630    Serratia spp. serological reagents. 866.3660    Shigella spp. serological reagents. 866.3680    Sporothrix  schenckii  serological  re-

agents.

866.3700 Staphylococcus aureus serological re- agents.

866.3720 Streptococcus spp. exoenzyme re- agents.

866.3740 Streptococcus spp. serological re- agents.

866.3780 Toxoplasma gondii serological re- agents.

866.3820 Treponema pallidum nontreponemal test   reagents.

866.3830 Treponema pallidum treponemal test reagents.

866.3850 Trichinella spiralis serological re- agents.

866.3860 Trichomonas vaginalis nucleic acid assay.

866.3870 Trypanosoma spp. serological re- agents.

866.3900 Varicella-zoster virus serological reagents.

866.3930   Vibrio cholerae serological reagents. 866.3940   West    Nile    virus    serological    re-

agents.

866.3945    Dengue virus serological reagents. 866.3946    Dengue  virus  nucleic  acid  amplifi-

cation test reagents.

866.3950 In vitro human immunodeficiency virus (HIV) drug resistance  genotype assay.

866.3980 Respiratory viral panel multiplex nucleic  acid  assay.

866.3990   Gastrointestinal	microorganism multiplex  nucleic  acid-based  assay.

#### Food and Drug Administration, HHS	Pt. 866

##### Subpart E—Immunology Laboratory Equipment and Reagents

866.4070    RNA Preanalytical Systems. 866.4100  Complement reagent.

866.4500    Immunoelectrophoresis equipment.

866.4520    Immunofluorometer equipment.

866.4540   Immunonephelometer equipment. 866.4600    Ouchterlony agar plate.

866.4700    Automated  fluorescence  in situ  hy- bridization   (FISH)   enumeration   systems.

866.4800    Radial immunodiffusion plate. 866.4830  Rocket	immunoelectrophoresis

equipment.

866.4900   Support gel.

##### Subpart F—Immunological Test Systems

866.5040    Albumin   immunological   test   sys- tem.

866.5060  Prealbumin	immunological	test system.

866.5065   Human  allotypic  marker  immuno- logical  test  system.

866.5080 Alpha-1-antichymotrypsin immuno- logical  test  system.

866.5090   Antimitochondrial	antibody immunological   test   system.

866.5100  Antinuclear	antibody	immuno- logical   test   system.

866.5110   Antiparietal	antibody	immuno- logical   test   system.

866.5120  Antismooth	muscle	antibody immunological   test   system.

866.5130    Alpha-1-antitrypsin    immunological test   system.

866.5150  Bence-Jones	proteins immunological   test   system.

866.5160    Beta-globulin    immunological     test system.

866.5170    Breast    milk    immunological    test system.

866.5180    Fecal    calprotectin    immunological test   system.

866.5200   Carbonic     anhydrase     B     and     C immunological   test   system.

866.5210    Ceruloplasmin   immunological   test system.

866.5220    Cohn fraction II immunological test system.

866.5230    Colostrum  immunological  test  sys- tem.

866.5240   Complement  components  immuno- logical   test   system.

866.5250    Complement    C1     inhibitor    (inact- ivator)   immunological   test   system.

866.5260 Complement	C3b	inactivator immunological   test   system.

866.5270    C-reactive    protein    immunological test   system.

866.5320    Properidin  factor  B  immunological test   system.

866.5330   Factor  XIII,  A,  S,  immunological test system.

866.5340    Ferritin    immunological    test    sys- tem.

866.5350    Fibrinopeptide     A     immunological test   system.

866.5360    Cohn    fraction    IV    immunological test   system.

866.5370    Cohn fraction V immunological test system.

866.5380    Free secretory component immuno- logical   test   system.

866.5400    Alpha-globulin   immunological   test system.

866.5420  Alpha-1-glycoproteins	immuno- logical  test  system.

866.5425  Alpha-2-glycoproteins	immuno- logical  test  system.

866.5430  Beta-2-glycoprotein	I	immuno- logical   test   system.

866.5440    Beta-2-glycoprotein     III     immuno- logical   test   system.

866.5460    Haptoglobin     immunological     test system.

866.5470    Hemoglobin     immunological     test system.

866.5490    Hemopexin immunological test sys- tem.

866.5500   Hypersensitivity	pneumonitis immunological   test   system.

866.5510   Immunoglobulins A, G, M, D, and E immunological test system.

866.5520    Immunoglobulin   G   (Fab   fragment specific)   immunological   test   system.

866.5530    Immunoglobulin    G    (Fc    fragment specific)   immunological   test   system.

866.5540    Immunoglobulin    G    (Fd    fragment specific)   immunological   test   system.

866.5550    Immunoglobulin   (light   chain   spe- cific)   immunological   test   system.

866.5560  Lactic	dehydrogenase	immuno- logical   test   system.

866.5570     Lactoferrin immunological test sys- tem.

866.5580    Alpha-1-lipoprotein    immunological test   system.

866.5590    Lipoprotein   X  immunological   test system.

866.5600    Low-density    lipoprotein    immuno- logical   test   system.

866.5620  Alpha-2-macroglobulin	immuno- logical  test  system.

866.5630  Beta-2-microglobulin	immuno- logical  test  system.

866.5640    Infectious  mononucleosis  immuno- logical   test   system.

866.5660    Multiple    autoantibodies    immuno- logical   test   system.

866.5680    Myoglobin  immunological  test  sys- tem.

866.5700   Whole   human   plasma   or   serum immunological   test   system.

866.5715    Plasminogen     immunological     test system.

866.5735    Prothrombin     immunological     test system.

866.5750  Radioallergosorbent	(RAST) immunological   test   system.

866.5760   Tryptase test system.

866.5765    Retinol-binding    protein    immuno- logical   test   system.

#### § 866.1

866.5775 Rheumatoid factor  immunological test   system.

866.5785  Anti-Saccharomyces	cerevisiae	(S. cerevisiae)  antibody  (ASCA)  test  systems. 866.5800   Seminal    fluid    (sperm)    immuno-

logical test system.

866.5820 Systemic lupus erythematosus immunological   test   system.

866.5860 Total  spinal  fluid  immunological test   system.

866.5870 Thyroid autoantibody  immuno- logical   test   system.

866.5880 Transferrin immunological test sys- tem.

866.5890   Inter-alpha       trypsin	inhibitor immunological   test   system.

866.5900 Cystic fibrosis transmembrane con- ductance  regulator  (CFTR)  gene   muta- tion   detection   system.

866.5910 Quality control material for cystic fibrosis nucleic acid assays.

866.5940 Autosomal recessive carrier screen- ing  gene  mutation  detection  system.

##### Subpart G—Tumor Associated Antigen Immunological Test Systems

866.6010  Tumor	associated	antigen immunological   test   system.

866.6020 Immunomagnetic  circulating  can- cer cell selection and  enumeration  sys- tem.

866.6030 AFP-L3% immunological  test  sys- tem.

866.6040 Gene expression profiling test  sys- tem  for  breast  cancer  prognosis.

866.6050 Ovarian adnexal mass assessment score  test  system.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  47  FR  50823,  Nov.  9,  1982,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 866 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 866.1   Scope.

- This part sets forth the classifica- tion of immunology and microbiology devices intended for human use that are in commercial distribution.

- The indentification  of  a  device  in a regulation in this part is not a  pre- cise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this  part,  but  shall  state  why  the  de-

21 CFR Ch. I (4–1–17 Edition)

vice     is     substantially     equivalent     to other devices, as required by § 807.87.

- To avoid duplicative listings, an immunology and microbiology device that has two or more types of uses (e.g., used both as a diagnostic device and as a microbiology device) is listed only in one subpart.

- References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

- Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 17733, May 11, 1987, as amended at 68

FR  5827,  Feb.  5,  2003;  79  FR  50552,  Aug.  25,

2014]

##### § 866.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (Premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraphs (b) and (c) of this sec- tion. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or  on the last day of the 30th full calendar month after the regulation that classi- fies the device into class III is effec- tive, whichever is later. See section 501(f)(2)(B) of the act. Accordingly, un- less an effective date of the require- ment for premarket approval is shown

#### Food and Drug Administration, HHS	§ 866.9

in the regulation for a device classified into class III in this part, the device may be commercially distributed with- out FDA’s issuance of an order approv- ing a PMA or declaring completed a PDP for the device.  If  FDA  promul- gates a regulation under section 515(b) of the act requiring premarket ap- proval for a device, section 501(f)(1)(A) of the act applies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

- A device identified in a regulation in this part that is classified into class III and that is subject to the transi- tional provisions of section 520(l) of the act is automatically classified by stat- ute into class III and must have an ap- proval under section 515 of the act be- fore being commercially distributed. Accordingly, the regulation for such a class III transitional device states that as of the enactment date of the amend- ments, May 28, 1976, the device must have an approval under section 515 of the act before commercial distribution.

[52 FR 17733, May 11, 1987; 52 FR 22577, June

12, 1987]

##### § 866.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent

that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

#### § 866.1620

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2311, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 866.1620   Antimicrobial  susceptibility test disc.

- Identification. An antimicrobial susceptibility test disc is a device that consists of antimicrobic-impregnated paper discs used to measure by a disc- agar diffusion technique or a disc-broth elution technique the in vitro suscepti- bility of most clinically important bac- terial pathogens to antimicrobial agents. In the disc-agar diffusion tech- nique, bacterial susceptibility is ascertained by directly measuring the magnitude of a zone of bacterial inhibi- tion around the disc on an agar sur- face. The disc-broth elution technique is associated with an automated rapid susceptibility test system and employs a fluid medium in which susceptibility is ascertained by photometrically measuring changes in bacterial growth resulting when antimicrobial material is eluted from the disc into the fluid medium. Test results are used to deter- mine the antimicrobial agent of choice in the treatment of bacterial diseases.

- Classification. Class II (perform- ance standards).

##### § 866.1640   Antimicrobial  susceptibility test powder.

- Identification. An antimicrobial susceptibility test powder is a device that consists of an antimicrobial drug powder packaged in vials in specified amounts and intended for use in clin- ical laboratories for determining in vitro susceptibility of bacterial patho- gens to these therapeutic agents. Test results are used to determine the anti-

21 CFR Ch. I (4–1–17 Edition)

microbial agent of choice in the treat- ment of bacterial diseases.

- Classification. Class II (perform- ance standards).

##### § 866.1645 Fully automated short-term incubation cycle antimicrobial sus- ceptibility system.

- Identification. A fully automated short-term incubation cycle anti- microbial susceptibility system is a de- vice  that  incorporates  concentrations of antimicrobial agents into a system for the purpose of determining in vitro susceptibility of bacterial pathogens isolated from clinical specimens. Test results obtained from short-term (less than 16 hours) incubation are used to determine the antimicrobial agent of choice to treat bacterial diseases.

- Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class  II Special  Controls  Guid- ance Document: Antimicrobial Suscep- tibility Test (AST) Systems; Guidance for Industry and FDA.’’

[68 FR 5827, Feb. 5, 2003]

##### § 866.1700 Culture medium for anti- microbial susceptibility tests.

- Identification. A culture medium for antimicrobial susceptibility tests is a device intended for medical purposes that consists of any medium capable of supporting the growth of many of the bacterial pathogens that are subject to antimicrobial susceptibility tests. The medium should be free of components known to be antagonistic to the com- mon agents for which susceptibility tests are performed in the treatment of disease.

- Classification. Class II (perform- ance standards).

## Subpart C—Microbiology Devices

##### § 866.2050   Staphylococcal	typing bacteriophage.

- Identification. A staphylococcal typing bacteriophage is a device con- sisting of a bacterial virus intended for medical purposes to identify patho- genic staphylococcal bacteria through use of the bacteria’s susceptibility to destruction by the virus. Test results are used principally for the collection of epidemiological information.

Food and Drug Administration, HHS	§ 866.2300

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 866.2120   Anaerobic chamber.

- Identification. An anaerobic cham- ber is a device intended for medical purposes to maintain an anaerobic (ox- ygen free) environment. It is used to isolate and cultivate anaerobic micro- organisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9. The device is also exempt from the good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements  concerning  records,  and

§ 820.198,    with    respect    to    complaint files.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38790, July 25, 2001]

##### § 866.2160   Coagulase plasma.

- Identification. Coagulase plasma is a device that consists of freeze-dried animal or human plasma that is in- tended for medical purposes to perform coagulase tests primarily on staphy- lococcal bacteria. When reconstituted, the fluid plasma is clotted by the ac- tion of the enzyme coagulase which is produced by pathogenic staphylococci. Test results are used primarily as an aid in the diagnosis of disease caused by pathogenic bacteria belonging to the genus Staphylococcus and provide epidemiological information on disease caused by these microorganisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  61

FR 1119,  Jan. 16,  1996;  66  FR 38790,  July 25,

2001]

##### § 866.2170   Automated colony counter.

- Identification. An automated col- ony counter is a mechanical device in- tended for medical purposes to deter- mine the number of bacterial colonies present on a bacteriological culture medium contained in a petri plate. The number of colonies counted is used in the diagnosis of disease as a measure of the degree of bacterial infection.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25045, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 866.2180   Manual colony counter.

- Identification. A manual colony counter is a device intended for med- ical purposes that consists of a printed grid system superimposed on an illumi- nated screen. Petri plates containing bacterial colonies to be counted are placed on the screen for better viewing and ease of  counting. The  number of colonies counted is used in the diag- nosis of disease as a measure of the de- gree of bacterial infection.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9. The device is also exempt from the good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements  concerning  records,  and

§ 820.198,    with    respect    to    complaint files.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38790, July 25, 2001]

##### § 866.2300   Multipurpose    culture    me- dium.

- Identification. A multipurpose cul- ture medium is a device that consists primarily of liquid or solid biological materials intended for medical pur- poses for the cultivation and identi- fication of several types of pathogenic microorganisms without the need of additional    nutritional    supplements.

#### § 866.2320

Test results aid in the diagnosis of dis- ease and also provide epidemiological information on diseases  caused  by these  microorganisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 866.2320   Differential	culture	me- dium.

- Identification. A differential cul- ture medium is a device that consists primarily of liquid biological materials intended for medical purposes to cul- tivate and identify different types of pathogenic microorganisms. The iden- tification of these microorganisms is accomplished by the addition of a spe- cific biochemical component(s) to the medium. Microorganisms are identified by a visible change (e.g., a color change) in a specific biochemical com- ponent(s) which indicates that specific metabolic reactions have occurred. Test results aid in the diagnosis of dis- ease and also provide epidemiological information on diseases caused by these  microorganisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38790, July 25,

2001]

##### § 866.2330   Enriched culture medium.

- Identification. An enriched culture medium is a device that consists pri- marily of liquid or solid biological ma- terials intended for medical purposes to cultivate and identify fastidious microorganisms (those having complex nutritional requirements). The device consists of a relatively simple basal medium enriched by the addition of such nutritional components as blood, blood serum, vitamins, and extracts of plant or animal tissues. The device is used in the diagnosis of disease caused by pathogenic microorganisms and also provides epidemiological information on these diseases.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2350   Microbiological   assay   cul- ture medium.

- Identification. A microbiological assay culture medium is a device that consists primarily of liquid or solid bi- ological materials intended for medical purposes to cultivate selected test microorganisms in order to measure by microbiological procedures the con- centration in a patient’s serum of cer- tain substances,  such  as  amino acids, antimicrobial agents, and vitamins. The concentration of these substances is measured by their ability to promote or inhibit the growth of the test orga- nism in the innoculated medium. Test results aid in the diagnosis of disease resulting from either deficient or ex- cessive amounts of these substances in a patient’s serum. Tests results may also be used to monitor the effects of the administration of certain anti- microbial drugs.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2360   Selective culture medium.

- Identification. A selective culture medium is a device that consists pri- marily of liquid or solid biological ma- terials intended for medical purposes to cultivate and identify certain patho- genic microorganisms. The device con- tains one or more components that suppress the growth of certain micro- organisms while either promoting or not affecting the growth of other microorganisms. The device aids in the diagnosis of disease caused by patho- genic microorganisms and also pro- vides epidemiological information on these diseases.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	§ 866.2450

premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2390   Transport culture medium.

  - Identification. A transport culture medium is a device that consists of a semisolid, usually non-nutrient, me- dium that maintains the viability of suspected pathogens contained in pa- tient specimens while in transit from the specimen collection area to the laboratory. The device aids in the diag- nosis of disease caused by pathogenic microorganisms and also provides epi- demiological information on these dis- eases.

  - Classification. Class I (general con- trols).

##### § 866.2410   Culture  medium  for  patho- genic Neisseria spp.

- Identification. A culture medium for pathogenic Neisseria spp. is a device that consists primarily  of liquid  or solid biological materials used to cul- tivate and identify pathogenic Neisseria spp. The identification aids in the diag- nosis of disease caused by bacteria be- longing to the genus Neisseria, such as epidemic cerebrospinal  meningitis, other meningococcal disease, and gon- orrhea, and also provides epidemiolog- ical information on these microorga- nisms.

- Classification. Class II (perform- ance standards).

##### § 866.2420   Oxidase  screening  test  for gonorrhea.

- Identification. An oxidase screen- ing test for gonorrhea is an in vitro de- vice that consists of the articles in- tended to identify by chemical reac- tion, cytochrome oxidase, an oxidizing enzyme that is associated with certain bacteria	including	Neisseria gonorrhoeae. A sample of a male’s urethral discharge is obtained on  a swab which is placed into a wetting agent containing an ingredient  that will react with cytochrome oxidase. When cytochrome oxidase is present, the swab turns a dark purple color within 3 minutes. Because it  is  un- likely   that   cytochrome   oxidase-posi-

tive organisms other than Neisseria gonorrhoeae are present in the urethral discharge of males, the  identification of cytochrome oxidase with this device indicates presumptive infection of the patient with the causative agent of gonorrhea.

- Classification. Class III (premarket approval) (transitional device).

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 866.3.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  52

FR 17734, May 11, 1987]

##### § 866.2440   Automated	medium	dis- pensing and stacking device.

- Identification. An automated me- dium dispensing and stacking device is a device intended for medical purposes to dispense a microbiological culture medium into petri dishes and then me- chanically stack the petri dishes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9. The device is also exempt from the good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements  concerning  records,  and

§ 820.198,  with  respect  to  complaint files.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38791, July 25, 2001]

##### § 866.2450   Supplement	for	culture media.

- Identification. A supplement for culture media is a device, such as a vi- tamin or sugar mixture,  that  is  added to a solid or liquid basal culture me- dium to produce a desired formulation and that is intended for medical pur- poses to enhance the growth of fas- tidious microorganisms (those having complex  nutritional  requirements). This device aids in the diagnosis of dis- eases caused by pathogenic microorga- nisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 866.2480

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2480   Quality  control  kit  for  cul- ture media.

  - Identification. A quality control kit for culture media is a device that consists of paper discs (or other suit- able materials), each impregnated with a specified, freeze-dried, viable micro- organism, intended for medical pur- poses to determine if a given culture medium is able to support the growth of that microorganism. The device aids in the diagnosis of disease caused by pathogenic microorganisms and also provides epidemiological information on these diseases.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2500   Microtiter diluting and dis- pensing device.

- Identification. A microtiter dilut- ing and dispensing device is a mechan- ical device intended for medical pur- poses to dispense or serially dilute very small quantities of biological or chem- ical reagents for use in various diag- nostic procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.2540   Microbiological incubator.

- Identification. A microbiological incubator is a device with various chambers or water-filled compartments in which controlled environmental con- ditions, particularly temperature, are maintained. It is intended for medical purposes to cultivate microorganisms and aid in the diagnosis of disease.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9. The device is also exempt from the good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements  concerning  records,  and

§ 820.198,    with    respect    to    complaint files.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38791, July 25, 2001]

##### § 866.2560   Microbial growth monitor.

- Identification. A microbial growth monitor is a device intended for med- ical purposes that measures the con- centration of bacteria suspended in a liquid medium by measuring changes in light scattering properties, optical density, electrical impedance, or by making direct bacterial counts. The device aids in the diagnosis of disease caused by pathogenic microorganisms.

- Classification. Class I. With the ex- ception of automated blood culturing system devices that are used in testing for bacteria, fungi, and other micro- organisms in blood and other normally sterile body fluids, this device is ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  60

FR 38482, July 27, 1995]

##### § 866.2580   Gas-generating device.

- Identification. A gas-generating device is a device intended for medical purposes that produces predetermined amounts of selected gases to be used in a closed chamber in order to establish suitable atmospheric conditions for cultivation of microorganisms with special atmospheric requirements. The device aids in the diagnosis of disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

#### Food and Drug Administration, HHS	§ 866.3020

##### § 866.2600   Wood’s fluorescent lamp.

  - Identification. A Wood’s fluores- cent lamp is a device intended for med- ical purposes to detect fluorescent ma- terials (e.g., fluorescein pigment pro- duced by certain microorganisms) as an aid in the identification of these microorganisms. The device aids in the diagnosis of disease.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9. The device is also exempt from the good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements  concerning  records,  and

§ 820.198,    with    respect    to    complaint files.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38791, July 25, 2001]

##### § 866.2660   Microorganism    differentia- tion and identification device.

- Identification. A microorganism differentiation and identification de- vice is a device intended for medical purposes that consists of one or more components, such as differential cul- ture media, biochemical reagents, and paper discs or paper strips impregnated with test reagents, that are usually contained in individual compartments and used to differentiate and identify selected microorganisms. The device aids in the diagnosis of disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.2850   Automated zone reader.

- Identification. An automated zone reader is a mechanical device intended for medical purposes to measure zone diameters of microbial growth inhibi- tion (or exhibition), such as those ob- served on the surface of certain culture media used in disc-agar diffusion anti- microbial susceptibility tests. The de- vice aids in decisionmaking respecting the treatment of disease.

- Classification. Class I (general con- trols).

##### § 866.2900   Microbiological	specimen collection and transport device.

- Identification. A microbiological specimen collection and transport de- vice is a specimen collecting chamber intended for medical purposes to pre- serve the viability or integrity of microorganisms in specimens during storage of specimens after their collec- tion and during their transport from the collecting area to the laboratory. The device may be labeled or otherwise represented as sterile. The device aids in the diagnosis of disease caused by pathogenic  microorganisms.

- Classification. Class I (general con- trols).

## Subpart D—Serological Reagents

§ 866.3010   Acinetobacter   calcoaceticus

##### serological reagents.

- Identification. Acinetobacter calcoaceticus serological reagents are devices that consist of Acinetobacter calcoaceticus antigens and antisera used to identify this bacterium from cul- tured isolates derived from clinical specimens. The identification aids in the diagnosis of disease caused by the bacterium Acinetobacter calcoaceticus and provides epidemiological informa- tion on disease caused by this micro- organism. This organism becomes pathogenic in patients with burns or with immunologic deficiency, and in- fection can result in sepsis (blood poi- soning).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.3020   Adenovirus   serological   re- agents.

- Identification. Adenovirus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to adenovirus in serum. Additionally, some of these reagents consist of adenovirus antisera conjugated with a

§ 866.3035

fluorescent dye and are used to identify adenoviruses directly  from  clinical specimens. The identification aids in the diagnosis of disease caused by adenoviruses and provides epidemiolog- ical information on these diseases. Adenovirus infections may cause phar- yngitis (inflammation of the throat), acute respiratory diseases, and certain external diseases of the eye (e.g., con- junctivitis).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.3035   Arizona  spp.  serological  re- agents.

- Identification. Arizona spp. sero- logical reagents are devices that con- sist of antisera and antigens used to identify Arizona spp. in cultured iso- lates derived from clinical specimens. The identification aids in the diagnosis of disease caused by bacteria belonging to the genus Arizona and provides epi- demiological information on diseases caused by these microorganisms. Ari- zona spp. can cause gastroenteritis (food poisoning) and sepsis (blood poi- soning).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

§ 866.3040   Aspergillus  spp.  serological reagents.

- Identification. Aspergillus spp. sero- logical reagents are devices that con- sist of antigens and antisera used in various serological tests to identify antibodies to Aspergillus spp. in serum. The identification aids in the diagnosis of aspergillosis caused by fungi belong- ing to the genus Aspergillus. Aspergillosis is a disease marked by in- flammatory granulomatous (tumor- like) lessions in the skin, ear, eyeball cavity, nasal sinuses, lungs, and occa- sionally the bones.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3050   Beta-glucan  serological  as- says.

- Identification. Beta-glucan sero- logical assays are devices that consist of antigens or proteases used in sero- logical assays. The device is intended for use for the presumptive diagnosis of fungal infection. The assay is indicated for use in patients with symptoms of, or medical conditions predisposing the patient to  invasive  fungal  infection. The device can be used as an aid in the diagnosis of deep seated mycoses and fungemias.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Serological Assays for the Detection of Beta-Glucan.’’ See § 866.1(e)  for the availability  of  this  guidance  document.

[69 FR 56936, Sept. 23, 2004]

§ 866.3060   Blastomyces dermatitidis se- rological reagents.

- Identification.	Blastomyces dermatitidis serological reagents are de- vices that consist of antigens and antisera used in serological tests to identify antibodies to Blastomyces determatitidis in serum. The identifica- tion aids in the diagnosis of blasto- mycosis caused by the fungus Blastomyces dermatitidis. Blastomycosis is a chronic granulomatous (tumor- like) disease, which may be limited to the skin or lung or may be widely dis- seminated in the body resulting in le- sions of the bones, liver, spleen, and kidneys.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

#### Food and Drug Administration, HHS	§ 866.3125

§ 866.3065   Bordetella   spp.   serological reagents.

  - Identification. Bordetella spp. sero- logical reagents are devices that con- sist of antigens and antisera, including antisera conjugated with a fluorescent dye, used in serological tests to iden- tify Bordetella spp. from cultured iso- lates or directly from clinical speci- mens. The identification aids in the di- agnosis of diseases caused by bacteria belonging to the genus Bordetella and provides epidemiological information on these diseases. Bordetella spp. cause whooping cough (Bordetella pertussis) and other similiarly contagious and acute respiratory infections character- ized by pneumonitis (inflammation of the lungs).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.3085   Brucella spp. serological re- agents.

- Identification. Brucella spp. sero- logical reagents are devices that con- sist of antigens and antisera used for serological identification  of  Brucella spp. from cultured isolates derived from clinical specimens or to identify antibodies to Brucella spp. in serum. Additionally, some of these reagents consist of antisera conjugated with a fluorescent dye (immunofluorescent re- agents) used to identify Brucella spp. directly from clinical specimens or cul- tured isolates derived from clinical specimens. The  identification  aids  in the diagnosis of brucellosis (e.g., undu- lant fever, Malta fever) caused by bac- teria belonging to the genus Brucella and provides epidemiological informa- tion on diseases caused by these micro- organisms.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

§ 866.3110   Campylobacter   fetus   sero- logical reagents.

- Identification. Campylobacter fetus serological reagents are devices that consist of antisera conjugated with a fluorescent dye used to identify Campylobacter fetus from clinical speci- mens or cultured isolates derived from clinical specimens. The identification aids in the diagnosis of diseases caused by this bacterium and provides epide- miological information on these dis- eases. Campylobacter fetus is a frequent cause of abortion in sheep and cattle and is sometimes responsible for endo- carditis (inflammation of certain mem- branes of the heart) and enteritis (in- flammation of the intestines) in hu- mans.

- Classification. Class I (general con- trols).

##### § 866.3120   Chlamydia    serological   re- agents.

- Identification. Chlamydia sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to chlamydia in  serum.  Additionally, some of these reagents consist of chlamydia antisera conjugated with a fluorescent dye used to identify chlamydia directly from clinical speci- mens or cultured isolates derived from clinical specimens. The identification aids in the diagnosis of disease caused by bacteria belonging to the genus Chlamydia and provides epidemiolog- ical information on these diseases. Chlamydia are the causative agents of psittacosis (a form of pneumonia), lymphogranuloma venereum (a vene- real disease), and trachoma (a chronic disease of the eye and eyelid).

- Classification. Class I (general con- trols).

§ 866.3125   Citrobacter  spp.  serological reagents.

- Identification. Citrobacter spp. sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify Citrobacter spp. from cultured isolates derived from clinical specimens. The identi- fication aids in the diagnosis of disease caused by bacteria belonging to the genus Citrobacter and provides epide- miological   information   on   diseases

#### § 866.3130

caused by  these microorganisms. Citrobacter spp. have occasionally been associated with urinary tract infec- tions.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.3130   Clostridium   difficile   toxin gene amplification assay.

- Identification. A Clostridium difficile toxin gene amplification assay is a device that consists of reagents for the amplification and detection of tar- get sequences in Clostridium difficile toxin genes in fecal specimens from pa- tients suspected of having Clostridium difficile infection (CDI). The detection of clostridial toxin genes, in conjunc- tion with other laboratory  tests,  aids in the clinical laboratory diagnosis of CDI caused by Clostridium difficile.

- Classification. Class II (special controls). The special controls are set forth in FDA’s guideline document en- titled: ‘‘Class II Special Controls Guideline: Toxin Gene Amplification Assays for the Detection of Clostridium difficile; Guideline for Industry and Food and Drug Administration Staff.’’ See § 866.1(e) for information on obtain- ing this document.

[80 FR 51939, Aug. 27, 2015]

§ 866.3135   Coccidioides   immitis   sero- logical reagents.

- Identification. Coccidioides immitis serological reagents are devices that consist of antigens and antisera used in serological tests to identify  antibodies to Coccidioides immitis in serum. The identification aids in the diagnosis of coccidioidomycosis caused by a fungus belonging to the genus Coccidioides and provides  epidemiological  information on diseases caused by this microorga- nism. An infection with Coccidioides immitis produces symptoms varying in severity from those accompanying the common cold to those of influenza.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

#### 21 CFR Ch. I (4–1–17 Edition)

in subpart E  of part  807  of  this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

§ 866.3140   Corynebacterium  spp.  sero- logical reagents.

- Identification. Corynebacterium spp. serological reagents are devices that consist of antisera conjugated with a fluorescent dye used to identify Corynebacterium spp. from clinical specimens. The identification aids in the diagnosis of disease caused by bac- teria belonging to the genus Corynebacterium and provides epidemio- logical information on diseases caused by these microorganisms. The principal human pathogen of this genus, Corynebacterium diphtheriae, causes diphtheria. However, many other types of corynebacteria form part of the nor- mal flora of the human respiratory tract, other mucus membranes, and skin, and are either nonpathogenic or have an uncertain role.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3145   Coxsackievirus	serological reagents.

- Identification. Coxsackievirus se- rological reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to coxsackievirus in serum. Additionally, some of these reagents consist of coxsackievirus antisera conjugated with a fluorescent dye that are used to identify coxsackievirus from clinical specimens or from tissue culture iso- lates derived from clinical specimens. The identification aids in the diagnosis of coxsackievirus infections and pro- vides epidemiological information on diseases caused by these viruses. Coxsackieviruses produce a variety of infections, including common colds, meningitis (inflammation of brain and spinal cord membranes), herpangina (brief fever accompanied by ulcerated lesions of the throat), and myopericarditis (inflammation of heart tissue).

#### Food and Drug Administration, HHS	§ 866.3205

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

§ 866.3165   Cryptococcus neoformans se- rological reagents.

- Identification. Cryptococcus neoformans serological reagents are de- vices that consist of antigens used in serological tests to identify  antibodies to Cryptococcus neoformans in serum. Additionally, some of these reagents consist of antisera conjugated with a fluorescent dye (immunofluorescent re- agents) and are used to identify Cryptococcus neoformans directly from clinical specimens or from cultured isolates derived from clinical speci- mens. The identification aids in the di- agnosis of cryptococcosis and provides epidemiological information on  this type of disease. Cryptococcosis infec- tions are found most often as chronic meningitis (inflammation of brain membranes) and, if not treated, are usually   fatal.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

##### § 866.3175   Cytomegalovirus serological reagents.

- Identification. Cytomegalovirus se- rological reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to cytomegalovirus in serum. The identi- fication aids in the diagnosis of dis- eases caused by cytomegaloviruses (principally cytomegalic inclusion dis- ease) and provides epidemiological in- formation	on	these	diseases. Cytomegalic inclusion disease is a gen- eralized infection of infants and is caused by intrauterine or early post- natal infection with the virus. The dis- ease may cause severe congenital ab- normalities, such as microcephaly (ab- normal smallness of the head), motor disability, and mental retardation. Cytomegalovirus   infection   has   also

been associated with  acquired hemo- lytic anemia, acute and chronic hepa- titis, and an infectious mononucleosis- like syndrome.

- Classification. Class II (perform- ance standards).

§ 866.3200   Echinococcus	spp.	sero- logical reagents.

- Identification. Echinococcus spp. se- rological reagents are devices that con- sist of Echinococcus spp. antigens and antisera used in serological tests to identify antibodies to Echinococcus spp. in serum. The identification aids in the diagnosis of echinococcosis, caused by parasitic tapeworms belonging to the genus Echinococcus and provides epide- miological information on this disease. Echinococcosis is characterized by the development of cysts in the liver, lung, kidneys, and other organs formed by the larva of the infecting organisms.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3205   Echovirus	serological	re- agents.

- Identification. Echovirus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to echovirus in serum. Additionally, some of these reagents consist of echovirus antisera conjugated with a fluorescent dye used to identify echoviruses from clinical specimens or from tissue cul- ture isolates derived from clinical specimens. The identification aids in the diagnosis of echovirus infections and provides epidemiological informa- tion on diseases caused by these vi- ruses. Echoviruses cause illnesses such as meningitis (inflammation of the brain and spinal cord membranes), feb- rile illnesses (accompanied by fever) with or without rash, and the common cold.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 866.3210

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

##### § 866.3210   Endotoxin assay.

  - Identification. An endotoxin assay is a device that uses serological tech- niques in whole blood. The device is in- tended for use in conjunction with other laboratory findings and clinical assessment of the patient to aid in the risk assessment of critically ill pa- tients for progression to severe sepsis.

  - Classification. Class II (special controls). The special control for this device is the FDA guidance entitled ‘‘Class II Special Controls Guidance Document:    Endotoxin    Assay.’’    See

§ 866.1(e)   for   the   availability   of   this guidance document.

[68 FR 62008, Oct. 31, 2003. Redesignated at 70

FR 53069, Sept. 7, 2005]

§ 866.3220   Entamoeba  histolytica  sero- logical reagents.

- Identification.	Entamoeba histolytica serological reagents are de- vices that consist of antigens and antisera used in serological tests to identify antibodies to Entamoeba histolytica in serum. Additionally, some of these  reagents  consist of antisera conjugated with a fluorescent dye (immunofluorescent reagents) used to identify Entamoeba histolytica directly from clinical specimens. The identi- fication aids in the diagnosis of amebi- asis caused by the microscopic proto- zoan parasite Entamoeba histolytica and provides  epidemiological  information on diseases caused  by  this  parasite. The parasite may invade the  skin, liver, intestines, lungs, and diaphragm, causing disease conditions such as in- dolent ulcers, an amebic hepatitis, amebic dysentery, and pulmonary le- sions.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47 FR 50823, Nov. 9, 1982; 47 FR 56846, Dec. 21,

1982, as amended at 63 FR 59226, Nov. 3, 1998]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 866.3225   Enterovirus	nucleic	acid assay.

- Identification. An enterovirus nu- cleic acid assay is a device that con- sists of primers, probes, enzymes, and controls for the amplification and de- tection of enterovirus ribonucleic acid (RNA) in cerebrospinal fluid (CSF) from individuals who have signs and symptoms consistent with meningitis or meningoencephalitis. The detection of enterovirus RNA, in conjunction with other laboratory tests, aids in the clinical laboratory diagnosis of viral meningitis caused by enterovirus.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Nucleic Acid  Amplification  Assay  for the Detection of  Enterovirus  RNA.’’ See § 866.1(e) for the availability of this guidance document.

[74 FR 8, Jan. 2, 2009]

##### § 866.3235   Epstein-Barr	virus	sero- logical reagents.

- Identification. Epstein-Barr virus serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to Epstein-Barr virus in serum. The identification aids in  the diagnosis  of Epstein-Barr virus infections and pro- vides epidemiological information on diseases caused by these viruses. Ep- stein-Barr viruses are thought to cause infectious mononucleosis and have been associated with Burkitt’s lymphoma (a tumor of the jaw in Afri- can children and young adults) and postnasal carcinoma (cancer).

- Classification. Class I (general con- trols).

##### § 866.3240   Equine	encephalomyelitis virus serological reagents.

- Identification.		Equine encephalomyelitis virus serological re- agents are devices that consist of anti- gens and antisera used in serological tests to identify antobodies to equine encephalomyelitis virus in serum. The identification aids in  the diagnosis  of diseases	caused	by		equine encephalomyelitis viruses and provides epidemiological  information  on  these

#### Food and Drug Administration, HHS	§ 866.3280

viruses. Equine encephalomyelitis vi- ruses are transmitted to humans by the bite of insects, such as mosquitos and ticks, and may cause encephalitis (inflammation of the brain), rash, acute arthritis, or hepatitis.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

§ 866.3250   Erysipelothrix rhusiopathiae

##### serological reagents.

- Identification. Erysipelothrix rhusiopathiae serological reagents are devices that consist of antigens and antisera used in serological tests to identify Erysipelothrix rhusiopathiae from cultured isolates derived from clinical specimens. The identification aids in the diagnosis of disease caused by this bacterium belonging to the genus Erysipelothrix. This organism is responsible for a variety of inflamma- tions of the skin following skin abra- sions from contact with fish, shellfish, or poultry.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

§ 866.3255   Escherichia  coli  serological reagents.

- Identification. Escherichia coli sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify Escherichia coli from cultured isolates derived from clinical specimens. Additionally, some of these reagents consist of Escherichia coli antisera conjugated with a fluores- cent dye used to identify Escherichia coli directly from clinical specimens or cultured isolates derived from clinical specimens. The identification aids in the diagnosis of diseases caused by this bacterium belonging to the genus Esch- erichia, and provides epidemiological information on diseases caused by this microorganism. Although  Escherichia coli constitutes the greater part of the

microorganisms found in the intestinal tract in humans and is usually non- pathogenic, those strains which are pathogenic may cause urinary tract in- fections or epidemic diarrheal disease, especially in children.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38791, July 25,

2001]

§ 866.3270   Flavobacterium   spp.   sero- logical reagents.

- Identification. Flavobacterium spp. serological reagents are devices that consist of antigens and antisera used in serological	tests	to	identify Flavobacteriuim spp. from cultured iso- lates derived from clinical specimens. The identification aids in the diagnosis of disease caused by bacteria belonging to the genus Flavobacterium and pro- vides epidemiological information on diseases caused by these microorga- nisms. Most members of this genus are found in soil and water and, under cer- tain conditions, may become patho- genic to humans. Flavobacterium meningosepticum is highly virulent for the newborn, in whom it may cause epidemics of septicemia (blood poi- soning) and meningitis (inflammation of the membranes of the brain) and is usually attributable to contaminated hospital equipment.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25046, June 12, 1989; 66 FR 38792, July 25,

2001]

§ 866.3280   Francisella  tularensis  sero- logical reagents.

- Identification. Francisella tularensis serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to Francisella tularensis in serum or to identify Francisella tularensis in cul- tured isolates derived from clinical specimens. Additionally, some of these reagents consist of antisera conjugated

#### § 866.3290

with	a	fluorescent	dye (immunofluorescent reagents) used to identify Francisella tularensis directly from clinical specimens. The identi- fication aids in the diagnosis of tula- remia caused by Francisella tularensis and provides epidemiological informa- tion on this disease. Tularemia is a desease principally of rodents, but may be transmitted to humans through handling of infected animals, animal products, or by the bites of fleas and ticks. The disease takes on several forms depending upon the site of infec- tion, such as skin lesions, lymph node enlargements, or pulmonary infection.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

##### § 866.3290   Gonococcal	antibody	test (GAT).

- Identification. A gonococcal anti- body test (GAT) is an in vitro device that consists of the reagents intended to identify by immunochemical tech- niques, such as latex agglutination, in- direct fluorescent antibody, or radioimmunoassay, antibodies to Neisseria gonorrhoeae in sera of asymp- tomatic females at low risk of infec- tion. Identification of antibodies with this device may indicate past  or present infection of the patient with Neisseria gonorrhoeae.

- Classification. Class III (premarket approval) (transitional device).

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 866.3.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  52

FR 17734, May 11, 1987]

§ 866.3300   Haemophilus	spp.	sero- logical reagents.

- Identification. Haemophilus spp. se- rological reagents are devices that con- sist of antigens and antisera, including antisera conjugated with a fluorescent dye, that are used in serological tests to identify Haemophilus spp. directly from clinical specimens or tissue cul- ture   isolates   derived   from   clinical

#### 21 CFR Ch. I (4–1–17 Edition)

specimens. The identification aids in the diagnosis of diseases caused by bac- teria belonging to the genus Haemophilus and provides epidemiolog- ical information on diseases cause by these microorganisms. Diseases most often caused by Haemophilus spp. in- clude pneumonia, pharyngitis, sinus- itis, vaginitis, chancroid venereal dis- ease, and a contagious form of conjunc- tivitis (inflammation of eyelid mem- branes).

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59226, Nov. 3, 1998]

##### § 866.3305   Herpes  simplex  virus  sero- logical assays.

- Identification. Herpes simplex virus serological assays are  devices that consist of antigens and antisera used in various serological tests to identify antibodies to herpes simplex virus in serum. Additionally, some of the assays consist of herpes simplex virus antisera conjugated with a fluo- rescent dye (immunofluorescent as- says) used to identify herpes simplex virus directly from clinical  specimens or tissue culture isolates derived from clinical specimens. The identification aids in the diagnosis of diseases caused by herpes simplex viruses and provides epidemiological information on these diseases. Herpes simplex viral infec- tions range from common and mild le- sions of the skin and mucous mem- branes to a severe form of encephalitis (inflammation of the brain). Neonatal herpes virus infections range from a mild infection to a severe generalized disease with a fatal outcome.

- Classification. Class II (special controls). The device is  classified  as class II (special controls). The special control for the device  is FDA’s revised guidance document entitled ‘‘Class II Special Controls Guidance Document: Herpes Simplex Virus Types 1 and 2 Se- rological Assays.’’  For  availability  of the   guidance   revised   document,   see

§ 866.1(e).

[72  FR  15830,  Apr.  3,  2007,  as  amended  at  74

FR  42775,  Aug.  25,  2009;  76  FR  48717,  Aug.  9,

2011]

#### Food and Drug Administration, HHS	§ 866.3328

##### § 866.3310   Hepatitis A virus (HAV) se- rological assays.

  - Identification. HAV serological as- says are devices that consist of anti- gens and antisera for the detection of hepatitis A  virus-specific IgM,  IgG, or total antibodies (IgM and IgG), in human serum or plasma. These devices are used for testing specimens from in- dividuals who have signs and symp- toms consistent with acute hepatitis to determine if an individual has been previously infected with HAV, or as an aid to identify HAV-susceptible indi- viduals. The detection of these anti- bodies aids in the clinical laboratory diagnosis of an acute or past infection by HAV in conjunction with other clin- ical laboratory findings. These devices are not intended for screening blood or solid or soft tissue donors.

  - Classification. Class II (special controls). The special control is ‘‘Guidance for Industry and FDA Staff: Class II Special Controls Guidance Document: Hepatitis A Virus Sero- logical Assays.’’ See § 866.1(e) for the availability  of  this  guidance  document.

[71 FR 6679, Feb. 9, 2006]

§ 866.3320   Histoplasma capsulatum se- rological reagents.

- Identification.	Histoplasma capsulatum serological reagents are de- vices that consist of antigens and antisera used in serological tests to identify antibodies to Histoplasma capsulatum in  serum.  Additionally, some of these reagents consist of Histoplasma capsulatum antisera con- jugated with a fluorescent dye (immunofluorescent reagents) used to identify Histoplasma capsulatum from clinical specimens or cultured isolates derived from clinical specimens. The identification aids in the diagnosis of histoplasmosis caused by this fungus belonging to the genus Histoplasma and provides  epidemiological  information on the diseases caused by this fungus. Histoplasmosis usually is a mild and often asymptomatic respiratory infec- tion, but in a small number of infected individuals the lesions may spread to practically all tissues and organs.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

in subpart E  of part  807  of  this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.3328   Influenza  virus  antigen  de- tection test system.

- Identification. An influenza virus antigen detection test system is a de- vice intended for the qualitative detec- tion of influenza viral antigens directly from clinical specimens in patients with signs and symptoms of res- piratory infection. The test aids in the diagnosis of influenza infection and provides epidemiological information on influenza. Due to the propensity of the virus to mutate, new strains emerge over time which may poten- tially affect the  performance of these devices. Because influenza is highly contagious and  may  lead  to  an  acute respiratory tract infection causing se- vere illness and even death, the accu- racy of these devices has serious public health implications.

- Classification. Class II (special controls). The special controls for this device are:

- The device’s sensitivity and speci- ficity performance characteristics or positive percent agreement and nega- tive percent agreement, for each speci- men type claimed in the intended use of the device, must meet one of the fol- lowing two minimum clinical perform- ance criteria:

  - For devices evaluated as compared to an FDA-cleared nucleic acid based- test or other currently appropriate and FDA accepted comparator method other than correctly performed viral culture method:

- The positive percent agreement estimate for the device when testing for influenza A and influenza B must be at the point estimate of at least 80 per- cent with a lower bound of the 95 per- cent confidence interval that is greater than or equal to 70 percent.

- The negative percent agreement estimate for the device when testing for influenza A and influenza B must be at the point estimate of at least 95 per- cent with a lower bound of the 95 per- cent confidence interval that is greater than or equal to 90 percent.

#### § 866.3328

  - For devices evaluated as com- pared to correctly performed viral cul- ture method as the comparator meth- od:

    - The sensitivity estimate for the device when testing for influenza  A must be at the point estimate of at least 90 percent with a lower bound of the 95 percent confidence interval that is greater than or equal to 80 percent. The sensitivity estimate for the device when testing for influenza B must be at the point estimate of at least 80 per- cent with a lower bound of the 95 per- cent confidence interval that is greater than or equal to 70 percent.

    - The specificity estimate for the device when testing for influenza A and influenza B must be at the point esti- mate of at least 95 percent with a lower bound of the 95 percent confidence in- terval that is greater than  or  equal  to 90 percent.

- When performing testing to dem- onstrate the device meets the require- ments in paragraph (b)(1) of this sec- tion, a currently appropriate and FDA accepted comparator method must be used to establish assay performance in clinical  studies.

- Annual analytical reactivity test- ing of the device must be performed with contemporary influenza strains. This annual analytical reactivity test- ing must meet the following criteria:

  - The appropriate strains to be test- ed will be identified by FDA in con- sultation with the Centers for Disease Control and Prevention (CDC) and sourced from CDC or an FDA-des- ignated source. If the annual strains are not available from CDC, FDA will identify an alternative source for ob- taining the requisite strains.

  - The testing must be conducted according to a standardized protocol considered and determined by FDA  to be acceptable and appropriate.

  - By July 31 of each calendar year, the results of the last 3 years of annual analytical reactivity testing must be included as part of the device’s label- ing. If a device has not been on the market long enough for 3 years of an- nual analytical reactivity testing to have been conducted since the device received marketing authorization from FDA, then the results of every annual analytical  reactivity  testing  since  the

#### 21 CFR Ch. I (4–1–17 Edition)

device received marketing authoriza- tion from FDA must be included. The results must be presented as part of the device’s labeling in a tabular format, which includes the detailed informa- tion for each virus tested as described in the certificate of authentication, ei- ther by:

- Placing the results directly in the device’s § 809.10(b) of this chapter com- pliant labeling that physically accom- panies the device in a separate section of the labeling where the analytical re- activity testing data can be found; or

- In the device’s label or in other labeling that physically accompanies the device, prominently providing a hyperlink to the manufacturer’s public Web site where the analytical reac- tivity testing data can be found. The manufacturer’s home page, as well as the primary part of the manufacturer’s Web site that discusses the device, must provide a prominently placed hyperlink to the Web page containing this information and must allow unre- stricted viewing access.

- If one of the actions listed at sec- tion 564(b)(1)(A)–(D) of the Federal Food, Drug, and Cosmetic Act occurs with respect to  an  influenza  viral strain,  or  if  the  Secretary  of  Health and Human Services (HHS) determines, under section 319(a) of the Public Health Service Act, that a disease or disorder presents a public health emer- gency, or that a public health emer- gency otherwise exists, with respect to an influenza viral strain:

  - Within 30 days from the date that FDA notifies manufacturers that char- acterized viral  samples  are  available for test evaluation, the manufacturer must have testing performed on the de- vice with those viral samples in accord- ance with a standardized protocol con- sidered and determined by FDA to be acceptable and appropriate. The proce- dure and location of testing may de- pend on the nature of the emerging virus.

  - Within 60 days from the date that FDA notifies manufacturers that char- acterized viral  samples  are  available for  test  evaluation  and  continuing until 3 years from that date, the re- sults of the influenza emergency ana- lytical  reactivity  testing,  including  the

#### Food and Drug Administration, HHS	§ 866.3336

detailed information for the virus test- ed as described in the certificate of au- thentication, must be included as part of the device’s labeling in a tabular for- mat, either by:

    - Placing the results directly in the device’s § 809.10(b) of this chapter com- pliant labeling that physically accom- panies the device in a separate section of the labeling where analytical reac- tivity testing data can be found, but separate from the annual analytical re- activity  testing  results;  or

    - In a section of the device’s label or in other labeling that physically ac- companies the device, prominently pro- viding a hyperlink to the manufactur- er’s public Web site where the analyt- ical reactivity testing data can be found. The manufacturer’s home page, as well as the primary part of the man- ufacturer’s Web site that discusses the device, must provide a prominently placed hyperlink to the Web page con- taining this information and must allow unrestricted viewing access.

[82 FR 3618, Jan. 12, 2017]

##### § 866.3330   Influenza   virus   serological reagents.

      - Identification. Influenza virus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to influenza in serum. The identification aids in the diagnosis of influenza (flu) and provides epidemiological informa- tion on influenza. Influenza is an acute respiratory tract disease, which is often epidemic.

      - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.3332   Reagents   for   detection   of specific novel influenza A viruses.

- Identification. Reagents for detec- tion of specific novel influenza A vi- ruses are devices that are intended for use in a nucleic acid amplification test to directly detect specific virus RNA in human respiratory specimens or viral cultures. Detection of specific virus RNA aids in the diagnosis of influenza

caused by specific novel influenza A vi- ruses in patients with clinical risk of infection with these viruses, and also aids in the presumptive laboratory identification of specific novel influ- enza A viruses to provide epidemiolog- ical information on influenza. These reagents include primers, probes, and specific influenza A virus controls.

- Classification. Class II (special controls). The special controls are:

- FDA’s guidance document entitled ‘‘Class II Special Controls  Guidance Document: Reagents for Detection of Specific Novel Influenza  A  Viruses.’’ See § 866.1(e) for information on obtain- ing this document.

- The distribution of these devices is limited to laboratories with experi- enced personnel who have training in standardized molecular testing proce- dures and expertise in viral diagnosis, and appropriate biosafety equipment and containment.

[71 FR 14379, Mar. 22, 2006]

##### § 866.3336   John Cunningham Virus se- rological reagents.

  - Identification. John Cunningham Virus serological reagents are devices that consist of antigens and antisera used in serological assays to identify antibodies to  John Cunningham  Virus in serum and plasma. The identifica- tion aids in the risk stratification  for the development of progressive multifocal leukoencephalopathy in multiple sclerosis and Crohn’s disease patients undergoing natalizumab ther- apy. These devices are for adjunctive use, in the context of other clinical risk factors for the development of pro- gressive multifocal leukoenceph- alopathy.

  - Classification. Class II (special controls). The special control for this device is the FDA guideline document entitled ‘‘Class II Special Controls Guideline: John Cunningham Virus Se- rological   Reagents.’’   For   availability of      the      guideline      document,      see

§ 866.1(e).

[79 FR 3740, Jan. 23, 2014]

#### § 866.3340

§ 866.3340   Klebsiella   spp.   serological reagents.

    - Identification. Klebsiella spp. sero- logical reagents are devices that con- sist of antigens and antisera, including antisera conjugated with a fluorescent dye  (immunofluorescent   reagents), that are used in serological tests to identify Klebsiella spp. from cultured isolates derived from clinical speci- mens. The identification aids in the di- agnosis of diseases caused by bacteria belonging to the genus Klebsiella and provides  epidemiological  information on these diseases. These organisms can cause serious urinary tract and pul- monary infections, particularly in hos- pitalized  patients.

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

§ 866.3350   Leptospira  spp.  serological reagents.

- Identification. Leptospira spp. sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to Leptospira spp. in serum or identify Leptospira spp. from cultured isolates derived from clinical specimens. Addi- tionally, some of these antisera are conjugated with a fluorescent dye (immunofluorescent reagents) and used to identify  Leptospira  spp.  directly from clinical specimens. The identi- fication aids in the diagnosis of lepto- spirosis caused by bacteria  belonging to the genus Leptospira and provides ep- idemiological information on this dis- ease. Leptospira infections range from mild fever-producing illnesses to severe liver and kidney involvement pro- ducing hemorrhage and dysfunction of these organs.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 866.3355   Listeria  spp.  serological  re- agents.

- Identification. Listeria spp. sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify Listeria spp. from cultured isolates derived from clinical specimens. Additionally, some of these reagents consist of Listeria spp. antisera conjugated with a fluorescent dye (immunofluorescent reagents) used to identify Listeria spp. directly from clinical specimens. The identification aids in the diagnosis of listeriosis, a disease caused by bacteria belonging to the genus Listeria, and provides epide- miological information on diseases caused by these microorganisms. Lis- teria monocytogenes, the most common human pathogen of this genus, causes meningitis (inflammation of the brain membranes) and meningoencephalitis (inflammation of the brain and brain membranes) and is often fatal if un- treated. A second form of human listeriosis is an intrauterine  infection in pregnant women that results in a high mortality  rate  for  infants  before or after birth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3360   Lymphocytic	choriomenin- gitis virus serological reagents.

- Identification. Lymphocytic cho- riomeningitis virus serological re- agents are devices that consist of anti- gens and antisera used in serological tests to identify antibodies to lymphocytic choriomeningitis virus in serum. The identification aids in the diagnosis of lymphocytic choriomenin- gitis virus infections and provides epi- demiological information on diseases caused by these viruses. Lymphocytic choriomeningitis viruses usually cause a mild cerebral meningitis (inflamma- tion of membranes that envelop the brain) and occasionally a mild pneu- monia, but in rare instances may produce severe and even fatal illnesses due to complications from cerebral meningitis and pneumonia.

#### Food and Drug Administration, HHS	§ 866.3373

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3365 Multiplex nucleic acid assay for identification of microorga- nisms and resistance markers from positive blood cultures.

- Identification. A multiplex nucleic acid assay for identification of micro- organisms and resistance markers from positive blood cultures is a qualitative in vitro device intended to simulta- neously detect and identify microorga- nism nucleic acids from blood cultures that test positive by Gram stain or other microbiological stains. The de- vice detects specific nucleic acid se- quences for microorganism identifica- tion as well as for antimicrobial resist- ance. This device aids in the diagnosis of bloodstream infections when used in conjunction with other clinical and laboratory findings. However, the de- vice does not replace traditional meth- ods for culture and susceptibility test- ing.

- Classification. Class II (special controls). The special control for this device is FDA’s guideline document en- titled ‘‘Class II Special Controls Guide- line: Multiplex Nucleic Acid Assay for Identification of Microorganisms and Resistance Markers  from  Positive Blood  Cultures.’’   For   availability   of the guideline document, see § 866.1(e).

[80 FR 30154, May 27, 2015]

§ 866.3370   Mycobacterium   tuberculosis

##### immunofluorescent reagents.

- Identification. Mycobacterium tu- berculosis immunofluorescent reagents are devices that consist of antisera conjugated with a fluorescent dye used to identify Mycobacterium tuberculosis directly from clinical specimens. The identification aids in the diagnosis of tuberculosis and provides epidemiolog- ical information on this disease. Mycobacterium tuberculosis is the com- mon causative organism in human tu- berculosis, a chronic infectious disease characterized by formation of tubercles (small rounded nodules) and tissue ne-

crosis (destruction), usually occurring in the lung.

- Classification. Class I (general con- trols).

##### § 866.3372 Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis com- plex in respiratory specimens.

- Identification. Nucleic acid-based in vitro diagnostic devices for the de- tection of Mycobacterium tuberculosis complex in respiratory specimens are qualitative nucleic acid-based in vitro diagnostic devices intended to detect Mycobacterium tuberculosis complex nu- cleic acids extracted from human res- piratory specimens. These devices are non-multiplexed and intended to be used as an aid in the diagnosis of pul- monary tuberculosis when used in con- junction with clinical and other labora- tory findings. These devices do not in- clude devices intended to detect the presence of organism mutations associ- ated with drug resistance. Respiratory specimens may include sputum (in- duced or expectorated), bronchial specimens (e.g., bronchoalveolar lavage or bronchial aspirate), or tracheal aspi- rates.

- Classification. Class II (special controls). The special control for this device is the FDA document entitled ‘‘Class II Special Controls Guideline: Nucleic Acid-Based In Vitro Diagnostic Devices for the Detection of Mycobacterium tuberculosis Complex in Respiratory Specimens.’’ For avail- ability  of  the  guideline  document,  see

§ 866.1(e).

[79 FR 31027, May 30, 2014]

##### § 866.3373 Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis com- plex (MTB-complex) and the genetic mutations associated with MTB- complex antibiotic resistance in respiratory specimens.

- Identification. Nucleic acid-based in vitro diagnostic devices for the de- tection of Mycobacterium tuberculosis complex (MTB-complex) and the ge- netic mutations associated with MTB- complex antibiotic resistance in res- piratory specimens are qualitative nu- cleic acid-based devices that detect the presence of MTB-complex-associated nucleic  acid  sequences  in  respiratory

#### § 866.3373

samples. These devices are intended to aid in the diagnosis of pulmonary tu- berculosis and the selection of an ini- tial treatment regimen when used in conjunction with clinical findings and other laboratory results. These devices do not provide confirmation of anti- biotic susceptibility since other mech- anisms of resistance may exist that may be associated with a lack of clin- ical response to treatment other than those detected by the device.

- Classification. Class II (special controls). The special controls for this device are:

  - The FDA document  entitled ‘‘Class II Special Controls Guideline: Nucleic Acid-Based In Vitro Diagnostic Devices for the Detection of Mycobacterium tuberculosis Complex and Genetic Mutations Associated with An- tibiotic Resistance in Respiratory Specimens,’’ which addresses the miti- gation of risks specific to the detection of MTB-complex.  For  availability  of the document, see § 866.1(e).

  - The following items, which ad- dress the mitigation of risks specific to the detection of the genetic mutations associated with antibiotic resistance of MTB-complex:

- The device must include an exter- nal positive assay control as appro- priate. Acceptable positive assay con- trols include MTB-complex isolates containing one or more antibiotic-re- sistance associated target sequences detected by the device.

- The device must include internal controls as appropriate. An acceptable internal control may include human nucleic acid co-extracted with MTB- complex containing nucleic acid se- quences associated with antibiotic re- sistance and primers amplifying human housekeeping genes (e.g., RNaseP, b- actin).

- The device’s intended use must include a description of the scope of antibiotic resistance targeted by the assay, i.e., the specific  drugs  and/or drug classes.

- The specific performance charac- teristics section of the device’s label- ing must include information regarding the specificity of the assay oligonucleotides	for	detecting mutations associated with antibiotic resistance  of  MTB-complex,  and  any

#### 21 CFR Ch. I (4–1–17 Edition)

information indicating the  potential for non-specific binding (e.g., BLAST search).

- In demonstrating device perform- ance you must perform:

  - Pre-analytical studies that evalu- ate:

    - Frozen samples. If there is use of any frozen samples in the device per- formance studies, or if there is a device claim for the use of frozen samples for testing, the effect of freezing samples prior to testing and the effect of mul- tiple freeze/thaw cycles on both anti- biotic susceptible and antibiotic resist- ant strains of MTB-complex.

    - Nucleic acid extraction methods. Ex- traction methods must parallel those used in devices for the detection of MTB-complex nucleic acid and confirm that the detection of the genetic mutations associated with antibiotic resistance is not affected.

  - Analytical studies that analyze:

    - Limit of Detection. Limit of Detec- tion must be determined in the most challenging matrix (e.g., sputum) claimed for use with the device. The Limit of Detection must be determined using both antibiotic susceptible and antibiotic resistant strains of MTB- complex. The antibiotic resistant strains must be those with well charac- terized genetic mutations associated with antibiotic resistance.

    - Analytical Reactivity (Inclusivity). Testing must be conducted to evaluate the ability of the device to detect ge- netic mutations associated with anti- biotic resistance in a diversity of MTB- complex strains. Isolates used in test- ing must be well characterized. Isolate strain characterization must be deter- mined using standardized reference methods recognized by a reputable sci- entific body and appropriate to the strain lineage.

    - Within-Laboratory (Repeatability) Precision Testing. Within-laboratory precision studies, if appropriate, must include at least one antibiotic resist- ant and one antibiotic susceptible strain of MTB-complex.

    - Between Laboratory Reproducibility Testing. The protocol for the reproduc- ibility study may vary slightly depend- ing  on  the  assay  format;  however,  the

#### Food and Drug Administration, HHS	§ 866.3390

panel must include at least one anti- biotic resistant and one antibiotic sus- ceptible strain of MTB-complex.

  - Clinical Studies. Clinical per- formance of the device must be estab- lished by conducting prospective clin- ical studies that include subjects with culture confirmed active tuberculosis. Studies must attempt to enroll sub- jects at risk for antibiotic-resistant MTB-complex; however, it may be nec- essary to include supplemental anti- biotic resistant retrospective and con- trived samples. Clinical studies must compare device results to both phenotypic drug susceptibility testing and genotypic reference methods. The genotypic reference method must be a polymerase chain reaction based meth- od that uses primers different from those in the experimental device and confirmed by bidirectional sequencing.

[79 FR 63036, Oct. 22, 2014]

§ 866.3375   Mycoplasma spp. serological reagents.

- Identification. Mycoplasma spp. se- rological reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to Mycoplasma spp. in serum. Addition- ally, some of these reagents consist of Mycoplasma spp. antisera conjugated with	a	fluorescent	dye (immunofluorescent reagents) used to identify Mycoplasma spp. directly from clinical specimens. The identification aids in the diagnosis of disease caused by bacteria belonging to the genus Mycoplasma and provides epidemiolog- ical information on diseases caused by these microorganisms. Mycoplasma spp. are associated with inflammatory con- ditions of the urinary and respiratory tracts, the  genitals,  and  the  mouth. The effects in humans of infection with Mycoplasma pneumoniae range from inapparent infection to mild or severe upper respiratory disease, ear infec- tion, and bronchial pneumonia.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3380   Mumps virus serological re- agents.

- Identification. Mumps virus sero- logical reagents consist of antigens and antisera used in serological tests to identify antibodies to mumps virus in serum. Additionally, some of these re- agents consist of antisera conjugated with	a	fluorescent	dye (immunofluorescent reagents) used in serological tests to identify mumps vi- ruses from tissue culture isolates de- rived from clinical specimens. The identification aids in the diagnosis of mumps and provides epidemiological information on mumps. Mumps is an acute  contagious  disease,  particularly in children, characterized by an en- largement of one or both of the parotid glands (glands situated near the ear), although other  organs  may  also be  in- volved.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2311, Jan. 14, 2000]

##### § 866.3390   Neisseria  spp.  direct  sero- logical test reagents.

- Identification. Neisseria spp. direct serological test reagents are  devices that consist of antigens and antisera used in serological tests to identify Neisseria spp. from cultured isolates. Additionally, some of these reagents consist of Neisseria spp. antisera con- jugated with a fluorescent dye (immunofluorescent reagents) which may be used to detect the presence of Neisseria spp. directly from  clinical specimens. The  identification  aids  in the diagnosis of disease caused by bac- teria belonging to the genus Neisseria, such as epidemic cerebrospinal menin- gitis, meningococcal disease, and gon- orrhea, and also provides epidemiolog- ical information on diseases caused by these microorganisms. The device does not include products for  the  detection of gonorrhea in humans by indirect methods, such as detection of anti- bodies or of oxidase produced by gono- coccal  organisms.

- Classification. Class II (perform- ance standards).

#### § 866.3395

##### § 866.3395   Norovirus	serological	re- agents.

  - Identification. Norovirus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to detect the presence of norovirus antigens in fecal samples. These devices aid in the diagnosis of norovirus infection in the setting of an individual patient with symptoms of acute gastroenteritis when the indi- vidual patient is epidemiologically linked to other patients with symp- toms of acute gastroenteritis and/or aid in the identification of norovirus as the etiology of an outbreak of acute gastroenteritis in the setting of epidemiologically linked patients with symptoms of acute gastroenteritis.

  - Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Norovirus   Serological   Reagents.’’   See

§ 866.1(e)   for   the   availability   of   this guidance document.

[76 FR 14274, Mar. 9, 2012]

##### § 866.3400   Parainfluenza    virus    sero- logical reagents.

- Identification. Parainfluenza virus serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to parainfluenza virus in serum. The identification aids in  the diagnosis  of parainfluenza virus infections and pro- vides epidemiological information on diseases caused by these viruses. Parainfluenza viruses cause a  variety of respiratory illnesses ranging from the common cold to pneumonia.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.3402   Plasmodium species antigen detection assays.

- Identification. A Plasmodium spe- cies antigen detection assay is a device that employs antibodies for the detec- tion of specific malaria parasite anti- gens, including histidine-rich protein-2

#### 21 CFR Ch. I (4–1–17 Edition)

(HRP2) specific antigens, and pan ma- larial antigens in human whole blood. These devices are used for testing specimens from individuals who have signs and symptoms consistent with malaria infection. The detection of these antigens aids in the clinical lab- oratory diagnosis of malaria caused by the four malaria species capable of in- fecting humans: Plasmodium falciparum, Plasmodium vivax, Plasmodium ovale, and Plasmodium malariae,  and  aids  in the differential diagnosis of Plasmodium falciparum infections from other less virulent Plasmodium species. The device is intended for use in conjunction with other clinical laboratory findings.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Plasmodium species Antigen Detection Assays.’’ See § 866.1(e) for the  avail- ability of this guidance document.

[73 FR 29054, May 20, 2008]

##### § 866.3405   Poliovirus	serological	re- agents.

- Identification. Poliovirus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to poliovirus in serum. Additionally, some of these reagents consist of poliovirus antisera conjugated with a fluorescent dye (immunofluorescent reagents) used to identify polioviruses from clinical specimens or from tissue culture iso- lates derived from clinical specimens. The identification aids in the diagnosis of poliomyelitis (polio) and provides epidemiological information on this disease. Poliomyelitis is an acute in- fectious disease which in its serious form affects the central nervous sys- tem resulting in atrophy  (wasting away) of groups of muscles, ending in contraction and permanent deformity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 866.3480

##### § 866.3410   Proteus spp. (Weil-Felix) se- rological reagents.

  - Identification. Proteus spp. (Weil- Felix) serological reagents are devices that consist of antigens and antisera, including antisera conjugated with a fluorescent dye (immunofluorescent re- agents), derived from the bacterium Proteus vulgaris used in agglutination tests (a specific type of antigen-anti- body reaction) for the detection of antibodies to rickettsia (virus-like  bac- teria) in serum. Test results aid in the diagnosis of diseases caused by bacteria belonging to the genus Rickettsiae  and provide epidemiological information on these diseases. Rickettsia are generally transmitted by arthropods (e.g., ticks and mosquitoes) and produce infections in humans characterized by rash and fever (e.g., typhus  fever,  spotted  fever, Q fever, and trench fever).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

§ 866.3415   Pseudomonas	spp.	sero- logical reagents.

- Identification. Pseudomonas spp. se- rological reagents are devices that con- sist of antigens and antisera, including antisera conjugated with a fluorescent dye  (immunofluorescent  reagents), used to identify Pseudomonas spp. from clinical specimens or from cultured isolates derived from clinical speci- mens. The identification aids in the di- agnosis of disease caused by bacteria belonging to the genus Pseudomonas. Pseudomonas aeruginosa is a  major cause of hospital-acquired infections, and has been associated with urinary tract infections, eye  infections,  burn and wound infections, blood poisoning, abscesses, and meningitis (inflamma- tion of brain membranes). Pseudomonas pseudomallei causes melioidosis, a chronic pneumonia.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

in subpart E  of part  807  of  this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.3460   Rabiesvirus	immuno- fluorescent reagents.

- Identification.	Rabiesvirus immunofluorescent reagents are de- vices that consist of  rabiesvirus antisera conjugated with a fluorescent dye used to identify rabiesvirus in specimens taken from suspected rabid animals. The identification aids in the diagnosis of rabies in patients exposed by animal bites and provides epidemio- logical information on rabies. Rabies is an acute infectious disease of the cen- tral nervous system which, if undiagnosed, may be fatal. The disease is commonly transmitted to humans by a bite from a rabid animal.

- Classification. Class II (perform- ance standards).

##### § 866.3470   Reovirus	serological	re- agents.

- Identification. Reovirus serological reagents are devices that consist of antigens and antisera used in sero- logical tests to identify antibodies to reovirus in serum. The identification aids in the diagnosis of reovirus infec- tions and provides epidemiological in- formation on diseases caused by these viruses. Reoviruses are thought to cause only mild respiratory and gastro- intestinal illnesses.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.3480   Respiratory  syncytial  virus serological reagents.

- Identification. Respiratory syncytial virus serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to respiratory syncytial virus in serum. Additionally, some of these reagents consist of res- piratory syncytial virus antisera con- jugated with a fluorescent dye (immunofluorescent reagents) and used

#### § 866.3490

to identify respiratory syncytial vi- ruses from clinical specimens or from tissue culture isolates derived from clinical specimens. The identification aids in the diagnosis of respiratory syncytial virus infections and provides epidemiological information on dis- eases caused by these viruses. Res- piratory syncytial viruses cause a number of respiratory tract infections, including the common cold, pharyn- gitis,  and  infantile  bronchopneumonia.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.3490   Rhinovirus   serological   re- agents.

- Identification. Rhinovirus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to rhinovirus in serum. The identification aids in the diagnosis of rhinovirus in- fections and provides epidemiological information on diseases caused by these viruses. Rhinoviruses cause com- mon colds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.3500   Rickettsia	serological	re- agents.

- Identification. Rickettsia sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to rickettsia in serum. Additionally,  some of these reagents consist of rickettsial antisera conjugated with a fluorescent dye (immunofluorescent reagents) used to identify rickettsia directly from clinical specimens. The identification aids in the diagnosis of diseases caused by virus-like bacteria belonging to the genus Rickettsiae and provides epide- miological information on these dis- eases. Rickettsia are generally trans- mitted  by  arthropods  (e.g.,  ticks  and

#### 21 CFR Ch. I (4–1–17 Edition)

mosquitoes) and produce infections in humans characterized  by  rash  and fever (e.g., typhus  fever,  spotted  fever, Q fever, and trench fever).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.3510   Rubella virus serological re- agents.

- Identification. Rubella virus sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to rubella virus in serum. The identifica- tion aids in the diagnosis of rubella (German measles) or confirmation of a person’s immune status from past in- fections or immunizations and provides epidemiological information on Ger- man measles. Newborns infected in the uterus with rubella virus may be born with multiple congenital defects (ru- bella syndrome).

- Classification. Class II. The special controls for this device are:

  - National Committee for Clinical Laboratory Standards’:

    - 1/LA6 ‘‘Detection and Quantita- tion of Rubella IgG Antibody: Evalua- tion and Performance Criteria for Mul- tiple Component  Test Products, Speciment Handling, and Use of the Test Products in the Clinical Labora- tory, October 1997,’’

    - 1/LA18 ‘‘Specifications for Immunological Testing for Infectious Diseases, December 1994,’’

    - D13 ‘‘Agglutination Characteris- tics, Methodology, Limitations, and Clinical Validation, October 1993,’’

    - EP5 ‘‘Evaluation of Precision Performance of Clinical Chemistry De- vices, February 1999,’’ and

    - EP10 ‘‘Preliminary Evaluation of the Linearity of Quantitive Clinical Laboratory Methods, May 1998,’’

  - Centers for Disease Control’s:

    - Low Titer Rubella Standard,

    - Reference Panel of Well Charac- terized Rubella Sera, and

#### Food and Drug Administration, HHS	§ 866.3630

  - World Health Organization’s International Rubella Standard.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  52

FR 17734,  May 11,  1987;  65  FR 17144,  Mar. 31,

2000]

##### § 866.3520   Rubeola  (measles)  virus  se- rological reagents.

- Identification. Rubeola (measles) virus serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to rubeola virus in serum. The identification aids in the diagnosis of measles and provides epidemiolog- ical information on the disease.  Mea- sles is an acute, highly infectious dis- ease of the respiratory and reticuloendothelial tissues, particu- larly in children, characterized by a confluent and blotchy rash.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

§ 866.3550   Salmonella  spp.  serological reagents.

- Identification. Salmonella spp. sero- logical  reagents  are  devices  that  con- sist of antigens and antisera used in se- rological   tests   to   identify   Salmonella spp.   from   cultured   isolates   derived from   clinical   specimens.   Additionally, some    of    these    reagents    consist    of antisera  conjugated  with  a  fluorescent dye (immunofluorescent reagents) used to    identify    Salmonella   spp.    directly from   clinical   specimens   or   cultured isolates   derived   from   clinical   speci- mens. The identification aids in the di- agnosis of salmonellosis caused by bac- teria belonging to the genus Salmonella and  provides  epidemiological  informa- tion  on  this  disease.  Salmonellosis  is characterized by high grade fever (‘‘en- teric    fever’’),    severe    diarrhea,    and cramps.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

in subpart E  of part  807  of  this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

§ 866.3600   Schistosoma spp. serological reagents.

- Identification. Schistosoma spp. se- rological reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify antibodies to Schistosoma spp. in serum. The identi- fication aids in the diagnosis of schis- tosomiasis caused by parasitic flatworms of the genus Schistosoma. Schistosomiasis is characterized by a variety of acute and chronic infections. Acute infection is marked by fever, al- lergic symptoms, and diarrhea. Chronic effects are usually severe and are caused by fibrous degeneration of tis- sue around deposited eggs of the para- site in the liver, lungs, and central nervous system. Schistosomes can also cause schistosome dermatitis (e.g., swimmer’s itch), a skin disease marked by intense itching.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.3630   Serratia spp. serological re- agents.

- Identification. Serratia spp. sero- logical reagents are devices that con- sist of antigens and antisera used in se- rological tests to identify Serratia spp. from cultured isolates. The identifica- tion aids in the diagnosis of disease caused by bacteria belonging to the genus Serratia and provides epidemio- logical information on these diseases. Serratia spp. are occasionally associ- ated with gastroenteritis (food  poi- soning) and wound infections.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

#### § 866.3660

##### § 866.3660   Shigella spp. serological re- agents.

  - Identification. Shigella spp. sero- logical reagents are devices that con- sist of antigens and antisera, including antisera conjugated with a fluorescent dye  (immunofluorescent  reagents), used in serological tests to identify Shigella spp. from  cultured  isolates. The identification aids in the diagnosis of shigellosis  caused by  bacteria be- longing to the genus Shigella and pro- vides epidemiological information on this disease. Shigellosis is  character- ized by abdominal pain, cramps, diar- rhea, and fever.

  - Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

§ 866.3680   Sporothrix   schenckii   sero- logical reagents.

- Identification. Sporothrix schenckii serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to Sporothrix schenckii in serum. The identification aids in  the diagnosis  of sporothrichosis caused by a fungus be- longing to the genus Sporothrix and provides epidemiological information on this disease. Sporothrichosis is a chronic tumorlike infection primarily of the skin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

§ 866.3700   Staphylococcus aureus sero- logical reagents.

- Identification. Staphylococcus aureus serological reagents are devices that consist of antigens and antisera used in serological tests to identify enterotoxin (toxin affecting the intes- tine) producing staphylococci from cul- tured isolates. The identification  aids in the diagnosis of disease caused by this bacterium belonging to the genus Staphylococcus and provides epidemio-

#### 21 CFR Ch. I (4–1–17 Edition)

logical information on these diseases. Certain strains of Staphylococcus aureus produce an enterotoxin while growing in meat, dairy, or bakery products. After ingestion, this enterotoxin is ab- sorbed in the gut and causes destruc- tion of the intestinal lining (gastroenteritis).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

§ 866.3720   Streptococcus	spp.	exo- enzyme reagents.

- Identification. Streptococcus spp. exoenzyme reagents are devices used to identify antibodies to Streptococcus spp. exoenzyme in serum. The identifica- tion aids in the diagnosis of disease caused by bacteria belonging to the genus Streptococcus and provides epide- miological information on these dis- eases. Pathogenic streptococci are as- sociated with infections, such as sore throat, impetigo (an infection charac- terized by small pustules on the skin), urinary tract infections, rheumatic fever, and kidney disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  61

FR 1119,  Jan. 16,  1996;  66  FR 38792,  July 25,

2001]

§ 866.3740   Streptococcus	spp.	sero- logical reagents.

- Identification. Streptococcus spp. serological reagents are devices that consist of antigens and antisera (ex- cluding streptococcal exoenzyme re- agents made from enzymes secreted by streptococci) used in serological tests to identify Streptococcus spp. from cul- tured isolates derived from clinical specimens. The identification aids in the diagnosis of diseases caused by bac- teria belonging to the genus Strepto- coccus and provides epidemiological in- formation on these diseases. Patho- genic streptococci are associated with

#### Food and Drug Administration, HHS	§ 866.3860

infections, such as sore throat, impe- tigo (an infection characterized by small pustules on the skin), urinary tract infections, rheumatic fever, and kidney disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

§ 866.3780   Toxoplasma	gondii	sero- logical reagents.

- Identification. Toxoplasma gondii serological reagents are devices that consist of antigens and antisera used in serological tests to identify  antibodies to Toxoplasma gondii in serum. Addi- tionally, some of these reagents consist of antisera conjugated with a fluores- cent dye (immunofluorescent reagents) used to identify Toxoplasma gondii from clinical specimens. The identification aids in the diagnosis of toxoplasmosis caused by the parasitic protozoan Toxoplasma gondii and provides epide- miological information on this disease. Congenital toxoplasmosis is character- ized by lesions of the central nervous system, which if undetected and un- treated may lead to brain defects, blindness, and death of an unborn fetus. The disease is characterized in children by inflammation of the brain and spinal cord.

- Classification. Class II (perform- ance standards).

§ 866.3820   Treponema   pallidum   non- treponemal test reagents.

- Identification. Treponema pallidum nontreponemal test reagents are de- vices that consist of antigens derived from nontreponemal sources (sources not	directly	associated	with treponemal organisms)  and  control sera (standardized sera with which test results are compared) used in sero- logical tests to identify reagin, an anti- body-like agent, which is produced from the reaction of treponema micro- organisms with body tissues. The iden- tification aids in the diagnosis  of syphilis caused by microorganisms be- longing to the genus Treponema and provides  epidemiological  information on  syphilis.

- Classification. Class II (perform- ance standards).

§ 866.3830   Treponema	pallidum	tre- ponemal test reagents.

- Identification. Treponema pallidum treponemal test reagents are devices that consist of the antigens, antisera and all control reagents (standardized reagents with which test results are compared) which are derived from treponemal sources and that are used in the fluorescent treponemal antibody absorption test (FTA-ABS), the Treponema pallidum immobilization test (T.P.I.), and other treponemal tests used to identify antibodies to Treponema pallidum directly from in- fecting treponemal organisms in serum. The identification aids in the diagnosis of syphilis caused by bacteria belonging to the genus Treponema and provides epidemiological information on syphilis.

- Classification. Class II (perform- ance standards).

§ 866.3850   Trichinella	spiralis	sero- logical reagents.

- Identification. Trichinella spiralis serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to Trichinella spiralis in serum. The identification aids in  the diagnosis  of trichinosis caused by parasitic roundworms belonging to the genus Trichinella and  provides  epidemiolog- ical information on trichinosis. Trichi- nosis is caused by ingestion of under- cooked, infested meat, especially pork, and characterized by fever, muscle weakness, and diarrhea.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

§ 866.3860   Trichomonas   vaginalis  nu- cleic acid assay.

- Identification. A Trichomonas vaginalis nucleic acid assay is a device that consists of primers, probes, en- zymes, and controls for the amplifi- cation and detection of trichomonas nucleic  acids  in  endocervical  swabs,

#### § 866.3870

vaginal swabs, and female urine speci- mens, from women symptomatic for vaginitis, cervicitis, or urethritis and/ or to aid in the diagnosis of trichomo- niasis in asymptomatic women. The de- tection of trichomonas nucleic  acids, in conjunction with other laboratory tests, aids in the clinical laboratory di- agnosis of trichomoniasis caused by Trichomonas vaginalis.

- Classification. Class II (special controls). The special controls are set forth in FDA’s guideline document en- titled: ‘‘Class II Special Controls Guideline: Nucleic Acid Amplification Assays for the Detection  of Trichomonas vaginalis; Guideline for In- dustry and Food and Drug Administra- tion Staff.’’ See § 866.1(e) for informa- tion on obtaining this document.

[80 FR 46192, Aug. 4, 2015]

§ 866.3870   Trypanosoma	spp.	sero- logical reagents.

- Identification. Trypanosoma spp. serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to Trypanosoma spp. in  serum. The identification aids in  the diagnosis  of trypanosomiasis, a disease caused by parasitic protozoans belonging to the genus Trypanosoma. Trypanosomiasis in adults is a chronic disease charac- terized by fever, chills, headache, and vomiting. Central nervous system in- volvement produces typical sleeping sickness syndrome: physical exhaus- tion, inability to eat, tissue wasting, and eventual death. Chagas disease, an acute form of trypanosomiasis in chil- dren, most seriously affects the central nervous system and heart muscle.

- Classification. Class I (general con- trols).

##### § 866.3900   Varicella-zoster  virus  sero- logical reagents.

- Identification. Varicella-zoster virus serological reagents are devices that consist of antigens and antisera used in serological tests to identify antibodies to varicella-zoster in serum. The identification aids in the diagnosis of diseases caused by varicella-zoster viruses and provides epidemiological information on these diseases. Varicella (chicken pox) is a mild, high- ly  infectious  disease,  chiefly  of  chil-

#### 21 CFR Ch. I (4–1–17 Edition)

dren. Zoster (shingles) is the recurrent form of the disease, occurring in adults who were previously infected with varicella-zoster viruses. Zoster is the response (characterized by a rash) of the partially immune host to a reac- tivation of varicella viruses present in latent form in the patient’s body.

- Classification. Class II (perform- ance standards).

§ 866.3930   Vibrio   cholerae   serological reagents.

- Identification. Vibrio cholerae sero- logical reagents are devices that are used in the agglutination (an antigen- antibody clumping reaction) test to identify Vibrio cholerae from cultured isolates derived from clinical speci- mens. The identification aids in the di- agnosis of cholera caused by the bac- terium Vibrio cholerae and provides epi- demiological information on cholera. Cholera is an acute infectious disease characterized by severe diarrhea with extreme fluid and electrolyte (salts) depletion, and by vomiting, muscle cramps, and prostration. If untreated, the severe dehydration may lead to shock, renal failure, cardiovascular collapse, and death.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.3940   West  Nile  virus  serological reagents.

- Identification. West Nile virus se- rological reagents are devices that con- sist of antigens and antisera for the de- tection of anti-West Nile virus IgM antibodies, in human serum, from indi- viduals who have signs and symptoms consistent with viral meningitis/en- cephalitis. The detection aids in the clinical laboratory diagnosis of viral meningitis/encephalitis caused by West Nile  virus.

- Classification. Class II (special controls). The special control is FDA’s guidance entitled ‘‘Class II Special Controls Guidance Document: Sero- logical Reagents for the Laboratory Di- agnosis    of    West    Nile    Virus.’’    See

#### Food and Drug Administration, HHS	§ 866.3980

§ 866.1(e)   for   the   availability   of   this guidance document.

[68 FR 61745, Oct. 30, 2003]

##### § 866.3945 Dengue virus serological re- agents.

  - Identification. Dengue virus sero- logical reagents are devices that con- sist of antigens and antibodies for the detection of dengue virus and dengue antibodies in individuals who have signs and symptoms of dengue fever or dengue hemorrhagic fever. The detec- tion aids in the clinical laboratory di- agnosis of dengue fever or dengue hem- orrhagic fever caused by dengue virus.

  - Classification. Class II (special controls). The special control is FDA’s guideline entitled ‘‘Class II Special Controls Guideline: Dengue Virus Sero- logical Reagents.’’  For  availability  of the guideline document, see § 866.1(e).

[79 FR 31023, May 30, 2014]

##### § 866.3946 Dengue virus nucleic acid amplification test reagents.

- Identification. Dengue virus nu- cleic acid  amplification  test  reagents are devices that consist of primers, probes, enzymes, and controls for the amplification and detection of dengue virus serotypes 1, 2, 3, or 4 from viral ribonucleic acid (RNA) in human serum and plasma from individuals who have signs and symptoms consistent with dengue (mild or severe). The identifica- tion of dengue virus serotypes 1, 2, 3, or 4 in human serum and plasma (sodium citrate) collected from human patients with dengue provides epidemiologic in- formation for surveillance of circu- lating dengue viruses.

- Classification. Class II (special controls). The special control is FDA’s guideline entitled ‘‘Class II Special Controls Guideline: Dengue Virus Nu- cleic Acid Amplification Test  Re- agents.’’ For availability of the guide- line document, see § 866.1(e).

[79 FR 53609, Sept. 10, 2014]

##### § 866.3950 In vitro human immuno- deficiency virus (HIV) drug resist- ance genotype assay.

- Identification. The in vitro HIV drug resistance genotype assay is a de- vice that consists of nucleic acid rea- gent primers and probes together with

software for predicting drug resistance/ susceptibility based on results obtained with these primers and probes. It is in- tended for use in detecting HIV genomic mutations that confer resist- ance to specific antiretroviral drugs, as an aid in monitoring and treating HIV infection.

- Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document: In Vitro HIV Drug Re- sistance Genotype Assay.’’ See § 866.1(e) for the availability of this guidance document.

[72 FR 44382, Aug. 8, 2007]

##### § 866.3980   Respiratory viral panel mul- tiplex nucleic acid assay.

- Identification. A respiratory viral panel multiplex nucleic acid assay is a qualitative in vitro diagnostic device intended to simultaneously detect and identify multiple viral nucleic acids ex- tracted from human respiratory speci- mens or viral culture. The detection and identification of a specific viral nucleic acid from individuals exhib- iting signs and symptoms of res- piratory infection aids in the diagnosis of respiratory viral infection when used in conjunction with other clinical and laboratory findings. The device is in- tended for detection and identification of a combination of the following vi- ruses:

  - Influenza A and Influenza B;

  - Influenza A subtype H1 and Influ- enza A subtype H3;

  - Respiratory Syncytial Virus subtype A and Respiratory Syncytial Virus subtype B;

  - Parainfluenza 1, Parainfluenza 2, and Parainfluenza 3 virus;

  - Human Metapneumovirus;

  - Rhinovirus; and

  - Adenovirus.

- Classification. Class II (special controls). The special controls are:

- FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Respiratory Viral Panel Multiplex Nucleic Acid Assay;’’

- For a device that detects and identifies Human Metapneumovirus, FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document:      Testing      for      Human

#### § 866.3990

Metapneumovirus  (hMPV)  Using  Nu- cleic Acid Assays;’’ and

- For a device that detects and dif- ferentiates Influenza A subtype H1 and subtype H3, FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Testing for Detec- tion and Differentiation of Influenza A Virus Subtypes Using Multiplex Nu- cleic Acid Assays.’’ See § 866.1(e) for the availability of these guidance docu- ments.

[74 FR 52138, Oct. 9, 2009]

##### § 866.3990 Gastrointestinal microorga- nism multiplex nucleic acid-based assay.

- Identification. A gastrointestinal microorganism multiplex nucleic acid- based assay is a qualitative in vitro di- agnostic device intended to simulta- neously detect and identify multiple gastrointestinal microbial  nucleic acids extracted from human stool specimens. The device detects specific nucleic acid sequences for organism identification as well as for deter- mining the presence of toxin genes. The detection and identification of a specific gastrointestinal microbial nu- cleic acid from individuals exhibiting signs and symptoms of gastrointestinal infection aids in the diagnosis of gas- trointestinal infection when used in conjunction with clinical evaluation and other laboratory findings. A gas- trointestinal microorganism multiplex nucleic acid-based assay also aids in the detection and identification of acute gastroenteritis in the context of outbreaks.

- Classification. Class II (special controls). The special controls are set forth in FDA’s guideline document en- titled: ‘‘Class II Special Controls Guideline: Gastrointestinal Microorga- nism Multiplex Nucleic Acid-Based As- says for Detection and Identification of Microorganisms and Toxin Genes from Human Stool Specimens.’’ For avail- ability of the guideline document, see

§ 866.1(e).

[80 FR 67314, Nov. 2, 2015]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart E—Immunology Labora- tory Equipment and  Re- agents

##### § 866.4070   RNA Preanalytical Systems.

- Identification. RNA Preanalytical Systems are devices intended to col- lect, store, and transport patient speci- mens, and stabilize intracellular RNA from the specimens, for subsequent iso- lation and purification of the intracellular RNA for RT–PCR used in in vitro molecular diagnostic testing.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: RNA Preanalytical Systems (RNA Col- lection, Stabilization and Purification System for RT–PCR Used in Molecular Diagnostic Testing).’’ See § 866.1(e) for the availability of this guidance docu- ment.

[70 FR 49863, Aug. 25, 2005]

##### § 866.4100   Complement reagent.

- Identification. A complement rea- gent is a device that consists of com- plement, a naturally occurring serum protein from any warm-blooded animal such as guinea pigs, that may be in- cluded as a component part of sero- logical test kits used in the diagnosis of disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  2001,  as  amended  at  66

FR 38792, July 25, 2001]

##### § 866.4500   Immunoelectrophoresis equipment.

- Identification.

Immunoelectrophoresis equipment for clinical use with its electrical power supply is a device used for separating protein	molecules.

Immunoelectrophoresis is a procedure in which a complex protein mixture  is placed in an agar gel and the various proteins are separated on the basis of their relative mobilities under the in- fluence of an electric current. The sep- arated proteins are then permitted to diffuse   through   the   agar   toward   a

#### Food and Drug Administration, HHS	§ 866.4700

multispecific antiserum, allowing pre- cipitation and visualization of the sep- arate complexes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.4520   Immunofluorometer    equip- ment.

- Identification. Immunofluorometer equipment for clinical use with its electrical power supply is a device used to measure  the fluorescence  of fluorochrome-labeled antigen-antibody complexes. The concentration of these complexes may be measured by means of reflected light. A beam of light is passed through a solution in which a fluorochrome has been selectively at- tached to serum protein antibody mol- ecules in suspension. The amount of light emitted  by  the  fluorochrome label is detected by a photodetector, which converts light energy into elec- trical energy. The amount of electrical energy registers on a readout system such as a digital voltmeter or a record- ing chart. This electrical readout is called the fluorescence value  and  is used to measure the concentration of antigen-antibody  complexes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.4540   Immunonephelometer equipment.

- Identification.

Immunonephelometer equipment for clinical use with its electrical power supply is a device that measures light scattering from antigen-antibody com- plexes. The concentration of these complexes may be measured by means of reflected light. A beam of light passed through a solution is  scattered by the particles in suspension. The amount of light is detected by a photodetector, which converts light en-

ergy into electrical  energy.  The amount of electrical energy registers on a readout system such as a digital voltmeter or a recording chart. This electrical readout is called the light- scattering value and is used to measure the concentration of antigen-antibody complexes. This generic type of device includes devices with various kinds of light sources, such as laser equipment.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.4600   Ouchterlony agar plate.

- Identification. An ouchterlony agar plate for clinical use is a device containing an agar gel used to examine antigen-antibody	reactions.	In immunodiffusion, antibodies and anti- gens migrate toward each other through gel which originally contained neither of these reagents. As the re- agents come in contact with each other, they combine to form a precipi- tate that is trapped in the gel matrix and is immobilized.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

##### § 866.4700 Automated fluorescence in situ hybridization (FISH) enumera- tion systems.

- Identification. An automated FISH enumeration system is a device that consists of an automated scanning mi- croscope, image analysis system, and customized software applications for FISH assays. This device is intended for in vitro diagnostic use with FISH assays as an aid in the detection, counting and classification of cells based on recognition of cellular color, size, and shape, and in the detection and enumeration of FISH signals in interphase nuclei of formalin-fixed, paraffin-embedded human tissue speci- mens.

#### § 866.4800

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Automated Fluorescence in situ Hy- bridization (FISH) Enumeration Sys- tems.’’ See § 866.1(e) for the availability of this guidance document.

[70 FR 14534, Mar. 23, 2005]

##### § 866.4800   Radial	immunodiffusion plate.

- Identification.        A	radial immunodiffusion plate for clinical use is a device that consists of a plastic plate to which agar gel containing antiserum is added. In radial immunodiffusion, antigens migrate through gel which originally contains specific antibodies. As the reagents come in contact with each other, they combine to form a precipitate that is trapped in  the  gel  matrix  and  immo- bilized.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  66

FR 38792, July 25, 2001]

##### § 866.4830   Rocket	immunoelectro- phoresis equipment.

- Identification.	Rocket immunoelectrophoresis equipment  for clinical use is a device used to perform a specific test on proteins by using a procedure		called	rocket immunoelectrophoresis. In this proce- dure, an electric current causes the protein in solution to migrate through agar gel containing specific  antisera. The protein precipitates with the antisera in a rocket-shaped pattern, giving the name to the device.  The height of the peak (or the  area  under the peak) is proportional to the con- centration of the protein.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 866.4900   Support gel.

- Identification. A support gel for clinical use is a device that consists of an agar or agarose preparation that is used while measuring various kinds of, or parts of, protein molecules by var- ious immunochemical techniques, such as	immunoelectrophoresis, immunodiffusion,  or  chromatography.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  54

FR 25047, June 12, 1989; 66 FR 38792, July 25,

2001]

## Subpart F—Immunological Test Systems

##### § 866.5040   Albumin immunological test system.

- Identification. An albumin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the albumin (a plasma protein) in serum and other body fluids. Meas- urement of albumin aids in the diag- nosis of kidney and intestinal diseases.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.5060   Prealbumin   immunological test system.

- Identification. A prealbumin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the prealbumin (a plasma pro- tein) in serum and other body fluids. Measurement of prealbumin levels in serum may aid in the assessment of the patient’s nutritional status.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

Food and Drug Administration, HHS	§ 866.5120

##### § 866.5065   Human	allotypic	marker immunological test system.

  - Identification. A human allotypic marker immunological test system is a device that consists of the reagents used to identify by immunochemical techniques the inherited human pro- tein allotypic markers (such as nGm, nA2 m, and Km allotypes) in serum and other body fluids. The identification may be used while studying population genetics.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5080   Alpha-1-antichymotrypsin immunological test system.

- Identification. An alpha-1- antichymotrypsin immunological test system is a device that consists of the reagents used to measure by immunochemical techniques alpha-1- antichymotrypsin (a protein) in serum, other body fluids, and tissues. Alpha-1- antichymotrypsin helps protect tissues against proteolytic (protein-splitting) enzymes released during infection.

- Classification. Class II (perform- ance standards).

##### § 866.5090   Antimitochondrial  antibody immunological test system.

- Identification.		An antimitochondrial	antibody immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the antimitochondrial anti- bodies in human serum. The measure- ments aid in the diagnosis of diseases that produce a spectrum of autoantibodies (antibodies produced against the body’s own tissue), such as primary biliary cirrhosis (degeneration of liver tissue) and chronic active hepa- titis (inflammation of the liver).

- Classification. Class II (perform- ance standards).

##### § 866.5100   Antinuclear	antibody immunological test system.

- Identification. An antinuclear antibody immunological test system is a device that consists of the reagents

used to measure by immunochemical techniques the autoimmune antibodies in serum, other body fluids, and tissues that react with cellular nuclear con- stituents (molecules present in the nu- cleus of a cell, such as ribonucleic acid, deoxyribonucleic acid, or nuclear pro- teins). The measurements aid in the di- agnosis	of	systemic	lupus erythematosus (a multisystem auto- immune disease in which antibodies at- tack the victim’s own tissues), hepa- titis (a liver disease), rheumatoid ar- thritis, Sjo¨ gren’s syndrome (arthritis with inflammation of the eye, eyelid, and salivary glands), and systemic scle- rosis (chronic hardening and shrinking of many body tissues).

- Classification. Class II (perform- ance standards).

##### § 866.5110   Antiparietal	antibody immunological test system.

- Identification. An antiparietal antibody immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the specific antibody for gastric parietal cells in serum and other body fluids. Gastric parietal cells are those cells located in the stomach that produce a protein that enables vi- tamin B12 to be absorbed by the body. The measurements aid in the diagnosis of vitamin B12 deficiency (or pernicious anemia), atrophic gastritis (inflamma- tion of the stomach), and autoimmune connective tissue diseases (diseases re- sulting when the body produces anti- bodies against its own tissues).

- Classification. Class II (perform- ance standards).

##### § 866.5120   Antismooth muscle antibody immunological test system.

- Identification. An antismooth mus- cle antibody immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the antismooth muscle antibodies (anti- bodies to nonstriated, involuntary muscle) in serum. The measurements aid in the diagnosis of chronic hepa- titis (inflammation of the liver) and autoimmune connective tissue diseases (diseases resulting from antibodies pro- duced against the body’s own tissues).

#### § 866.5130

- Classification Class II (perform- ance standards).

##### § 866.5130   Alpha-1-antitrypsin immunological test system.

- Identification. An alpha-1- antitrypsin immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the alpha-1-antitrypsin (a plasma protein) in serum, other body fluids, and tissues. The measurements aid in the diagnosis of several condi- tions including juvenile and adult cir- rhosis of the liver. In addition, alpha-1- antitrypsin deficiency has been associ- ated with pulmonary emphysema.

- Classification. Class II (perform- ance standards).

##### § 866.5150   Bence-Jones	proteins immunological test system.

- Identification. A Bence-Jones pro- teins immunological test system is a device that consists of  the  reagents used to measure by immunochemical techniques the Bence-Jones proteins in urine and plasma. Immunoglobulin molecules normally consist of pairs of polypeptide chains (subunits) of un- equal size (light chains and heavy chains) bound together by several di- sulfide bridges. In some cancerous con- ditions, there is a proliferation of one plasma cell (antibody-producing cell) with excess production of  light  chains of one specific kind (monoclonal light chains). These free homogeneous light chains not associated with an immunoglobulin molecule can be found in urine and plasma, and have been called Bence-Jones proteins. Measure- ment of Bence-Jones proteins and de- termination that they are monoclonal aid in the diagnosis of multiple myeloma (malignant proliferation of plasma	cells),	Waldenstrom’s macroglobulinemia (increased produc- tion of large immunoglobulins by spleen and bone marrow cells), leu- kemia (cancer of the blood-forming or- gans), and lymphoma (cancer of the lymphoid  tissue).

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 866.5160   Beta-globulin	immunolog- ical test system.

- Identification. A beta-globulin immunological test system is a device that consists of reagents used to meas- ure by  immunochemical  techniques beta globulins (serum protein) in serum and other body fluids. Beta-globulin proteins include beta-lipoprotein, transferrin, glycoproteins, and com- plement, and are rarely associated with specific pathologic disorders.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5170   Breast  milk  immunological test system.

- Identification. A breast milk immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the breast milk proteins.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5180   Fecal	calprotectin immunological test system.

- Identification. A fecal calprotectin immunological test system is an in vitro diagnostic device that consists of reagents used to quantitatively meas- ure, by immunochemical techniques, fecal calprotectin in human stool specimens. The device is intended forin vitro diagnostic use as an aid in the di- agnosis of inflammatory bowel diseases (IBD), specifically Crohn’s disease and ulcerative colitis, and as an aid in dif- ferentiation of IBD from  irritable bowel syndrome.

- Classification. Class II (special controls). The special control for these devices is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance	Document:	Fecal

#### Food and Drug Administration, HHS	§ 866.5250

Calprotectin Immunological Test Sys- tems.’’ For the availability of this guidance document, see § 866.1(e).

[71 FR 42598, July 27, 2006]

##### § 866.5200   Carbonic  anhydrase  B  and C immunological test system.

  - Identification. A carbonic anhy- drase B and C immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques specific carbonic anhydrase protein molecules in serum and other body fluids. Meas- urements of carbonic anhydrase B and C aid in the diagnosis of abnormal he- moglobin metabolism.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5210   Ceruloplasmin    immunolog- ical test system.

- Identification. A ceruloplasmin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the ceruloplasmin (copper- transporting serum protein) in serum, other body fluids, or tissues. Measure- ments of ceruloplasmin aid in the diag- nosis of copper metabolism disorders.

- Classification. Class II (perform- ance standards).

##### § 866.5220   Cohn fraction II immunolog- ical test system.

- Identification. A Cohn fraction II immunological test system is a device that consists of the reagents that con- tain or are used to measure that frac- tion of plasma containing protein gamma globulins, predominantly of the IgG class. The device may be used as a coprecipitant in radioimmunoassay methods, as raw material for the puri- fication of IgG subclasses, and to re- duce nonspecific adsorption of plasma proteins in immunoassay techniques. Measurement of these proteins aids in the diagnosis of any disease concerned with abnormal levels of IgG gamma globulins such as agammaglobulinemia or multiple myeloma.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5230   Colostrum	immunological test system.

- Identification. A colostrum immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the specific proteins in colos- trum. Colostrum is a substance ex- creted by the mammary glands during pregnancy and until production  of breast milk begins 1 to 5 days after childbirth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5240   Complement	components immunological test system.

- Identification. A complement com- ponents immunological test system is a device that consists of the reagents used to measure by immunochemical techniques complement components C1q, C1r, C1s, C2, C3, C4, C5, C6, C7, C8, and C9, in serum, other body fluids, and tis- sues. Complement is a group of serum proteins which destroy infectious agents. Measurements of these proteins aids in the diagnosis of immunologic disorders, especially those associated with deficiencies of complement com- ponents.

- Classification. Class II (perform- ance standards).

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  53

FR 11253, Apr. 6, 1988]

##### § 866.5250 Complement C 2 inhibitor (inactivator) immunological test system.

- Identification. A complement C1 inhibitor (inactivator) immunological test system is a device that consists of the   reagents   used   to   measure   by

#### § 866.5260

immunochemical techniques the com- plement C1 inhibitor (a plasma protein) in serum. Complement C1 inhibitor oc- curs normally in plasma and blocks the action of the C1 component of com- plement (a group of serum proteins which destroy infectious agents). Meas- urement of complement C1 inhibitor aids in the diagnosis of hereditary angioneurotic edema (increased blood vessel permeability causing swelling of tissues) and a rare form of angioedema associated with lymphoma (lymph node cancer).

- Classification. Class II (perform- ance standards).

##### § 866.5260   Complement C3b   inactivator immunological test system.

- Identification. A complement C3b inactivator immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the complement C3b inactivator (a plasma protein) in serum. Complement is a group of serum proteins that destroy infectious agents. Measurement of complement C3b inactivator aids in the diagnosis of in- herited antibody dysfunction.

- Classification. Class II (perform- ance standards).

##### § 866.5270   C-reactive  protein  immuno- logical test system.

- Identification. A C-reactive protein immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the C-reactive protein in serum and other body fluids. Measurement of C-reactive protein aids in evaluation of the amount of injury to body tissues.

- Classification. Class II (perform- ance standards).

##### § 866.5320   Properdin factor B immuno- logical test system.

- Identification. A properdin factor B immunological test system is a de- vice that consists of the reagents used to measure by immunochemical tech- niques properdin factor B in serum and other body fluids. The deposition of properdin factor B in body tissues or a corresponding depression in the amount of properdin factor B in serum and other body fluids is evidence of the involvement  of  the  alternative  to  the

#### 21 CFR Ch. I (4–1–17 Edition)

classical pathway of activation of com- plement (a group of plasma proteins which cause the destruction of cells which are foreign to the body). Meas- urement of properdin factor B aids in the diagnosis of several kidney dis- eases, e.g., chronic glomerulonephritis (inflammation of the glomeruli of the kidney), lupus nephritis (kidney dis- ease associated with a multisystem autoimmune disease, systemic lupus erythematosus), as well as several skin diseases, e.g., dermititis herpetiformis (presence of vesicles on the skin that burn and itch), and pemphigus vulgaris (large vesicles on the skin). Other dis- eases in which the alternate pathway of complement activation has been im- plicated include rheumatoid arthritis, sickle cell anemia, and gram-negative bacteremia.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.5330   Factor  XIII,  A,  S,  immuno- logical test system.

- Identification. A factor XIII, A, S, immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the factor XIII (a bloodclotting factor), in platelets (A) or serum (S). Measurements of factor XIII, A, S, aid in the diagnosis and treatment of cer- tain bleeding disorders resulting from a deficiency of this factor.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9. This exemption does not apply to factor deficiency tests classified under § 864.7290 of this chap- ter.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5340   Ferritin  immunological  test system.

- Identification. A ferritin immunological test system is a device that  consists  of  the  reagents  used  to

#### Food and Drug Administration, HHS	§ 866.5400

measure by immunochemical tech- niques the ferritin (an iron-storing pro- tein) in serum and other body fluids. Measurements of ferritin aid in the di- agnosis of diseases affecting iron me- tabolism, such as hemochromatosis (iron overload) and iron deficiency amemia.

- Classification. Class II (perform- ance standards).

##### § 866.5350   Fibrinopeptide   A   immuno- logical test system.

- Identification. A fibrinopeptide A immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the fibrinopeptide A (a blood- clotting factor) in plasma and other body	fluids.	Measurement	of fibrinopeptide A may aid in the diag- nosis and treatment of certain blood- clotting disorders.

- Classification. Class II (perform- ance standards).

##### § 866.5360   Cohn  fraction  IV  immuno- logical test system.

- Identification. A Cohn fraction IV immunological test system is a device that consists of or measures that frac- tion of plasma proteins, predominantly alpha- and beta- globulins, used as a raw material for the production of pure alpha- or beta- globulins. Measurement of specific alpha- or beta- globulins aids in the diagnosis of many diseases, such as Wilson’s disease (an inherited dis- ease affecting the liver and brain), Tangier’s disease (absence of alpha-1- lipoprotein), malnutrition, iron  defi- ciency anemia, red blood cell disorders, and kidney disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47 FR 50823, Nov. 9, 1982; 47 FR 56846, Dec. 21,

1982, as amended at 59 FR 63007, Dec. 7, 1994;

66 FR 38793, July 25, 2001]

##### § 866.5370   Cohn   fraction   V   immuno- logical test system.

- Identification. A Cohn fraction V immunological test system is a device that consists of or measures that frac- tion of plasma containing predomi- nantly   albumin   (a   plasma   protein).

This test aids in the diagnosis of dis- eases where albumin levels may be de- pressed, e.g., nephrosis (disease of the kidney), proteinuria (protein in the urine), gastroenteropathy (disease of the stomach and small intestine), rheu- matoid arthritis, and viral hepatitis.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5380   Free   secretory   component immunological test system.

- Identification. A free secretory component immunological test system is a device that consists of the reagents used to measure by immunochemical techniques free secretory component (normally a portion of the secretory IgA antibody molecule) in body fluids. Measurement of free secretory compo- nent (protein molecules) aids in the di- agnosis or repetitive lung infections and other hypogammaglobulinemic conditions (low antibody levels).

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.5400   Alpha-globulin	immuno- logical test system.

- Identification. An alpha-globulin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the alpha-globulin (a serum pro- tein) in serum and other body fluids. Measurement of alpha-globulin  may aid in the diagnosis of inflammatory lesions, infections, severe burns, and a variety of other conditions.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

#### § 866.5420

##### § 866.5420   Alpha-1-glycoproteins immunological test system.

  - Identification. An alpha-1- glycoproteins immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques alpha-1- glycoproteins (a group of plasma pro- teins found in the alpha-1 group when subjected to electrophoresis) in serum and other body fluids. Measurement of specific alpha-1-glycoproteins may aid in the diagnosis of collagen (connective tissue) disorders, tuberculosis, infec- tions, extensive malignancy, and diabe- tes.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5425   Alpha-2-glycoproteins immunological test system.

- Identification. An alpha-2- glycoproteins immunolgical test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the alpha- 2-glycoproteins (a group of plasma pro- teins found in the alpha-2 group when subjected to electrophoresis) in serum and other body fluids. Measurement of alpha-2-glycoproteins aids in the diag- nosis of some cancers and genetically inherited deficiencies of these plasma proteins.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5430   Beta-2-glycoprotein	I immunological test system.

- Identification.         A	beta-2- glycoprotein I immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the beta-2- glycoprotein I (a serum protein) in serum and other body fluids. Measure- ment  of  beta-2-glycoprotein  I  aids  in

#### 21 CFR Ch. I (4–1–17 Edition)

the diagnosis of an inherited deficiency of this serum protein.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5440   Beta-2-glycoprotein	III

immunological test system.

- Identification.         A	beta-2- glycoprotein III immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the beta-2- glycoprotein III (a serum protein) in serum and other body fluids. Measure- ment of beta-2-glycoprotein III aids in the diagnosis of an inherited deficiency of this serum protein and a variety of other  conditions.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5460   Haptoglobin  immunological test system.

- Identification. A haptoglobin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the haptoglobin (a protein that binds hemoglobin, the oxygen-carrying pigment in red blood cells) in serum. Measurement of haptoglobin may aid in the diagnosis of hemolytic diseases (diseases in which the red blood cells rupture and release hemoglobin) re- lated to the formation of hemoglobin- haptoglobin complexes and certain kid- ney diseases.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

#### Food and Drug Administration, HHS	§ 866.5520

##### § 866.5470   Hemoglobin   immunological test system.

  - Indentification. A hemoglobin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the different types of free hemo- globin (the oxygen-carrying pigment in red blood cells) in blood, urine, plasma, or other body fluids. Measurements of free hemoglobin aid in the diagnosis of various hematologic disorders, such as sickle cell anemia, Fanconi’s anemia (a rare inherited disease), aplastic anemia (bone marrow does not produce enough blood cells), and leukemia (cancer of the blood-forming organs).

  - Classification. Class II (perform- ance standards).

##### § 866.5490   Hemopexin	immunological test system.

- Indentification. A hemopexin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the hemopexin (a serum protein that binds heme, a component of hemo- globin) in serum. Measurement of hemopexin aids in the diagnosis of var- ious hematologic disorders, such as he- molytic anemia (anemia due to short- ened in vivo survival of mature red blood cells and inability of the bone marrow to compensate for their de- creased life span) and sickle cell ane- mia.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  63

FR 59227, Nov. 3, 1998]

##### § 866.5500   Hypersensitivity	pneumo- nitis immunological test system.

- Identification. A hypersensitivity pneumonitis immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the immunoglobulin antibodies in serum which react specifically with organic dust derived from fungal or animal pro- tein sources. When these antibodies react with such dusts in the lung, im- mune complexes precipitate and trig- ger an inflammatory reaction (hyper-

sensitivity pneumonitis). Measurement of these immunoglobulin G antibodies aids in the diagnosis of hyper- sensitivity pneumonitis and other al- lergic respiratory disorders.

- Classification. Class II (perform- ance standards).

##### § 866.5510 Immunoglobulins A, G, M, D, and E immunological test system.

- Identification.		An immunoglobulins A, G, M, D, and E immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the immunoglobulins A,  G,  M, D, an E (serum antibodies) in serum. Measurement		of	these immunoglobulins aids in the diagnosis of abnormal protein metabolism and the body’s lack of ability to resist in- fectious agents.

- Classification. Class II (perform- ance standards).

##### § 866.5520 Immunoglobulin G (Fab fragment specific) immunological test system.

- Identification. An immunoglobulin G	(Fab	fragment		specific) immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the Fab antigen-binding frag- ment resulting from breakdown of immunoglobulin G antibodies in urine, serum, and other body fluids. Measure- ment of Fab fragments of immunoglobulin G aids in the diag- nosis of lymphoproliferative disorders, such as multiple myeloma (tumor of bone marrow cells), Waldenstrom’s macroglobulinemia	(increased immunoglobulin production by the spleen and bone marrow cells), and lymphoma (tumor of the lymphoid tis- sues).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

#### § 866.5530

##### § 866.5530 Immunoglobulin G (Fc frag- ment specific) immunological test system.

  - Identification. An immunoglobulin G	(Fc	fragment	specific) immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the Fc  (carbohydrate con- taining) fragment of  immunoglobulin G (resulting from breakdown of immunoglobulin G antibodies) in urine, serum, and other body fluids. Measure- ment of immunoglobulin G Fc frag- ments aids in the diagnosis of plasma cell antibody-forming abnormalities, e.g., gamma heavy chain disease.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

##### § 866.5540 Immunoglobulin G (Fd frag- ment specific) immunological test system.

- Identification. An immunoglobulin G	(Fd	fragment	specific) immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the amino terminal (antigen- binding) end (Fd fragment) of the heavy chain (a subunit) of the immunoglobulin antibody molecule in serum.		Measurement		of immunoglobulin G Fd fragments  aids in the diagnosis of plasma antibody- forming cell abnormalities.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5550   Immunoglobulin	(light chain specific) immunological test system.

- Identification. An immunoglobulin (light chain specific) immunological test system is a device that consists of the reagents used to measure by immunochemical      techniques      both

#### 21 CFR Ch. I (4–1–17 Edition)

kappa and lambda types of light chain portions of immunoglobulin molecules in serum, other body fluids, and tis- sues. In some disease states, an excess of light chains are produced by the antibody-forming cells. These  free light chains, unassociated with gamma globulin molecules, can be found in a patient’s body fluids and tissues. Meas- urement of the various amounts of the different types of light chains aids in the diagnosis of multiple myeloma (cancer of antibody-forming cells), lymphocytic neoplasms (cancer of lymphoid tissue), Waldenstrom’s macroglobulinemia (increased produc- tion of large immunoglobulins), and connective tissue diseases such as rheumatoid arthritis or systemic lupus erythematosus.

- Classification. Class II (perform- ance standards).

##### § 866.5560   Lactic	dehydrogenase immunological test system.

- Identification. A lactic dehydro- genase immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the activity of the lactic dehydrogenase enzyme in serum. In- creased levels of lactic dehydrogenase are found in a variety of conditions, in- cluding megaloblastic anemia (de- crease in the number of mature red blood cells), myocardial infarction (heart disease), and some forms of leu- kemia (cancer of the blood-forming or- gans). However, the diagnostic useful- ness of this device is limited because of the many conditions known to cause increased lactic dehydrogenase levels.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5570   Lactoferrin	immunological test system.

- Identification. A lactoferrin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the lactoferrin (an iron-binding protein with the ability to inhibit the growth  of  bacteria)  in  serum,  breast

#### Food and Drug Administration, HHS	§ 866.5660

milk, other body fluids, and tissues. Measurement of lactoferrin may aid in the diagnosis of an inherited deficiency of this protein.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2312, Jan. 14, 2000]

##### § 866.5580   Alpha-1-lipoprotein immunological test system.

- Identification. An alpha-1- lipoprotein immunological test  system is a device that consists of the reagents used to measure by immunochemical techniques the alpha-1-lipoprotein (high-density lipoprotein) in serum and plasma. Measurement of alpha-1- lipoprotein may aid in the diagnosis of Tangier disease (a hereditary  disorder of fat metabolism).

- Classification. Class II (perform- ance standards).

##### § 866.5590   Lipoprotein   X   immunolog- ical test system.

- Identification. A lipoprotein X immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques lipoprotein X (a high-density lipoprotein) in serum and other body fluids. Measurement of lipoprotein X aids in the diagnosis of obstructive liver disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2313, Jan. 14, 2000]

##### § 866.5600   Low-density	lipoprotein immunological test system.

- Identification. A low-density lipoprotein immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the low-density lipoprotein in serum and other body fluids. Meas- urement of low-density lipoprotein in serum may aid in the diagnosis of dis- orders of lipid (fat) metabolism and help to identify young persons at risk from cardiovascular diseases.

- Classification. Class II (perform- ance standards).

##### § 866.5620   Alpha-2-macroglobulin immunological test system.

- Identification. An alpha-2- macroglobulin immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the alpha- 2-macroglobulin (a serum protein) in plasma. Measurement of alpha-2- macroglobulin may aid in the diagnosis of blood-clotting or clot lysis disorders.

- Classification. Class II (perform- ance standards).

##### § 866.5630   Beta-2-microglobulin immunological test system.

- Identification. A beta-2-microglob- ulin immunological test system is a de- vice that consists of the reagents used to measure by immunochemical tech- niques beta-2-microglobulin (a protein molecule) in serum, urine, and other body fluids. Measurement of beta-2- microglobulin aids in the diagnosis of active rheumatoid arthritis and kidney disease.

- Classification. Class II (perform- ance standards).

##### § 866.5640   Infectious	mononucleosis immunological test system.

- Identification. An infectious mono- nucleosis immunological test system is a device that consists of the reagents used to measure by immunochemical techniques heterophile antibodies fre- quently associated with infectious mononucleosis in serum, plasma, and other body fluids. Measurements of these antibodies aid in the diagnosis of infectious mononucleosis.

- Classification. Class II (perform- ance standards).

[47 FR 50823, Nov. 9, 1982; 47 FR 56846, Dec. 21,

1982]

##### § 866.5660   Multiple	autoantibodies immunological test system.

- Identification. A multiple autoantibodies immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the autoantibodies (antibodies produced against   the   body’s   own   tissues)   in

#### § 866.5680

serum and other body fluids. Measure- ment  of  multiple  autoantibodies   aids in the diagnosis of autoimmune dis- orders (disease produced when the body’s own tissues are injured by autoantibodies).

- Classification. Class II (perform- ance standards).

##### § 866.5680   Myoglobin	immunological test system.

- Identification. A myoglobin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the myoglobin (an oxygen stor- age protein found in muscle) in serum and other body fluids. Measurement of myoglobin aids in the rapid diagnosis of heart or renal disease.

- Classification. Class II (perform- ance standards).

##### § 866.5700   Whole   human   plasma   or serum immunological test system.

- Identification. A whole human plasma or serum immunological test system is a device that consists of re- agents used to measure by immunochemical techniques the pro- teins in plasma or serum. Measure- ments of proteins in plasma or serum aid in the diagnosis of any disease con- cerned with abnormal levels of plasma or serum proteins, e.g., agammaglobu- linemia, allergies, multiple myeloma, rheumatoid vasculitis, or hereditary angioneurotic edema.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  59

FR 63007,  Dec. 7,  1994;  66  FR 38793,  July 25,

2001]

##### § 866.5715   Plasminogen  immunological test system.

- Identification. A plasminogen immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the plasminogen (an inactive substance from which plasmin, a blood- clotting factor, is formed) in serum, other body fluids, and tissues. Measure- ment of plasminogen levels may aid in

#### 21 CFR Ch. I (4–1–17 Edition)

the   diagnosis   of   fibrinolytic   (blood- clotting)  disorders.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2313, Jan. 14, 2000]

##### § 866.5735   Prothrombin immunological test system.

- Identification. A prothrombin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the prothrombin (clotting factor

II) in serum. Measurements of the amount of antigenically competent (ability to react with protein anti- bodies) prothrombin aid in the diag- nosis of blood-clotting disorders.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9. This exemption does not apply to multipurpose  systems  for in vitro coagulation studies classified under § 864.5425 of this chapter or pro- thrombin    time    tests    classified    under

§ 864.7750 of this chapter.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2313, Jan. 14, 2000]

##### § 866.5750   Radioallergosorbent  (RAST) immunological test system.

- Identification.	A radioallergosorbent	immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the aller- gen antibodies (antibodies which cause an allergic reaction) specific for a given allergen. Measurement of specific allergen antibodies may aid in the di- agnosis of asthma, allergies, and other pulmonary disorders.

- Classification. Class II (perform- ance standards).

##### § 866.5760   Tryptase test system.

- Identification. A tryptase test sys- tem is a device that aids in the diag- nosis of systemic mastocytosis. It is in- tended for in vitro diagnostic use as an aid in the clinical diagnosis of patients with      a      suspicion      of      systemic

#### Food and Drug Administration, HHS	§ 866.5860

mastocytosis in conjunction with other clinical and laboratory findings.

- Classification. Class II (special controls). The special control is FDA’s guideline entitled ‘‘Class II Special Controls Guideline: Tryptase Test Sys- tem as an Aid in the Diagnosis of Sys- temic Mastocytosis.’’ For availability of the document, see § 866.1(e).

[79 FR 56010, Sept. 18, 2014]

##### § 866.5765   Retinol-binding	protein immunological test system.

- Identification. A retinol-binding protein immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the retinol-binding protein that binds and transports vitamin A in serum and urine. Measurement of this protein may aid in the diagnosis of kid- ney disease and in monitoring patients with kidney transplants.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  65

FR 2313, Jan. 14, 2000]

##### § 866.5775   Rheumatoid factor immuno- logical test system.

- Identification. A rheumatoid factor immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the rheumatoid factor (anti- bodies to immunoglobulins) in serum, other body fluids, and tissues. Measure- ment of rheumatoid factor may aid in the diagnosis of rheumatoid arthritis.

- Classification. Class II (perform- ance standards).

§ 866.5785   Anti-Saccharomyces cerevisiae  (S.  cerevisiae)  antibody (ASCA) test systems.

- Identification. The Anti-Saccharo- myces cerevisiae (S. cerevisiae) antibody (ASCA) test system is an in vitro diag- nostic device that consists of the re- agents used to measure, by immunochemical techniques, anti- bodies to S. cerevisiae (baker’s or brew- er’s yeast) in human serum or plasma. Detection of S. cerevisiae antibodies may aid in the diagnosis of Crohn’s dis- ease.

- Classification. Class II (special controls). The special control is FDA’s ‘‘Guidance for Industry and FDA Re- viewers: Class II Special Control Guid- ance Document for Anti-Saccharomyces cerevisiae (S. cerevisiae) Antibody (ASCA) Premarket Notifications.’’

[65 FR 70307, Nov. 22, 2000]

##### § 866.5800   Seminal	fluid	(sperm) immunological test system.

- Identification. A seminal fluid (sperm) immunological test system is a device that consists of the reagents used for legal purposes to identify and differentiate animal and human semen. The test results may be used as court evidence in alleged instances of rape and other sex-related crimes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[54 FR 25047, June 12, 1989, as amended at 66

FR 38793, July 25, 2001]

##### § 866.5820   Systemic   lupus   erythema- tosus immunological test system.

- Identification. A systemic lupus erythematosus  (SLE)  immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the auto- immune antibodies in serum and other body fluids that react with cellular nu- clear double-stranded deoxyribonucleic acid (DNA) or other nuclear constitu- ents that are specifically diagnostic of SLE. Measurement of nuclear double- stranded DNA antibodies aids in the di- agnosis of SLE (a multisystem auto- immune disease in which tissues are attacked by the person’s own anti- bodies).

- Classification. Class II (perform- ance standards).

##### § 866.5860   Total  spinal  fluid  immuno- logical test system.

- Identification. A total spinal fluid immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the total protein in cerebro- spinal fluid. Measurement of spinal fluid proteins may aid in the diagnosis

#### § 866.5870

of multiple sclerosis and other diseases of the nervous system.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

##### § 866.5870   Thyroid	autoantibody immunological test system.

- Identification. A thyroid autoantibody immunological test sys- tem is a device that consists of the re- agents used to measure by immunochemical techniques the thy- roid autoantibodies (antibodies pro- duced against the body’s own tissues). Measurement of thyroid autoantibodies may aid in the diagnosis of certain thy- roid disorders, such as Hashimoto’s dis- ease (chronic lymphocytic thyroiditis), nontoxic goiter (enlargement of thy- roid gland), Grave’s disease (enlarge- ment of the thyroid gland with protru- sion of the eyeballs), and cancer of the thyroid.

- Classification. Class II (perform- ance standards).

##### § 866.5880   Transferrin    immunological test system.

- Identification. A transferrin immunological test system is a device that consists of the reagents used to measure by immunochemical tech- niques the transferrin (an iron-binding and transporting serum  protein) in serum, plasma, and other body fluids. Measurement of transferrin levels aids in the diagnosis of malnutrition, acute inflammation, infection, and red blood cell disorders, such as iron deficiency anemia.

- Classification. Class II (perform- ance standards).

##### § 866.5890   Inter-alpha   trypsin   inhib- itor immunological test system.

- Identification. An inter-alpha trypsin inhibitor immunological test system is a device that consists of the reagents used to measure by immunochemical techniques the inter- alpha trypsin inhibitor (a protein) in serum and other body fluids. Measure- ment  of  inter-alpha  trypsin  inhibitor

#### 21 CFR Ch. I (4–1–17 Edition)

may aid in the diagnosis of acute bac- terial infection and inflammation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 866.9.

[47  FR  50823,  Nov.  9,  1982,  as  amended  at  53

FR  11253,  Apr.  6,  1988;  65  FR  2313,  Jan.  14,

2000]

##### § 866.5900   Cystic	fibrosis transmembrane conductance regu- lator (CFTR) gene mutation detec- tion system.

- Identification. The CFTR gene mu- tation detection system is a device used to simultaneously detect and identify a panel of mutations and variants in the CFTR gene. It is in- tended as an aid in confirmatory diag- nostic testing of individuals with sus- pected cystic fibrosis (CF), carrier identification, and newborn screening. This device is not intended for stand- alone diagnostic purposes, prenatal di- agnostic, pre-implantation, or popu- lation screening.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: CFTR Gene Mutation Detection Sys- tem.’’ See § 866.1(e) for the availability of this guidance document.

[70 FR 61738, Oct. 26, 2005]

##### § 866.5910   Quality control material for cystic fibrosis nucleic acid assays.

- Identification. Quality control ma- terial for cystic fibrosis nucleic acid assays. A quality control material for cystic fibrosis nucleic acid assays is a device intended  to help  monitor reli- ability of a test system by detecting analytical deviations such as  those that may arise from reagent or instru- ment variation in genetic testing. This type of device includes recombinant, synthetic, and cell line-based DNA con- trols.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Quality Control Material for Cystic Fi- brosis    Nucleic    Acid    Assays.’’    See

#### Food and Drug Administration, HHS	§ 866.5940

§ 866.1(e)   for   the   availability   of   this guidance document.

[72 FR 1176, Jan. 10, 2007]

##### § 866.5940 Autosomal recessive carrier screening gene mutation detection system.

  - Identification. Autosomal reces- sive carrier screening gene mutation detection system is a  qualitative  in vitro molecular diagnostic system used for genotyping of clinically relevant variants in genomic DNA isolated from human specimens intended for pre- scription use or over-the-counter use. The device is intended for autosomal recessive disease carrier screening in adults of reproductive  age.  The  device is not intended for copy number vari- ation, cytogenetic, or biochemical test- ing.

  - Classification. Class II (special controls). Autosomal recessive carrier screening gene mutation detection sys- tem must comply with the following special  controls:

- If the device is offered over-the- counter, the device manufacturer must provide information to a potential pur- chaser or actual test report recipient about how to obtain access to a board- certified clinical molecular geneticist or equivalent to assist in pre- and post- test counseling.

- The device must use a collection device that is FDA cleared, approved, or classified as 510(k) exempt, with an indication for in vitro diagnostic use in DNA testing.

- The device’s labeling must include a prominent hyperlink to the manufac- turer’s public Web site where the man- ufacturer shall make the information identified in this section publicly available. The manufacturer’s home page, as well as the primary part of the manufacturer’s Web site that discusses the device, must provide a prominently placed hyperlink to the Web page con- taining this information and must allow unrestricted viewing access. If the device can be purchased from the Web site or testing using the device can be ordered from the Web site, the same information must be found on the Web page for ordering the device or provided in a prominently placed and publicly accessible hyperlink on the Web page for ordering the device. Any

changes to the device that could sig- nificantly affect safety or effectiveness would require new data or information in support of such changes, which would also have to be posted on the manufacturer’s Web site. The informa- tion must include:

- A detailed device description in- cluding:

  - Gene (or list of the genes if more than one) and variants the test detects (using standardized nomenclature, Human Genome Organization (HUGO) nomenclature, and coordinates).

  - Scientifically established clinical validity of each variant detected and reported by the test, which must be well-established in peer-reviewed jour- nal  articles,  authoritative  summaries of the  literature  such  as  Genetics Home  Reference

GeneReviews	(or similar summaries of valid scientific evidence,   and/or   professional   society recommendations,  including:

    - Genotype-phenotype information for the reported mutations.

    - Relevant American College of Medical Genetics (ACMG) or American Congress of Obstetricians and Gyne- cologists (ACOG) guideline recom- mending testing of the specific gene(s) and variants the test detects and rec- ommended populations, if available. If not available, a statement stating that professional guidelines currently do not recommend testing for this specific gene(s) and variants.

    - Table of expected prevalence of carrier status in major ethnic and ra- cial populations and the general popu- lation.

  - The specimen type (e.g., saliva, whole blood), matrix, and volume.

  - Assay steps and technology used.

  - Specification of required ancil- lary reagents, instrumentation, and equipment.

  - Specification of the specimen col- lection, processing, storage, and prepa- ration methods.

  - Specification of risk mitigation elements and description of all addi- tional procedures, methods, and prac- tices incorporated into the directions for use that mitigate risks associated with testing.

#### § 866.5940

  - Information pertaining to the probability of test failure (e.g., failed quality control) based on data from clinical samples, description of sce- narios in which a test can fail (i.e., low sample volume, low DNA concentra- tion, etc.), how customers will be noti- fied, and followup actions to be taken.

(I) Specification of the criteria for test result interpretation and report- ing.

- Information that demonstrates the performance characteristics of the device, including:

- Accuracy (method comparison) of study results for each claimed speci- men type.

  - Accuracy of the device shall be evaluated with fresh clinical specimens collected and processed in a manner consistent with the device’s instruc- tions for use. If this is impractical, fresh clinical samples may be sub- stituted or supplemented with archived clinical samples. Archived  samples shall have been collected previously in accordance with the device’s instruc- tions for use, stored appropriately, and randomly selected. In some instances, use of contrived samples or human cell line samples may also be appropriate; the contrived or human cell line sam- ples shall mimic clinical specimens as much as is feasible and provide an un- biased evaluation of the device’s accu- racy.

  - Accuracy must be evaluated as compared to  bidirectional  sequencing or other methods identified as appro- priate by  FDA.  Performance  criteria for both the comparator method and device must be predefined and appro- priate to the test’s intended use. De- tailed  appropriate  study  protocols must be provided.

  - Information provided shall in- clude the number and type of speci- mens, broken down by clinically rel- evant variants, that were compared to bidirectional sequencing or other methods identified as appropriate by FDA. The accuracy, defined as positive percent agreement  (PPA) and  negative percent agreement (NPA), must be measured; accuracy point estimates must be greater than 99 percent (both per reported variant and overall) and uncertainty of the point estimate must be presented using the 95 percent con-

#### 21 CFR Ch. I (4–1–17 Edition)

fidence interval. Clinical specimens must include both homozygous wild type and heterozygous genotypes. The number of clinical specimens for each variant reported that must be included in the accuracy study must be based on the variant prevalence. Common variants (greater than 0.1 percent allele frequency in ethnically relevant popu- lation) must have at least 20 unique heterozygous clinical specimens tested. Rare variants (less than or equal to 0.1 percent allele frequency in ethnically relevant population) shall have at least three unique  mutant  heterozygous specimens tested. Any no calls (i.e., ab- sence of a result) or invalid calls (e.g., failed quality control) in the  study must be included in accuracy study re- sults and reported separately. Variants that have a point estimate for PPA or NPA of less than 99 percent (incorrect test results as compared to bidirectional sequencing or other methods identified as appropriate by FDA) must not be incorporated into test claims and reports. Accuracy measures generated from clinical speci- mens versus contrived samples or cell lines must be presented separately. Re- sults must be summarized and pre- sented in  tabular  format,  by  sample and by genotype.  Point  estimate  of PPA should be calculated as the  num- ber of positive results divided by the number of specimens known to harbor variants (mutations)  without  ‘‘no calls’’ or invalid calls. The point esti- mate of NPA should be calculated  as the number of negative results divided by the number of wild type specimens tested without ‘‘no calls’’ or  invalid calls, for each variant that is being re- ported. Point estimates should be cal- culated along with 95 percent two-sided confidence  intervals.

  - Information shall be reported on the clinical positive predictive value (PPV) and negative predictive value (NPV) for carrier status (and where possible, for each variant) in each pop- ulation. Specifically, to  calculate  PPV and NPV, estimate test coverage (TC) and the percent of persons with vari- ant(s) included in the device among all carriers: PPV = (PPA * TC * )/(PPA * TC * + (1 ¥ NPA) * (1 ¥ )) and NPV

= (NPA * (1 ¥ ))/(NPA *(1 ¥ ) + (1 ¥

#### Food and Drug Administration, HHS	§ 866.5940

PPA*TC) * ) where PPA and NPA de- scribed	either	in	paragraph (b)(3)(ii)(A)(4)(i)     or     in     paragraph

(b)(3)(ii)(A)(4)(ii) of this section and is prevalence of carriers in the population (pre-test risk to be a carrier for the dis- ease).

    - For the point estimates of PPA and NPA less than 100 percent, use the calculated estimates in the PPV and NPV  calculations.

    - Point estimates of 100 percent may have high uncertainty. If these variants are measured using highly multiplexed technology, calculate the random error rate for the overall de- vice and incorporate that rate in the estimation of the PPA and NPA as cal- culated previously. Then use these cal- culated estimates in the PPV and NPV calculations. This type of accuracy study is helpful in determining that there is no systematic error in such de- vices.

- Precision (reproducibility): Preci- sion data must be generated using mul- tiple instruments and multiple opera- tors, on multiple non-consecutive days, and using multiple reagent  lots. The sample panel must include specimens with claimed sample type (e.g. saliva samples)	representing	different genotypes		(i.e.,	wild		type, heterozygous). Performance criteria must be predefined. A detailed study protocol must be created in advance of the study and  then  followed.  The ‘‘failed quality control’’ rate must be indicated. It must be clearly docu- mented whether results were generated from clinical specimens, contrived samples, or cell lines. The study results shall state, in a tabular format, the variants tested in the study and the number of replicates for each variant, and what testing conditions were stud- ied (i.e., number of runs, days, instru- ments, reagent lots, operators, speci- mens/type, etc). The study must in- clude all nucleic acid extraction steps from the claimed specimen type or ma- trix, unless a separate extraction study for the claimed sample type is per- formed. If the device is to be used at more than one laboratory, different laboratories must be included in the precision study (and reproducibility must be evaluated). The percentage of ‘‘no  calls’’  or  invalid  calls,  if  any,  in

the study must be provided as a part of the precision (reproducibility) study results.

- Analytical specificity data: Data must be generated evaluating the ef- fect on test performance of potential endogenous and exogenous interfering substances relevant to the specimen type, evaluation of cross-reactivity of known cross-reactive alleles and pseudogenes, and  assessment  of  cross- contamination.

- Analytical sensitivity data: Data must be generated demonstrating the minimum amount of DNA that will en- able the test to perform  accurately  in 95 percent of runs.

- Device stability data: The manu- facturer must establish upper and lower limits of input nucleic acid and sample stability that will achieve the claimed accuracy and reproducibility. Data supporting such claims must be described.

- Specimen type and matrix com- parison data: Specimen type and ma- trix comparison data must be gen- erated if more than one specimen type or anticoagulant can be tested with the device, including failure rates for the different specimen types.

    - If the device is offered over-the- counter, including cases in which the test results are provided direct-to-con- sumer, the manufacturer must conduct a study that assesses user comprehen- sion of the device’s labeling and test process and provide a concise summary of the results of the study. The fol- lowing items must be included in the user study:

      - The test manufacturer must per- form pre- and post-test user com- prehension studies to assess user abil- ity to understand the possible results of a carrier test and their clinical meaning. The comprehension test ques- tions must directly evaluate the mate- rial being presented to the user in the test reports.

      - The test manufacturer must pro- vide a carrier testing education module to potential and actual test report re- cipients. The module must define terms that are used in the test reports and explain the significance of carrier status.

      - The user study must meet the fol- lowing criteria:

#### § 866.5940

        - The study participants must be comprised of a statistically justified and demographically diverse popu- lation (determined using methods such as quota-based sampling) that is rep- resentative of the intended user popu- lation. Furthermore, the users must be comprised of a diverse range of age and educational levels that have no prior experience with the test or its manu- facturer. These factors shall be well-de- fined in the inclusion and exclusion criteria.

        - All sources of bias (e.g., non-re- sponders) must be predefined and ac- counted for in the study results with regard to both responders and non-re- sponders.

        - The testing must follow a format where users have limited time to com- plete the studies (such as an onsite sur- vey format and a one-time visit with a cap on the maximum amount of time that a participant has to complete the tests).

        - Users must be randomly assigned to study arms. Test reports given to users must: Define the condition being tested and related symptoms; explain the intended use and limitations of the test; explain the relevant ethnicities regarding the variant tested; explain carrier status and relevance to the user’s ethnicity; and provide links to additional information pertaining to situations where the user is concerned about their test results or would like followup information as indicated in test labeling. The study shall assess participants’ ability to understand the following comprehension concepts: The test’s limitations, purpose, and results.

        - Study participants must be un- trained, naive to the test subject of the study, and be provided only the mate- rials that will be available to them when the test is marketed.

        - The user comprehension study must meet the predefined primary end- point criteria,  including  a  minimum  of a 90 percent or greater overall com- prehension rate (i.e. selection of the correct answer) for each comprehen- sion concept to demonstrate that the education module and test reports are adequate for over-the-counter use.

      - A summary of the user com- prehension study must be provided and include the following:

#### 21 CFR Ch. I (4–1–17 Edition)

- Results regarding reports that are provided for each gene/variant/eth- nicity  tested.

- Statistical methods used to ana- lyze all data sets.

- Completion rate, non-responder rate, and reasons for non-response/data exclusion, as well as  a  summary  table of comprehension rates regarding com- prehension concepts (purpose of test, test results, test limitations, ethnicity relevance for the test results, etc.) for each study report.

- Your 21 CFR 809.10 compliant la- beling and any test report generated must include the  following  warning and limitation statements, as applica- ble:

  - A warning that reads ‘‘The test is intended only for autosomal recessive carrier screening in adults of reproduc- tive age.’’

  - A statement accurately dis- closing the genetic coverage of the test in lay terms, including, as applicable, information on variants not queried by the test, and the proportion of incident disease that is not related to the gene(s) tested. For example, where ap- plicable, the statement would have to include a warning that the test does not or may not detect all genetic variants related to the genetic disease, and that the absence of a variant test- ed does not rule out the presence of other genetic variants that may be dis- ease-related. Or, where applicable, the statement would have to include a warning that the basis for the  disease for which the genetic carrier status is being tested is unknown or believed to be non-heritable in a substantial num- ber of people who have the disease, and that a negative test result cannot  rule out the possibility that any offspring may be affected with the disease. The statement would have to include any other warnings needed to accurately convey to consumers the degree to which the test is informative for car- rier  status.

  - For prescription use tests, the following warnings that read:

- ‘‘The results of this test are in- tended to be interpreted by a board- certified clinical molecular geneticist or equivalent and should be used in conjunction with other available lab- oratory and clinical information.’’

#### Food and Drug Administration, HHS	§ 866.6010

- ‘‘This device is not intended for disease diagnosis, prenatal testing of fetuses, risk assessment, prognosis or pre-symptomatic testing, suscepti- bility testing, or newborn screening.’’

  - For over-the-counter tests, a statement that reads ‘‘This test is not intended to diagnose a disease, or tell you anything about your risk for devel- oping a disease in the future. On its own, this test is also not  intended  to tell you anything about the health of your fetus, or  your  newborn  child’s risk of developing a particular disease later on in life.’’

  - For over-the-counter tests, the following warnings that read:

    - ‘‘This test is not a substitute for visits to a healthcare provider. It is recommended that you consult with a healthcare provider if you have any questions or concerns about your re- sults.’’

    - ‘‘The test does not diagnose any health conditions. Results should be used along with other clinical informa- tion for any medical purposes.’’

    - ‘‘The laboratory may not be able to process your sample. The prob- ability that the laboratory cannot process your saliva sample can be up to [actual probability percentage].’’

    - ‘‘Your ethnicity may affect how your genetic health results are inter- preted.’’

  - For a positive result in an over- the-counter test when the positive pre- dictive value for  a  specific  population is less than 50 percent and more than 5 percent, a warning that reads ‘‘The positive result you obtained may false- ly identify you as a carrier. Consider genetic counseling and followup test- ing.’’

  - For a positive result in an over- the-counter test when the positive pre- dictive value for  a  specific  population is less than 5 percent, a warning that reads ‘‘The positive result you obtained is very likely to be incorrect due to the rarity of this variant. Consider genetic counseling and followup testing.’’

- The testing done to comply with paragraph (b)(3) of this section must show the device meets or exceeds each of the following performance specifica- tions:

  - The accuracy must be shown to be equal to or greater than 99 percent for

both PPA  and  NPA.  Variants that  have a point estimate for PPA or NPA  of less than 99 percent (incorrect test re- sults as compared to bidirectional se- quencing or other methods  identified as appropriate by FDA) must not be in- corporated into test claims and re- ports.

  - Precision (reproducibility) per- formance must meet or exceed 99 per- cent for both positive and negative re- sults.

  - The user comprehension study must obtain values of 90 percent or greater user comprehension for each comprehension concept.

- The distribution of this device, ex- cluding the collection device described in paragraph (b)(2) of this section, shall be limited to the manufacturer, the manufacturer’s subsidiaries, and lab- oratories regulated under the Clinical Laboratory Improvement Amend- ments.

[80 FR 65630, Oct. 27, 2015]

## Subpart G—Tumor Associated Antigen immunological Test Systems

##### § 866.6010 Tumor-associated antigen immunological test system.

- Identification. A tumor-associated antigen immunological test system is a device that consists of reagents used to qualitatively or quantitatively meas- ure, by immunochemical techniques, tumor-associated antigens in serum, plasma, urine, or  other  body  fluids. This device is intended as an aid in monitoring patients  for  disease progress or response to therapy or for the detection of recurrent or residual disease.

- Classification.   Class    II    (special controls). Tumor markers must comply with the following special controls: (1) A  guidance  document  entitled  ‘‘Guid- ance  Document  for  the  Submission  of Tumor   Associated   Antigen   Premarket Notifications (510(k)s) to FDA,’’ and (2) voluntary assay performance standards issued  by  the  National  Committee  on Clinical  Laboratory  Standards.

[62 FR 66005, Dec. 17, 1997]

§ 866.6020

##### § 866.6020 Immunomagnetic  circu- lating cancer cell selection and enu- meration system.

  - Identification.	An immunomagnetic circulating cancer cell selection and enumeration system is a device that consists of biological probes, fluorochromes, and other re- agents; preservation and preparation devices; and a semiautomated analyt- ical instrument to select and count cir- culating cancer  cells  in  a  prepared sample of whole blood. This device is intended for adjunctive use in moni- toring or predicting cancer disease pro- gression, response to therapy, and for the detection of recurrent disease.

  - Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document: Immunomagnetic Cir- culating Cancer Cell Selection and Enumeration System.’’ See § 866.1(e) for availability  of  this  guidance  document.

[69 FR 26038, May 11, 2004]

##### § 866.6030 AFP-L3% immunological test system.

- Identification. An AFP-L3% immunological test system is an in vitro device that consists of reagents and an automated instrument used to quantitatively	measure,	by immunochemical techniques, AFP and AFP-L3 subfraction in human serum. The device is intended for in vitro diag- nostic use as an aid in the risk assess- ment of patients with chronic liver dis- ease for development of hepatocellular carcinoma, in conjunction with other laboratory findings, imaging studies, and clinical assessment.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: AFP-L3% Immunological Test Sys- tems.’’ See § 866.1(e) for the availability of this guidance document.

[70 FR 57749, Oct. 4, 2005]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 866.6040 Gene expression profiling test system for breast cancer prog- nosis.

- Identification. A gene  expression profiling test system for breast cancer prognosis is a device that measures the ribonucleic acid (RNA) expression level of multiple genes and combines this in- formation to yield a signature (pattern or classifier or index) to aid in prog- nosis of previously diagnosed breast cancer.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Gene Expression Profiling Test System for    Breast    Cancer    Prognosis.’’    See

§ 866.1(e)   for   the   availability   of   this guidance document.

[72 FR 26291, May 9, 2007]

##### § 866.6050 Ovarian adnexal mass as- sessment score test system.

- Identification. An ovarian/adnexal mass assessment test system is a de- vice that measures one or more pro- teins in serum or plasma. It yields a single result for the likelihood that an adnexal pelvic mass in a woman, for whom surgery is planned, is malignant. The test is for adjunctive use, in the context of a negative primary clinical and radiological evaluation, to aug- ment the identification of patients whose gynecologic surgery requires on- cology expertise and resources.

- Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document:  Ovarian  Adnexal Mass Assessment Score  Test  System.’’  For the availability of this guidance docu- ment, see § 866.1(e).

- Black box warning. Under section 520(e) of the Federal Food, Drug, and Cosmetic Act these devices are subject to the following restriction: A warning statement must be placed in a black box and must appear in all advertising, labeling, and promotional material for these devices. That warning statement must read:

#### Food and Drug Administration, HHS	Pt. 868

[76 FR 16294, Mar. 23, 2011, as amended at 76

FR 82131, Dec. 30, 2011]

## PART 868—ANESTHESIOLOGY DEVICES

##### Subpart A—General Provisions

Sec.

868.1    Scope.

868.3   Effective dates of requirement for pre- market  approval.

868.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

868.1030   Manual algesimeter.

868.1040   Powered algesimeter. 868.1075   Argon gas analyzer.

868.1100  Arterial blood sampling kit.

868.1120   Indwelling     blood     oxyhemoglobin concentration    analyzer.

868.1150   Indwelling    blood    carbon    dioxide partial  pressure  (PCO2)  analyzer.

868.1170   Indwelling blood hydrogen ion con- centration  (pH)  analyzer.

868.1200    Indwelling    blood    oxygen    partial pressure  (PO2)  analyzer.

868.1400   Carbon dioxide gas analyzer. 868.1430   Carbon monoxide gas analyzer. 868.1500   Enflurane gas analyzer.

868.1575   Gas collection vessel. 868.1620   Halothane gas analyzer. 868.1640   Helium gas analyzer.

868.1670   Neon gas analyzer. 868.1690    Nitrogen gas analyzer.

868.1700    Nitrous oxide gas analyzer. 868.1720    Oxygen gas analyzer.

868.1730   Oxygen uptake computer. 868.1750   Pressure plethysmograph.

868.1760   Volume plethysmograph.

868.1780    Inspiratory airway pressure meter. 868.1800  Rhinoanemometer.

868.1840   Diagnostic spirometer.

868.1850   Monitoring spirometer.

868.1860    Peak-flow meter for spirometry. 868.1870   Gas volume calibrator.

868.1880    Pulmonary-function      data      calcu- lator.

868.1890  Predictive	pulmonary-function value    calculator.

868.1900    Diagnostic   pulmonary-function   in- terpretation     calculator.

868.1910   Esophageal stethoscope.

868.1920    Esophageal  stethoscope  with  elec- trical    conductors.

868.1930  Stethoscope head. 868.1965   Switching valve (ploss). 868.1975   Water vapor analyzer.

##### Subpart C—Monitoring Devices

868.2025    Ultrasonic air embolism monitor. 868.2300   Bourdon gauge flowmeter.

868.2320   Uncompensated  thorpe   tube  flow- meter.

868.2340   Compensated    thorpe    tube    flow- meter.

868.2350    Gas calibration flowmeter. 868.2375    Breathing frequency monitor. 868.2377  Apnea monitor.

868.2380    Nitric oxide analyzer. 868.2385    Nitrogen dioxide analyzer. 868.2450    Lung water monitor.

868.2480   Cutaneous  carbon  dioxide  (PcCO2) monitor.

868.2500    Cutaneous oxygen (PcO2) monitor. 868.2550   Pneumotachometer.

868.2600    Airway pressure monitor. 868.2610   Gas pressure gauge.

868.2620   Gas pressure calibrator. 868.2700   Pressure regulator.

868.2775   Electrical  peripheral  nerve  stimu- lator.

868.2875    Differential pressure transducer. 868.2885   Gas flow transducer.

868.2900   Gas pressure transducer.

##### Subparts D–E [Reserved]

Subpart F—Therapeutic Devices

868.5090   Emergency airway needle. 868.5100   Nasopharyngeal airway.

868.5110   Oropharyngeal airway.

868.5115    Device  to  relieve  acute  upper  air- way   obstruction.

868.5120   Anesthesia conduction catheter. 868.5130   Anesthesia conduction filter.

868.5140    Anesthesia conduction kit. 868.5150    Anesthesia conduction needle.

§ 868.1

21 CFR Ch. I (4–1–17 Edition)

| 868.5160 Gas machine for anesthesia or anal- | 868.5895 | Continuous ventilator. |
| --- | --- | --- |
| gesia. | 868.5905 | Noncontinuous ventilator (IPPB). |
| 868.5165 Nitric oxide administration appa- | 868.5915 | Manual emergency ventilator. |
| ratus. | 868.5925 | Powered emergency ventilator. |
| 868.5170 Laryngotracheal topical anesthesia | 868.5935 | External negative pressure venti- |

applicator.

868.5180    Rocking bed.

868.5220    Blow bottle.

868.5240   Anesthesia breathing circuit. 868.5250   Breathing circuit circulator. 868.5260    Breathing circuit bacterial filter. 868.5270  Breathing system heater.

868.5280   Breathing tube support. 868.5300   Carbon dioxide absorbent. 868.5310  Carbon dioxide absorber. 868.5320  Reservoir bag.

868.5330   Breathing gas mixer. 868.5340   Nasal oxygen cannula. 868.5350    Nasal oxygen catheter.

868.5365    Posture   chair   for   cardiac   or   pul- monary    treatment.

868.5375   Heat and moisture condenser (arti- ficial   nose).

868.5400    Electroanesthesia apparatus.

868.5420   Ether hook.

868.5430   Gas-scavenging apparatus. 868.5440    Portable oxygen generator. 868.5450    Respiratory gas humidifier.

868.5460   Therapeutic   humidifier   for   home use.

868.5470  Hyperbaric chamber.

868.5530   Flexible laryngoscope.

868.5540   Rigid laryngoscope. 868.5550    Anesthetic gas mask. 868.5560   Gas mask head strap. 868.5570   Nonrebreathing mask.

868.5580   Oxygen mask.

868.5590   Scavenging mask.

868.5600    Venturi mask.

868.5620   Breathing mouthpiece.

868.5630   Nebulizer.

868.5640    Medicinal  nonventilatory  nebulizer (atomizer).

868.5650   Esophageal obturator. 868.5655    Portable liquid oxygen unit. 868.5665  Powered percussor.

868.5675   Rebreathing device.

868.5690   Incentive spirometer. 868.5700   Nonpowered oxygen tent.

868.5710   Electrically powered oxygen tent. 868.5720  Bronchial tube.

868.5730   Tracheal tube.

868.5740    Tracheal/bronchial differential ven- tilation    tube.

868.5750    Inflatable tracheal tube cuff. 868.5760  Cuff spreader.

868.5770    Tracheal tube fixation device. 868.5780   Tube introduction forceps.

868.5790   Tracheal tube stylet.

868.5795    Tracheal tube cleaning brush. 868.5800    Tracheostomy tube and tube cuff. 868.5810   Airway connector.

868.5820   Dental protector.

868.5830    Autotransfusion apparatus. 868.5860   Pressure tubing and accessories. 868.5870   Nonrebreathing valve.

868.5880   Anesthetic vaporizer.

lator.

868.5955    Intermittent     mandatory     ventila- tion    attachment.

868.5965   Positive   end   expiratory   pressure breathing     attachment.

868.5975    Ventilator tubing. 868.5995   Tee drain (water trap).

##### Subpart G—Miscellaneous

868.6100 Anesthetic cabinet, table, or tray. 868.6175 Cardiopulmonary emergency cart. 868.6225  Nose clip.

868.6250   Portable air compressor. 868.6400   Calibration gas.

868.6700   Anesthesia stool.

868.6810     Tracheobronchial suction catheter. 868.6820    Patient position support.

868.6885    Medical gas yoke assembly.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE: 47 FR 31142, July 16, 1982, unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 868 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 868.1   Scope.

- This part sets forth the classifica- tion of anesthesiology devices intended for human use that are in commercial distribution.

- The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

- To avoid duplicative listings, an anesthesiology device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed only in one subpart.

- References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

Food and Drug Administration, HHS	§ 868.9

- Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 17734, May 11, 1987, as amended at 67

FR 76681,  Dec. 13,  2002;  78  FR 18233,  Mar. 26,

2013]

##### § 868.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer-

cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17734, May 11, 1987]

##### § 868.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

#### § 868.1030

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2313, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 868.1030   Manual algesimeter.

    - Identification. A manual algesim- eter is a mechanical device intended to determine a patient’s sensitivity to pain after administration of an anes- thetic agent, e.g., by pricking with a sharp point.

21 CFR Ch. I (4–1–17 Edition)

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[54 FR 25048, June 12, 1989, as amended at 66

FR 38793, July 25, 2001]

##### § 868.1040   Powered algesimeter.

- Identification. A powered algesim- eter is a device using electrical stimu- lation intended to determine a pa- tient’s sensitivity to pain after admin- istration of an anesthetic agent.

- Classification. Class II (perform- ance standards).

##### § 868.1075   Argon gas analyzer.

- Identification. An  argon  gas ana- lyzer is a device intended to measure the concentration of argon in a gas mixture to aid in determining the pa- tient’s ventilatory status. The device may use techniques such as mass spec- trometry or thermal conductivity.

- Classification. Class II (perform- ance standards).

##### § 868.1100   Arterial blood sampling kit.

- Identification. An arterial blood sampling kit is a device, in kit form, used to obtain arterial blood samples from a patient for blood gas determina- tions. The kit may include a syringe, needle, cork, and heparin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

##### §           868.1120              Indwelling	blood oxyhemoglobin   concentration  ana- lyzer.

- Identification. An indwelling blood oxyhemoglobin concentration analyzer

#### Food and Drug Administration, HHS	§ 868.1430

is a photoelectric device used to meas- ure, in vivo, the oxygen-carrying ca- pacity of hemoglobin in blood to aid in determining the patient’s physiological status.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before September 21, 2004, for	any	indwelling	blood oxyhemoglobin concentration analyzer that was in commercial distribution before May 28, 1976, or that has, on or before September 21, 2004, been found to be substantially equivalent to an in- dwelling blood oxyhemoglobin con- centration analyzer that was in com- mercial distribution before May 28, 1976. Any other indwelling blood oxyhemoglobin concentration analyzer shall have an approved PMA or de- clared completed PDP in effect before being placed in commercial distribu- tion.

[47 FR 31142, July 16, 1982, as amended at 52

FR 17735,  May 11,  1987;  52  FR 22577,  June 12,

1987; 69 FR 34920, June 23, 2004]

##### § 868.1150 Indwelling blood carbon di- oxide partial pressure (PCO2) ana- lyzer.

- Identification. An indwelling blood carbon dioxide partial pressure PCO2 analyzer is a device that consists of a catheter-tip PCO2 transducer (e.g., PCO2 electrode) and that is used to measure, in vivo, the partial pressure of carbon dioxide in blood to aid in determining the patient’s circulatory, ventilatory, and metabolic status.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Indwelling Blood Gas Analyzers; Final Guidance for Industry and FDA.’’

[47 FR 31142, July 16, 1982; 47 FR 40410, Sept.

14, 1982, as amended at 52 FR 17735, May 11,

1987; 66 FR 57368, Nov. 15, 2001]

##### § 868.1170   Indwelling  blood  hydrogen ion concentration (pH) analyzer.

- Identification. An indwelling blood hydrogen ion concentration (pH) ana- lyzer is a device that consists of a cath- eter-tip pH electrode and that is used

to measure, in vivo, the hydrogen ion concentration (pH) in blood to aid in determining the patient’s acid-base balance.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Indwelling Blood Gas Analyzers; Final Guidance for Industry and FDA.’’

[47 FR 31142, July 16, 1982, as amended at 52

FR  17735,  May  11,  1987;  66  FR  57368,  Nov.  15,

2001]

##### § 868.1200   Indwelling	blood	oxygen partial pressure (PO2) analyzer.

- Identification. An indwelling blood oxygen partial pressure (PO2) analyzer is a device that consists of a catheter- tip PO2 transducer (e.g., PO2 electrode) and that is used to measure, in vivo, the partial pressure of oxygen in blood to aid in determining the patient’s cir- culatory, ventilatory, and metabolic status.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Indwelling Blood Gas Analyzers; Final Guidance for Industry and FDA.’’

[47 FR 31142, July 16, 1982; 47 FR 40410, Sept.

14, 1982, as amended at 52 FR 17735, May 11,

1987; 66 FR 57368, Nov. 15, 2001]

##### § 868.1400   Carbon   dioxide   gas   ana- lyzer.

- Identification. A carbon dioxide gas analyzer is a device intended to measure the concentration of carbon dioxide in a gas mixture to aid in de- termining the patient’s ventilatory, circulatory, and metabolic status. The device may use techniques such as chemical titration, absorption of infra- red radiation, gas chromatography, or mass spectrometry.

- Classification. Class II (perform- ance standards).

##### § 868.1430   Carbon  monoxide  gas  ana- lyzer.

- Identification. A carbon monoxide gas analyzer is a device intended to measure the concentration of carbon monoxide in a gas mixture to aid in de- termining the patient’s ventilatory status. The device may use techniques

#### § 868.1500

such as infrared absorption or gas chro- matography.

- Classification. Class II (perform- ance standards).

##### § 868.1500   Enflurane gas analyzer.

- Identification. An enflurane gas analyzer is a device intended to meas- ure the concentration of enflurane an- esthetic in a gas mixture.

- Classification. Class II (perform- ance standards).

##### § 868.1575   Gas collection vessel.

- Identification. A gas collection vessel is a container-like device in- tended to collect a patient’s exhaled gases for subsequent analysis. It does not include a sampling pump.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

##### § 868.1620   Halothane gas analyzer.

- Identification. A halothane gas an- alyzer is a device intended to measure the concentration of halothane anes- thetic in a gas mixture. The device may use techniques such as mass spec- trometry or absorption of infrared or ultraviolet radiation.

- Classification. Class II (perform- ance standards).

##### § 868.1640   Helium gas analyzer.

- Identification. A helium gas ana- lyzer is a device intended to measure the concentration of helium in a gas mixture during pulmonary function testing. The device may use techniques such as thermal conductivity, gas chro- matography, or mass spectrometry.

- Classification. Class II (perform- ance standards).

##### § 868.1670   Neon gas analyzer.

- Identification. A neon gas analyzer is a device intended to measure the concentration of neon in a gas mixture exhaled by a patient. The device may use techniques such as mass spectrom- etry or thermal conductivity.

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 868.1690   Nitrogen gas analyzer.

- Identification. A nitrogen gas ana- lyzer is a device intended to measure the concentration of nitrogen in res- piratory gases to aid in determining a patient’s ventilatory status. The device may use techniques such as gas chro- matography or mass spectrometry.

- Classification. Class II (perform- ance standards).

##### § 868.1700   Nitrous oxide gas analyzer.

- Identification. A nitrous oxide gas analyzer is a device intended to meas- ure the concentration of nitrous oxide anesthetic in a gas mixture. The device may use techniques such as infrared absorption or mass spectrometry.

- Classification. Class II (perform- ance standards).

##### § 868.1720   Oxygen gas analyzer.

- Identification. An oxygen gas ana- lyzer is a device intended to measure the concentration of oxygen in res- piratory gases by techniques such as mass spectrometry, polarography, thermal conductivity, or gas chroma- tography. This generic type  of device also includes paramagnetic analyzers.

- Classification. Class II (perform- ance standards).

##### § 868.1730   Oxygen uptake computer.

- Identification. An oxygen uptake computer is a device intended to com- pute the amount of oxygen  consumed by a patient and may include compo- nents for determining expired gas vol- ume and composition.

- Classification. Class II (perform- ance standards).

##### § 868.1750   Pressure plethysmograph.

- Identification. A pressure plethysmograph is a device used to de- termine a patient’s airway resistance and lung volumes by measuring pres- sure changes while the patient is in an airtight  box.

- Classification. Class II (perform- ance standards).

##### § 868.1760   Volume plethysmograph.

- Identification. A volume plethysmograph is an airtight box, in

#### Food and Drug Administration, HHS	§ 868.1910

which a patient sits, that is used to de- termine the patient’s lung volume changes.

- Classification. Class II (perform- ance standards).

##### § 868.1780   Inspiratory airway pressure meter.

- Identification. An inspiratory air- way pressure meter is a device used to measure the amount of pressure pro- duced in a patient’s airway during maximal inspiration.

- Classification. Class II (perform- ance standards).

##### § 868.1800   Rhinoanemometer.

- Identification. A rhinoanemometer is a device used to quantify the amount of nasal congestion by measuring the airflow through, and differential pres- sure across, a patient’s nasal passages.

- Classification. Class II (perform- ance standards).

##### § 868.1840   Diagnostic spirometer.

- Identification. A diagnostic spi- rometer is a device used in pulmonary function testing to measure the volume of gas moving in or out of a patient’s lungs.

- Classification. Class II (perform- ance standards).

##### § 868.1850   Monitoring spirometer.

- Identification. A monitoring spi- rometer is a device used to measure continuously a patient’s tidal volume (volume of gas inhaled by the patient during each respiration cycle) or minute volume (the tidal volume mul- tiplied by  the  rate of  respiration  for 1 minute) for the evaluation of the pa- tient’s  ventilatory  status.

- Classification. Class II (perform- ance standards).

##### § 868.1860   Peak-flow	meter	for spirometry.

- Identification. A peak-flow meter for spirometry is a device used to measure a patient’s maximum venti- latory flow rate.

- Classification. Class II (perform- ance standards).

##### § 868.1870   Gas volume calibrator.

- Identification. A gas volume cali- brator is a device that is intended for

medical purposes and that is used to calibrate the output of gas volume measurement instruments by deliv- ering a known gas volume.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

##### § 868.1880   Pulmonary-function	data calculator.

- Identification. A pulmonary-func- tion data calculator is a device used to calculate pulmonary-function values based on actual physical data obtained during pulmonary-function testing.

- Classification. Class II (perform- ance standards).

##### § 868.1890   Predictive   pulmonary-func- tion value calculator.

- Identification. A predictive pul- monary-function value calculator is a device used to calculate normal pul- monary-function values based on em- pirical equations.

- Classification. Class II (perform- ance standards).

##### § 868.1900   Diagnostic  pulmonary-func- tion interpretation calculator.

- Identification. A diagnostic pul- monary-function interpretation calcu- lator is a device that interprets pul- monary study data to determine clin- ical significance of pulmonary-function values.

- Classification. Class II (perform- ance standards).

##### § 868.1910   Esophageal stethoscope.

- Identification. An esophageal stethoscope is a nonpowered device that is inserted into a patient’s esoph- agus to enable the user to listen to heart and breath sounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2313, Jan. 14, 2000]

#### § 868.1920

##### § 868.1920   Esophageal	stethoscope with electrical conductors.

  - Identification. An esophageal stethoscope with electrical conductors is a device that is inserted into the esophagus to listen to a patient’s heart and breath sounds and to monitor electrophysiological signals. The de- vice may also incorporate a thermistor for temperature measurement.

  - Classification. Class II (perform- ance standards).

##### § 868.1930   Stethoscope head.

- Identification. A stethoscope head is a weighted chest piece used during anesthesia to listen to a patient’s heart, breath, and other physiological sounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38793, July 25,

2001]

##### § 868.1965   Switching valve (ploss).

- Identification. A switching valve (ploss) is a three-way valve located be- tween a stethoscope placed over the heart, a blood pressure cuff, and an ear- piece. The valve allows the user to eliminate one sound channel and listen only to a patient’s heart or korotkoff (blood pressure) sounds through the other channel.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38793, July 25,

2001]

##### § 868.1975   Water vapor analyzer.

- Identification. A water vapor ana- lyzer is a device intended to measure

#### 21 CFR Ch. I (4–1–17 Edition)

the concentration of water vapor in a patient’s expired gases by using tech- niques such as mass spectrometry.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38793,  July 25,

2001]

## Subpart C—Monitoring Devices

##### § 868.2025   Ultrasonic	air	embolism monitor.

- Identification. An ultrasonic air embolism monitor is a device used to detect air bubbles in a patient’s blood stream. It may use Doppler or other ul- trasonic principles.

- Classification. Class II (perform- ance standards).

##### § 868.2300   Bourdon gauge flowmeter.

- Identification. A bourdon gauge flowmeter is a device intended for med- ical purposes that is used in conjunc- tion with respiratory equipment to sense gas pressure. The device is cali- brated to indicate gas flow rate when the outflow is open to the atmosphere.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2320   Uncompensated thorpe tube flowmeter.

- Identification. An uncompensated thorpe tube flowmeter is a device in- tended for medical purposes that is used to indicate and control gas flow rate accurately. The device includes a vertically mounted tube and is cali- brated when the outlet of the flow- meter is open to the atmosphere.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

Food and Drug Administration, HHS	§ 868.2450

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2340   Compensated   thorpe   tube flowmeter.

  - Identification. A compensated thorpe tube flowmeter is a device in- tended for medical purposes that is used to control and measure gas flow rate accurately. The device includes a vertically mounted tube, with the out- let of the flowmeter calibrated to a ref- erence pressure.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2350   Gas calibration flowmeter.

- Identification. A gas calibration flowmeter is a device intended for med- ical purposes that is used to calibrate flowmeters and accurately measure gas flow.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2375   Breathing   frequency   mon- itor.

- Identification. A breathing (venti- latory) frequency monitor is a device intended to measure or monitor a pa- tient’s  respiratory  rate.  The   device may provide an audible or visible alarm when the respiratory rate, aver- aged over time, is outside operator settable alarm limits. This device does not include the apnea monitor classi- fied in § 868.2377.

- Classification. Class II (perform- ance standards).

[47 FR 31142, July 16, 1982, as amended at 67

FR 46852, July 17, 2002]

##### § 868.2377   Apnea monitor.

- Identification. An apnea monitor is a complete system intended to alarm primarily upon the cessation of breath- ing timed from the last detected breath. The apnea monitor also in- cludes indirect methods of apnea detec- tion such as monitoring of heart rate and other physiological parameters linked to the presence or absence of adequate respiration.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Apnea Monitors; Guidance for Industry and FDA.’’

[67 FR 46852, July 17, 2002]

##### § 868.2380   Nitric oxide analyzer.

- Identification. The nitric oxide an- alyzer is a device intended to measure the concentration of nitric oxide in res- piratory gas mixtures during adminis- tration of nitric oxide.

- Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance  Document  for  Premarket  Notifi- cation  Submissions  for  Nitric  Oxide Administration	Apparatus,	Nitric Oxide Analyzer, and Nitrogen Dioxide Analyzer.’’

[65 FR 14465, Mar. 3, 2000]

##### § 868.2385   Nitrogen dioxide analyzer.

- Identification. The nitrogen diox- ide analyzer is a device intended to measure the concentration of nitrogen dioxide in respiratory gas mixtures during administration of nitric oxide.

- Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance  Document  for  Premarket  Notifi- cation  Submissions  for  Nitric  Oxide Administration	Apparatus,	Nitric Oxide Analyzer, and Nitrogen Dioxide Analyzer.’’

[65 FR 11465, Mar. 3, 2000]

##### § 868.2450   Lung water monitor.

- Identification. A lung water mon- itor is a device used to monitor the trend of fluid volume changes in a pa- tient’s lung by measuring changes in thoracic electrical impedance (resist- ance to alternating current) by means

#### § 868.2480

of  electrodes  placed  on  the  patient’s chest.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP for a device is re- quired to be filed with the Food and Drug Administration on or before July 12, 2000, for any lung water monitor that was in commercial distribution before May 28, 1976, or that has, on or before July 12, 2000, been found to be substantially equivalent to a  lung water monitor that was in commercial distribution before May 28, 1976. Any other lung water monitor device shall have an approved PMA or declared completed PDP in effect before being placed in commercial distribution.

[47 FR 31142, July 16, 1982, as amended at 52

FR  17735,  May  11,  1987;  65  FR  19834,  Apr.  13,

2000]

##### § 868.2480   Cutaneous   carbon   dioxide (PcCO2) monitor.

- Identification. A cutaneous carbon dioxide (PcCO2) monitor is a noninvasive heated sensor  and a  pH- sensitive glass electrode placed on a patient’s skin, which is intended to monitor relative changes in a hemodynamically stable patient’s cu- taneous carbon dioxide tension as an adjunct to arterial carbon dioxide ten- sion measurement.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Cutaneous Carbon Dioxide (PcCO2) and Oxygen (PcO2) Monitors; Guidance for Industry and FDA.’’ See § 868.1(e) for the avail- ability of this guidance document.

[54 FR 27160, June 28, 1989, as amended at 67

FR 76681, Dec. 13, 2002]

##### § 868.2500   Cutaneous   oxygen   (PcO2) monitor.

- Identification. A cutaneous oxygen (PcO2) monitor is a noninvasive, heat- ed sensor (e.g., a Clark-type polargraphic electrode) placed on the patient’s skin that is intended to mon- itor relative changes in the cutaneous oxygen tension.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con-

#### 21 CFR Ch. I (4–1–17 Edition)

trols Guidance Document: Cutaneous Carbon Dioxide (PcCO2) and Oxygen (PcO2) Monitors; Guidance for Industry and FDA.’’ See § 868.1(e) for the avail- ability of this guidance document.

[67 FR 76681, Dec. 13, 2002]

##### § 868.2550   Pneumotachometer.

- Identification.	A pneumotachometer is a device intended for medical purposes that is used to de- termine gas flow by measuring the pressure differential across a known re- sistance. The device may use a set of capillaries or a metal screen for the re- sistive  element.

- Classification. Class II (perform- ance standards).

##### § 868.2600   Airway pressure monitor.

- Identification. An airway pressure monitor is a device used to  measure the pressure in a patient’s upper air- way. The device may include a pressure gauge and an alarm.

- Classification. Class II (perform- ance standards).

##### § 868.2610   Gas pressure gauge.

- Identification. A gas pressure gauge (e.g., bourdon tube pressure gauge) is a device intended for medical purposes that is used to measure gas pressure in a medical gas delivery sys- tem.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2620   Gas pressure calibrator.

- Identification. A gas pressure cali- brator is a device intended for medical purposes that is used to calibrate pres- sure-measuring instruments by gener- ating a known gas pressure.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 868.5110

##### § 868.2700   Pressure regulator.

  - Identification. A pressure regu- lator is a device, often called a pres- sure-reducing valve, that is intended for medical purposes and that is used to convert a medical gas pressure from a high variable pressure to a lower, more constant working pressure. This device includes mechanical oxygen reg- ulators.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2775   Electrical  peripheral  nerve stimulator.

- Identification. An electrical pe- ripheral nerve stimulator (neuro- muscular blockade monitor) is a device used to apply an electrical current to a patient to test the level of pharma- cological effect of anesthetic drugs and gases.

- Classification. Class II (perform- ance standards).

##### § 868.2875   Differential  pressure  trans- ducer.

- Identification. A differential pres- sure transducer is a two-chambered de- vice intended for medical purposes that is often used during pulmonary func- tion testing. It generates an electrical signal for subsequent display or proc- essing that is proportional to the dif- ference in gas pressures in the two chambers.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2885   Gas flow transducer.

- Identification. A gas flow trans- ducer is a device intended for medical purposes that is used to convert gas flow rate into an electrical signal for subsequent display or processing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1119,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.2900   Gas pressure transducer.

- Identification. A gas pressure transducer is a device intended for medical purposes that is used to con- vert gas pressure into an electrical sig- nal for subsequent display or proc- essing.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120, Jan. 16, 1996]

## Subparts D–E [Reserved]

Subpart F—Therapeutic Devices

##### § 868.5090   Emergency airway needle.

- Identification. An emergency air- way needle is a device intended to puncture a patient’s cricothyroid mem- brane to provide an emergency airway during upper airway obstruction.

- Classification. Class II (perform- ance standards).

##### § 868.5100   Nasopharyngeal airway.

- Identification. A nasopharyngeal airway is a device used to aid breathing by means of a tube inserted into a pa- tient’s pharynx through the nose to provide a patent airway.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5110   Oropharyngeal airway.

- Identification. An oropharyngeal airway is a device inserted into a pa- tient’s pharynx through the mouth to provide a patent airway.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

§ 868.5115

premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5115   Device	to	relieve	acute upper airway obstruction.

  - Identification. The device is a raised, rounded pad that, in the event of choking on a foreign body, can be applied to the abdomen and pushed up- ward to generate expulsion pressure to remove the obstruction to relieve acute upper airway obstruction.

  - Classification. Class II (special controls) (‘‘Class II Special Control Guidance Document for Acute Upper Airway Obstruction Devices’’). The de- vice is exempt from the premarket no- tification procedures in subpart E of part   807   of   this   chapter,   subject   to

§ 868.9.

[65 FR 39099, June 23, 2000; 65 FR 47669, Aug.

3, 2000]

##### § 868.5120   Anesthesia conduction cath- eter.

- Identification. An anesthesia con- duction catheter is a flexible tubular device used to inject local anesthetics into a patient and to provide contin- uous regional anesthesia.

- Classification. Class II (perform- ance standards).

##### § 868.5130   Anesthesia   conduction   fil- ter.

- Identification. An anesthesia con- duction filter is a microporous filter used while administering to a patient injections of local anesthetics to mini- mize particulate (foreign material) contamination of the injected fluid.

- Classification. Class II (perform- ance standards).

##### § 868.5140   Anesthesia conduction kit.

- Identification. An anesthesia con- duction kit is a device used to admin- ister to a patient conduction, regional, or local anesthesia. The device may contain syringes, needles, and drugs.

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 868.5150   Anesthesia  conduction  nee- dle.

- Identification. An anesthesia con- duction needle is a device used to in- ject local anesthetics into a patient to provide regional anesthesia.

- Classification. Class II (perform- ance standards).

##### § 868.5160   Gas machine for anesthesia or analgesia.

- Gas machine for anesthesia—(1) Identification. A gas machine for anes- thesia is a device used to administer to a patient, continuously or intermit- tently, a general inhalation anesthetic and to maintain a patient’s ventila- tion. The device may include a gas flowmeter, vaporizer, ventilator, breathing circuit with bag, and emer- gency air supply.

(2) Classification. Class II (perform- ance standards).

- Gas machine for analgesia—(1) Identification. A gas machine for anal- gesia is a device used to administer to a patient an analgesic agent, such as a nitrous oxide-oxygen mixture (max- imum concentration of 70 percent ni- trous oxide).

- Classification. Class II (perform- ance standards).

##### § 868.5165   Nitric  oxide  administration apparatus.

  - Identification. The nitric oxide ad- ministration apparatus is a device used to add nitric oxide to gases that are to be breathed by a patient. The nitric oxide administration apparatus is to be used in conjunction with a  ventilator or other breathing gas administration system.

  - Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance  Document  for  Premarket  Notifi- cation  Submissions  for  Nitric  Oxide Administration	Apparatus,	Nitric Oxide Analyzer, and Nitrogen Dioxide Analyzer.’’

[65 FR 11465, Mar. 3, 2000]

##### § 868.5170   Laryngotracheal topical an- esthesia applicator.

- Identification. A laryngotracheal topical anesthesia applicator is a de- vice used to apply topical anesthetics to a patient’s laryngotracheal area.

#### Food and Drug Administration, HHS	§ 868.5300

- Classification. Class II (perform- ance standards).

##### § 868.5180   Rocking bed.

- Identification. A rocking bed is a device intended for temporary use to help patient ventilation (breathing) by repeatedly tilting the patient, thereby using the weight of the abdominal con- tents to move the diaphragm.

- Classification. Class II (perform- ance standards).

##### § 868.5220   Blow bottle.

- Identification. A blow bottle is a device that is intended for medical pur- poses to induce a forced expiration from a patient. The patient blows into the device to move a column of water from one bottle to another.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38794, July 25,

2001]

##### § 868.5240   Anesthesia    breathing    cir- cuit.

- Identification. An anesthesia breathing circuit is a device that is in- tended to administer medical gases to a patient during anesthesia. It provides both an inhalation and exhalation route and may include a connector, adaptor, and Y-piece.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5250   Breathing circuit circulator.

- Identification. A breathing circuit circulator is a turbine device that is attached to a closed breathing circuit and that is intended to circulate anes- thetic gases continuously by maintain- ing the unidirectional valves in  an open position and reducing mechanical dead space and resistance in the breathing circuit.

- Classification. Class II (perform- ance standards).

##### § 868.5260   Breathing  circuit  bacterial filter.

- Identification. A breathing circuit bacterial filter is a device that is in- tended to remove microbiological and particulate matter from the gases in the breathing circuit.

- Classification. Class II (perform- ance standards).

##### § 868.5270   Breathing system heater.

- Identification. A breathing system heater is a device that is intended to warm breathing gases before they enter a patient’s airway. The device may in- clude a temperature controller.

- Classification. Class II (perform- ance standards).

##### § 868.5280   Breathing tube support.

- Identification. A breathing tube support is a device that is intended to support and anchor a patient’s breath- ing tube(s).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38794, July 25,

2001]

##### § 868.5300   Carbon dioxide absorbent.

- Identification. A carbon dioxide ab- sorbent is a device intended for med- ical purposes that consists of an ab- sorbent material (e.g., soda lime) that is intended to remove carbon dioxide from the gases in the breathing circuit.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 868.5310

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5310   Carbon dioxide absorber.

  - Identification. A carbon dioxide ab- sorber is a device that is intended for medical purposes and that is used in a breathing circuit as a container for carbon dioxide absorbent. It may in- clude a canister and water drain.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5320   Reservoir bag.

- Identification. A reservoir bag is a device, usually made of conductive rub- ber, intended for use in a breathing cir- cuit as a reservoir for breathing gas and to assist, control, or monitor a pa- tient’s  ventilation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38794,  July 25,

2001]

##### § 868.5330   Breathing gas mixer.

- Identification. A breathing gas mixer is a device intended for use in conjunction with a respiratory support apparatus to control the mixing of gases that are to be breathed by a pa- tient.

- Classification. Class II (perform- ance standards).

##### § 868.5340   Nasal oxygen cannula.

- Identification. A nasal oxygen cannula is a two-pronged device used to administer oxygen to a patient through both nostrils.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38794,  July 25,

2001]

##### § 868.5350   Nasal oxygen catheter.

- Identification. A nasal oxygen catheter is a device intended to be in- serted through a patient’s nostril to administer oxygen.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38794,  July 25,

2001]

##### § 868.5365   Posture chair for cardiac or pulmonary treatment.

- Identification. A posture chair for cardiac or pulmonary treatment is a device intended to assist in the reha- bilitation and mobilization of patients with chronic heart or lung disease.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38794, July 25,

2001]

##### § 868.5375   Heat	and	moisture	con- denser (artificial nose).

- Identification. A heat and moisture condenser (artificial nose) is a device intended to be positioned over a tra- cheotomy (a surgically created opening in the throat) or tracheal tube (a tube inserted into the trachea) to warm and humidify gases breathed in by a pa- tient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 868.5460

##### § 868.5400   Electroanesthesia	appa- ratus.

  - Identification.	An electroanesthesia apparatus is a device used for the induction and mainte- nance of anesthesia during surgical procedures by means of an alternating or pulsed electric current that is passed through electrodes fixed to a pa- tient’s head.

  - Classification. Class III (premarket approval).

  - Date PMA or notice of completion of a PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any  electroanesthesia  apparatus that was in commercial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equivalent to an electroanesthesia  apparatus  that  was in commercial distribution before May 28, 1976. Any other electroanesthesia apparatus shall have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[47 FR 31142, July 16, 1982, as amended at 52

FR 17735, May 11, 1987; 61 FR 50706, Sept. 27,

1996]

##### § 868.5420   Ether hook.

- Identification. An ether hook is a device that fits inside a patient’s mouth and that is intended to deliver vaporized ether.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38795, July 25,

2001]

##### § 868.5430   Gas-scavenging apparatus.

- Identification. A gas-scavenging apparatus is a device intended to col- lect excess anesthetic, analgesic, or trace gases or vapors from a patient’s breathing system, ventilator, or extracorporeal pump-oxygenator,  and to conduct these gases out of the area by means of an exhaust system.

- Classification. Class II (perform- ance standards).

##### § 868.5440   Portable oxygen generator.

- Identification. A portable oxygen generator is a device that is intended to release oxygen for respiratory ther- apy by means of either a chemical re- action or physical means (e.g., a molec- ular sieve).

- Classification. Class II (perform- ance standards).

##### § 868.5450   Respiratory gas humidifier.

- Identification. A respiratory gas humidifier is a device that is intended to add moisture to, and sometimes to warm, the breathing gases for adminis- tration to a patient. Cascade, gas, heated, and prefilled humidifiers are included in this generic type of device.

- Classification. Class II (perform- ance standards).

##### § 868.5460   Therapeutic  humidifier  for home use.

- Identification. A therapeutic hu- midifier for home use is a device that adds water vapor to breathing gases and that is intended for respiratory therapy or other medical purposes. The vapor produced by the device pervades the area surrounding the patient, who breathes the vapor during normal res- piration.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982; 47 FR 40410, Sept.

14, 1982, as amended at 61 FR 1120, Jan. 16,

1996; 66 FR 38795, July 25, 2001]

#### § 868.5470

##### § 868.5470   Hyperbaric chamber.

  - Identification. A hyperbaric cham- ber is a device that is intended to in- crease the environmental oxygen pres- sure to promote the movement of oxy- gen from the environment to a  pa- tient’s tissue by means of pressuriza- tion that is greater than atmospheric pressure. This device does not include topical oxygen chambers  for  extrem- ities (§ 878.5650).

  - Classification. Class II (perform- ance standards).

##### § 868.5530   Flexible laryngoscope.

- Identification. A flexible laryn- goscope is a fiberoptic device used to examine and visualize a patient’s upper airway and aid placement of a tracheal tube.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 41107, Sept. 17, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5540   Rigid laryngoscope.

- Identification. A rigid laryn- goscope is a device used to examine and visualize a patient’s upper airway and aid placement of a tracheal tube.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9

[47 FR 41107, Sept. 17, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5550   Anesthetic gas mask.

- Identification. An anesthetic gas mask is a device, usually made of con- ductive rubber, that is positioned over a patient’s nose or mouth to direct an- esthetic gases to the upper airway.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 41107, Sept. 17, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 868.5560   Gas mask head strap.

- Identification. A gas mask head strap is a device used to hold an anes- thetic gas mask in position on a pa- tient’s face.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 41107, Sept. 17, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38795, July 25,

2001]

##### § 868.5570   Nonrebreathing mask.

- Identification. A nonrebreathing mask is a device fitting over a pa- tient’s face to administer oxygen. It utilizes one-way valves to prevent the patient from rebreathing previously ex- haled gases.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5580   Oxygen mask.

- Identification. An oxygen mask is a device placed over a patient’s nose, mouth, or tracheostomy to administer oxygen or aerosols.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5590   Scavenging mask.

- Identification. A scavenging mask is a device positioned over a patient’s nose to deliver anesthetic or analgesic gases to the upper airway and to re- move excess and exhaled gas. It is usu- ally used during dentistry.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 868.5675

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5600   Venturi mask.

  - Identification. A venturi mask is a device containing an air-oxygen mixing mechanism that dilutes 100 percent ox- ygen to a predetermined concentration and delivers the mixed gases to a pa- tient.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5620   Breathing mouthpiece.

- Identification. A breathing mouth- piece is a rigid device that is inserted into a patient’s mouth and that con- nects with diagnostic or therapeutic respiratory devices.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2313, Jan. 14, 2000]

##### § 868.5630   Nebulizer.

- Identification. A nebulizer is a de- vice intended to spray liquids in aer- osol form into gases that are delivered directly to the patient for breathing. Heated, ultrasonic, gas, venturi, and refillable nebulizers are included in this generic type of device.

- Classification. Class II (perform- ance standards).

##### § 868.5640   Medicinal	nonventilatory nebulizer (atomizer).

- Identification. A medicinal non- ventilatory nebulizer (atomizer) is a device that is intended to spray liquid medication in aerosol form into the air that a patient will breathe.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

subpart  E  of  part  807  of  this  chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2313, Jan. 14, 2000]

##### § 868.5650   Esophageal obturator.

- Identification. An esophageal obtu- rator is a device inserted through a pa- tient’s mouth to aid ventilation of the patient during emergency resuscitation by occluding (blocking) the esophagus, thereby permitting positive pressure ventilation through the trachea. The device consists of a closed-end semirigid esophageal tube that is at- tached to a face mask.

- Classification. Class II (perform- ance standards).

##### § 868.5655   Portable liquid oxygen unit.

- Identification. A portable liquid oxygen unit is a portable, thermally in- sulated container of liquid oxygen that is intended to supplement gases to be inhaled by a patient, is sometimes ac- companied by tubing and an oxygen mask. An empty portable liquid oxygen unit is a device, while the oxygen con- tained therein is a drug.

- Classification. Class II (perform- ance standards).

##### § 868.5665   Powered percussor.

- Identification. A powered percussor is a device that is intended to transmit vibration through a patient’s  chest wall to aid in freeing mucus deposits in the lung in order to improve bronchial drainage and that may be powered by electricity or compressed gas.

- Classification. Class II (perform- ance standards).

##### § 868.5675   Rebreathing device.

- Identification. A rebreathing de- vice is a device that enables a patient to rebreathe exhaled gases. It may be used in conjunction with pulmonary function testing or for increasing minute ventilation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2313, Jan. 14, 2000]

#### § 868.5690

##### § 868.5690   Incentive spirometer.

  - Identification. An incentive spi- rometer is a device that indicates a pa- tient’s breathing volume or flow and that provides an incentive to the pa- tient to improve his or her ventilation.

  - Classification. Class II (perform- ance standards).

##### § 868.5700   Nonpowered oxygen tent.

- Identification. A nonpowered oxy- gen tent is a device that encloses a pa- tient’s head and upper body to contain oxygen delivered to the patient for breathing. This generic type of device includes infant oxygen hoods.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2313, Jan. 14, 2000]

##### § 868.5710   Electrically powered oxygen tent.

- Identification. An electrically pow- ered oxygen tent is a device that en- closes a patient’s head and, by means of an electrically powered unit, admin- isters breathing oxygen and controls the temperature and humidity of the breathing gases. This generic type de- vice includes the pediatric aerosol tent.

- Classification. Class II (perform- ance standards).

##### § 868.5720   Bronchial tube.

- Identification. A bronchial tube is a device used to differentially intubate a patient’s bronchus (one of the two main branches of the trachea leading directly to the lung) in order to isolate a portion of lung distal to the tube.

- Classification. Class II (perform- ance standards).

##### § 868.5730   Tracheal tube.

- Identification. A tracheal tube is a device inserted into a patient’s trachea via the nose or mouth and used to maintain an open airway.

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 868.5740   Tracheal/bronchial  differen- tial ventilation tube.

- Identification. A tracheal/bron- chial differential ventilation tube is a device used to isolate the left or the right lung  of a  patient  for anesthesia or pulmonary function testing.

- Classification. Class II (perform- ance standards).

##### § 868.5750   Inflatable	tracheal	tube cuff.

- Identification. An inflatable tra- cheal tube cuff is a device used to pro- vide an airtight seal between a tra- cheal tube and a patient’s trachea.

- Classification. Class II (perform- ance standards).

##### § 868.5760   Cuff spreader.

- Identification. A cuff spreader is a device used to install tracheal tube cuffs on tracheal or tracheostomy tubes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38795, July 25,

2001]

##### § 868.5770   Tracheal  tube  fixation  de- vice.

- Identification. A tracheal tube fix- ation device is a device used to hold a tracheal tube in place, usually by means of straps or pinch rings.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 868.5860

##### § 868.5780   Tube introduction forceps.

  - Identification. Tube introduction forceps (e.g., Magill forceps) are a right-angled device used to grasp a tra- cheal tube  and  place  it  in  a  patient’s trachea.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5790   Tracheal tube stylet.

- Identification. A tracheal tube sty- let is a device used temporarily to make rigid a flexible tracheal tube to aid its insertion into a patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5795   Tracheal	tube	cleaning brush.

- Identification. A tracheal tube cleaning brush is a device consisting of a brush with plastic bristles intended to clean tracheal cannula devices after their removal from patients.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[51  FR  40388,  Nov.  6,  1986,  as  amended  at  66

FR 38795, July 25, 2001]

##### § 868.5800   Tracheostomy tube and tube cuff.

- Identification. A tracheostomy tube and tube cuff is a device intended to be placed into a surgical opening of

the trachea to facilitate ventilation to the lungs. The cuff may be a separate or integral part of the tracheostomy tube and is, when inflated, intended to establish a seal between the tracheal wall and the tracheostomy tube. The cuff is used to prevent the patient’s as- piration of substances, such as blood or vomit, or to provide a means for posi- tive-pressure ventilation of the pa- tient. This device is made of either stainless steel or plastic.

- Classification. Class II.

[51 FR 40389, Nov. 6, 1986]

##### § 868.5810   Airway connector.

- Identification. An airway con- nector is a device intended to connect a breathing gas source to a tracheal tube, tracheostomy tube, or mask.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5820   Dental protector.

- Identification. A dental protector is a device intended to protect a pa- tient’s teeth during manipulative pro- cedures within a patient’s oral cavity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38795,  July 25,

2001]

##### § 868.5830   Autotransfusion apparatus.

- Identification. An autotransfusion apparatus is a device used to collect and reinfuse the blood lost by a patient due to surgery or trauma.

- Classification. Class II (perform- ance standards).

##### § 868.5860   Pressure  tubing  and  acces- sories.

- Identification. Pressure tubing and accessories are flexible or rigid devices intended to deliver pressurized medical gases.

#### § 868.5870

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

##### § 868.5870   Nonrebreathing valve.

- Identification. A nonrebreathing valve is a one-way valve that directs breathing gas flow to the patient and vents exhaled gases into the atmos- phere.

- Classification. Class II (perform- ance standards).

##### § 868.5880   Anesthetic vaporizer.

- Identification. An anesthetic va- porizer is a device used to vaporize liq- uid anesthetic and deliver a controlled amount of the vapor to the patient.

- Classification. Class II (perform- ance standards).

##### § 868.5895   Continuous ventilator.

- Identification. A continuous venti- lator (respirator) is a device intended to mechanically control or assist pa- tient breathing by delivering a pre- determined percentage of oxygen in the breathing gas. Adult, pediatric, and neonatal ventilators are included in this generic type of device.

- Classification. Class II (perform- ance standards).

##### § 868.5905   Noncontinuous	ventilator (IPPB).

- Identification. A noncontinuous ventilator (intermittent positive pres- sure breathing-IPPB) is a device in- tended to deliver intermittently an aerosol to a patient’s lungs or to assist a patient’s breathing.

- Classification. Class II (perform- ance standards).

##### § 868.5915   Manual    emergency    venti- lator.

- Identification. A manual emer- gency ventilator is a device, usually in- corporating a bag and valve, intended to provide emergency respiratory sup- port by means of a face mask or a tube inserted into a patient’s airway.

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 868.5925   Powered   emergency   venti- lator.

- Identification. A powered emer- gency ventilator is a demand valve or inhalator intended to provide emer- gency respiratory support by means of a face mask or a tube inserted into a patient’s airway.

- Classification. Class II (perform- ance standards).

##### § 868.5935   External  negative  pressure ventilator.

- Identification. An external nega- tive pressure ventilator (e.g., iron lung, cuirass) is a device chamber that is in- tended to support a patient’s ventila- tion by alternately applying and re- leasing external negative pressure over the diaphragm and upper trunk of the patient.

- Classification. Class II (perform- ance standards).

##### § 868.5955   Intermittent	mandatory ventilation attachment.

- Identification. An intermittent mandatory ventilation (IMV) attach- ment is a device attached to a mechan- ical ventilator that allows spontaneous breathing by a patient while providing mechanical ventilation at a preset rate.

- Classification. Class II (perform- ance standards).

##### § 868.5965   Positive	end	expiratory pressure breathing attachment.

- Identification. A positive end ex- piratory pressure (PEEP) breathing at- tachment is a device attached to a ven- tilator that is used to elevate pressure in a patient’s lungs above atmospheric pressure at the end of exhalation.

- Classification. Class II (perform- ance standards).

##### § 868.5975   Ventilator tubing.

- Identification. Ventilator tubing is a device intended for use as a conduit for gases between a ventilator and a patient during ventilation of the pa- tient.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 868.6700

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

##### § 868.5995   Tee drain (water trap).

  - Identification. A tee drain (water trap) is a device intended to trap and drain water that collects in ventilator tubing during respiratory therapy, thereby preventing an increase in breathing resistance.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1120,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

## Subpart G—Miscellaneous

##### § 868.6100   Anesthetic cabinet, table, or tray.

- Identification. An anesthetic cabi- net, table, or tray is a device intended to store anesthetic equipment and drugs. The device is usually con- structed to eliminate build-up of static electrical charges.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38796, July 25,

2001]

##### § 868.6175   Cardiopulmonary	emer- gency cart.

- Identification. A cardiopulmonary emergency cart is a device intended to store and transport resuscitation sup- plies for emergency treatment. The de- vice does not include any equipment used  in  cardiopulmonary  resuscitation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex-

ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38796, July 25,

2001]

##### § 868.6225   Nose clip.

- Identification. A nose clip is a de- vice intended to close a patient’s exter- nal nares (nostrils) during  diagnostic or therapeutic procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25048, June 12, 1989; 66 FR 38796, July 25,

2001]

##### § 868.6250   Portable air compressor.

- Identification. A portable air com- pressor is a device intended to provide compressed air for medical purposes, e.g., to drive ventilators and other res- piratory devices.

- Classification. Class II (perform- ance standards).

##### § 868.6400   Calibration gas.

- Identification. A calibration gas is a device consisting of a container of gas of known concentration  intended to calibrate medical gas concentration measurement devices.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

##### § 868.6700   Anesthesia stool.

- Identification. An anesthesia stool is a device intended for use as a stool

§ 868.6810

for  the  anesthesiologist  in  the  oper- ating room.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 54

FR 25049, June 12, 1989; 66 FR 38796, July 25,

2001]

##### § 868.6810   Tracheobronchial	suction catheter.

| ient’s upper airway. 870.1100 Blood pressure alarm. |
| --- |
| (b) Classification. Class 1 (general con- | 870.1110 870.1120 | Blood pressure computer. Blood pressure cuff. |
| rols). The device is exempt from the 870.1130 Noninvasive blood pressure meas- remarket notification procedures in urement system. |

(a) Identification. A tracheobronchial suction catheter is a device used to as- pirate liquids or semisolids from a pa- t

t p

subpart  E  of  part  807  of  this  chapter subject to § 868.9.

[47 FR 31142, July 16, 1982, as amended at 65

FR 2314, Jan. 14, 2000]

##### § 868.6820   Patient position support.

- Identification. A patient position support is a device intended to main- tain the position of an anesthetized pa- tient during surgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

##### § 868.6885   Medical gas yoke assembly.

- Identification. A medical gas yoke assembly is a device intended to con- nect medical gas cylinders to regu- lators or needle valves to supply gases for anesthesia or respiratory therapy. The device may include a particulate filter.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 868.9.

[47 FR 31142, July 16, 1982, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38796,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

PART 870—CARDIOVASCULAR DEVICES

##### Subpart A—General Provisions

Sec.

870.1    Scope.

870.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Cardiovascular Diagnostic Devices

870.1025 Arrhythmia detector and alarm (in- cluding ST-segment measurement and alarm).

870.1140 Venous blood pressure manometer. 870.1200 Diagnostic intravascular catheter. 870.1210    Continuous flush catheter.

870.1220    Electrode    recording    catheter    or electrode  recording  probe.

870.1230    Fiberoptic oximeter catheter. 870.1240   Flow-directed catheter.

870.1250    Percutaneous catheter.

870.1270    Intracavitary    phonocatheter    sys- tem.

870.1280   Steerable catheter.

870.1290     Steerable catheter control system. 870.1300   Catheter cannula.

870.1310   Vessel    dilator    for    percutaneous catheterization.

870.1330   Catheter guide wire. 870.1340   Catheter introducer. 870.1350    Catheter balloon repair kit. 870.1360  Trace microsphere.

870.1370    Catheter tip occluder. 870.1380    Catheter stylet.

870.1390   Trocar.

870.1415    Coronary  vascular  physiologic  sim- ulation  software  device.

870.1425  Programmable	diagnostic	com- puter.

870.1435    Single-function, preprogrammed di- agnostic    computer.

870.1450    Densitometer.

870.1650    Angiographic injector and syringe. 870.1660   Indicator injector.

870.1670    Syringe actuator for an injector. 870.1750    External  programmable  pacemaker

pulse generator.

870.1800   Withdrawal-infusion pump.

870.1875    Stethoscope.

870.1915   Thermodilution probe.

##### Subpart C—Cardiovascular Monitoring Devices

870.2050    Biopotential   amplifier   and   signal conditioner.

#### Food and Drug Administration, HHS	Pt. 870

870.2060    Transducer signal amplifier and sig- nal   conditioner.

870.2100   Cardiovascular blood flowmeter. 870.2120   Extravascular blood flow probe. 870.2300 Cardiac	monitor	(including

cardiotachometer and rate alarm).

870.2310  Apex	cardiograph (vibrocardiograph).

870.2320      Ballistocardiograph.

870.2330     Echocardiograph.

870.2340      Electrocardiograph.

870.2350    Electrocardiograph   lead   switching adaptor.

870.2360    Electrocardiograph electrode. 870.2370    Electrocardiograph     surface     elec-

trode tester.

870.2390     Phonocardiograph.

870.2400      Vectorcardiograph.

870.2450    Medical cathode-ray tube display. 870.2600    Signal isolation system.

870.2620   Line isolation monitor.

870.2640    Portable leakage current alarm. 870.2675  Oscillometer.

870.2700   Oximeter.

870.2710   Ear oximeter.

870.2750  Impedance phlebograph.

870.2770   Impedance plethysmograph.

870.2780    Hydraulic,    pneumatic,    or    photo- electric     plethysmographs.

870.2800    Medical magnetic tape recorder. 870.2810    Paper chart recorder.

870.2840    Apex cardiographic transducer. 870.2850    Extravascular blood pressure trans-

ducer.

870.2855    Implantable   Intra-aneurysm   Pres- sure   Measurement   System.

870.2860   Heart sound transducer.

870.2870    Catheter tip pressure transducer. 870.2880    Ultrasonic transducer.

870.2890    Vessel occlusion transducer.  870.2900    Patient   transducer   and   electrode

cable (including connector).

870.2910    Radiofrequency    physiological    sig- nal  transmitter  and  receiver.

870.2920   Telephone	electrocardiograph transmitter   and   receiver.

##### Subpart D—Cardiovascular Prosthetic Devices

870.3250    Vascular clip. 870.3260    Vena cava clip.

870.3300    Vascular embolization device. 870.3375     Cardiovascular intravascular filter. 870.3450    Vascular graft prosthesis.

870.3460    Endovascular suturing system. 870.3470    Intracardiac patch or pledget made

of	polypropylene,		polyethylene terephthalate,	or	polytetrafluoro- ethylene.

870.3535    Intra-aortic    balloon    and    control system.

870.3545    Ventricular bypass (assist) device. 870.3600    External   pacemaker   pulse   gener-

ator.

870.3605   Pacing system analyzer.

870.3610    Implantable  pacemaker  pulse  gen- erator.

870.3620   Pacemaker lead adaptor.

870.3630    Pacemaker  generator  function  ana- lyzer.

870.3640    Indirect pacemaker generator func- tion   analyzer.

870.3650    Pacemaker polymeric mesh bag. 870.3670   Pacemaker charger.

870.3680    Cardiovascular  permanent  or  tem- porary  pacemaker  electrode.

870.3690    Pacemaker test magnet. 870.3700   Pacemaker programmers.

870.3710    Pacemaker   repair   or   replacement material.

870.3720    Pacemaker  electrode  function  test- er.

870.3730    Pacemaker service tools. 870.3800   Annuloplasty ring.

870.3850    Carotid sinus nerve stimulator. 870.3925    Replacement heart valve.

870.3935    Prosthetic heart valve holder. 870.3945    Prosthetic heart valve sizer.

##### Subpart E—Cardiovascular Surgical Devices

870.4075   Endomyocardial biopsy device. 870.4100    Extracorporeal   circuit   and   acces-

sories	for	long-term	respiratory/ cardiopulmonary  failure.

870.4200   Cardiopulmonary  bypass  accessory equipment.

870.4205   Cardiopulmonary bypass bubble de- tector.

870.4210    Cardiopulmonary   bypass   vascular catheter,  cannula,  or   tubing.

870.4220    Cardiopulmonary bypass heart-lung machine   console.

870.4230   Cardiopulmonary bypass defoamer. 870.4240   Cardiopulmonary   bypass   heat   ex-

changer.

870.4250   Cardiopulmonary   bypass   tempera- ture   controller.

870.4260    Cardiopulmonary    bypass    arterial line  blood  filter.

870.4270  Cardiopulmonary	bypass	cardi- otomy  suction  line  blood  filter.

870.4280    Cardiopulmonary prebypass filter. 870.4290   Cardiopulmonary    bypass    adaptor,

stopcock, manifold, or fitting.

870.4300   Cardiopulmonary   bypass   gas   con- trol   unit.

870.4310   Cardiopulmonary   bypass   coronary pressure  gauge.

870.4320    Cardiopulmonary   bypass   pulsatile flow  generator.

870.4330   Cardiopulmonary     bypass     on-line blood  gas  monitor.

870.4340   Cardiopulmonary bypass level sens- ing  monitor  and/or  control.

870.4350  Cardiopulmonary	bypass oxygenator.

870.4360   Nonroller-type blood pump.

870.4370    Roller-type    cardiopulmonary    by- pass blood pump.

#### § 870.1

870.4380  Cardiopulmonary	bypass	pump speed  control.

870.4390   Cardiopulmonary bypass pump tub- ing.

870.4400   Cardiopulmonary bypass blood res- ervoir.

870.4410  Cardiopulmonary	bypass	in-line blood gas sensor.

870.4420  Cardiopulmonary	bypass	cardi- otomy  return  sucker.

870.4430  Cardiopulmonary	bypass intracardiac   suction   control.

870.4450   Vascular clamp. 870.4475   Surgical vessel dilator.

870.4500  Cardiovascular	surgical	instru- ments.

870.4510   Apical closure device. 870.4875    Intraluminal artery stripper. 870.4885   External vein stripper.

##### Subpart F—Cardiovascular Therapeutic Devices

870.5050    Patient care suction apparatus. 870.5100    Percutaneous    Transluminal    Coro-

nary Angioplasty (PTCA) Catheter.

870.5150    Embolectomy catheter.

870.5175    Septostomy catheter.  870.5200    External cardiac compressor.

870.5210  Cardiopulmonary	resuscitation (CPR)  aid.

870.5225    External counter-pulsating device. 870.5300    DC-defibrillator (including paddles). 870.5310    Automated     external     defibrillator

system.

870.5325    Defibrillator tester.

870.5550    External     transcutaneous     cardiac pacemaker    (noninvasive).

870.5700    Steerable  cardiac  ablation  catheter remote  control  system.

870.5800   Compressible limb sleeve. 870.5900    Thermal regulating system.

870.5910    Esophageal  thermal  regulation  de- vice.

870.5925     Automatic rotating tourniquet.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE: 45 FR 7907, Feb. 5, 1980, unless oth-

erwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 870 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 870.1   Scope.

    - This part sets forth the classifica- tion of cardiovascular devices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer   who   submits   a   pre-

21 CFR Ch. I (4–1–17 Edition)

market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - To avoid duplicative listings, a cardiovascular device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed only in one subpart.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 17735, May 11, 1987, as amended at 68

FR 61344,  Oct. 28,  2003;  78  FR 18233,  Mar. 26,

2013]

##### § 870.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a device, has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except as provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the

#### Food and Drug Administration, HHS	§ 870.9

device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17735, May 11, 1987]

##### § 870.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II

device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the

#### § 870.1025

assay  is  intended  for  use  in  matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2314, Jan. 14, 2000]

## Subpart B—Cardiovascular Diagnostic Devices

##### § 870.1025 Arrhythmia detector and alarm (including ST-segment meas- urement and alarm).

    - Identification. The arrhythmia de- tector and alarm device monitors an electrocardiogram and is designed to produce a visible or audible signal or alarm when atrial or ventricular ar- rhythmia, such as premature contrac- tion or ventricular fibrillation, occurs.

    - Classification. Class II (special controls). The guidance document enti- tled ‘‘Class II Special  Controls  Guid- ance Document: Arrhythmia Detector and Alarm’’ will serve as the special control. See § 870.1 for  the  availability of this guidance document.

[68 FR 61344, Oct. 28, 2003]

##### § 870.1100   Blood pressure alarm.

- Identification. A blood pressure alarm is a device that accepts the sig- nal from a blood pressure transducer amplifier, processes the signal, and emits an alarm when the blood pres- sure falls outside a pre-set upper or lower limit.

- Classification. Class II (perform- ance standards).

##### § 870.1110   Blood pressure computer.

- Identification. A blood pressure computer is a device that accepts the electrical signal from a blood pressure transducer amplifier and indicates the systolic, diastolic, or mean pressure based on the input signal.

- Classification. Class II (perform- ance standards).

##### § 870.1120   Blood pressure cuff.

- Identification. A blood pressure cuff is a device that has an inflatable bladder in an inelastic sleeve (cuff) with a mechanism for inflating and de- flating the bladder. The cuff is used in

21 CFR Ch. I (4–1–17 Edition)

conjunction with another device to de- termine a subject’s blood pressure.

- Classification. Class II (perform- ance standards).

##### § 870.1130   Noninvasive  blood  pressure measurement system.

- Identification. A noninvasive blood pressure measurement system is a de- vice that provides a signal from which systolic, diastolic, mean, or any com- bination of the three pressures can be derived through the use of tranducers placed on the surface of the body.

- Classification. Class II (perform- ance standards).

##### § 870.1140   Venous  blood  pressure  ma- nometer.

- Identification. A venous blood pressure manometer is a device at- tached to a venous catheter to indicate manometrically the central or periph- eral venous pressure.

- Classification. Class II (perform- ance standards).

##### § 870.1200   Diagnostic	intravascular catheter.

- Identification. An intravascular di- agnostic catheter is a device used to record intracardiac pressures, to sam- ple blood, and to introduce substances into the heart and vessels. Included in this generic device are right-heart catheters, left-heart catheters, and angiographic catheters, among others.

- Classification. Class II (perform- ance standards).

##### § 870.1210   Continuous flush catheter.

- Identification. A continuous flush catheter is an attachment to a cath- eter-transducer system that permits continuous intravascular flushing at a slow infusion rate for the purpose of eliminating  clotting,  back-leakage, and waveform damping.

- Classification. Class II (perform- ance standards).

##### § 870.1220   Electrode   recording   cath- eter or electrode recording probe.

- Identification. An electrode record- ing catheter or an electrode recording probe is a device used to detect an intracardiac electrocardiogram, or to detect cardiac output or left-to-right heart   shunts.   The   device   may   be

#### Food and Drug Administration, HHS	§ 870.1350

unipolar or multipolar for electro- cardiogram detection, or may be a platinum-tipped catheter which senses the presence of a special indicator for cardiac output or left-to-right heart shunt  determinations.

- Classification. Class II (perform- ance standards).

##### § 870.1230   Fiberoptic    oximeter    cath- eter.

- Identification. A fiberoptic oxim- eter catheter is a device used to esti- mate the oxygen saturation of the blood. It consists of two fiberoptic bun- dles that conduct light at a desired wavelength through blood and detect the reflected and scattered light at the distal end of the catheter.

- Classification. Class II (perform- ance standards).

##### § 870.1240   Flow-directed catheter.

- Identification. A flow-directed catheter is a device that incorporates a gas-filled balloon to help direct the catheter to the desired position.

- Classification. Class II (perform- ance standards).

##### § 870.1250   Percutaneous catheter.

- Identification. A percutaneous catheter is a device that is introduced into a vein or artery through the skin using a dilator and a sheath (intro- ducer) or guide wire.

- Classification. Class II (perform- ance standards).

##### § 870.1270   Intracavitary phonocatheter system.

- Identification. An intracavitary phonocatheter system is a system that includes a catheter with an acoustic transducer and the associated device that processes the signal from the transducer; this device records bio- acoustic phenomena from a transducer placed within the heart, blood vessels, or body cavities.

- Classification. Class II (perform- ance standards).

##### § 870.1280   Steerable catheter.

- Identification. A steerable catheter is a catheter used for diagnostic and monitoring purposes whose movements are directed by a steering control unit.

- Classification. Class II (perform- ance standards).

##### § 870.1290   Steerable   catheter   control system.

- Identification. A steerable catheter control system is a device that is con- nected to the proximal end of a steer- able guide wire that controls the mo- tion of the steerable catheter.

- Classification. Class II (perform- ance standards).

##### § 870.1300   Catheter cannula.

- Identification. A catheter cannula is a hollow tube which is inserted into a vessel or cavity; this device provides a rigid or semirigid structure which can be connected to a tube or con- nector.

- Classification. Class II (perform- ance standards).

##### § 870.1310   Vessel	dilator	for percutaneous catheterization.

- Identification. A vessel dilator for percutaneous catheterization is a de- vice which is placed over the guide wire to enlarge the opening in the vessel, and which is then removed before slid- ing the catheter over the guide wire.

- Classification. Class II (perform- ance standards).

##### § 870.1330   Catheter guide wire.

- Identification. A catheter guide wire is a coiled wire that is designed to fit inside a percutaneous catheter for the purpose of directing the catheter through a blood vessel.

- Classification. Class II (perform- ance standards).

##### § 870.1340   Catheter introducer.

- Identification. A catheter intro- ducer is a sheath used to facilitate placing a catheter through the skin into a vein or artery.

- Classification. Class II (perform- ance standards).

##### § 870.1350   Catheter balloon repair kit.

- Identification. A catheter balloon repair kit is a device used to repair or replace the balloon of a balloon cath- eter. The kit contains the materials, such as glue and balloons, necessary to effect the repair or replacement.

#### § 870.1360

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any catheter balloon repair kit that was in commercial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be sub- stantially equivalent to a catheter bal- loon repair kit that was in commercial distribution before May 28, 1976. Any other catheter balloon repair kit shall have an approved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 61 FR 50706, Sept. 27, 1996]

##### § 870.1360   Trace microsphere.

- Identification. A trace microsphere is a radioactively tagged nonbiodegrad- able particle that is intended to be in- jected into an artery or vein and trapped in the capillary bed for the purpose of studying blood flow within or to an organ.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any trace microsphere that was in commercial distribution before May 28, 1976, or that has, on or before  Decem- ber 26, 1996 been found to be substan- tially equivalent to a  trace  micro- sphere that was in commercial dis- tribution before May 28, 1976. Any other trace microsphere shall have an approved PMA or a declared completed PDP in effect before being placed in commercial     distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 61 FR 50706, Sept. 27, 1996]

##### § 870.1370   Catheter tip occluder.

- Identification. A catheter tip occluder is a device that is inserted into certain catheters to prevent flow through one or more orifices.

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 870.1380   Catheter stylet.

- Identification. A catheter stylet is a wire that is run through a catheter or cannula to render it stiff.

- Classification. Class II (perform- ance standards).

##### § 870.1390   Trocar.

- Identification. A trocar is a sharp- pointed instrument used with a cannula for piercing a vessel or cham- ber to facilitate insertion of the cannula.

- Classification. Class II (perform- ance standards).

##### § 870.1415   Coronary   vascular   physio- logic simulation software device.

- Identification. A coronary vascular physiologic simulation software device is a prescription device that provides simulated functional assessment of blood flow in the coronary vascular system using data extracted from med- ical device imaging to solve algorithms and yield simulated metrics of physio- logic information (e.g., blood flow, cor- onary flow reserve, fractional flow re- serve, myocardial perfusion). A coro- nary vascular physiologic simulation software device is intended to generate results for use and review by a quali- fied clinician.

- Classification. Class II (special controls). The special controls for this device are:

  - Adequate software verification and validation based on comprehensive hazard analysis, with identification of appropriate mitigations, must be per- formed,  including:

- Full characterization of the tech- nical parameters of the software, in- cluding:

- Any proprietary algorithm(s) used to model the vascular anatomy; and

- Adequate description of the ex- pected impact of all applicable image acquisition hardware features and characteristics on performance and any associated minimum specifications;

- Adequate consideration of pri- vacy and security issues in the system design; and

#### Food and Drug Administration, HHS	§ 870.1415

- Adequate mitigation of the im- pact of failure of any subsystem com- ponents (e.g., signal detection and anal- ysis, data storage, system communica- tions and cybersecurity) with respect to incorrect patient reports and oper- ator failures.

  - Adequate non-clinical perform- ance testing must be provided to dem- onstrate the validity of computational modeling methods for flow measure- ment; and

  - Clinical data supporting the pro- posed intended use must be provided, including the  following:

    - Output measure(s) must be com- pared to a clinically acceptable method and must adequately represent the simulated measure(s) the device pro- vides in an accurate and reproducible manner;

    - Clinical utility of the device measurement accuracy must be dem- onstrated by comparison to that of other available diagnostic tests (e.g., from  literature  analysis);

    - Statistical performance of the device within clinical risk strata (e.g., age, relevant comorbidities, disease stability) must be reported;

    - The dataset must be adequately representative of the intended use pop- ulation for the device (e.g., patients, range of vessel sizes, imaging device models). Any selection criteria or limi- tations of the samples must be fully de- scribed and justified;

    - Statistical methods must consider the predefined endpoints:

- Estimates of probabilities of in- correct results must be provided for each endpoint,

- Where multiple samples from the same patient are used, statistical anal- ysis must not assume statistical inde- pendence without adequate justifica- tion, and

- The report must provide appro- priate confidence intervals for each performance metric;

    - Sensitivity and  specificity  must be characterized across the range of available  measurements;

    - Agreement of the simulated measure(s) with clinically acceptable measure(s) must be assessed across the full range of measurements;

    - Comparison of the measure- ment  performance  must  be  provided

across the range of intended image ac- quisition hardware; and

    - If the device uses a cutoff thresh- old or operates across a spectrum of disease, it must be established prior to validation, and it must be justified as to how it was determined and clinically validated;

  - Adequate validation must be per- formed and controls implemented to characterize and ensure consistency (i.e., repeatability and reproducibility) of measurement outputs:

    - Acceptable incoming image qual- ity control measures and the resulting image rejection rate for the clinical data must be specified, and

    - Data must be provided within the clinical validation study or using equivalent datasets demonstrating the consistency (i.e., repeatability and re- producibility) of the output that is rep- resentative of the range of data quality likely to be encountered in the in- tended use population and relevant use conditions in the intended use environ- ment;

      - Testing must be performed using multiple operators meeting planned qualification criteria and using the procedure that will be implemented in the production use of the device, and

      - The factors (e.g., medical imaging dataset, operator) must be identified regarding which were held constant and which were varied during the eval- uation, and a description must be pro- vided for the computations and statis- tical analyses used to evaluate  the data;

  - Human factors evaluation and val- idation must be provided to dem- onstrate adequate performance of the user interface to allow for users to ac- curately measure intended parameters, particularly where parameter settings that have impact on measurements re- quire significant user intervention; and

  - Device labeling must be provided that adequately describes the fol- lowing:

    - The device’s intended use, includ- ing the type of imaging data used, what the device measures and outputs to the user, whether the measure is qualitative or quantitative, the clin- ical indications for which it is to be used, and the specific population for which the device use is intended;

#### § 870.1425

    - Appropriate warnings specifying the intended patient population, iden- tifying anatomy and image acquisition factors that may impact measurement results, and providing cautionary guid- ance for interpretation of the provided measurements;

    - Key assumptions made in the calculation and determination of simu- lated measurements;

    - The measurement performance of the device for all presented param- eters, with appropriate confidence in- tervals, and the supporting evidence for this performance. Per-vessel clin- ical performance, including where ap- plicable localized performance accord- ing to vessel and segment, must be in- cluded as well as a characterization of the measurement error across the ex- pected range of measurement for key parameters based on the clinical data;

    - A detailed description of the pa- tients studied in the clinical validation (e.g., age, gender, race or ethnicity, clinical stability, current treatment regimen) as well as  procedural  details of the clinical study (e.g., scanner rep- resentation, calcium scores, use of beta-blockers or nitrates); and

    - Where significant human inter- face is necessary for accurate analysis, adequately detailed description of the analysis procedure using the device and any data features that could affect ac- curacy of results.

[80 FR 63673, Oct. 21, 2015]

##### § 870.1425   Programmable	diagnostic computer.

- Identification. A programmable di- agnostic computer is a device that can be programmed to compute various physiologic or blood flow parameters based on the output from one or more electrodes, transducers, or measuring devices; this device includes any asso- ciated commercially supplied pro- grams.

- Classification. Class II (perform- ance standards).

##### § 870.1435   Single-function, preprogrammed	diagnostic	com- puter.

- Identification. A single-function, preprogrammed diagnostic computer is a hard-wired computer that calculates a  specific  physiological  or  blood-flow

#### 21 CFR Ch. I (4–1–17 Edition)

parameter based on information ob- tained from one or more electrodes, transducers, or measuring devices.

- Classification. Class II (perform- ance standards).

##### § 870.1450   Densitometer.

- Identification. A densitometer is a device used to measure the trans- mission of light through  an  indicator in a sample of blood.

- Classification. Class II (perform- ance standards).

##### § 870.1650   Angiographic   injector   and syringe.

- Identification. An angiographic in- jector and syringe is a device that con- sists of a syringe and a high-pressure injector which are used to inject con- trast material into the heart, great vessels, and coronary arteries to study the heart and vessels by x-ray photog- raphy.

- Classification. Class II (perform- ance standards).

##### § 870.1660   Indicator injector.

- Identification. An  indicator injec- tor is an electrically or gas-powered device designed to inject accurately an indicator solution into the blood stream. This device may be used in conjuction with a densitometer or thermodilution device to determine cardiac output.

- Classification. Class II (perform- ance standards).

##### § 870.1670   Syringe  actuator  for  an  in- jector.

- Identification. A syringe actuator for an injector is an electrical device that controls the timing of an injection by an angiographic or indicator injec- tor and synchronizes the injection with the electrocardiograph signal.

- Classification. Class II (perform- ance standards).

##### § 870.1750   External	programmable pacemaker pulse generator.

- Identification. An external pro- grammable pacemaker pulse genera- tors is a device that can be pro- grammed to produce one or more pulses at preselected intervals; this de- vice is used in electrophysiological studies.

#### Food and Drug Administration, HHS	§ 870.2310

- Classification. Class II (perform- ance standards).

##### § 870.1800   Withdrawal-infusion pump.

- Identification. A withdrawal-infu- sion pump is a device designed to inject accurately drugs into the bloodstream and to withdraw blood samples for use in determining cardiac output.

- Classification. Class II (perform- ance standards).

##### § 870.1875   Stethoscope.

- Manual stethoscope—(1) Identifica- tion. A manual stethoscope is a me- chanical device used to project the sounds associated with the heart, arte- ries, and veins and other internal or- gans.

(2) Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

- Electronic stethoscope—(1) Identi- fication. An electronic stethoscope is an electrically amplified device used to project the sounds associated with the heart, arteries, and veins and other in- ternal  organs.

- Classification. Class II (perform- ance standards).

[45 FR 7907, Feb. 5, 1980, as amended at 59 FR

63007, Dec. 7, 1994; 66 FR 38796, July 25, 2001]

##### § 870.1915   Thermodilution probe.

  - Identification. A thermodilution probe is a device that monitors cardiac output by use of thermodilution tech- niques; this device is commonly at- tached to a catheter that may have one or more probes.

  - Classification. Class II (perform- ance standards).

## Subpart C—Cardiovascular Monitoring Devices

##### § 870.2050   Biopotential   amplifier   and signal conditioner.

- Identification. A biopotential am- plifier and signal conditioner is a de- vice used to amplify or condition an electrical signal of biologic origin.

- Classification. Class II (perform- ance standards).

##### § 870.2060   Transducer signal amplifier and conditioner.

- Identification. A transducer signal amplifier and conditioner is a device used to provide the excitation energy for the transducer and to amplify or condition the signal emitted by the transducer.

- Classification. Class II (perform- ance standards).

##### § 870.2100   Cardiovascular  blood  flow- meter.

- Identification. A cardiovascular blood flowmeter is a device that is con- nected to a flow transducer that ener- gizes the transducer and processes and displays the blood flow signal.

- Classification. Class II (perform- ance standards).

##### § 870.2120   Extravascular	blood	flow probe.

- Identification. An extravascular blood flow probe is an extravascular ul- trasonic or electromagnetic probe used in conjunction with a blood flowmeter to measure blood flow in a chamber or vessel.

- Classification. Class II (perform- ance standards).

##### § 870.2300   Cardiac  monitor  (including cardiotachometer and rate alarm).

- Identification. A cardiac monitor (including cardiotachometer and rate alarm) is a device used to measure the heart rate from an analog signal pro- duced by an electrocardiograph, vectorcardiograph, or blood pressure monitor. This device may sound an alarm when the heart rate falls outside preset upper and lower limits.

- Classification. Class II (perform- ance standards).

##### § 870.2310   Apex	cardiograph (vibrocardiograph).

- Identification. An apex cardio- graph (vibrocardiograph) is a device used to amplify or condition the signal from an apex cardiographic transducer and to produce a visual display of the motion of the heart; this device also provides any excitation energy re- quired by the transducer.

- Classification. Class II (perform- ance standards).

§ 870.2320

##### § 870.2320   Ballistocardiograph.

  - Identification.	A ballistocardiograph is a device, includ- ing a supporting structure on which the patient is placed, that moves in re- sponse to blood ejection from the heart. The device often provides a vis- ual display.

  - Classification. Class II (perform- ance standards).

##### § 870.2330   Echocardiograph.

- Identification. An echocardiograph is a device that uses ultrasonic energy to create images of cardiovascular structures. It includes  phased  arrays and two-dimensional scanners.

- Classification. Class II (perform- ance standards).

##### § 870.2340   Electrocardiograph.

- Identification. An electrocardio- graph is a device used to process the electrical signal transmitted through two or more electrocardiograph elec- trodes and to produce a visual display of the electrical signal produced by the heart.

- Classification. Class II (perform- ance standards).

##### § 870.2350   Electrocardiograph	lead switching adaptor.

- Identification. An electrocardio- graph lead switching adaptor is a pas- sive switching device to which electro- cardiograph limb and chest leads may be attached. This device is used to con- nect various combinations of limb and chest leads to the output terminals in order to create standard lead combina- tions such as leads I, II, and III.

- Classification. Class II (perform- ance standards).

##### § 870.2360   Electrocardiograph	elec- trode.

- Identification. An electrocardio- graph electrode is the electrical con- ductor which is applied to the surface of the body to transmit the electrical signal at the body surface to a proc- essor that produces an electrocardio- gram or vectorcardiogram.

- Classification. Class II (special controls). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9. The

#### 21 CFR Ch. I (4–1–17 Edition)

special control for this device is the FDA  guidance  document   entitled ‘‘Class II Special  Controls Guidance Document: Electrocardiograph Elec- trodes.’’ See § 870.1(e) for availability information of guidance documents.

[45 FR 7907, Feb. 5, 1980, as amended at 76 FR

43585, July 21, 2011]

##### § 870.2370   Electrocardiograph   surface electrode tester.

- Identification. An electrocardio- graph surface electrode tester is a de- vice used to test the function and ap- plication of electrocardiograph elec- trodes.

- Classification. Class II (perform- ance standards).

##### § 870.2390   Phonocardiograph.

- Identification. A phonocardiograph is a device used to amplify or condition the signal from a heart sound trans- ducer. This device furnishes the exci- tation energy for the transducer and provides a visual or  audible display  of the heart sounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.2400   Vectorcardiograph.

- Identification. A vectorcardiograph is a device used to process the elec- trical signal transmitted through elec- trocardiograph electrodes and to produce a visual display of the mag- nitude and direction of the electrical signal produced by the heart.

- Classification. Class II (perform- ance standards).

##### § 870.2450   Medical    cathode-ray    tube display.

- Identification. A medical cathode- ray tube display is a device designed primarily to display selected biological signals. This device often incorporates special display features unique to a specific biological signal.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 870.2800

##### § 870.2600   Signal isolation system.

  - Identification. A signal isolation system is a device that electrically iso- lates the patient from equipment con- nected to the commercial power supply received from a utility company. This isolation may be accomplished, for ex- ample, by transformer coupling, acous- tic coupling, or optical coupling.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.2620   Line isolation monitor.

- Identification. A line isolation monitor is a device used to monitor the electrical leakage current from a power supply electrically isolated from the commercial power supply received from a utility company.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.2640   Portable	leakage	current alarm.

- Identification. A portable leakage current alarm is a device used to meas- ure the electrical leakage current be- tween any two points of an electrical system and to sound an alarm if the current exceeds a certain threshold.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.2675   Oscillometer.

- Identification. An oscillometer is a device used to measure physiological oscillations of any kind,  e.g.,  changes in the volume of arteries.

- Classification. Class II (perform- ance standards).

##### § 870.2700   Oximeter.

- Identification. An oximeter is a de- vice used to transmit radiation at a known wavelength(s) through blood and to measure the blood oxygen satu- ration based on the amount of reflected or scattered radiation. It may be used alone or in conjunction with a fiberoptic oximeter catheter.

- Classification. Class II (perform- ance standards).

##### § 870.2710   Ear oximeter.

- Identification. An ear oximeter is an extravascular device used to trans- mit light at a known wavelength(s) through blood in the ear. The amount of reflected or scattered light as indi- cated by this device is used to measure the blood oxygen saturation.

- Classification. Class II (perform- ance standards).

##### § 870.2750   Impedance phlebograph.

- Identification. An impedance phlebograph is a device used to provide a visual display of the venous pulse or drainage by measuring electrical im- pedance changes in a region of the body.

- Classification. Class II (perform- ance standards).

##### § 870.2770   Impedance plethysmograph.

- Identification. An impedance plethysmograph is a device used to es- timate peripheral blood flow by meas- uring electrical  impedance  changes  in a region of the body such as the arms and legs.

- Classification. Class II (perform- ance standards).

##### § 870.2780   Hydraulic,	pneumatic,	or photoelectric plethysmographs.

- Identification. A hydraulic, pneu- matic,	or	photoelectric plethysmograph is a device used to es- timate blood flow in a region of the body using hydraulic, pneumatic, or photoelectric  measurement  techniques.

- Classification. Class II (perform- ance standards).

##### § 870.2800   Medical  magnetic  tape  re- corder.

- Identification. A medical magnetic tape recorder is a device used to record

#### § 870.2810

and play back signals from, for exam- ple, physiological amplifiers, signal conditioners, or computers.

- Classification. Class II (perform- ance standards).

##### § 870.2810   Paper chart recorder.

- Identification. A paper chart re- corder is a device used to print on paper, and create a permanent record of the signal from, for example, a phys- iological amplifier, signal conditioner, or computer.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.2840   Apex   cardiographic   trans- ducer.

- Identification. An apex cardio- graphic transducer is a device used to detect motion of the heart (accelera- tion, velocity, or displacement) by changes in the mechanical or electrical properties of the device.

- Classification. Class II (perform- ance standards).

##### § 870.2850   Extravascular   blood   pres- sure transducer.

- Identification. An extravascular blood pressure transducer is a device used to measure blood pressure by changes in the mechanical or electrical properties of the device. The proximal end of the transducer is connected to a pressure monitor that produces an ana- log or digital electrical  signal  related to the electrical or mechanical changes produced in the transducer.

- Classification. Class II (perform- ance standards).

##### § 870.2855   Implantable Intra-aneurysm Pressure Measurement System.

- Identification. Implantable intra- aneurysm pressure measurement sys- tem is a device used to measure the intra-sac pressure in a vascular aneu- rysm. The device consists of a pressure transducer that is implanted into the aneurysm and a monitor that reads the pressure from the transducer.

- Classification. Class II (special controls).  The  special  control  is  FDA’s

#### 21 CFR Ch. I (4–1–17 Edition)

guidance document entitled ‘‘Class II Special Controls Guidance Document: Implantable Intra-Aneurysm Pressure Measurement System.’’ See § 870.1 (e) for the availability of this guidance document.

[71 FR 7871, Feb. 15, 2006]

##### § 870.2860   Heart sound transducer.

- Identification. A heart sound transducer is an external  transducer that exhibits a change  in  mechanical or electrical properties in relation to sounds produced by the heart. This de- vice may be used in conjunction with a phonocardiograph to record heart sounds.

- Classification. Class II (perform- ance standards).

##### § 870.2870   Catheter tip pressure trans- ducer.

- Identification. A catheter tip pres- sure transducer is a device incor- porated into the distal end of a cath- eter. When placed in the bloodstream, its mechanical or electrical properties change in relation to changes in blood pressure. These changes are trans- mitted to accessory equipment for processing.

- Classification. Class II (perform- ance standards).

##### § 870.2880   Ultrasonic transducer.

- Identification. An ultrasonic trans- ducer is a device applied to the skin to transmit and receive ultrasonic energy that is used in conjunction with an echocardiograph to provide imaging of cardiovascular structures. This device includes phased arrays and two-dimen- sional scanning transducers.

- Classification. Class II (perform- ance standards).

##### § 870.2890   Vessel occlusion transducer.

- Identification. A vessel occlusion transducer is a device used to provide an electrical signal corresponding to sounds produced in a partially occluded vessel. This device includes motion, sound, and ultrasonic transducers.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 870.3375

##### § 870.2900 Patient transducer and elec- trode cable (including connector).

  - Identification. A patient trans- ducer and electrode cable (including connector) is an electrical conductor used to transmit signals from, or power or excitation signals to, patient-con- nected electrodes or transducers.

  - Classification. Class II (perform- ance standards).

##### § 870.2910 Radiofrequency physio- logical signal transmitter and re- ceiver.

- Identification. A radiofrequency physiological signal transmitter and receiver is a device used to condition a physiological signal so that it can be transmitted via radiofrequency from one location to another, e.g., a central monitoring station. The received signal is reconditioned by the device into its original format so that it can be dis- played.

- Classification. Class II (perform- ance standards).

##### § 870.2920 Telephone electrocardio- graph transmitter and receiver.

- Identification. A telephone electro- cardiograph transmitter and receiver is a device used to condition an electro- cardiograph signal so that it can be transmitted via a telephone line to an- other location. This  device  also  in- cludes a receiver that reconditions the received signal into its original format so that it can be displayed. The device includes devices used to transmit and receive pacemaker signals.

- Classification. Class II (perform- ance standards).

## Subpart D—Cardiovascular Prosthetic Devices

##### § 870.3250   Vascular clip.

- Identification. A vascular clip is an implanted extravascular device de- signed to occlude, by compression, blood flow in small blood vessels other than intracranial vessels.

- Classification. Class II (perform- ance standards).

##### § 870.3260   Vena cava clip.

- Identification. A vena cava clip is an implanted extravascular device de-

signed to occlude partially  the  vena cava for the purpose of inhibiting the flow of thromboemboli through that vessel.

- Classification. Class II (perform- ance standards).

##### § 870.3300   Vascular   embolization   de- vice.

- Identification. A vascular embolization device is an intravascular implant intended to control hem- orrhaging due to aneurysms, certain types of tumors (e.g., nephroma, hepatoma, uterine fibroids), and arteriovenous malformations. This does not include cyanoacrylates and other embolic agents, which act by po- lymerization	or	precipitation. Embolization devices used in neuro- vascular applications are also not in- cluded    in    this    classification,    see

§ 882.5950 of this chapter.

- Classification. Class II (special controls.) The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Vascular and Neurovascular Embolization Devices.’’ For availability of this guidance docu- ment, see § 870.1(e).

[69 FR 77899, Dec. 29, 2004]

##### § 870.3375   Cardiovascular intravascular filter.

- Identification. A cardiovascular intravascular filter is  an  implant  that is placed in the inferior vena cava for the purpose of preventing pulmonary thromboemboli (blood clots generated in the lower limbs and broken loose into the blood stream) from flowing into the right side of the heart and the pulmonary circulation.

- Classification. Class II. The special controls for this device are:

  - ‘‘Use of International Standards Organization’s ISO 10993 ‘Biological Evaluation  of  Medical  Devices  Part  I:

Evaluation and Testing,’ ’’ and

  - FDA’s:

    - ‘‘510(k) Sterility Review Guidance and Revision of 2/12/90 (K90–1)’’ and

    - ‘‘Guidance for Cardiovascular Intravascular Filter 510(k) Submis- sions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 65 FR 17144, Mar. 31, 2000]

§ 870.3450

##### § 870.3450   Vascular graft prosthesis.

      - Identification. A vascular graft prosthesis is an implanted device in- tended to repair, replace, or bypass sec- tions of native or artificial vessels, ex- cluding coronary or cerebral vasculature, and to provide vascular access. It is commonly constructed of materials such as polyethylene terephthalate and polytetrafluoro- ethylene, and it may be coated with a biological coating, such as albumin or collagen, or a synthetic  coating,  such as silicone. The graft structure itself is not made of materials of animal origin, including human umbilical cords.

      - Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Guidance Document for Vas- cular Prostheses 510(k) Submissions.’’

[66 FR 18542, Apr. 10, 2001]

##### § 870.3460   Endovascular Suturing Sys- tem.

- Identification. An endovascular su- turing system is a medical device in- tended to provide fixation and sealing between an endovascular graft and the native artery. The system is comprised of the implant device and an endovascular delivery device used to implant the endovascular suture.

- Classification. Class II (special controls). The special controls for this device are:

  - The device should be dem- onstrated to be biocompatible;

  - Sterility and shelf life testing should demonstrate the sterility of pa- tient-contacting components and the shelf-life of these components;

  - Non-clinical and clinical perform- ance testing should demonstrate sub- stantial equivalence in safety and ef- fectiveness, including durability, com- patibility, migration resistance, corro- sion resistance, and delivery and de- ployment;

  - Non-clinical testing should evalu- ate the compatibility of the device in an magnetic resonance (MR) environ- ment;

  - Appropriate analysis and non-clin- ical testing should validate electro- magnetic compatibility (EMC) and electrical safety;

#### 21 CFR Ch. I (4–1–17 Edition)

  - The sale, distribution, and use of the device are restricted to prescrip- tion  use  in  accordance  with  21  CFR

801.109 of this chapter; and

  - Labeling must bear all informa- tion required for the safe and effective use   of   the   device   as   outlined   in

§ 801.109(c) of this chapter, including a detailed summary of  the  non-clinical and clinical evaluations  pertinent  to use of the device.

[77 FR 8119, Feb. 14, 2012]

##### § 870.3470 Intracardiac patch or pledg- et made of polypropylene, poly- ethylene terephthalate, or poly- tetrafluoroethylene.

- Identification. An intracardiac patch or pledget made of poly- propylene, polyethylene terephthalate, or polytetrafluoroethylene is a fabric device placed in the heart that is used to repair septal defects, for patch graft- ing, to repair tissue, and to buttress su- tures.

- Classification. Class II (perform- ance standards).

##### § 870.3535 Intra-aortic balloon and control system.

- Identification. An intra-aortic bal- loon and control system is a prescrip- tion device that consists of an inflat- able balloon, which is placed in the aorta to improve cardiovascular func- tioning during certain life-threatening emergencies, and a control system for regulating the inflation and  deflation of the balloon. The control system, which monitors and is synchronized with the electrocardiogram, provides a means for setting the inflation and de- flation of the balloon with the cardiac cycle.

- Classification. (1) Class II (special controls) when the device is indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complica- tions of heart failure. The special con- trols for this device are:

- Appropriate analysis and non-clin- ical testing must be conducted to vali- date electromagnetic  compatibility and electrical safety of the device;

- Software verification, validation, and hazard analysis must be performed;

- The device must be dem- onstrated to be biocompatible;

#### Food and Drug Administration, HHS	§ 870.3600

- Sterility and shelf-life testing must demonstrate the  sterility of  pa- tient-contacting components and the shelf life of these components;

- Non-clinical performance evalua- tion of the device must demonstrate mechanical integrity,  durability, and reliability to support its intended pur- pose; and

- Labeling must include a detailed summary of the device- and procedure- related complications pertinent to use of the device.

(2) Class III (premarket approval) when the device is indicated for septic shock and pulsatile flow generation.

- Date premarket approval application (PMA) or notice of completion of product development protocol (PDP) is required. A PMA or notice of completion of a PDP is required to be filed with the  Food and Drug Administration on or before March 31, 2014, for any intra-aortic bal- loon and control system indicated for septic shock or pulsatile flow genera- tion that was in commercial distribu- tion before May 28, 1976, or that has, on or before March 31, 2014, been found to be substantially equivalent to any intra-aortic balloon and control system indicated for septic shock or pulsatile flow generation that was in commer- cial distribution before May 28, 1976. Any other intra-aortic balloon and con- trol system indicated for  septic  shock or pulsatile flow generation shall  have an approved PMA or declared com- pleted PDP in  effect  before  being placed in commercial distribution.

[78 FR 79303, Dec. 31, 2013]

##### § 870.3545   Ventricular   bypass   (assist) device.

- Identification. A ventricular by- pass (assist) device is a device that as- sists the left or right ventricle in main- taining circulatory blood flow. The de- vice is either totally or partially im- planted in the body.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before November 21, 2011, for any ventricular bypass (assist) de- vice that was in commercial distribu- tion before May 28, 1976, or that has, on

or before November 21, 2011, been found to be substantially equivalent to any ventricular bypass (assist) device that was in commercial distribution before May 28, 1976. Any other ventricular by- pass (assist) device shall have an ap- proved PMA or declared  completed PDP in effect before being placed in commercial   distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 76 FR 50666, Aug. 16, 2011]

##### § 870.3600   External   pacemaker   pulse generator.

    - Identification. An external pace- maker pulse generator (EPPG) is a pre- scription device that has a power sup- ply and electronic circuits  that produce a periodic electrical pulse to stimulate the heart. This device, which is used outside the body, is used as a temporary substitute for the heart’s intrinsic pacing system until a perma- nent pacemaker can be implanted, or to control  irregular  heartbeats  in  pa- tients following cardiac surgery or a myocardial infarction. The device may have adjustments for impulse strength, duration, R-wave sensitivity, and other pacing variables.

    - Classification. Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing must validate electromagnetic compatibility (EMC) within a hospital environment.

- Electrical bench testing must demonstrate device safety during in- tended use. This must include testing with the specific power source (i.e., bat- tery power, AC mains connections, or both).

- Non-clinical performance testing data must demonstrate the perform- ance characteristics of the device. Testing must include the following:

  - Testing must demonstrate the ac- curacy  of   monitoring   functions, alarms, measurement features, thera- peutic features, and all adjustable or programmable parameters as identified in  labeling;

  - Mechanical bench testing of ma- terial strength must demonstrate that the device and connection cables will withstand forces or conditions encoun- tered during use;

#### § 870.3605

  - Simulated use analysis/testing must demonstrate adequate user inter- face for adjustable parameters, per- formance of alarms, display screens, interface with external devices  (e.g. data storage, printing), and  indi- cator(s) functionality under intended use conditions; and

  - Methods and instructions for cleaning the pulse generator and con- nection cables must be validated.

- Appropriate software verification, validation,  and  hazard  analysis  must be performed.

- Labeling must include the fol- lowing:

  - The labeling must clearly  state that these devices are intended for use in a hospital environment and under the supervision of a clinician trained in their use;

  - Connector terminals should be clearly, unambiguously marked on the outside of the EPPG device. The mark- ings should identify positive (+) and negative (¥) polarities. Dual chamber devices should clearly  identify  atrial and  ventricular  terminals;

  - The labeling must list all pacing modes available in the device;

  - Labeling must include a detailed description of any special capabilities (e.g., overdrive pacing or automatic mode switching); and

  - Appropriate electromagnetic com- patibility information must be in- cluded.

[81 FR 22529, Apr. 18, 2016]

##### § 870.3605   Pacing system analyzer.

- Identification. A pacing system an- alyzer (PSA) is a prescription device that combines the functionality of a pacemaker   electrode   function   tester (§ 870.3720) and an external pacemaker pulse generator (EPPG) (§ 870.3600). It is connected to a pacemaker  lead  and uses a power supply and electronic cir- cuits to supply an accurately cali- brated, variable pacing pulse for meas- uring the patient’s  pacing  threshold and intracardiac R-wave potential.  A PSA may be a single, dual, or triple chamber system and can simulta- neously deliver pacing therapy while testing one or more implanted pacing leads.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing must validate electromagnetic compatibility (EMC) within a hospital environment.

- Electrical bench testing must demonstrate device safety during in- tended use. This must include testing with the specific power source (i.e., bat- tery power, AC mains connections, or both).

- Non-clinical performance testing data must demonstrate the perform- ance characteristics of the device. Testing must include the following:

  - Testing must demonstrate the ac- curacy  of   monitoring   functions, alarms, measurement features, thera- peutic features, and all adjustable or programmable parameters as identified in  labeling;

  - Mechanical bench testing of ma- terial strength must demonstrate that the device and connection cables will withstand forces or conditions encoun- tered during use;

  - Simulated use analysis/testing must demonstrate adequate user inter- face for adjustable parameters, per- formance of alarms, display screens, interface with external devices  (e.g. data storage, printing), and  indi- cator(s) functionality under intended use conditions; and

  - Methods and instructions for cleaning the pulse generator and con- nection cables must be validated.

- Appropriate software verification, validation,  and  hazard  analysis  must be performed.

- Labeling must include the fol- lowing:

  - The labeling must clearly  state that these devices are intended for use in a hospital environment and under the supervision of a clinician trained in their use;

  - Connector terminals should be clearly, unambiguously marked on the outside of the PSA. The markings should identify positive (+) and nega- tive (¥) polarities. Dual chamber de- vices should clearly identify atrial and ventricular terminals. Triple chamber devices should clearly identify atrial, right ventricular, and left ventricular terminals;

#### Food and Drug Administration, HHS	§ 870.3670

  - The labeling must list all pacing modes available in the device;

  - Labeling must include a detailed description of any special capabilities (e.g., overdrive pacing or automatic mode switching);

  - Labeling must limit the use of ex- ternal pacing to the implant procedure; and

  - Appropriate electromagnetic compatibility information must be in- cluded.

[81 FR 22350, Apr. 18, 2016]

##### § 870.3610   Implantable	pacemaker pulse generator.

- Identification. An implantable pacemaker pulse generator is a device that has a power supply and electronic circuits that produce a periodic elec- trical pulse to stimulate  the  heart. This device is used as a substitute for the heart’s intrinsic pacing system to correct both intermittent and contin- uous cardiac rhythm disorders. This device may include triggered, inhib- ited, and asynchronous modes and is implanted in the human body.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before September 20, 2012, for any implantable pacemaker pulse generator device that was in commer- cial distribution before May 28, 1976, or that has, on or before September 20, 2012, been found to be substantially equivalent to any implantable pace- maker pulse generator device that was in commercial distribution before May 28, 1976. Any other implantable pace- maker pulse generator device shall have an approved PMA or declared completed PDP in effect before being placed in commercial distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 77 FR 37576, June 22, 2012]

##### § 870.3620   Pacemaker lead adaptor.

- Identification. A pacemaker lead adaptor is a device used to adapt a pacemaker lead so that it can be con- nected to a pacemaker pulse generator produced by a different manufacturer.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Guidance for the  Submission of Research and Marketing Applica- tions for Permanent Pacemaker Leads and for Pacemaker Lead Adaptor 510(k) Submissions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 66 FR 18542, Apr. 10, 2001]

##### § 870.3630   Pacemaker  generator  func- tion analyzer.

- Identification. A pacemaker gener- ator function analyzer is a device that is connected to a pacemaker pulse gen- erator to test any or all of the genera- tor’s parameters, including pulse dura- tion, pulse amplitude, pulse rate, and sensing threshold.

- Classification. Class II (perform- ance standards).

##### § 870.3640   Indirect   pacemaker   gener- ator function analyzer.

- Identification. An indirect pace- maker generator function analyzer is an electrically powered device that is used to determine pacemaker function or pacemaker battery function by peri- odically monitoring an implanted pace- maker’s pulse rate and pulse width. The device is noninvasive, and it de- tects pacemaker pulse rate and width via external electrodes in contact with the patient’s skin.

- Classification. Class II (perform- ance standards).

##### § 870.3650   Pacemaker  polymeric  mesh bag.

- Identification. A pacemaker poly- meric mesh bag is an implanted device used to hold a pacemaker pulse gener- ator. The bag is designed to create a stable implant environment for the pulse generator.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.3670   Pacemaker charger.

- Identification. A pacemaker charg- er is a device used transcutaneously to

#### § 870.3680

recharge  the  batteries  of  a  recharge- able pacemaker.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.3680   Cardiovascular	permanent or temporary pacemaker electrode.

- Temporary pacemaker electrode—(1) Identification. A temporary pacemaker electrode is a device consisting of flexi- ble insulated electrical  conductors with one end connected to an external pacemaker pulse generator and the other end applied to the heart. The de- vice is used to transmit a pacing elec- trical stimulus from the pulse gener- ator to the heart and/or to  transmit the electrical signal of the heart to the pulse generator.

(2) Classification. Class II (perform- ance standards).

- Permanent pacemaker electrode—(1) Identification. A permanent pacemaker electrode is a device consisting of flexi- ble insulated electrical  conductors with one end connected to an implantable pacemaker pulse gener- ator and the other end applied to the heart. The device is used to transmit a pacing electrical stimulus from the pulse generator to the heart and/or to transmit the electrical signal of the heart to the pulse generator.

(2) Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before October 4, 2012, for any permanent pacemaker electrode device that was in commercial dis- tribution before May 28, 1976, or that has, on or before October 4, 2012, been found to be substantially equivalent to any permanent pacemaker electrode device that was in commercial dis- tribution before May 28, 1976. Any other pacemaker repair or replacement material device shall have an approved PMA or declared completed PDP in ef-

#### 21 CFR Ch. I (4–1–17 Edition)

fect before being placed in commercial distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 77 FR 39927, July 6, 2012]

##### § 870.3690   Pacemaker test magnet.

- Identification. A pacemaker test magnet is a device used to test an in- hibited or triggered type of pacemaker pulse generator and cause an inhibited or triggered generator to revert to asynchronous operation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38796, July 25, 2001]

##### § 870.3700   Pacemaker programmers.

- Identification. A pacemaker pro- grammer is a device used to noninvasively change one or more of the electrical operating characteristics of a pacemaker.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before September 20, 2012, for any pacemaker programmer  that was in commercial distribution before May 28, 1976, or that has, on or before September 20, 2012, been found to be substantially equivalent to any pace- maker programmer that was in com- mercial distribution before May 28, 1976. Any other pacemaker programmer shall have an approved PMA or de- clared completed PDP in effect before being placed in commercial distribu- tion.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 77 FR 37573, June 22, 2012]

##### § 870.3710   Pacemaker   repair   or   re- placement material.

- Identification. A pacemaker repair or replacement material is an adhesive, a sealant, a screw, a crimp, or any other material used to repair a pace- maker lead or to reconnect a pace- maker lead to a pacemaker pulse gen- erator.

#### Food and Drug Administration, HHS	§ 870.3925

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before November 21, 2011, for any pacemaker repair or replace- ment material device that was in com- mercial distribution before May 28, 1976, or that has, on or before Novem- ber 21, 2011, been found to be substan- tially equivalent to any pacemaker re- pair or replacement material  device that was in commercial distribution before May 28, 1976. Any other pace- maker repair or replacement material device shall have an approved PMA or declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 76 FR 50666, Aug. 16, 2011]

##### § 870.3720   Pacemaker  electrode  func- tion tester.

- Identification. A  pacemaker  elec- trode function tester is a device which is connected to an implanted pace- maker lead that supplies an accurately calibrated, variable pacing pulse for measuring the patient’s pacing thresh- old and intracardiac R-wave potential.

- Classification. Class II (perform- ance standards).

##### § 870.3730   Pacemaker service tools.

- Identification. Pacemaker service tools are devices such as screwdrivers and Allen wrenches, used to repair a pacemaker lead or to reconnect a pace- maker lead to a pacemaker generator.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 54 FR

25049, June 12, 1989; 66 FR 38797, July 25, 2001]

##### § 870.3800   Annuloplasty ring.

- Identification. An annuloplasty ring is a rigid or flexible ring im- planted around the mitral or tricuspid heart valve for reconstructive treat- ment of valvular insufficiency.

- Classification. Class II (special controls).  The  special  control  for  this

device is the FDA guidance document entitled ‘‘Guidance for Annuloplasty Rings 510(k) Submissions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 66 FR 18542, Apr. 10, 2001]

##### § 870.3850   Carotid  sinus  nerve  stimu- lator.

- Identification. A carotid sinus nerve stimulator is an implantable de- vice used to decrease arterial pressure by stimulating Hering’s nerve at the carotid sinus.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any carotid sinus nerve stimulator that was in commercial  distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to a carotid sinus nerve stimulator that was in commercial distribution before May 28, 1976. Any other carotid sinus nerve stimulator   shall   have   an   approved PMA or a declared completed PDP in effect before being placed in commer- cial    distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17736, May 11, 1987; 61 FR 50706, Sept. 27, 1996]

##### § 870.3925   Replacement heart valve.

- Identification. A replacement heart valve is a device intended to per- form the function of any of the heart’s natural valves. This device includes valves constructed of prosthetic mate- rials, biologic valves (e.g., porcine valves), or valves constructed of a com- bination of prosthetic and biologic ma- terials.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before December 9, 1987 for any replacement heart valve  that  was in  commercial  distribution  before  May

#### § 870.3935

28, 1976, or that has on or before De- cember 9, 1987 been found to be sub- stantially equivalent to a replacement heart valve that was in commercial distribution before May 28, 1976. Any other  replacement  heart  valve  shall have an approved PMA or a declared completed PDP in effect before being placed  in  commercial  distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

18163, May 13, 1987; 52 FR 23137, June 17, 1987]

##### § 870.3935   Prosthetic heart valve hold- er.

  - Identification. A prosthetic heart valve holder is a device used to hold a replacement heart valve while it is being sutured into place.

  - Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996]

##### § 870.3945   Prosthetic heart valve sizer.

- Identification. A prosthetic heart valve sizer is a device used to measure the size of the natural valve opening to determine the size of the appropriate replacement heart valve.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 61 FR

1121, Jan. 16, 1996; 66 FR 38797, July 25, 2001]

## Subpart E—Cardiovascular Surgical Devices

##### § 870.4075   Endomyocardial  biopsy  de- vice.

- Identification. An endomyocardial biopsy device is a device used in a cath- eterization procedure to remove sam- ples of tissue from the inner wall of the heart.

- Classification. Class II (perform- ance standards).

##### § 870.4100 Extracorporeal circuit and accessories for long-term res- piratory/cardiopulmonary failure.

- Identification. An extracorporeal circuit and accessories for long-term respiratory/cardiopulmonary     support

21 CFR Ch. I (4–1–17 Edition)

(>6 hours) is a system of devices and accessories that provides assisted extracorporeal circulation and physio- logic gas exchange of the patient’s blood in patients with acute res- piratory	failure	or	acute cardiopulmonary failure, where other available  treatment  options   have failed, and continued clinical deterio- ration is expected or the risk of death is imminent. The main devices and ac- cessories of the system include, but are not limited to, the console (hardware), software, and disposables, including, but not  limited to,  an oxygenator, blood pump, heat exchanger, cannulae, tubing, filters, and other accessories (e.g., monitors, detectors, sensors, con- nectors).

- Classification—Class II (special controls). The special controls for this device are:

- The technological characteristics of the device must ensure that the ge- ometry and design parameters are con- sistent with the intended use, and that the devices and accessories in the cir- cuit are compatible;

- The devices and accessories in the circuit must be demonstrated to be bio- compatible;

- Sterility and shelf-life testing must demonstrate the sterility of any patient-contacting devices and acces- sories in the circuit and the shelf life of these devices and accessories;

- Non-clinical performance evalua- tion of the devices and accessories in the circuit must demonstrate substan- tial equivalence of the performance characteristics on the bench, mechan- ical integrity, electromagnetic com- patibility (where applicable), software, durability, and reliability;

- In vivo evaluation of the devices and accessories in the circuit must demonstrate their performance  over the intended duration of use, including a detailed summary of the clinical evaluation pertinent to the use of the devices and accessories to demonstrate their effectiveness if a specific indica- tion (patient population and/or condi- tion) is identified; and

- Labeling must include a detailed summary of the non-clinical and in vivo evaluations pertinent to use of the devices and accessories in the circuit and adequate instructions with respect

#### Food and Drug Administration, HHS	§ 870.4250

to anticoagulation, circuit setup, per- formance characteristics  with  respect to compatibility among different de- vices and accessories in the circuit, and maintenance during a procedure.

[81 FR 7451, Feb. 12, 2016]

##### § 870.4200 Cardiopulmonary bypass ac- cessory equipment.

  - Identification. Cardiopulmonary bypass accessory equipment is a device that has  no  contact  with  blood  and that is used in the cardiopulmonary bypass circuit to support, adjoin, or connect components, or to aid in the setup of the extracorporeal line,  e.g., an oxygenator mounting bracket or system-priming  equipment.

  - Classification. (1) Class I. The de- vice is classified as class I if it does not involve an electrical connection to the patient. The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 870.9.

- Class II (special controls). The de- vice is classified as class II if it in- volves an electrical connection to the patient. The special controls are as fol- lows:

  - The performance standard under part 898 of this chapter, and

  - The guidance document entitled ‘‘Guidance on the Performance Stand- ard for Electrode Lead Wires and Pa- tient Cables.’’ The device is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 870.9.

[65 FR 19319, Apr. 11, 2000]

##### § 870.4205 Cardiopulmonary bypass bubble detector.

- Identification. A cardiopulmonary bypass bubble detector is a device used to detect bubbles in the arterial return line of the cardiopulmonary bypass cir- cuit.

- Classification. Class II (perform- ance standards).

##### § 870.4210 Cardiopulmonary bypass vascular catheter, cannula, or tub- ing.

- Identification. A cardiopulmonary bypass vascular catheter, cannula, or tubing is a device used in cardiopulmonary surgery to cannulate

the vessels, perfuse the coronary arte- ries, and to interconnect the catheters and cannulas with an oxygenator. The device includes accessory bypass equip- ment.

- Classification. Class II (perform- ance standards).

##### § 870.4220   Cardiopulmonary	bypass heart-lung machine console.

- Identification. A cardiopulmonary bypass heart-lung machine console is a device that consists of a control panel and the electrical power and control circuitry for a heart-lung machine. The console is designed to interface with the basic units used in a gas exchange system, including the pumps, oxygenator, and heat exchanger.

- Classification. Class II (perform- ance standards).

##### § 870.4230   Cardiopulmonary	bypass defoamer.

- Identification. A cardiopulmonary bypass defoamer is a device used in conjunction with an oxygenator during cardiopulmonary bypass surgery to re- move gas bubbles from the blood.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Guidance for Extracorporeal Blood Circuit Defoamer 510(k) Submis- sions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 66 FR 18542, Apr. 10, 2001]

##### § 870.4240   Cardiopulmonary	bypass heat exchanger.

- Identification. A cardiopulmonary bypass heat exchanger is a device, con- sisting of a heat exchange system used in extracorporeal circulation to  warm or cool the blood or perfusion fluid flowing through the device.

- Classification. Class II (perform- ance standards).

##### § 870.4250   Cardiopulmonary	bypass temperature controller.

- Identification. A cardiopulmonary bypass temperature controller is a de- vice used to control the temperature of the fluid entering and leaving a heat exchanger.

- Classification. Class II (perform- ance standards).

#### § 870.4260

##### § 870.4260 Cardiopulmonary bypass ar- terial line blood filter.

  - Identification. A cardiopulmonary bypass arterial line blood filter is a de- vice used as part of a gas exchange (oxygenator) system to filter nonbio- logic particles and emboli  (blood  clots or pieces of foreign material flowing in the bloodstream which will obstruct circulation by blocking a vessel) out of the blood. It is used in the arterial re- turn  line.

  - Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled	‘‘Guidance	for Cardiopulmonary Bypass Arterial Line Blood Filter 510(k) Submissions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 66 FR 18542, Apr. 10, 2001]

##### § 870.4270 Cardiopulmonary bypass cardiotomy suction line blood filter.

- Identification. A cardiopulmonary bypass cardiotomy suction line blood filter is a device used as part of a gas exchange (oxygenator) system to filter nonbiologic particles and emboli (a blood clot or a piece of foreign  mate- rial flowing in the bloodstream which will obstruct circulation by blocking a vessel) out of the blood. This device is intended for use in the cardiotomy suc- tion  line.

- Classification. Class II (perform- ance standards).

##### § 870.4280 Cardiopulmonary prebypass filter.

- Identification. A cardiopulmonary prebypass filter is a device used during priming of the oxygenator circuit to remove particulates or other debris from the circuit prior to initiating by- pass. The device is not used to filter blood.

- Classification. Class II (perform- ance standards).

##### § 870.4290 Cardiopulmonary bypass adaptor, stopcock, manifold, or fit- ting.

- Identification. A cardiopulmonary bypass adaptor, stopcock, manifold, or fitting is a device used in cardio- vascular diagnostic, surgical, and therapeutic applications to inter- connect tubing, catheters, or other de- vices.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class II (perform- ance standards).

##### § 870.4300   Cardiopulmonary	bypass gas control unit.

- Identification. A cardiopulmonary bypass gas control unit is a device used to control and measure the flow of gas into the oxygenator. The device is cali- brated for a specific gas.

- Classification. Class II (perform- ance standards).

##### § 870.4310   Cardiopulmonary	bypass coronary pressure gauge.

- Identification. A cardiopulmonary bypass coronary pressure gauge is a de- vice used in cardiopulmonary bypass surgery to measure the pressure of the blood perfusing the coronary arteries.

- Classification. Class II (perform- ance standards).

##### § 870.4320   Cardiopulmonary	bypass pulsatile flow generator.

- Identification. A cardiopulmonary bypass pulsatile flow generator is an electrically and pneumatically  oper- ated device used to create pulsatile blood flow. The device is placed in a cardiopulmonary bypass circuit down- stream from the oxygenator.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before September 21, 2004, for any cardiopulmonary bypass pulsatile flow generator that was in commercial distribution before May 28, 1976, or that has, on or before Sep- tember 21, 2004, been found to be sub- stantially		equivalent	to		any cardiopulmonary bypass pulsatile flow generator that was in commercial dis- tribution before May 28, 1976. Any other	cardiopulmonary		bypass pulsatile flow generator shall have an approved PMA or declared completed PDP in effect before being placed in commercial   distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 69 FR 34920, June 23, 2004]

#### Food and Drug Administration, HHS	§ 870.4360

##### § 870.4330 Cardiopulmonary bypass on-line blood gas monitor.

  - Identification. A cardiopulmonary bypass on-line blood gas monitor is a device used in conjunction with a blood gas sensor to measure the level of gases in the blood.

  - Classification. Class II (perform- ance standards).

##### § 870.4340 Cardiopulmonary bypass level sensing monitor and/or con- trol.

- Identification. A cardiopulmonary bypass level sensing monitor and/or control is a device used to monitor and/ or control the level of blood in the blood reservoir and to sound an alarm when the level falls below a predeter- mined value.

- Classification. Class II (perform- ance standards).

##### § 870.4350 Cardiopulmonary bypass oxygenator.

- Identification. A cardiopulmonary bypass oxygenator is a device used to exchange gases between blood and a gaseous environment to satisfy the gas exchange needs of a patient during open-heart surgery.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled	‘‘Guidance	for Cardiopulmonary Bypass Oxygenators 510(k) Submissions.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 66 FR 18542, Apr. 10, 2001]

##### § 870.4360   Nonroller-type blood pump.

- Nonroller-type cardiopulmonary and circulatory bypass blood pump—(1) Iden- tification.	A	nonroller-type cardiopulmonary and circulatory by- pass blood pump is a prescription de- vice that uses a method other than re- volving rollers to pump the blood through an extracorporeal circuit for periods lasting less than 6 hours for the purpose of providing either:

- Full or partial cardiopulmonary bypass (i.e., circuit includes an oxygenator) during open surgical pro- cedures on the heart  or  great  vessels; or

- Temporary circulatory bypass for diversion of flow around a planned dis- ruption    of    the    circulatory    pathway

necessary for open surgical procedures on the aorta or vena cava.

(2) Classification—Class II (special controls). The special controls for this device are:

- Non-clinical performance testing must perform as intended over the in- tended duration of use and dem- onstrate the following: Operating pa- rameters, dynamic blood damage, heat generation, air entrapment, mechan- ical integrity, and durability/reli- ability;

- The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible;

- Sterility and shelf life testing must demonstrate the  sterility of  pa- tient-contacting components and the shelf life of these components; and

- Labeling must include informa- tion regarding the duration of use, and a detailed summary of the device- and procedure-related complications perti- nent to use of the device.

- Nonroller-type temporary ventric- ular support blood pump—(1) Identifica- tion. A nonroller-type temporary ven- tricular support blood pump is a pre- scription device that uses any method resulting in blood propulsion to provide the temporary ventricular assistance required for support of the systemic and/or pulmonary circulations during periods when there is ongoing or an- ticipated hemodynamic instability due to  immediately  reversible  alterations in ventricular myocardial function re- sulting from mechanical or physiologic causes. Duration of use would be less than 6 hours.

(2) Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of product development protocol (PDP) is required. A PMA or notice of completion of a PDP is required to be filed with FDA on or before September 8, 2015, for any non- roller-type temporary ventricular sup- port blood pump that was in commer- cial distribution before May 28, 1976, or that has, on or before September 8, 2015, been found to be substantially equivalent to any nonroller-type tem- porary ventricular support blood pump that   was   in   commercial   distribution

#### § 870.4370

before May 28, 1976. Any other non- roller-type temporary ventricular sup- port blood pump shall have an ap- proved PMA or declared  completed PDP in effect before being placed in commercial   distribution.

[80 FR 32311, June 8, 2015]

##### § 870.4370   Roller-type cardiopulmonary	bypass	blood pump.

  - Identification. A roller-type cardiopulmonary bypass blood pump is a device that uses a revolving roller mechanism to pump the blood through the cardiopulmonary bypass circuit during bypass surgery.

  - Classification. Class II (perform- ance standards).

##### § 870.4380   Cardiopulmonary	bypass pump speed control.

- Identification. A cardiopulmonary bypass pump speed control is a device used that incorporates an electrical system or a mechanical system,  or both, and  is used  to  control the  speed of	blood	pumps	used	in cardiopulmonary bypass surgery.

- Classification. Class II (perform- ance standards).

##### § 870.4390   Cardiopulmonary	bypass pump tubing.

- Identification. A cardiopulmonary bypass pump tubing is polymeric tub- ing which is used in the blood pump head and which  is cyclically com- pressed by the pump to cause the blood to flow through the cardiopulmonary bypass  circuit.

- Classification. Class II (perform- ance standards).

##### § 870.4400   Cardiopulmonary	bypass blood reservoir.

- Identification. A cardiopulmonary bypass blood reservoir is a device used in conjunction with short-term extracorporeal circulation devices to hold a reserve supply of blood in the bypass  circulation.

- Classification. Class II (perform- ance standards), except that a reservoir that contains a defoamer or filter is classified into the same class as the defoamer or filter.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 870.4410   Cardiopulmonary bypass in- line blood gas sensor.

- Identification. A cardiopulmonary bypass in-line blood gas sensor is a transducer that measures the level of gases in the blood.

- Classification. Class II (perform- ance standards).

##### § 870.4420   Cardiopulmonary	bypass cardiotomy return sucker.

- Identification. A cardiopulmonary bypass cardiotomy return sucker is a device that consists of tubing, a con- nector, and a probe or tip that is used to remove blood from the  chest  or heart during cardiopulmonary bypass surgery.

- Classification. Class II (perform- ance standards).

##### § 870.4430   Cardiopulmonary	bypass intracardiac suction control.

- Identification. A cardiopulmonary bypass intracardiac suction control is a device which provides the vacuum and control for a cardiotomy return sucker.

- Classification. Class II (perform- ance standards).

##### § 870.4450   Vascular clamp.

- Identification. A vascular clamp is a surgical instrument used to occlude a blood vessel temporarily.

- Classification. Class II (perform- ance standards).

##### § 870.4475   Surgical vessel dilator.

- Identification. A surgical vessel di- lator is a device used to enlarge or cali- brate a vessel.

- Classification. Class II (perform- ance standards).

##### § 870.4500   Cardiovascular  surgical  in- struments.

- Identification. Cardiovascular sur- gical instruments are surgical instru- ments that have special features for use in cardiovascular surgery. These devices include, e.g., forceps, retrac- tors, and scissors.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 870.5100

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 870.9.

[45 FR 7907, Feb. 5, 1980, as amended at 54 FR

25049, June 12, 1989; 66 FR 38797, July 25, 2001]

##### § 870.4510   Apical closure device.

  - Identification. An apical closure device is a prescription device con- sisting of a delivery system and im- plant component that is used for soft tissue approximation of cardiac apical tissue during transcatheter valve re- placement procedures.

  - Classification. Class II (special controls). The special controls for this device are:

- The patient  contacting  materials must be evaluated to be biocompatible.

- Performance data must validate the sterility of the patient-contacting components of the device.

- Performance data must support the shelf life of the device by dem- onstrating continued sterility, package integrity, and device  functionality over the labeled shelf life.

- Non-clinical performance testing data must demonstrate that the device performs as intended under anticipated conditions of use. The following per- formance characteristics must be test- ed:

  - Consistent and reliable implant deployment;

  - Assessment of implant pull-out force; and

  - Sheath size compatibility with implant.

- In vivo evaluation of the device must demonstrate device performance, including device operation resulting in closure of the myocardial wound.

- Labeling must include the fol- lowing:

  - Detailed information explaining how the device operates;

  - Sheath size that device can ac- commodate;

  - Identification of the minimum myocardial wall thickness to ensure optimal device function; and

  - A shelf life.

[81 FR 71371, Oct. 17, 2016]

##### § 870.4875   Intraluminal   artery   strip- per.

- Identification. An intraluminal ar- tery stripper is a device used to per-

form an endarterectomy (removal of plaque deposits from arterisclerotic ar- teries.)

- Classification. Class II (perform- ance standards).

##### § 870.4885   External vein stripper.

- Identification. An external vein stripper is an extravascular device used to remove a section of a vein.

- Classification. Class II (perform- ance standards).

## Subpart F—Cardiovascular Therapeutic Devices

##### § 870.5050 Patient care suction appa- ratus.

- Identification. A patient care suc- tion apparatus is a device used with an intrathoracic catheter to withdraw fluid from the chest during the recov- ery period following surgery.

- Classification. Class II (perform- ance standards).

##### § 870.5100 Percutaneous Transluminal Coronary Angioplasty (PTCA) Cath- eter.

- Standard PTCA Catheter—(1) Iden- tification. A PTCA catheter is a device that operates on the principle of hy- draulic pressurization applied through an inflatable balloon attached to the distal end. A PTCA balloon catheter has a single or double lumen shaft. The catheter features a balloon of appro- priate compliance for the clinical ap- plication, constructed from a polymer. The balloon is designed to uniformly expand to a specified diameter and length at a specific pressure as labeled, with well characterized rates of infla- tion and deflation and a defined burst pressure. The device generally features a type of radiographic marker to facili- tate fluoroscopic visualization of the balloon during use. A PTCA catheter is intended for balloon dilatation of a hemodynamically significant coronary artery or bypass graft stenosis in pa- tients evidencing coronary  ischemia for the purpose of improving myocar- dial perfusion. A PTCA catheter may also be intended for the treatment of acute myocardial infarction; treatment of in-stent restenosis (ISR) and/or post- deployment stent expansion.

§ 870.5150

(2) Classification. Class II (special con- trols). The special control for this de- vice is ‘‘Class II Special Controls Guid- ance	Document	for	Certain Percutaneous Transluminal Coronary Angioplasty    (PTCA)    Catheters.’’    See

§ 870.1(e)   for   the   availability   of   this guidance document.

- Cutting/scoring PTCA Catheter—(1) Identification. A cutting/scoring PTCA catheter is a balloon-tipped catheter with cutting/scoring elements at- tached, which is used in those cir- cumstances where a high pressure bal- loon resistant lesion is encountered. A cutting/scoring PTCA catheter is in- tended for the treatment of hemodynamically significant coronary artery stenosis for the purpose of im- proving myocardial perfusion. A cut- ting/scoring PTCA catheter may also be indicated for use in complex type C lesions or for the treatment of in-stent restenosis.

(2) Classification. Class III (premarket approval). As of May 28, 1976, an ap- proval under section 515 of the act is required before this device may be commercially  distributed.  See  § 870.3.

[75 FR 54496, Sept. 8, 2010]

##### § 870.5150   Embolectomy catheter.

- Identification. An embolectomy catheter is a balloon-tipped catheter that is used to remove thromboemboli, i.e., blood clots which have migrated in blood vessels from one site in the vas- cular tree to another.

- Classification. Class II (perform- ance standards).

##### § 870.5175   Septostomy catheter.

- Identification. A septostomy cath- eter is a special balloon catheter that is used to create or enlarge the atrial septal defect found in the heart of cer- tain  infants.

- Classification. Class II (perform- ance standards).

##### § 870.5200   External	cardiac	com- pressor.

- Identification. An external cardiac compressor is an externally applied prescription device that is electrically, pneumatically, or manually powered and is used to compress the chest peri- odically in the region  of the heart  to provide blood flow during cardiac ar-

#### 21 CFR Ch. I (4–1–17 Edition)

rest. External cardiac compressor de- vices are used as an adjunct to manual cardiopulmonary resuscitation (CPR) when effective manual CPR is not pos- sible (e.g., during patient transport or extended CPR when fatigue may pro- hibit the delivery of effective/con- sistent compressions to the victim, or when insufficient EMS personnel are available to provide effective CPR).

- Classification. Class II (special controls). The special controls for this device are:

  - Nonclinical performance testing under simulated physiological condi- tions must demonstrate the reliability of the delivery of specific compression depth and rate over the intended dura- tion of use.

  - Labeling must include the fol- lowing:

- The clinical training necessary for the safe use of this device;

- Adjunctive use only indication prominently displayed on labels phys- ically placed on the device and in any device manuals or other labeling;

- Information on the patient popu- lation for which the device has been demonstrated to be effective (including patient size and/or age limitations, e.g., adult, pediatric and/or infant); and

- Information on the time nec- essary to deploy the device as dem- onstrated in the performance testing.

  - For devices that incorporate elec- trical components, appropriate anal- ysis and testing must demonstrate that the device is electrically safe and elec- tromagnetically compatible in its in- tended use environment.

  - Human factors testing and anal- ysis must validate that the device de- sign and labeling are sufficient for ef- fective use by the intended user, in- cluding an evaluation for the time nec- essary to deploy the device.

  - For devices containing software, software verification, validation, and hazard analysis must be performed.

  - Components of the device that come into human contact must be dem- onstrated to be biocompatible.

[81 FR 33133, May 25, 2016]

##### § 870.5210   Cardiopulmonary  resuscita- tion (CPR) aid.

    - CPR aid without feedback—(1) Iden- tification. A CPR aid without feedback

#### Food and Drug Administration, HHS	§ 870.5225

is a device that performs a simple func- tion such  as  proper hand placement and/or simple prompting for rate and/or timing of compressions/breathing  for the professionally trained rescuer, but offers no feedback related to the qual- ity of the CPR being provided. These devices are intended for use by persons professionally trained in CPR to assure proper use and the delivery of optimal CPR  to  the  victim.

(2) Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 870.9.

    - CPR aid with feedback—(1) Identi- fication. A CPR Aid device with feed- back is a device that provides real-time feedback to the rescuer regarding the quality of CPR being delivered to the victim, and provides either audio and/ or visual information to encourage the rescuer to continue the consistent ap- plication of effective manual CPR in accordance with current accepted CPR guidelines (to include, but not be lim- ited to, parameters such as compres- sion rate, compression depth, ventila- tion, recoil, instruction for one or mul- tiple rescuers, etc.). These devices may also perform a coaching function to aid rescuers in the sequence of steps nec- essary to perform effective CPR on a victim.

- Classification. Class II (special con- trols). The special controls for this de- vice are:

  - Nonclinical performance testing under simulated physiological or use conditions must demonstrate the accu- racy and reliability of the feedback to the user on specific compression rate, depth and/or respiration over the in- tended duration, and environment of use.

  - Labeling must include the clin- ical training, if needed, for the safe use of this device and information on the patient population for which the device has been demonstrated to be effective (including patient size and/or age limi- tations, e.g., adult, pediatric and/or in- fant).

  - For devices that incorporate electrical components, appropriate analysis and testing must demonstrate that  the  device  is  electrically  safe  and

electromagnetically  compatible  in  its intended use environment.

  - For devices containing software, software verification, validation, and hazard analysis must be performed.

  - Components of the device that come into human contact must be dem- onstrated to be biocompatible.

  - Human factors testing and anal- ysis must validate that the device de- sign and labeling are sufficient for ef- fective use by the intended user.

- Premarket notification. The  CPR Aid with feedback device is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter if it does not contain software (e.g., is mechanical or electro-mechan- ical) and is in compliance with the spe- cial controls under paragraph (b)(2) of this section, subject to  the  limitations of exemptions in § 870.9.

[81 FR 33134, May 25, 2016]

##### § 870.5225   External    counter-pulsating device.

- Identification. An external counter-pulsating device is a noninvasive, prescription device used to assist the heart by applying positive or negative pressure to one or more of the body’s limbs in synchrony with the heart cycle.

- Classification. (1) Class II (special controls) when the device is intended for the treatment of chronic stable an- gina that is refractory to optimal anti- anginal medical therapy and without options for revascularization. The spe- cial controls for this device are:

- Nonclinical performance evalua- tion of the device must demonstrate a reasonable assurance of safety and ef- fectiveness for applied pressure, syn- chronization of therapy with the appro- priate phase of the cardiac cycle, and functionality of alarms during a device malfunction or an abnormal patient condition;

- Reliabilities of the mechanical and electrical systems must be estab- lished through bench testing under simulated use conditions and matched by appropriate maintenance schedules;

- Software design and verification and validation must be appropriately documented;

#### § 870.5300

- The skin-contacting components of the device must be demonstrated to be biocompatible;

- Appropriate analysis and testing must be conducted to verify electrical safety and electromagnetic compat- ibility of the device; and

- Labeling must include a detailed summary of the device-related and pro- cedure-related complications pertinent to use of the device.

(2) Class III (premarket approval) for the following intended uses: Unstable angina pectoris; acute myocardial in- farction; cardiogenic shock; congestive heart failure; postoperative treatment of patients who have undergone coro- nary artery bypass surgery; peripheral arterial disease  associated  with ischemic ulcers rest pain or claudication, threatened gangrene, in- sufficient blood supply at an amputa- tion site, persisting ischemia after embolectomy or bypass surgery, and/or pre- and post-arterial reconstruction to improve runoff;  diabetes  complicated by peripheral arterial disease or other conditions possibly related to arterial insufficiency including nocturnal leg cramps	and/or	necrobiosis diabeticorum; venous diseases, includ- ing prophylaxis of deep vein thrombophlebitis, edema (e.g., chronic lymphedema) and/or induration (e.g., stasis dermatitis) associated with chronic venous stasis, venous stasis ul- cers, and/or thrombophlebitis; athletic injuries, including Charley  horses, pulled muscles and/or edematous mus- cles;  necrotizing  cellulitis.

- Date premarket approval application (PMA) or notice of completion of product development protocol (PDP) is required. A PMA or notice of completion of a PDP is required to be filed with FDA on or before March 31, 2014, for any external counter-pulsating device, with an in- tended use described in paragraph (b)(2) of this section, that was in commercial distribution before May 28, 1976, or that has, on or before March 31, 2014, been found to be substantially equivalent to any external counter-pulsating device, with an intended use described in para- graph (b)(2) of this section, that was in commercial distribution before May 28, 1976. Any other external counter-pul- sating device with an intended use de- scribed in paragraph (b)(2) of this sec-

#### 21 CFR Ch. I (4–1–17 Edition)

tion shall have an approved PMA or de- clared completed PDP in effect before being placed in commercial distribu- tion.

[78 FR 79307, Dec. 30, 2013]

##### § 870.5300   DC-defibrillator	(including paddles).

- Low-energy DC-defibrillator—(1) Identification. A low-energy DC- defibrillator is a device that delivers into a 50 ohm test load an electrical shock of a maximum of 360 joules of en- ergy used for defibrillating (restoring normal heart rhythm) the atria or ven- tricles of the heart or  to  terminate other cardiac arrhythmias. This ge- neric type of device includes low en- ergy defibrillators with a maximum electrical output of less than 360 joules of energy that are used in pediatric defibrillation or in cardiac  surgery. The device may either synchronize the shock with the proper phase of the electrocardiogram or may operate asynchronously. The  device   delivers the electrical shock through paddles placed either directly across  the  heart or on the surface of the body.

(2) Classification. Class II (perform- ance standards).

- High-energy DC-defibrillator—(1) Identification. A high-energy DC- defibrillator is a device that delivers into a 50 ohm test load an electrical shock of greater than 360 joules of en- ergy used for defibrillating the atria or ventricles of the heart or to terminate other cardiac arrhythmias. The device may either synchronize the shock with the proper phase of the electrocardio- gram or may operate asynchronously. The device delivers the electrical shock through paddles placed either directly across the heart or on the  surface  of the body.

(2) Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any DC-defibrillator (including pad- dles) described in paragraph (b)(1) of this section that was in commercial distribution before May 28, 1976, or that has,  on  or  before  December  26,  1996

#### Food and Drug Administration, HHS	§ 870.5700

been found to be substantially equiva- lent to a DC-defibrillator (including paddles) described in paragraph (b)(1) of this section that was in commercial distribution before May 28, 1976. Any other DC-defibrillator (including pad- dles) described in paragraph (b)(1) of this section shall have  an  approved PMA or declared completed PDP in ef- fect before being placed in commercial distribution.

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 61 FR 50706, Sept. 27, 1996]

##### § 870.5310   Automated	external defibrillator system.

  - Identification. An automated ex- ternal defibrillator (AED) system con- sists of an AED and those accessories necessary for the AED to detect and in- terpret an electrocardiogram and de- liver an electrical shock (e.g., battery, pad electrode, adapter, and hardware key for pediatric use). An AED system analyzes the  patient’s  electrocardio- gram, interprets the cardiac rhythm, and automatically delivers an elec- trical shock (fully automated AED), or advises the user to deliver the shock (semi-automated or shock advisory AED) to treat ventricular fibrillation or pulseless ventricular tachycardia.

  - Classification. Class III (premarket approval)

  - Date PMA or notice of completion of PDP is required. A PMA will be required to be submitted to the Food and Drug Administration by April 29, 2015, for any AED that was in commercial dis- tribution before May 28, 1976, or that has, by April 29, 2015, been found to be substantially equivalent to any  AED that was in commercial distribution before May 28, 1976. A PMA will be re- quired to be submitted to the Food and Drug Administration by April 29, 2015, for any AED accessory described in paragraph (a) that was in commercial distribution before May 28, 1976, or that has, by April 29, 2015, been found to be substantially equivalent to any AED accessory described in paragraph (a) that was in commercial distribution before May 28, 1976. Any other AED and AED accessory described in paragraph (a), shall have an approved PMA or de- clared  completed  PDP  in  effect  before

being  placed  in  commercial  distribu- tion.

[68  FR 61344,  Oct. 28,  2003;  69  FR 10615,  Mar.

8, 2004, as amended at 80 FR 5682, Feb. 3, 2015]

##### § 870.5325   Defibrillator tester.

- Identification. A defibrillator test- er is a device that is connected to the output of a defibrillator and is used to measure the energy delivered by the defibrillator into a standard resistive load. Some testers also provide wave- form information.

- Classification. Class II (perform- ance standards).

##### § 870.5550   External	transcutaneous cardiac pacemaker (noninvasive).

- Identification. An external trans- cutaneous	cardiac	pacemaker (noninvasive) is a device used to supply a periodic electrical pulse intended to pace the heart. The pulse from the de- vice is usually applied to the surface of the chest through electrodes such as defibrillator paddles.

- Classification. Class II. The special controls for this device are:

  - ‘‘American National Standards In- stitute/American Association for Med- ical Instrumentation’s DF–21 ‘Cardiac Defibrillator Devices’ ’’ 2d ed., 1996, and

  - ‘‘The maximum pulse amplitude should not exceed 200  milliamperes. The maximum pulse duration should not exceed 50 milliseconds.’’

[45 FR 7907, Feb. 5, 1980, as amended at 52 FR

17737, May 11, 1987; 65 FR 17144, Mar. 31, 2000]

##### § 870.5700   Steerable   cardiac   ablation catheter remote control system.

    - Identification. A steerable cardiac ablation catheter remote control sys- tem is a prescription device that is ex- ternal to the body and interacts with the manual handle of a steerable car- diac ablation catheter to remotely con- trol the advancement, retraction, rota- tion, and deflection of a compatible, steerable ablation catheter used for the treatment of cardiac arrhythmias in the right side of the heart. The device allows reversion to manual control of the steerable cardiac ablation catheter without withdrawal of the catheter and interruption of the procedure.

#### § 870.5700

    - Classification. Class II (special controls). The special controls for this device are:

- Non-clinical mechanical perform- ance testing  must  demonstrate  that the device performs as intended under anticipated conditions of use. The fol- lowing performance testing must be performed:

  - Mechanical performance of the system (without catheter connected);

  - Mechanical performance of the system with compatible catheters con- nected to verify that the system does not impact catheter function or per- formance. Assessments must include the following:

    - Side-by-side remote control and manual comparisons of catheter ma- nipulation (including all ranges of mo- tion of catheter deflection and tip curl) for all compatible catheters; must in- clude testing for worst-case conditions, and

    - Evaluation of the accuracy and function of all device control safety features; and

  - Simulated-use testing in a bench anatomic model or animal model.

- Non-clinical electrical testing must include validation of electro- magnetic compatibility (EMC), elec- trical safety, thermal safety, and elec- trical system performance. The fol- lowing performance testing must be performed:

- Electrical performance of the sys- tem with compatible catheters con- nected to verify that the system does not impact catheter function or per- formance. Assessments must include the following:

- Side-by-side remote control and manual comparisons of catheter ma- nipulation (including all ranges of mo- tion of catheter deflection and tip curl) for all compatible catheters; must in- clude testing for worst-case conditions, and

- Evaluation of the accuracy and function of all device control safety features; and

- Electrical safety between the de- vice and ablation catheter system and with other electrical equipment ex- pected in the catheter lab or operating room.

- In vivo testing must demonstrate that  the  device  performs  as  intended

#### 21 CFR Ch. I (4–1–17 Edition)

under anticipated conditions of use, in- cluding an assessment of the system impact on the functionality and per- formance of compatible catheters, and documentation of the adverse event profile associated with clinical use. Evidence must be submitted to address the following:

- Manipulation and Positioning: Ability to manipulate compatible cath- eters to pre-specified cardiac locations and confirm proper anatomic place- ment and tissue contact, in accordance with the system indications for use and the compatible catheter indications for use;

- Safety: Assess device-related complication rate and major proce- dural complication rate (regardless of device relatedness) in comparison to literature and/or a manual comparison group for compatible ablation cath- eters to support the indications for use;

- Efficacy: Assess ablation success in comparison to literature and/or a manual comparison group for compat- ible ablation catheters to support the indications for use; and

- User assessment of device remote controls and safety features.

- Post-market surveillance (PMS) must be conducted and completed in accordance with FDA agreed upon PMS protocol.

- A training program must be in- cluded with sufficient educational ele- ments that, upon completion of the training program, the clinician and supporting staff can:

  - Identify the safe environments for device use,

  - Use all safety features of device, and

  - Operate the device in  simulated or actual use environments representa- tive of indicated environments and use for the indication of compatible cath- eters.

- Performance data must dem- onstrate the sterility of the sterile dis- posable components of the system.

- Performance data must support shelf life by demonstrating continued sterility of the device (of the sterile disposable components), package integ- rity, and device functionality over the requested shelf life.

- Labeling must include the fol- lowing:

#### Food and Drug Administration, HHS	§ 870.5910

  - Appropriate instructions, warn- ings, cautions, limitations, and infor- mation related to the intended patient population, compatible ablation cath- eters, and the device safeguards for the safe use of the device;

  - Specific instructions and the clin- ical training needed for the safe use of the device, which includes:

- Instructions on assembling the device in all available configurations, including installation and removal of compatible catheters;

- Instructions and explanation of all controls, inputs, and outputs;

- Instructions on all available modes or states of the device;

- Instructions on all safety fea- tures of the device; and

- Validated methods and instruc- tions for reprocessing/disinfecting any reusable components;

  - A detailed summary of the me- chanical compatibility testing includ- ing:

    - A table with a complete list of compatible catheters tested (manufac- turer trade name and model number), and

    - A table with detailed test results, including type of test, acceptance cri- teria, and test results (i.e., pass for meeting acceptance criteria);

  - A detailed summary of the in vivo testing including:

    - A table with a complete list of compatible catheters used during test- ing (manufacturer trade name  and model number);

    - Adverse events encountered perti- nent to use of the device under use con- ditions;

    - A detailed summary of the device- and procedure-related com- plications;  and

    - A summary of study outcomes and  endpoints.  Information  pertinent to the fluoroscopy times/exposure for the procedure, patient, and operator fluoroscopic exposure;

  - Other labeling items:

    - A detailed summary of pertinent non-clinical testing information: EMC, mechanical, electrical, and steriliza- tion of device and components;

    - A detailed summary of the device technical parameters; and

    - An expiration date/shelf life and storage conditions for the sterile acces- sories; and

  - When available, and according to the timeframe included in the PMS protocol agreed upon with FDA, pro- vide a detailed summary of the PMS data including:

    - Updates to the labeling to accu- rately reflect outcomes or necessary modifications based upon data col- lected during the PMS experience, and

    - Inclusion of results and adverse events associated with utilization of the device during the PMS.

[80 FR 58606, Sept. 30, 2015]

##### § 870.5800   Compressible limb sleeve.

- Identification. A compressible limb sleeve is a device that is used to pre- vent pooling of blood in a limb by in- flating periodically a sleeve around the limb.

- Classification. Class II (perform- ance standards).

##### § 870.5900   Thermal regulating system.

- Identification. A thermal regu- lating system is an external system consisting of a device that is placed in contact with the patient and a tem- perature controller for the device. The system is used to regulate patient tem- perature.

- Classification. Class II (perform- ance standards).

##### § 870.5910   Esophageal  thermal  regula- tion device.

- Identification. An esophageal ther- mal regulation device is a prescription device used to apply a specified tem- perature to the endoluminal surface of the esophagus via an external con- troller. This device may incorporate a mechanism for gastric decompression and suctioning. The device is used to regulate patient temperature.

- Classification. Class II (special controls). The special controls for this device are:

- The patient contacting materials must be demonstrated to be biocompat- ible.

- Non-clinical performance evalua- tion must demonstrate that the device performs as intended under anticipated

#### § 870.5925

conditions of use. The following per- formance characteristics must be test- ed:

  - Mechanical integrity testing.

  - Testing to determine tempera- ture change rate(s).

  - Testing to demonstrate compat- ibility with the indicated external con- troller.

  - Shelf life testing.

- Animal testing must demonstrate that the device does not cause esopha- geal injury and that body temperature remains within appropriate boundaries under anticipated conditions of use.

- Labeling must include the fol- lowing:

- Detailed insertion instructions.

- Warning against attaching the device to unintended connections, such as external controllers for which the device is not indicated, or pressurized air outlets instead of vacuum outlets for those devices, including gastric suc- tion.

- The operating parameters, name, and model number of the indicated ex- ternal  controller.

- The intended duration of use.

[80 FR 49896, Aug. 18, 2015]

##### § 870.5925   Automatic	rotating	tour- niquet.

- Identification. An automatic rotat- ing tourniquet is a device that prevents blood flow in one limb at a time, which temporarily reduces the total blood volume, thereby reducing the normal workload of the heart.

- Classification. Class II (perform- ance standards).

## PART 872—DENTAL DEVICES

##### Subpart A—General Provisions

Sec.

872.1    Scope.

872.3   Effective dates of requirement for pre- market  approval.

872.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

872.1500    Gingival fluid measurer. 872.1720  Pulp tester.

872.1730    Electrode gel for pulp testers. 872.1740   Caries detection device.

21 CFR Ch. I (4–1–17 Edition)

872.1745 Laser fluorescence caries detection device.

872.1800    Extraoral source x-ray system. 872.1810    Intraoral source x-ray system. 872.1820    Dental   x-ray   exposure   alignment

device.

872.1830    Cephalometer.

872.1840 Dental x-ray position indicating de- vice.

872.1850    Lead-lined position indicator. 872.1870    Sulfide detection device.

872.1905    Dental x-ray film holder. 872.2050   Dental sonography device. 872.2060    Jaw tracking device.

##### Subpart C [Reserved]

Subpart D—Prosthetic Devices

872.3060   Noble metal alloy.

872.3070 Dental amalgam, mercury, and amalgam   alloy.

872.3080   Mercury and alloy dispenser. 872.3100    Dental amalgamator.

872.3110    Dental amalgam capsule. 872.3130  Preformed anchor.

872.3140   Resin applicator.

872.3150   Articulator.

872.3165    Precision attachment. 872.3200    Resin tooth bonding agent. 872.3220 Facebow.

872.3240   Dental bur.

872.3250    Calcium hydroxide cavity liner. 872.3260   Cavity varnish.

872.3275    Dental cement.

872.3285  Preformed clasp.

872.3300 Hydrophilic resin coating for den- tures.

872.3310   Coating material for resin fillings. 872.3330  Preformed crown.

872.3350   Gold or stainless steel cusp. 872.3360  Preformed cusp.

872.3400 Karaya and sodium borate with or without acacia denture adhesive.

872.3410 Ethylene oxide homopolymer and/or carboxymethylcellulose sodium denture adhesive.

872.3420    Carboxymethylcellulose	sodium and cationic polyacrylamide polymer denture adhesive.

872.3450 Ethylene oxide homopolymer and/or karaya  denture  adhesive.

872.3480 Polyacrylamide polymer (modified cationic)  denture   adhesive.

872.3490    Carboxymethylcellulose	sodium and/or polyvinylmethylether maleic acid calcium-sodium double salt denture ad- hesive.

872.3500 Polyvinylmethylether maleic anhy- dride (PVM-MA), acid copolymer, and carboxymethylcellulose	sodium (NACMC) denture adhesive.

872.3520   OTC denture cleanser. 872.3530    Mechanical denture cleaner. 872.3540   OTC denture cushion or pad. 872.3560   OTC denture reliner.

872.3570  OTC denture repair kit.

#### Food and Drug Administration, HHS	§ 872.1

872.3580  Preformed gold denture tooth. 872.3590  Preformed plastic denture tooth. 872.3600    Partially fabricated denture kit. 872.3630   Endosseous  dental  implant  abut-

ment.

872.3640   Endosseous dental implant. 872.3645     Subperiosteal implant material. 872.3660   Impression material.

872.3661    Optical    Impression    Systems    for CAD/CAM.

872.3670  Resin impression tray material. 872.3680 Polytetrafluoroethylene	(PTFE)

vitreous carbon materials.

872.3690    Tooth shade resin material. 872.3710    Base metal alloy.

872.3730    Pantograph.

872.3740    Retentive and splinting pin.

872.3750    Bracket   adhesive   resin   and   tooth conditioner.

872.3760    Denture  relining,  repairing,  or  re- basing  resin.

872.3765   Pit  and  fissure  sealant  and  condi- tioner.

872.3770   Temporary crown and bridge resin. 872.3810   Root canal post.

872.3820 Root canal filling resin. 872.3830 Endodontic paper point. 872.3840 Endodontic silver point. 872.3850   Gutta percha.

872.3890   Endodontic stabilizing splint. 872.3900   Posterior   artificial   tooth   with   a

metal insert.

872.3910   Backing and facing for an artificial tooth.

872.3920   Porcelain tooth. 872.3930    Bone grafting material.

872.3940 Total	temporomandibular	joint prosthesis.

872.3950   Glenoid fossa prosthesis. 872.3960   Mandibular condyle prosthesis.

872.3970   Interarticular	disc	prosthesis (interpositional     implant).

872.3980    Endosseous   dental   implant   acces- sories.

##### Subpart E—Surgical Devices

872.4120    Bone cutting instrument and acces- sories.

872.4130   Intraoral dental drill.

872.4200   Dental handpiece and accessories. 872.4465   Gas-powered jet injector.

872.4475 Spring-powered jet injector. 872.4535 Dental diamond instrument. 872.4565    Dental hand instrument.

872.4600    Intraoral ligature and wire lock. 872.4620    Fiber optic dental light.

872.4630    Dental operating light. 872.4730    Dental injecting needle. 872.4760  Bone plate.

872.4770   Temporary  mandibular  condyle  re- construction    plate.

872.4840    Rotary scaler.

872.4850   Ultrasonic scaler.

872.4880   Intraosseous fixation screw or wire. 872.4920    Dental  electrosurgical  unit  and  ac-

cessories.

##### Subpart F—Therapeutic Devices

872.5410   Orthodontic   appliance   and   acces- sories.

872.5470    Orthodontic plastic bracket. 872.5500    Extraoral orthodontic headgear. 872.5525   Preformed tooth positioner.

872.5550   Teething ring.

872.5560    Electrical salivary stimulatory sys- tem.

872.5570 Intraoral devices for snoring and intraoral devices for snoring and obstruc- tive sleep apnea.

872.5580   Oral rinse to reduce the adhesion of dental plaque.

##### Subpart G—Miscellaneous Devices

872.6010   Abrasive device and accessories. 872.6030    Oral     cavity     abrasive     polishing

agent.

872.6050    Saliva absorber.

872.6070   Ultraviolet   activator   for   polym- erization.

872.6080   Airbrush.

872.6100   Anesthetic warmer.

872.6140  Articulation paper. 872.6200  Base plate shellac.

872.6250    Dental chair and accessories. 872.6290   Prophylaxis cup.

872.6300   Rubber dam and accessories. 872.6350    Ultraviolet detector.

872.6390   Dental floss.

872.6475    Heat source for bleaching teeth. 872.6510    Oral irrigation unit.

872.6570  Impression tube.

872.6640   Dental  operative  unit  and  acces- sories.

872.6650    Massaging  pick  or  tip  for  oral  hy- giene.

872.6660   Porcelain powder for clinical use. 872.6670   Silicate protector.

872.6710    Boiling water sterilizer. 872.6730    Endodontic dry heat sterilizer. 872.6770   Cartridge syringe.

872.6855   Manual toothbrush.

872.6865  Powered toothbrush. 872.6870   Disposable fluoride tray. 872.6880   Preformed impression tray. 872.6890   Intraoral dental wax.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  52  FR  30097,  Aug.  12,  1987,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 872 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 872.1   Scope.

- This part sets forth the classifica- tion of dental devices intended for human use that are in commercial dis- tribution.

§ 872.3

- The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

- To avoid duplicative listings, a dental device that has two or more types of uses (e.g., used both as a diag- nostic device and as a therapeutic de- vice) is listed in one subpart only.

- References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

- Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 30097, Aug. 12, 1987, as amended at 68

FR 19737,  Apr. 22,  2003;  79  FR 50552,  Aug. 25,

2014]

##### § 872.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a device, has an approval under section 515 of the act, FDA must promulgate a regula- tion under section 515(b) of the act re- quiring such approval, except as pro- vided in paragraphs (b) and (c) of this section.  Such  a  regulation  under  sec-

#### 21 CFR Ch. I (4–1–17 Edition)

tion 515(b) of the act shall not be effec- tive during the grace period ending on the 90th day  after  its  promulgation or on the last day of the 30th full calendar month after the regulation that classi- fies the device into class III is effec- tive, whichever is later. See section 501(f)(2)(B) of the act. Accordingly, un- less an effective date of the require- ment for premarket approval is shown in the regulation for a device classified into class III in this part, the device may be commercially distributed with- out FDA’s issuance of an order approv- ing a PMA or declaring completed a PDP for the device.  If  FDA  promul- gates a regulation under section 515(b) of the act requiring premarket ap- proval for a device, section 501(f)(1)(A) of the act applies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

- A device identified in a regulation in this part that is classified into class III and that is subject to the transi- tional provisions of section 520(1) of the act is automatically classified by stat- ute into class III and must have an ap- proval under section 515 of the act be- fore being commercially distributed. Accordingly, the regulation for such a class III transitional device states that as of the enactment date of the amend- ments, May 28, 1976, the device must have an approval under section 515 of the act before commercial distribution.

#### Food and Drug Administration, HHS	§ 872.1730

##### § 872.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening,   diagnosis,   or   monitoring

life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2314, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 872.1500   Gingival fluid measurer.

- Identification. A gingival fluid measurer is a gauge device intended to measure the amount of fluid in the gin- gival sulcus (depression between the tooth and gums) to determine if there is a gingivitis condition.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63007,  Dec. 7,  1994;  66  FR 38797,  July 25,

2001]

##### § 872.1720   Pulp tester.

- Identification. A pulp tester is an AC or battery powered device intended to evaluate the pulpal vitality of teeth by employing high frequency current transmitted by an electrode to stimu- late the nerve tissue in the dental pulp.

- Classification. Class II.

##### § 872.1730   Electrode  gel  for  pulp  test- ers.

- Identification. An electrode gel for pulp testers is a device intended to be applied to the surface of a tooth before use of a pulp tester to aid conduction of electrical current.

§ 872.1740

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.1740   Caries detection device.

- Identification. The caries detection device is a device intended to show the existence of decay in a patient’s tooth by use of electrical current.

- Classification. Class II.

##### § 872.1745   Laser    fluorescence    caries detection device.

- Identification. A laser fluorescence caries detection device is a laser, a flu- orescence detector housed in a dental handpiece, and a control console that performs device calibration, as well as variable tone emitting and fluores- cence measurement functions. The in- tended use of the device is to aid in the detection of tooth decay by measuring increased laser induced fluorescence.

- Classification. Class II, subject to the following special controls:

  - Sale, distribution, and use of this device  are  restricted   to   prescription use in accordance with § 801.109 of this chapter;

  - Premarket notifications must in- clude clinical studies, or other relevant information,  that  demonstrates   that the device aids in the  detection  of tooth decay by measuring increased laser induced fluorescence; and

  - The labeling must include de- tailed use instructions with pre- cautions that urge users to:

- Read and understand all directions before using the device,

- Store probe tips under proper conditions,

- Properly sterilize the emitter-de- tector handpick before each use, and

- Properly maintain and handle the instrument in the specified manner and condition.

[65 FR 18235, Apr. 7, 2000]

##### § 872.1800   Extraoral  source  x-ray  sys- tem.

- Identification. An extraoral source x-ray system is an AC-powered device that  produces  x-rays  and  is  intended

#### 21 CFR Ch. I (4–1–17 Edition)

for dental  radiographic  examination and diagnosis of diseases of the teeth, jaw, and oral structures. The x-ray source (a tube) is located outside the mouth. This generic type of device may include patient and equipment sup- ports and component parts.

- Classification. Class II.

##### § 872.1810   Intraoral  source  x-ray  sys- tem.

- Identification. An intraoral source x-ray system is an electrically powered device that produces x-rays and is in- tended for dental radiographic exam- ination and diagnosis of diseases of the teeth, jaw, and oral structures. The x- ray source (a tube) is located inside the mouth. This generic type of device may include patient and equipment sup- ports and component parts.

- Classification. Class II.

##### § 872.1820   Dental x-ray exposure align- ment device.

- Identification. A dental x-ray expo- sure alignment device is a device in- tended to position x-ray film and to align the examination site with the x- ray beam.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38797,  July 25,

2001]

##### § 872.1830   Cephalometer.

- Identification. A cephalometer is a device used in dentistry during x-ray procedures. The device is intended to place and to hold a patient’s head in a standard position during dental x-rays.

- Classification. Class II.

##### § 872.1840   Dental  x-ray  position  indi- cating device.

- Identification. A dental x-ray posi- tion indicating device is a device, such as a collimator, cone, or aperture, that is used in dental radiographic examina- tion. The device is intended to  align the examination site with the x-ray beam and to restrict the dimensions of the dental x-ray field by limiting the size of the primary x-ray beam.

#### Food and Drug Administration, HHS	§ 872.2050

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38797,  July 25,

2001]

##### § 872.1850   Lead-lined	position	indi- cator.

- Identification. A lead-lined posi- tion indicator is a cone-shaped device lined with lead that is attached to a dental x-ray tube and intended to aid in positioning the tube, to prevent the misfocusing of the x-rays by absorbing divergent radiation, and to prevent leakage of radiation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38797,  July 25,

2001]

##### § 872.1870   Sulfide detection device.

- Identification. A sulfide detection device is a device consisting of an AC- powered control unit, probe handle, probe tips, cables, and accessories. This device is intended to be used in vivo, to manually measure periodontal pocket probing depths, detect the presence or absence of bleeding on probing, and de- tect the presence of sulfides in peri- odontal pockets, as an adjunct in the diagnosis of periodontal diseases in adult patients.

- Classification. Class II (special controls) prescription use in accord- ance with § 801.109 of this chapter; con- formance with recognized standards of biocompatibility, electrical safety, and sterility; clinical and analytical per- formance  testing,  and  proper  labeling.

[63 FR 59717, Nov. 5, 1998]

##### § 872.1905   Dental x-ray film holder.

- Identification. A dental x-ray film holder is a device intended to position and to hold x-ray film inside the mouth.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter,     with     the     exceptions     of

§ 820.180, with respect to general re- quirements     concerning     records,     and

§ 820.198,  with  respect  to  complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.2050   Dental sonography device.

- Dental sonography device for moni- toring—(1) Identification. A dental sonography device for monitoring is an electrically  powered  device,  intended to	be	used	to	monitor temporomandibular joint sounds. The device detects and records sounds made by the temporomandibular joint.

(2) Classification. Class  I.  The  device is exempt from the premarket notifica- tion provisions of subpart E of part 807 of this chapter subject to § 872.9.

- Dental sonography device for inter- pretation and diagnosis—(1) Identifica- tion. A dental sonography device for in- terpretation and diagnosis is an elec- trically powered device, intended to in- terpret temporomandibular  joint sounds for the diagnosis of temporomandibular joint disorders and associated orofacial pain. The device detects, records, displays, and stores sounds	made	by	the temporomandibular joint during jaw movement. The device interprets these sounds to generate meaningful output, either directly or by connection to a personal computer. The device may be part of a system of devices, contrib- uting joint sound information to be considered with data from other diag- nostic components.

- Classification. Class II (special con- trols). The special control for this de- vice is FDA’s guidance document enti- tled ‘‘Class II Special Controls Guid- ance Document: Dental Sonography and Jaw Tracking Devices.’’

[68 FR 67367, Dec. 2, 2003]

#### § 872.2060

##### § 872.2060   Jaw tracking device.

  - Jaw tracking device for monitoring mandibular jaw positions relative to the maxilla—(1) Identification. A jaw track- ing device for monitoring mandibular jaw positions relative to the maxilla is a nonpowered or electrically powered device that measures and records ana- tomical distances and angles in three dimensional space, to determine the relative position of the mandible with respect to the location and position of the maxilla, while at rest and during jaw movement.

(2) Classification. Class I (general con- trols). The device is exempt from the premarket notification provisions of subpart E of part 807 of this chapter subject to § 872.9.

  - Jaw tracking device for interpreta- tion of mandibular jaw positions for the diagnosis—(1) Identification. A jaw tracking device for interpretation of mandibular jaw  positions  relative  to the maxilla for the diagnosis of temporomandibular joint disorders and associated orofacial pain is a nonpow- ered or electrically powered  device that measures and records anatomical distances and angles to determine the relative position of the mandible  in three dimensional space, with  respect to the location and position of the maxilla, while at rest and during jaw movement. The device records, dis- plays, and stores  information  about jaw position. The device interprets jaw position to generate meaningful  out- put, either directly or by connection to a personal computer. The device  may be a part of a system of devices, con- tributing jaw position information to be considered with data from other di- agnostic  components.

(2) Classification. Class II (special con- trols). The special control for this de- vice is FDA’s guidance document enti- tled ‘‘Class II Special Controls Guid- ance Document: Dental Sonography and Jaw Tracking Devices.’’

[68 FR 67367, Dec. 2, 2003]

## Subpart C [Reserved]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart D—Prosthetic Devices

##### § 872.3060   Noble metal alloy.

- Identification. A noble metal alloy is a device composed primarily of noble metals, such as gold, palladium, plat- inum, or silver, that is intended for use in the fabrication of cast or porcelain- fused-to-metal crown and bridge res- torations.

- Classification. Class II (special controls). The special control for these devices is FDA’s ‘‘Class II Special Con- trols  Guidance  Document:   Dental Noble Metal Alloys.’’ The devices are exempt from the premarket  notifica- tion procedures in subpart E of part 807 of this chapter subject to the limita- tions in § 872.9. See § 872.1(e) for avail- ability  of  guidance  information.

[69 FR 51766, Aug. 23, 2004]

##### § 872.3070   Dental   amalgam,   mercury, and amalgam alloy.

- Identification. Dental amalgam  is a device that consists of a combination of elemental mercury, supplied as a liq- uid in bulk, sachet, or predosed capsule form, and amalgam alloy composed pri- marily of silver, tin, and copper, sup- plied as a powder in bulk, tablet, or predosed capsule form, for the direct filling of carious lesions or structural defects in teeth. This device also in- cludes the individual component de- vices, mercury and amalgam alloy, when intended to be combined with each other to form dental amalgam.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Dental Amalgam, Mercury, and Amalgam Alloy.’’ See § 872.1(e) for the availability of this guidance document.

[74 FR 38714, Aug. 4, 2009]

##### § 872.3080   Mercury	and	alloy	dis- penser.

- Identification. A mercury and alloy dispenser is a device with a spring-activated valve intended to measure and dispense into a mixing capsule a predetermined amount of dental mercury in droplet form and a premeasured amount of alloy pellets.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	§ 872.3165

premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.3100   Dental amalgamator.

  - Identification. A dental amalga- mator is a device, usually AC-powered, intended to mix, by shaking, amalgam capsules containing mercury and den- tal alloy particles, such as silver, tin, zinc, and copper. The mixed dental amalgam material is intended for fill- ing dental caries.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38797,  July 25,

2001]

##### § 872.3110   Dental amalgam capsule.

- Identification. A dental amalgam capsule is a container device in which silver alloy is intended to be mixed with mercury to form dental amalgam.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.3130   Preformed anchor.

- Identification. A preformed anchor is a device made of austenitic alloys or alloys containing 75 percent or greater gold or metals of the platinum group intended to be incorporated into a den- tal appliance, such  as  a  denture,  to help stabilize the appliance in the pa- tient’s  mouth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38797,  July 25,

2001]

##### § 872.3140   Resin applicator.

- Identification. A resin applicator is a brushlike device intended for use in spreading dental resin on a tooth dur- ing application of tooth shade mate- rial.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, the device is ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this   chapter,   with   the   exceptions   of

§ 820.180,   with   respect   to   general   re- quirements     concerning     records,     and

§ 820.198,    with    respect    to    complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.3150   Articulator.

- Identification. An articulator is a mechanical device intended to simu- late movements of a patient’s upper and lower jaws. Plaster casts of the pa- tient’s teeth and gums are placed  in the device to reproduce the occlusion (bite) and articulation of the patient’s jaws. An articulator is intended to fit dentures or provide orthodontic treat- ment.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, the device is ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this   chapter,   with   the   exceptions   of

§ 820.180,   with   respect   to   general   re- quirements     concerning     records,     and

§ 820.198,    with    respect    to    complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.3165   Precision attachment.

- Identification. A precision attach- ment or preformed bar is a device made

#### § 872.3200

of austenitic alloys or  alloys  con- taining 75 percent or greater gold and metals of the platinum group intended for use in prosthetic dentistry in con- junction with removable partial  den- tures. Various forms of the device are intended to connect a lower partial denture with another lower partial denture, to connect an upper partial denture with another upper partial denture, to connect either an upper or lower partial denture to a tooth or a crown, or to connect a fixed bridge to a partial  denture.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38797,  July 25,

2001]

##### § 872.3200   Resin tooth bonding agent.

- Identification. A resin tooth bond- ing agent is a device material, such as methylmethacrylate, intended to be painted on the interior of a prepared cavity of a tooth to improve retention of a restoration, such as a filling.

- Classification. Class II.

##### § 872.3220   Facebow.

- Identification. A facebow is a de- vice intended for use in denture fab- rication to determine the spatial rela- tionship between the upper  and  lower jaws. This  determination  is  intended for use in placing denture casts accu- rately into an articulator (§ 872.3150) and thereby aiding correct  placement of  artificial  teeth  into  a  denture  base.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, the device is ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this   chapter,   with   the   exceptions   of

§ 820.180,   with   respect   to   general   re- quirements     concerning     records,     and

#### 21 CFR Ch. I (4–1–17 Edition)

§ 820.198,    with    respect    to    complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38797,  July 25,

2001]

##### § 872.3240   Dental bur.

- Identification. A dental bur is a ro- tary cutting device made from carbon steel or tungsten carbide intended to cut hard structures in the mouth, such as teeth or bone. It is also intended to cut hard metals, plastics, porcelains, and similar materials intended for use in the fabrication of dental devices.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3250   Calcium   hydroxide   cavity liner.

- Identification. A calcium hydrox- ide cavity liner is a device material in- tended to be applied to the interior of a prepared cavity before insertion of restorative material, such as amalgam, to protect the pulp of a tooth.

- Classification. Class II.

##### § 872.3260   Cavity varnish.

- Identification. Cavity varnish is a device that consists of a compound in- tended to coat a prepared cavity of a tooth before insertion of restorative materials. The device is intended to prevent penetration of restorative ma- terials, such as amalgam, into the dentinal tissue.

- Classification. Class II.

##### § 872.3275   Dental cement.

- Zinc oxide-eugenol—(1) Identifica- tion. Zinc oxide-eugenol is a device composed of zinc oxide-eugenol in- tended to serve as a temporary tooth filling or as a base cement to affix a temporary tooth filling, to affix dental devices such as crowns or bridges, or to be applied to a tooth to protect the tooth pulp.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 872.3360

subpart  E  of  part  807  of  this  chapter subject to § 872.9.

- Dental cement other than zinc oxide-eugenol—(1) Identification. Dental cement other than zinc oxide-eugenol is a device composed of various mate- rials other than zinc oxide-eugenol in- tended to serve as a temporary tooth filling or as a base cement to affix a temporary tooth filling, to affix dental devices such as crowns or bridges, or to be applied to a tooth to protect the tooth pulp.

(2) Classification. Class II.

[52 FR 30097, Aug. 12, 1987, as amended at 65

FR 2314, Jan. 14, 2000]

##### § 872.3285   Preformed clasp.

- Identification. A preformed clasp or a preformed wire clasp is a prefab- ricated device made of austenitic al- loys or alloys containing 75 percent or greater gold and metals of the plat- inum group intended to be incor- porated into a dental  appliance,  such as a partial denture, to help  stabilize the appliance in the patient’s mouth by fastening the appliance to an adjacent tooth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3300   Hydrophilic   resin   coating for dentures.

- Identification. A hydrophilic resin coating for dentures is a device that consists of a water-retaining polymer that is intended to be applied to the base of a denture before the denture is inserted into the patient’s mouth to improve denture retention and com- fort.

- Classification. Class II.

##### § 872.3310   Coating  material  for  resin fillings.

- Identification. A coating material for resin fillings is a device intended to be applied to the surface of a restora- tive resin  dental  filling  to  attain  a smooth, glaze-like finish on the surface of the filling.

- Classification. Class II.

##### § 872.3330   Preformed crown.

  - Identification. A preformed crown is a prefabricated device made of plas- tic or austenitic alloys or alloys con- taining 75 percent or greater gold and metals of the platinum group intended to be affixed temporarily to a tooth after removal of, or breakage of, the natural crown (that portion of  the tooth that  normally  protrudes  above the gums). It is intended for use as a functional restoration until a perma- nent crown is constructed. The device also may be intended for use as a func- tional restoration for a badly decayed deciduous (baby) tooth until the adult tooth  erupts.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3350   Gold or stainless steel cusp.

- Identification. A gold or stainless steel cusp is a prefabricated device made of austenitic alloys or alloys con- taining 75 percent or greater gold and metals of the platinum group or stain- less steel intended to provide a perma- nent cusp (a projection on the chewing surface of a tooth) to achieve occlusal harmony (a proper bite) between the teeth and a removable denture.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3360   Preformed cusp.

- Identification. A performed cusp is a prefabricated device made  of  plastic or austenitic alloys or  alloys  con- taining 75 percent or greater gold and metals of the platinum group intended to be used as a temporary cusp (a pro- jection on the chewing surface of  a tooth) to achieve occlusal harmony (a proper bite) before permanent restora- tion of a tooth.

#### § 872.3400

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3400 Karaya and sodium borate with or without acacia denture ad- hesive.

- Identification. A karaya and so- dium borate with or without acacia denture adhesive is a device composed of karaya and sodium borate with or without acacia intended to be applied to the base of a denture before the den- ture is inserted into patient’s mouth to improve denture retention and com- fort.

- Classification. (1) Class I (general controls) if the device  contains  less than 12 percent by weight of sodium borate. The class I device is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 872.9.

(2) Class III if the device contains 12 percent or more by weight of sodium borate.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any karaya and  sodium  borate  with or  without  acacia  denture   adhesive that was in commercial distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to a karaya and sodium borate with or without aca- cia denture adhesive that was in com- mercial distribution before  May  28, 1976. Any other karaya and sodium bo- rate with or without acacia denture ad- hesive shall  have  an  approved  PMA  or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50706,  Sept. 27,  1996;  65  FR 2315,  Jan. 14,

2000]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 872.3410 Ethylene	oxide homopolymer	and/or carboxymethylcellulose sodium denture adhesive.

- Identification. An ethylene oxide homopolymer	and/or carboxymethylcellulose sodium den- ture adhesive is a device containing ethylene oxide homopolymer and/or carboxymethylcellulose sodium in- tended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to improve den- ture retention and comfort.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3420 Carboxymethylcellulose so- dium and cationic polyacrylamide polymer denture adhesive.

- Identification.	A carboxymethylcellulose sodium and cationic polyacrylamide polymer den- ture adhesive is a device composed of carboxymethylcellulose sodium and cationic polyacrylamide polymer in- tended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to improve den- ture retention and comfort.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any carboxymethylcellulose sodium and cationic polyacrylamide polymer denture adhesive that was in commer- cial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a carboxymethylcellulose so- dium and cationic polyacrylamide polymer denture adhesive that was in commercial distribution before May 28, 1976. Any other carboxymethylcellulose sodium and cationic polyacrylamide polymer denture adhesive shall have an approved PMA or a declared completed

#### Food and Drug Administration, HHS	§ 872.3500

PDP  in  effect  before  being  placed  in commercial distribution.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50707, Sept. 27, 1996]

##### § 872.3450   Ethylene	oxide homopolymer and/or karaya den- ture adhesive.

  - Identification. Ethylene oxide homopolymer and/or karaya denture adhesive is a device composed of ethyl- ene oxide homopolymer and/or karaya intended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to improve den- ture retention and comfort.

  - Classification. (1) Class I if the de- vice is made of wax-impregnated cot- ton cloth that the patient applies  to the base or inner surface of a denture before inserting the denture into the mouth. The device is intended to be discarded following 1 day’s use. The class I device is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter sub- ject to § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR  63008,  Dec.  7,  1994;  65  FR  2315,  Jan.  14,

2000]

##### § 872.3480 Polyacrylamide polymer (modified cationic) denture adhe- sive.

- Identification. A polyacrylamide polymer (modified cationic) denture adhesive is a device composed of polyacrylamide polymer (modified cat- ionic) intended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to im- prove denture retention and comfort.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any polyacrylamide polymer (modi- fied  cationic)  denture  adhesive   that was in commercial distribution  before May 28,  1976, or  that has,  on or  before December 26, 1996 been found to be sub- stantially	equivalent	to	a polyacrylamide polymer (modified cat- ionic) denture adhesive that was in commercial distribution before May 28, 1976. Any other polyacrylamide  poly- mer   (modified   cationic)   denture   adhe-

sive shall have an approved PMA or a declared completed PDP in effect be- fore being place in commercial dis- tribution.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50707, Sept. 27, 1996]

##### § 872.3490 Carboxymethylcellulose so- dium and/or polyvinylmethylether maleic acid calcium-sodium double salt denture adhesive.

- Identification.	A carboxymethylcellulose sodium and/or polyvinylmethylether maleic acid cal- cium-sodium double salt denture adhe- sive is a device composed of carboxymethylcellulose sodium and/or polyvinylmethylether maleic acid cal- cium-sodium double salt intended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to improve denture retention and comfort.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3500 Polyvinylmethylether  ma- leic anhydride (PVM-MA), acid co- polymer,	and

carboxymethylcellulose	sodium (NACMC) denture adhesive.

- Identification.

Polyvinylmethylether maleic anhy- dride (PVM-MA), acid copolymer, and carboxymethylcellulose	sodium (NACMC) denture adhesive is a device composed of polyvinylmethylether ma- leic anhydride, acid copolymer, and carboxymethylcellulose sodium in- tended to be applied to the base of a denture before the denture is inserted in a patient’s mouth to improve den- ture retention and comfort.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any polyvinylmethylether maleic anhydride (PVM-MA), acid copolymer, and carboxymethylcellulose sodium (NACMC) denture adhesive that was in

#### § 872.3520

commercial distribution before May 28, 1976, or that has, on or before Decem- ber 26, 1996 been found to be substan- tially	equivalent	to		a polyvinylmethylether maleic  anhy- dride (PVM-MA), acid copolymer, and carboxymethylcellulose		sodium (NACMC) denture adhesive that was in commercial distribution before May 28, 1976. Any other polyvinylmethylether maleic anhydride (PVM-MA), acid co- polymer, and carboxymethylcellulose sodium (NACMC) denture  adhesive shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50707, Sept. 27, 1996]

##### § 872.3520   OTC denture cleanser.

  - Identification. An OTC denture cleanser is a device that consists of material in the form of a powder, tab- let, or paste that is intended to remove debris from removable prosthetic den- tal appliances, such as bridges or den- tures. The dental appliance is removed from the patient’s mouth when the ap- pliance is cleaned.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3530   Mechanical denture cleaner.

- Identification. A  mechanical den- ture cleaner is a device, usually AC- powered, that consists of a container for mechanically agitating a denture cleansing solution. The device is in- tended to clean a denture by submer- sion in the agitating cleansing solution in the container.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 872.3540   OTC   denture   cushion   or pad.

- Identification. An OTC denture cushion or pad is a prefabricated or noncustom made disposable device that is intended to improve the fit of a loose or uncomfortable denture, and may be available for purchase over-the- counter.

- Classification. (1) Class I if the de- vice is made of wax-impregnated cot- ton cloth that the patient applies  to the base or inner surface of a denture before inserting the denture into the mouth. The device is intended to be discarded following 1 day’s use. The class I device is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter sub- ject to § 872.9.

- Class II if the OTC denture cush- ion or pad is made of a material other than wax-impregnated cotton cloth or if the intended use of the device differs from that described in paragraph (b)(1) of this section. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical—Devices  Part  I:  Evaluation and Testing,’ ’’ and

  - ‘‘OTC Denture Reliners, Repair Kits, and Partially Fabricated Denture Kits.’’

[52 FR 30097, Aug. 12, 1987, as amended at 65

FR 2315, 2000; 65 FR 17144, Mar. 31, 2000]

##### § 872.3560   OTC denture reliner.

- Identification. An OTC denture reliner is a device consisting of a mate- rial such as plastic resin that is in- tended to be applied as a permanent coating or lining on the base or tissue- contacting surface of a denture. The device is intended to replace a worn denture lining and may be available for purchase over the counter.

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and Testing,’ ’’ and

#### Food and Drug Administration, HHS	§ 872.3640

  - ‘‘OTC Denture Reliners, Repair Kits, and Partially Fabricated Denture Kits.’’

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50707, Sept. 27, 1996; 65 FR 17144, Mar. 31,

2000]

##### § 872.3570   OTC denture repair kit.

    - Identification. An OTC denture re- pair kit is a device consisting of a ma- terial, such as a resin monomer system of powder and liquid glues, that is in- tended to be applied permanently to a denture to mend cracks or breaks. The device may be available for purchase over-the counter.

    - Classification. Class II. The special controls for this device are FDA’s:

      - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and Testing,’ ’’ and

      - ‘‘OTC Denture Reliners, Repair Kits, and Partially Fabricated Denture Kits.’’

[52 FR 30097, Aug. 12, 1987, as amended at 65

FR 17144, Mar. 31, 2000]

##### § 872.3580   Preformed	gold	denture tooth.

        - Identification. A preformed gold denture tooth is a device composed of austenitic  alloys  or  alloys  containing 75 percent or greater gold and metals of the platinum group intended for use as a tooth or a portion of a tooth in a fixed or removable partial denture.

        - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3590   Preformed   plastic   denture tooth.

- Identification. A preformed plastic denture tooth is a prefabricated device, composed of materials such as methyl methacrylate, that is intended for use as a tooth in a denture.

- Classification. Class II.

##### § 872.3600   Partially fabricated denture kit.

- Identification. A partially fab- ricated denture kit is a device com- posed of connected preformed teeth that is intended for use in construction of a denture. A denture base is con- structed using the patient’s mouth as a mold, by partially polymerizing the resin denture base materials while the materials are in contact with the oral tissues. After the denture base is con- structed, the  connected preformed teeth are chemically bonded to the base.

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and Testing,’ ’’ and

  - ‘‘OTC Denture Reliners, Repair Kits, and Partially Fabricated Denture Kits.’’

[52 FR 30097, Aug. 12, 1987, as amended at 65

FR 17144, Mar. 31, 2000]

##### § 872.3630   Endosseous  dental  implant abutment.

    - Identification. An endosseous den- tal implant abutment is a premanufactured prosthetic component directly connected to the endosseous dental implant and is intended for use as an aid in prosthetic rehabilitation.

    - Classification. Class II (special controls). The guidance document enti- tled ‘‘Class II Special  Controls  Guid- ance	Document:	Root-Form Endosseous Dental Implants and Endosseous Dental Implant Abut- ments’’ will serve as the special con- trol. (See § 872.1(e) for  the  availability of this guidance document.)

[69 FR 26304, May 12, 2004]

##### § 872.3640   Endosseous dental implant.

- Identification. An endosseous den- tal implant is a prescription device made of a material such as titanium or titanium alloy that is intended to be surgically placed in the bone of the upper or lower jaw arches to provide support for prosthetic devices, such as artificial teeth, in order to restore a patient’s chewing function.

- Classification. (1) Class II (special controls).   The   device   is   classified   as

#### § 872.3645

class II if  it is  a root-form endosseous dental implant. The root-form endosseous dental implant is charac- terized by four geometrically distinct types: Basket, screw,  solid  cylinder, and hollow cylinder. The guidance doc- ument entitled ‘‘Class II Special Con- trols Guidance Document: Root-Form Endosseous Dental Implants and Endosseous Dental Implant Abut- ments’’ will serve as the special con- trol. (See § 872.1(e) for  the  availability of this guidance document.)

- Classification. Class II (special con- trols). The device is classified as class II if it is a blade-form endosseous den- tal implant. The special controls for this device are:

  - The design characteristics of the device must ensure that the geometry and material composition are con- sistent with the intended use;

  - Mechanical performance (fatigue) testing under simulated physiological conditions to demonstrate maximum load (endurance limit) when the device is subjected to compressive and shear loads;

  - Corrosion testing under simu- lated physiological conditions to dem- onstrate corrosion potential of each metal or alloy, couple potential for an assembled dissimilar metal implant system, and corrosion rate for an as- sembled dissimilar metal implant sys- tem;

  - The device must be demonstrated to be biocompatible;

  - Sterility testing must dem- onstrate the sterility of the device;

  - Performance testing to evaluate the compatibility of the device in a magnetic resonance (MR) environment;

  - Labeling must include a clear description of the technological fea- tures, how the device should be used in patients, detailed surgical protocol and restoration procedures, relevant pre- cautions and warnings based on the clinical use of the device, and quali- fications and training requirements for device users including technicians and clinicians;

  - Patient labeling must contain a description of how the device works, how the device is placed, how the pa- tient needs to care for the implant, possible adverse events and how to re- port any complications; and

#### 21 CFR Ch. I (4–1–17 Edition)

  - Documented clinical experience must demonstrate safe  and  effective use and capture any adverse events ob- served during clinical use.

[69 FR 26304, May 12, 2004, as amended at 79

FR 34625, June 18, 2014]

##### § 872.3645   Subperiosteal implant mate- rial.

- Identification. Subperiosteal im- plant material is a device composed of titanium or cobalt chrome molyb- denum intended to construct custom prosthetic devices which are surgically implanted into the lower or upper jaw between the periosteum (connective tissue covering the bone) and sup- porting bony structures. The device is intended to provide support for pros- theses, such as dentures.

- Classification. Class II.

##### § 872.3660   Impression material.

- Identification. Impression material is a device composed of materials such as alginate or polysulfide intended to be placed on a preformed impression tray and used to reproduce the struc- ture of a patient’s teeth and gums. The device is intended to provide models for study and for production of restora- tive prosthetic devices, such as gold in- lays and dentures.

- Classification. Class II (Special Controls).

[52 FR 30097, Aug. 12, 1987, as amended at 68

FR 19738, Apr. 22, 2003]

##### § 872.3661   Optical  Impression Systems for CAD/CAM.

- Identification. An optical impres- sion system for computer assisted de- sign and  manufacturing  (CAD/CAM)  is a device used to record the topo- graphical characteristics of teeth, den- tal impressions, or stone models by analog or digital methods for use in the computer-assisted design and manufac- turing of dental restorative prosthetic devices. Such systems may consist of a camera, scanner, or equivalent type of sensor and a computer with software.

- Classification. Class II (Special Controls). The device is exempt from the premarket notification procedures in subpart E of part 807 of the chapter subject to the limitations in § 872.9. The special control for these devices is the

#### Food and Drug Administration, HHS	§ 872.3730

FDA guidance document  entitled ‘‘Class II Special Controls Guidance Document: Optical Impression Systems for Computer Assisted Design and Man- ufacturing (CAD/CAM) of Dental Res- torations; Guidance for Industry and FDA.’’ For the availability of this guidance document, see § 872.1(e).

[68 FR 19738, Apr. 22, 2003]

##### § 872.3670   Resin impression tray mate- rial.

  - Identification. Resin impression tray material is a device intended for use in a two-step dental mold fabri- cating process. The device consists of a resin material, such as methyl meth- acrylate, and is used to form a custom impression tray for use in cases in which a preformed impression tray is not suitable, such as the fabrication of crowns, bridges, or full dentures. A pre- liminary plaster or stone model of the patient’s teeth and gums is made. The resin impression tray material is ap- plied to this preliminary study model to form a custom tray. This tray is then filled with impression material and inserted into the patient’s mouth to make an impression, from which a final, more precise, model of the pa- tient’s mouth is cast.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3680   Polytetrafluoroethylene (PTFE) vitreous carbon materials.

- Identification. Polytetrafluoro- ethylene (PTFE) vitreous carbon mate- rial is a device composed of polytetra- fluoroethylene (PTFE) vitreous carbon intended for use in maxillofacial alve- olar ridge augmentation (building up the upper or lower jaw area that con-

tains the sockets in which teeth are rooted) or  intended  to coat metal sur- gical implants to be placed in the alveoli (sockets  in  which the  teeth  are rooted) or the temporomandibular joints (the joint between the upper and lower jaws).

- Classification. Class II.

[52 FR 30097, Aug. 12, 1987; 52 FR 34456, Sept.

11, 1987]

##### § 872.3690   Tooth shade resin material.

- Identification. Tooth shade resin material is a device composed of mate- rials such as bisphenol-A  glycidyl methacrylate (Bis-GMA) intended to restore carious lesions or structural de- fects in teeth.

- Classification. Class II.

##### § 872.3710   Base metal alloy.

- Identification. A base  metal  alloy is a device composed primarily of base metals, such as nickel, chromium, or cobalt, that is intended for use in fab- rication of cast or porcelain-fused-to- metal crown and bridge restorations.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Dental Base Metal Alloys.’’ The device is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter  subject  to  the  limitations  in

§ 872.9.  See  § 872.1(e)  for  availability  of guidance    information.

[69 FR 51766, Aug. 23, 2004]

##### § 872.3730   Pantograph.

- Identification. A pantograph is a device intended to be attached to a pa- tient’s head to duplicate lower jaw movements to aid in construction of restorative and prosthetic dental de- vices. A marking pen is attached to the lower jaw component of the device and, as the patient’s mouth opens, the pen records on graph paper the angle be- tween the upper and the lower jaw.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from  the  current  good  manufacturing

#### § 872.3740

practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[52 FR 30097, Aug. 12, 1987, as amended at 66

FR 38798, July 25, 2001]

##### § 872.3740   Retentive and splinting pin.

  - Identification. A retentive and splinting pin is a device made of aus- tenitic alloys or alloys containing 75 percent or greater gold and metals  of the platinum group intended to be placed permanently in a tooth to pro- vide retention and stabilization for a restoration, such as a crown, or to join two or more teeth together.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9

[52 FR 30097, Aug. 12, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38798, July 25,

2001]

##### § 872.3750   Bracket adhesive resin and tooth conditioner.

- Identification. A bracket adhesive resin and tooth conditioner is a device composed of an adhesive compound, such as polymethylmethacrylate, in- tended to cement an orthodontic bracket to a tooth surface.

- Classification. Class II.

##### § 872.3760   Denture relining, repairing, or rebasing resin.

- Identification. A denture relining, repairing, or rebasing resin is a device composed of materials such as methylmethacrylate, intended to re- line a denture surface that contacts tissue, to repair a fractured denture, or to form a new denture base. This device is not available for over-the-counter (OTC) use.

- Classification. Class II.

##### § 872.3765   Pit and fissure sealant and conditioner.

- Identification. A pit and fissure sealant and conditioner is a device composed of resin, such as polymethylmethacrylate, intended for use primarily in young children to seal pit  and  fissure  depressions  (faults  in

#### 21 CFR Ch. I (4–1–17 Edition)

the  enamel)  in  the  biting  surfaces  of teeth to prevent cavities.

- Classification. Class II.

##### § 872.3770   Temporary	crown	and bridge resin.

- Identification. A temporary crown and bridge resin is a device composed of	a	material,	such	as polymethylmethacrylate, intended to make a temporary prosthesis, such as a crown or bridge, for use until a perma- nent restoration is fabricated.

- Classification. Class II.

##### § 872.3810   Root canal post.

- Identification. A root canal post is a device made of austenitic alloys or alloys containing 75 percent or greater gold and metals of the platinum group intended to  be  cemented into the root canal of a tooth to stabilize and sup- port a restoration.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38798, July 25,

2001]

##### § 872.3820   Root canal filling resin.

- Identification. A root canal filling resin is a device composed of material, such as methylmethacrylate, intended for use during endodontic therapy to fill the root canal of a tooth.

- Classification. (1) Class II if chloro- form is not used as an ingredient in the device.

(2) Class III if chloroform is used as an ingredient in the device.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any  root  canal  filling  resin  de- scribed in paragraph (b)(2) of this sec- tion that was  in  commercial  distribu- tion before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equivalent to a root canal filling resin described in para- graph (b)(2) of this section that was in commercial distribution before May 28, 1976.  Any  other  root  canal  filling  resin

#### Food and Drug Administration, HHS	§ 872.3920

shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 50707, Sept. 27, 1996]

##### § 872.3830   Endodontic paper point.

  - Identification. An endodontic paper point is a device made of paper intended for use during endodontic therapy to dry, or apply medication to, the root canal of a tooth.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38798,  July 25,

2001]

##### § 872.3840   Endodontic silver point.

- Identification. An endodontic sil- ver point is a device made of silver in- tended for use during endodontic ther- apy to fill permanently the root canal of a tooth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38798,  July 25,

2001]

##### § 872.3850   Gutta percha.

- Identification. Gutta percha is a device made from coagulated sap of certain tropical trees intended to fill the root canal of a tooth. The gutta percha is softened by heat and inserted into the root canal, where it hardens as it cools.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38798,  July 25,

2001]

##### § 872.3890   Endodontic	stabilizing splint.

- Identification. An endodontic sta- bilizing splint is a device made of a ma-

terial, such as titanium, intended to be inserted through the root canal into the upper or lower jaw bone to stabilize a tooth.

- Classification. Class II.

##### § 872.3900   Posterior	artificial	tooth with a metal insert.

- Identification. A posterior artifi- cial tooth with a metal insert is a por- celain device with an insert made of austenitic  alloys  or  alloys  containing 75 percent or greater gold and metals of the platinum group intended to replace a natural tooth. The device is attached to surrounding teeth by a bridge and is intended to provide both an improve- ment in appearance and functional oc- clusion (bite).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38798,  July 25,

2001]

##### § 872.3910   Backing  and  facing  for  an artificial tooth.

- Identification. A backing and fac- ing for an artificial tooth is a device intended for use in fabrication of a fixed or removable dental appliance, such as a crown or bridge. The backing, which is made of gold, is attached to the dental appliance and supports the tooth-colored facing, which is made of porcelain or plastic.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.3920   Porcelain tooth.

- Identification. A  porcelain  tooth  is a prefabricated device made of por- celain      powder      for      clinical       use (§ 872.6660) intended for use in construc- tion of fixed or removable prostheses, such as crowns and partial dentures.

- Classification. Class II.

#### § 872.3930

##### § 872.3930   Bone grafting material.

  - Identification. Bone grafting mate- rial is a material such as hydroxyapatite, tricalcium phosphate, polylactic and polyglycolic acids, or collagen, that is intended to fill, aug- ment, or reconstruct periodontal or bony defects of the oral and maxillo- facial region.

  - Classification. (1) Class II (special controls) for bone grafting materials that do not contain a drug that is a therapeutic biologic. The special con- trol is FDA’s ‘‘Class II Special Controls Guidance Document: Dental Bone Grafting     Material     Devices.’’     (See

§ 872.1(e)   for   the   availability   of   this guidance document.)

(2) Class III (premarket approval) for bone grafting materials that contain a drug that is a therapeutic  biologic. Bone grafting materials that contain a drug that is a  therapeutic  biologic, such as biological response modifiers, require premarket approval.

  - Date premarket approval application (PMA) or notice of product development protocol (PDP) is required. Devices de- scribed in paragraph (b)(2) of this sec- tion shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[70 FR 21949, Apr. 28, 2005]

##### § 872.3940   Total	temporomandibular joint prosthesis.

- Identification.	A	total temporomandibular joint  prosthesis  is a device that is  intended to  be im- planted in the human jaw to  replace the mandibular condyle and augment the glenoid fossa to functionally recon- struct  the  temporomandibular  joint.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before March 30, 1999, for any total temporomandibular joint prosthesis that was in commercial dis- tribution before May 28, 1976, or that has, on or before March 30, 1999, been found to be substantially equivalent to a total temporomandibular joint pros- thesis that was in commercial distribu- tion  before  May  28,  1976.  Any  other

#### 21 CFR Ch. I (4–1–17 Edition)

total temporomandibular joint pros- thesis shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[59 FR 65478, Dec. 20, 1994, as amended at 63

FR 71746, Dec. 30, 1998]

##### § 872.3950   Glenoid fossa prosthesis.

- Identification. A glenoid fossa prosthesis is a device  that  is  intended to	be	implanted	in	the temporomandibular joint to augment a glenoid fossa or to provide an articula- tion surface for the head of a man- dibular  condyle.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before March 30, 1999, for any glenoid fossa prosthesis  that  was in commercial distribution before May 28, 1976, or that has on or before March 30, 1999, been found to be substantially equivalent to a glenoid fossa prosthesis that was in commercial distribution before May 28, 1976. Any other glenoid fossa prosthesis shall have an approved PMA or a declared completed PDP in effect before being placed in commer- cial   distribution.

[59 FR 65478, Dec. 20, 1994, as amended at 63

FR 71746, Dec. 30, 1998]

##### § 872.3960   Mandibular   condyle   pros- thesis.

- Identification. A mandibular condyle prosthesis is a device that is intended to be implanted in the human jaw to replace the mandibular condyle and to articulate within a glenoid fossa.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before March 30, 1999, for any mandibular  condyle  prosthesis that was in commercial distribution before May 28, 1976, or that has, on or before March 30, 1999, been found to be substantially equivalent to a man- dibular condyle prosthesis that was in commercial distribution before May 28, 1976.   Any   other   mandibular   condyle

#### Food and Drug Administration, HHS	§ 872.4465

prosthesis shall have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[59 FR 65478, Dec. 20, 1994, as amended at 63

FR 71746,  Dec. 30,  1998;  78  FR 79310,  Dec. 30,

2013]

##### § 872.3970   Interarticular	disc	pros- thesis (interpositional implant).

  - Identification. An interarticular disc prosthesis (interpositional im- plant) is a device that is intended to be an interface between the natural ar- ticulating surface of the mandibular condyle and glenoid fossa.

  - Classification. Class III.

  - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before March 30, 1999, for any interarticular disc prosthesis (interpositional implant) that was in commercial distribution before May 28, 1976, or that has on or before March 30, 1999, been found to be substantially equivalent to an interarticular disc prosthesis (interpositional  implant) that was in commercial distribution before May 28, 1976. Any other inter- articular disc prosthesis (interposi- tional implant) shall have an approved PMA or a declared completed PDP in effect before being placed in commer- cial   distribution.

[59 FR 65478, Dec. 20, 1994, as amended at 63

FR 71746, Dec. 30, 1998]

##### § 872.3980   Endosseous  dental  implant accessories.

- Identification. Endosseous dental implant accessories are manually pow- ered devices intended to aid in the placement or removal of endosseous dental implants and abutments, pre- pare the site for placement of endosseous dental implants or abut- ments, aid in the fitting of endosseous dental implants or abutments, aid in the fabrication of dental  prosthetics, and be used as an accessory with endosseous dental implants when tis- sue contact will last less than 1 hour. These devices include drill bits, screw- drivers, countertorque devices, place- ment and removal tools, laboratory pieces used  for  fabrication of dental prosthetics,  and  trial  abutments.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[65 FR 60099, Oct. 10, 2000]

## Subpart E—Surgical Devices

##### § 872.4120   Bone	cutting	instrument and accessories.

- Identification. A bone cutting in- strument and accessories is a metal de- vice intended for use in reconstructive oral surgery to drill or cut into the upper or lower jaw and may be used to prepare bone to insert a wire, pin, or screw. The device includes the manual bone drill and wire driver, powered bone drill, rotary bone cutting hand- piece, and AC-powered bone saw.

- Classification. Class II.

##### § 872.4130   Intraoral dental drill.

- Identification. An intraoral dental drill is a rotary device intended to be attached to a dental handpiece to drill holes in teeth to secure cast or preformed pins to retain operative den- tal appliances.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.4200   Dental handpiece and acces- sories.

- Identification. A dental handpiece and accessories is an AC-powered, water-powered, air-powered, or belt- driven, hand-held device that may in- clude a foot controller for regulation of speed and direction of rotation or a contra-angle attachment for difficult to reach areas intended to prepare den- tal cavities for restorations, such as fillings, and for cleaning teeth.

- Classification. Class I.

[55 FR 48439, Nov. 20, 1990]

##### § 872.4465   Gas-powered jet injector.

- Identification. A gas-powered jet injector is a syringe device intended to

#### § 872.4475

administer a local anesthetic. The sy- ringe is powered by a cartridge con- taining pressurized carbon dioxide which provides  the  pressure to force the anesthetic out of the syringe.

- Classification. Class II.

##### § 872.4475   Spring-powered jet injector.

- Identification. A spring-powered jet injector is a syringe device intended to administer a local anesthetic. The syringe is powered by a spring mecha- nism which provides the pressure to force the anesthetic out of the syringe.

- Classification. Class II.

##### § 872.4535   Dental diamond instrument.

- Identification. A dental diamond instrument is an abrasive device in- tended to smooth tooth surfaces during the fitting of crowns or bridges. The device consists of a shaft which is in- serted into a handpiece and a head which has diamond chips imbedded into it. Rotation of the diamond instrument provides an abrasive action when it contacts a tooth.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008, Dec. 7, 1994]

##### § 872.4565   Dental hand instrument.

- Identification. A dental hand in- strument is a hand-held device in- tended to perform various tasks in gen- eral dentistry and oral surgery proce- dures. The device includes the opera- tive burnisher, operative amalgam car- rier, operative dental amalgam carver, surgical bone chisel, operative amal- gam and foil condenser, endodontic cu- rette, operative curette, periodontic curette, surgical curette, dental sur- gical elevator, operative dental exca- vator, operative explorer surgical bone file, operative margin finishing file, periodontic file, periodontic probe, sur- gical rongeur forceps, surgical tooth extractor forceps, surgical hemostat, periodontic hoe, operative matrix contouring instrument, operative cut- ting instrument, operative margin fin- ishing periodontic knife, periodontic marker, operative pliers, endodontic root canal plugger, endodontic root canal preparer, surgical biopsy punch,

#### 21 CFR Ch. I (4–1–17 Edition)

endodontic pulp canal reamer, crown remover, periodontic scaler, collar and crown scissors, endodontic pulp canal filling material spreader, surgical osteotome chisel, endodontic broach, dental wax carver, endodontic pulp canal file, hand instrument for calculus removal, dental depth gauge instru- ment, plastic dental filling instrument, dental instrument handle, surgical tis- sue scissors, mouth mirror, ortho- dontic band driver, orthodontic band pusher, orthodontic band setter, ortho- dontic bracket aligner, orthodontic pliers, orthodontic ligature tucking in- strument, forceps, for articulation paper, forceps for dental dressing, den- tal matrix band, matrix retainer, den- tal retractor, dental retractor acces- sories, periodontic or endodontic irri- gating syringe, and restorative or im- pression material syringe.

- Classification. Class I (general con- trols). If the device is made of the same materials that were used in the device before May 28, 1976, it is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38799,  July 25,

2001]

##### § 872.4600   Intraoral  ligature  and  wire lock.

- Identification. An intraoral liga- ture and wire lock is a metal device in- tended to constrict fractured bone seg- ments in the oral cavity. The bone seg- ments are stabilized by wrapping the ligature (wire) around the fractured bone segments and locking the ends to- gether.

- Classification. Class II.

##### § 872.4620   Fiber optic dental light.

- Identification. A fiber optic dental light is a device that is a light, usually AC-powered, that consists of glass or plastic fibers which have special opti- cal properties. The device is usually at- tached to a dental handpiece and is in- tended to illuminate a patient’s oral structures.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 872.4920

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.4630   Dental operating light.

  - Identification. A dental operating light, including the surgical headlight, is an AC-powered device intended to il- luminate oral structures and operating areas.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR 1121,  Jan. 16,  1996;  66  FR 38799,  July 25,

2001]

##### § 872.4730   Dental injecting needle.

- Identification. A dental injecting needle is a slender, hollow metal device with a sharp point intended to be at- tached to a syringe to inject local anes- thetics and other drugs.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63008,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.4760   Bone plate.

- Identification. A bone plate is a metal device intended to stabilize frac- tured bone structures in the oral cav- ity. The bone segments are attached to the plate with screws to prevent move- ment of the segments.

- Classification. Class II.

##### § 872.4770   Temporary	mandibular condyle reconstruction plate.

- Identification. A temporary man- dibular condyle reconstruction plate is a device that is intended to stabilize mandibular bone and provide for tem- porary reconstruction of the man- dibular condyle until permanent recon- struction is completed in patients who have undergone resective surgical pro- cedures requiring removal of the man- dibular condyle and mandibular bone. This device is not intended for treat-

ment  of  temporomandibular  joint  dis- orders.

- Classification. Class II (special controls). The special controls for this device is FDA’s guideline entitled ‘‘Temporary Mandibular Condyle Re- construction Plate Class II Special Controls Guideline.’’ See § 872.1(e) for the availability of this guidance docu- ment.

[78 FR 79310, Dec. 30, 2013]

##### § 872.4840   Rotary scaler.

- Identification. A rotary scaler is an abrasive device intended to be at- tached to a powered handpiece to re- move calculus deposits from teeth dur- ing dental cleaning and periodontal (gum) therapy.

- Classification. Class II.

##### § 872.4850   Ultrasonic scaler.

- Identification. An ultrasonic scaler is a device intended for use during den- tal cleaning and periodontal (gum) therapy to remove calculus deposits from teeth by application of an ultra- sonic vibrating scaler tip to the teeth.

- Classification. Class II.

##### § 872.4880   Intraosseous  fixation  screw or wire.

- Identification. An intraosseous fix- ation screw or wire is a metal device intended to be inserted into fractured jaw bone segments to prevent their movement.

- Classification. Class II.

##### § 872.4920   Dental  electrosurgical  unit and accessories.

- Identification.       A	dental electrosurgical unit and accessories is an AC-powered device consisting of a controlled power source and a set of cutting and coagulating electrodes. This device is intended to cut or re- move soft tissue or to control bleeding during surgical procedures in the oral cavity. An electrical current passes through the tip of the electrode  into the tissue and, depending upon the op- erating mode selected, cuts through soft tissue or coagulates the tissue.

- Classification. Class II.

#### § 872.5410

Subpart F—Therapeutic Devices

##### § 872.5410   Orthodontic  appliance  and accessories.

  - Identification. An orthodontic ap- pliance and accessories is a device in- tended for use in orthodontic treat- ment. The device is affixed to a tooth so that pressure can be exerted on the teeth. This device includes the preformed orthodontic band, ortho- dontic band material, orthodontic elas- tic band, orthodontic metal bracket, orthodontic wire clamp, preformed or- thodontic space maintainer, ortho- dontic expansion screw retainer, ortho- dontic spring, orthodontic tube, and or- thodontic wire.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.5470   Orthodontic plastic bracket.

- Identification. An orthodontic plastic bracket is a plastic device in- tended to be bonded to a tooth to apply pressure to a tooth from a flexible or- thodontic wire to alter its position.

- Classification. Class II.

##### § 872.5500   Extraoral orthodontic head- gear.

- Identification. An extraoral ortho- dontic headgear is a device intended for use with an orthodontic appliance to exert pressure on the teeth from outside the mouth. The headgear has a strap intended to wrap around the pa- tient’s neck or head and an inner bow portion intended to be fastened to the orthodontic appliance in the patient’s mouth.

- Classification. Class II.

##### § 872.5525   Preformed tooth positioner.

- Identification. A preformed tooth positioner is a plastic device that is an impression of a perfected bite intended to prevent a patient’s teeth from shift- ing position or to move teeth to a final position after orthodontic appliances (braces) have been removed. The pa- tient bites down on the device for sev- eral hours a day to force the teeth into

#### 21 CFR Ch. I (4–1–17 Edition)

a  final  position  or  to  maintain  the teeth in their corrected position.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.5550   Teething ring.

(a) Identification. A teething ring is a divice intended for use by infants for medical purposes to soothe gums dur- ing the teething process.

(b)(1) Classification. Class I if the teething ring does not contain a fluid, such as water. The device is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter.

- Class II if the teething ring con- tains a fluid, such as water.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009, Dec. 7, 1994]

##### § 872.5560   Electrical	salivary	stimu- latory system.

  - Identification. An electrical sali- vary stimulatory system is a prescrip- tion intraoral device that  is  intended to electrically stimulate a relative in- crease in saliva production.

  - Classification—Class II (special controls). The special controls for this device are:

- The design characteristics of the device must ensure that the device de- sign, material composition, and elec- trical output characteristics are con- sistent with the intended use;

- Any element of the device that contacts the patient must be dem- onstrated to be biocompatible;

- Appropriate analysis and/or test- ing must validate electromagnetic compatibility and electrical safety, in- cluding the safety of any battery used in the device;

- Software validation, verification, and hazard testing must be performed; and

- Documented clinical experience must demonstrate safe  and  effective use for stimulating saliva  production by addressing the risks of damage to intraoral    tissue    and    of    ineffective

#### Food and Drug Administration, HHS	§ 872.6050

treatment and must capture any ad- verse events observed during clinical use.

[80 FR 72586, Nov. 22, 2015]

##### § 872.5570 Intraoral devices for snor- ing and intraoral devices for snor- ing and obstructive sleep apnea.

  - Identification. Intraoral devices for snoring and intraoral devices for snor- ing and obstructive sleep apnea are de- vices that are worn during sleep to re- duce the incidence of snoring and to treat obstructive sleep apnea. The de- vices are designed to increase the pa- tency of the airway and to decrease air turbulence and  airway  obstruction. The classification includes palatal lift- ing devices, tongue retaining devices, and mandibular repositioning devices.

  - Classification. Class II (special controls). The special control for these devices is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Intraoral Devices for Snoring and/or Obstructive Sleep Apnea; Guidance for  Industry  and FDA.’’

[67 FR 68512, Nov. 12, 2002]

##### § 872.5580 Oral rinse to reduce the ad- hesion of dental plaque.

- Identification. The device is as- signed the generic name oral rinse to reduce the adhesion of dental plaque and is identified as a device intended to reduce the presence of bacterial plaque on teeth and oral mucosal surfaces by physical means. The device type in- cludes those devices that act by reduc- ing the attachment and inhibiting the growth of bacterial plaque.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Oral Rinse to Reduce the Adhesion of Dental Plaque.’’ See § 872.1(e) for the availability  of  this  guidance  document.

[70 FR 55028, Sept. 20, 2005]

## Subpart G—Miscellaneous Devices

##### § 872.6010 Abrasive device and acces- sories.

- Identification. An abrasive device and accessories is a device constructed

of various abrasives, such as diamond chips, that are glued to shellac-based paper. The device is intended to re- move excessive restorative materials, such as gold, and to smooth rough sur- faces from oral restorations, such as crowns. The device is attached to a shank that is held by a handpiece. The device includes the abrasive  disk, guard for an abrasive disk, abrasive point, polishing agent strip, and polishing wheel.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38799,  July 25,

2001]

##### § 872.6030   Oral	cavity	abrasive polishing agent.

- Identification. An oral cavity abra- sive polishing agent is a device in paste or powder form that contains an abra- sive material, such as silica pumice, in- tended to remove debris from the teeth. The abrasive polish is applied to the teeth by a handpiece attachment (prophylaxis cup).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.6050   Saliva absorber.

- Identification. A saliva absorber is a device made of paper or cotton in- tended to absorb moisture from the oral cavity during dental procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject  to  the  limitations  in  § 872.9.  If

§ 872.6070

the device is not labeled or otherwise represented  as  sterile,  it  is   exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38799,  July 25,

2001]

##### § 872.6070   Ultraviolet activator for po- lymerization.

  - Identification. An ultraviolet acti- vator for polymerization is a device that produces ultraviolet radiation in- tended to polymerize (set) resinous dental pit and fissure sealants or re- storative materials by transmission of light through a rod.

  - Classification. Class II.

##### § 872.6080   Airbrush.

- Identification. An airbrush is an AC-powered device intended for use in conjunction with articulation paper. The device uses air-driven particles to roughen the surfaces of dental restora- tions. Uneven areas of the restorations are then identified by use of articula- tion paper.

- Classification. Class II. The special control for this device is International Electrotechnical Commission’s IEC 60601–1–AM2 (1995–03), Amendment 2, ‘‘Medical     Electrical     Equipment—Part 1:  General  Requirements  for  Safety.’’

[52 FR 30097, Aug. 12, 1987; 52 FR 49250, Dec.

30, 1987, as amended at 71 FR 17144, Mar. 31,

2006]

##### § 872.6100   Anesthetic warmer.

- Identification. An anesthetic warmer is an AC-powered device into which tubes containing anesthetic so- lution are intended to be placed to warm them prior to administration of the anesthetic.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 60

FR 38900,  July 28,  1995;  66 FR 38799, July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 872.6140   Articulation paper.

- Identification. Articulation paper is a device composed of paper coated with an ink dye intended to be placed between the patient’s upper and lower teeth when the teeth are in the bite po- sition to locate uneven or high areas.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.6200   Base plate shellac.

- Identification. Base plant shellac is a device composed of shellac intended to rebuild the occlusal rim of full or partial dentures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13830,  Apr. 5,  1989;  66  FR 38799,  July 25,

2001]

##### § 872.6250   Dental    chair    and    acces- sories.

- Identification. A dental chair and accessories is a device, usually AC- powered, in which a patient sits. The device is intended to properly position a patient to perform dental procedures. A dental operative unit may be at- tached.

#### Food and Drug Administration, HHS	§ 872.6510

- Classification. Class I. The dental chair without the operative unit device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63009, Dec. 7, 1994]

##### § 872.6290   Prophylaxis cup.

- Identification. A prophylaxis cup is a device made of rubber intended to be held by a dental handpiece and used to apply polishing agents during prophy- laxis (cleaning). The dental handpiece spins the rubber cup holding the polishing agent and the user applies it to the teeth to remove debris.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38799,  July 25,

2001]

##### § 872.6300   Rubber	dam	and	acces- sories.

- Identification. A rubber dam and accessories is a device composed of a thin sheet of latex with a hole in the center intended to isolate a tooth from fluids in the mouth during dental pro- cedures, such as filling a cavity prepa- ration. The  device  is  stretched  around a tooth by inserting a tooth through a hole in the center. The device  includes the rubber dam, rubber dam clamp, rubber dam frame, and forceps for a rubber dam clamp. This classification does not  include  devices  intended  for use in preventing transmission of sexu- ally transmitted diseases  through  oral sex; those devices are classified as condoms in § 884.5300 of this chapter.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 872.9. If the device is not la-

beled or otherwise represented as ster- ile, it is exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,   with   respect   to   complaint files.

[65 FR 2315, Jan. 14, 2000]

##### § 872.6350   Ultraviolet detector.

- Identification. An ultraviolet de- tector is a device intended to provide a source of ultraviolet light  which  is used to identify  otherwise invisible material, such as dental plaque, present in or on teeth.

- Classification. Class II.

##### § 872.6390   Dental floss.

- Identification. Dental floss is a string-like device made of cotton or other fibers intended to remove plaque and food particles from between the teeth to reduce tooth decay. The fibers of the device may be coated with wax for easier use.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 872.9.

[52 FR 30097, Aug. 12, 1987, as amended at 61

FR  1121,  Jan.  16,  1996;  65  FR  2315,  Jan.  14,

2000]

##### § 872.6475   Heat  source  for  bleaching teeth.

- Identification. A heat source for bleaching teeth is an AC-powered de- vice that consists of a light or an elec- tric heater intended to apply heat to a tooth after it is treated with a bleach- ing agent.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38799,  July 25,

2001]

##### § 872.6510   Oral irrigation unit.

- Identification. An oral irrigation unit is an AC-powered device intended

#### § 872.6570

to provide a pressurized stream of water to remove food particles from be- tween the teeth and promote good peri- odontal (gum) condition.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38800,  July 25,

2001]

##### § 872.6570   Impression tube.

- Identification. An impression tube is a device consisting of a hollow cop- per tube intended to take an impres- sion of a single tooth. The hollow tube is filled with impression material. One end of the tube is sealed with a soft- ened material, such as wax, the re- maining end is slipped over the tooth to make the impression.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38800,  July 25,

2001]

##### § 872.6640   Dental  operative  unit  and accessories.

- Identification. A dental operative unit and accessories is an AC-powered device that is intended to supply power to and serve as a base for other dental devices, such as a dental handpiece, a dental operating light, an air or water syringe unit, and oral  cavity evacuator, a  suction  operative unit, and other dental devices and acces- sories. The device may be attached to a dental chair.

- Classification. Class I (general con- trols). Except for dental operative unit, accessories are exempt from premarket notification procedures in subpart E of

#### 21 CFR Ch. I (4–1–17 Edition)

part   807   of   this   chapter   subject   to

§ 872.9.

[55 FR 48439, Nov. 20, 1990, as amended at 59

FR  63009,  Dec.  7,  1994;  65  FR  2315,  Jan.  14,

2000]

##### § 872.6650   Massaging  pick  or  tip  for oral hygiene.

- Identification. A massaging pick or tip for oral hygiene is a rigid, pointed device intended to be used manually to stimulate and massage the gums to promote good periodontal (gum) condi- tion.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38800,  July 25,

2001]

##### § 872.6660   Porcelain  powder  for  clin- ical use.

- Identification. Porcelain powder for clinical use is a device consisting of a mixture of kaolin, felspar, quartz, or other substances intended for use in the production of artificial teeth in fixed or removable dentures, of jacket crowns, facings, and veneers. The de- vice is used in prosthetic dentistry by heating the powder mixture to a high temperature in an oven to produce a hard prosthesis with a glass-like finish.

- Classification. Class II.

##### § 872.6670   Silicate protector.

- Identification. A silicate protector is a device made of silicone intended to be applied with an absorbent tipped ap- plicator to the surface of a new res- toration to exclude temporarily fluids from its surface.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject  to  the  limitations  in  § 872.9.  If

#### Food and Drug Administration, HHS	§ 872.6865

the device is not labeled or otherwise represented  as  sterile,  it  is   exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38800,  July 25,

2001]

##### § 872.6710   Boiling water sterilizer.

  - Identification. A boiling water sterilizer is an AC-powered device that consists of a container for boiling water. The device is intended to steri- lize dental and surgical instruments by submersion in the boiling water in the container.

  - Classification. Class I (general con- trols).

[55 FR 48439, Nov. 20, 1990, as amended at 66

FR 46952, Sept. 10, 2001]

##### § 872.6730   Endodontic  dry  heat  steri- lizer.

- Identification. An endodontic dry heat sterilizer is a device intended to sterilize endodontic and other dental instruments by the application of dry heat. The heat is supplied  through glass beads which have been heated by electricity.

- Classification. Class III.

- Date premarket approval application (PMA) or notice of completion of product development protocol (PDP) is required. A PMA or notice of completion of a PDP is required to be filed with the  Food and Drug Administration on or before April 21, 1997, for any endodontic dry heat sterilizer that was in commercial distribution before May 28, 1976, or that has on or before April 21, 1997, been found to be substantially equivalent to the endodontic dry heat sterilizer that was in commercial distribution before May 28, 1976. Any other endodontic dry heat sterilizer shall have an approved PMA or declared completed PDP in ef- fect before being placed in commercial distribution.

[52 FR 30097, Aug. 12, 1987, as amended at 62

FR 2902,  Jan. 21,  1997;  62  FR 31512,  June 10,

1997]

##### § 872.6770   Cartridge syringe.

- Identification. A cartridge syringe is a device intended to inject anes- thetic agents subcutaneously or intramuscularly. The device consists of a metal syringe body into which a dis- posable, previously filled, glass carpule (a cylindrical cartridge) containing an- esthetic is placed. After attaching a needle to the syringe body and acti- vating the carpule by partially insert- ing the plunger on the syringe, the de- vice is used to administer an injection to the patient.

- Classification. Class II.

##### § 872.6855   Manual toothbrush.

- Identification. A manual tooth- brush is a device composed of a shaft with either natural or synthetic bris- tles at one end intended to remove ad- herent plaque and food debris from the teeth to reduce tooth decay.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38800,  July 25,

2001]

##### § 872.6865   Powered toothbrush.

- Identification. A powered tooth- brush is an AC-powered or battery-pow- ered device  that  consists  of  a  handle containing a motor that provides me- chanical movement to a brush intended to be applied to the teeth. The device is intended to remove adherent plaque and food debris from the teeth to re- duce tooth decay.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 872.6870

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 872.9.

[55 FR 48440, Nov. 20, 1990, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38800,  July 25,

2001]

##### § 872.6870   Disposable fluoride tray.

  - Identification. A disposable fluo- ride tray is a device made of styrofoam intended to apply fluoride topically to the teeth. To use the tray, the patient bites down on the tray which has been filled with a fluoride solution.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 54

FR 13831,  Apr. 5,  1989;  66  FR 38800,  July 25,

2001]

##### § 872.6880   Preformed impression tray.

21 CFR Ch. I (4–1–17 Edition)

§ 872.6890   Intraoral dental wax.

- Identification. Intraoral  dental wax is a device made of wax intended to construct patterns from which cus- tom made metal dental prostheses, such as crowns and bridges, are cast. In orthodontic dentistry, the device is in- tended to make a pattern of a patient’s bite to make study models of the teeth.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 30097, Aug. 12, 1987, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38800,  July 25,

2001]

## PART 874—EAR, NOSE, AND THROAT DEVICES

##### Subpart A—General Provisions

  - Identification. A preformed im- pression tray is a metal or plastic de- vice intended to hold impression mate- rial, such as alginate, to make an im- pression of a patient’s teeth or alveolar process (bony tooth sockets) to repro- duce the structure of a patient’s teeth and gums.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 872.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice   requirements   of   the   quality

Sec.

874.1    Scope.

874.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

874.1050    Audiometer.

874.1060    Acoustic  chamber  for  audiometric testing.

874.1070    Short   increment   sensitivity   index (SISI)   adapter.

874.1080    Audiometer calibration set. 874.1090    Auditory impedance tester.

874.1100   Earphone  cushion  for  audiometric testing.

874.1120   Electronic     noise     generator     for

| system regulation in part 820 of this audiometric testing. |
| --- |
| chapter, with the exception of § 820.180, | 874.1325 | Electroglottograph. |
| with respect to general requirements | 874.1500 | Gustometer. |
| concerning records, and § 820.198, with | 874.1600 | Olfactory test device. |
| respect to complaint files. 874.1800 Air or water caloric stimulator. |
|  | 874.1820 | Surgical nerve stimulator/locator. |
| [52 FR 30097, Aug. 12, 1987, as amended at 54 | 874.1925 | Toynbee diagnostic tube. |
| FR 13832, Apr. 5, 1989; 66 FR 38800, July 25, |  |  |
| 2001] |  | Subpart C [Reserved] |

Food and Drug Administration, HHS	§ 874.1

##### Subpart D—Prosthetic Devices

874.3300    Hearing aid.

874.3305    Wireless air-conduction hearing aid. 874.3310    Hearing aid  calibrator and  analysis

system.

874.3315   Tympanic membrane contact hear- ing aid.

874.3320   Group  hearing  aid  or  group  audi- tory  trainer.

874.3330   Master hearing aid.

874.3375     Battery-powered artificial larynx. 874.3400   Tinnitus masker.

874.3430    Middle ear mold.

874.3450   Partial ossicular replacement pros- thesis.

874.3495    Total   ossicular   replacement   pros- thesis.

874.3540    Prosthesis  modification  instrument for  ossicular  replacement  surgery.

874.3620    Ear,   nose,   and   throat   synthetic polymer  material.

874.3695    Mandibular    implant    facial    pros- thesis.

874.3730  Laryngeal prosthesis (Taub design). 874.3760   Sacculotomy tack (Cody tack).

874.3820    Endolymphatic shunt.

874.3850    Endolymphatic    shunt    tube    with valve.

874.3880   Tympanostomy tube.

874.3900   Nasal dilator.

874.3930  Tympanostomy	tube	with semipermeable   membrane.

874.3950  Transcutaneous	air	conduction hearing  aid  system.

##### Subpart E—Surgical Devices

874.4100   Epistaxis balloon.  874.4140   Ear, nose, and throat bur. 874.4175    Nasopharyngeal catheter.

874.4180    Eustachian   tube   balloon   dilation system.

874.4250   Ear,  nose,  and  throat  electric  or pneumatic surgical drill.

874.4350   Ear,  nose,  and  throat  fiberoptic light source and carrier.

874.4420    Ear,  nose,  and  throat  manual  sur- gical    instrument.

874.4490    Argon  laser  for  otology,  rhinology, and   laryngology.

874.4500    Ear, nose, and throat microsurgical carbon  dioxide  laser.

874.4680   Bronchoscope (flexible or rigid) and accessories.

874.4710   Esophagoscope   (flexible   or   rigid) and  accessories.

874.4720   Mediastinoscope and accessories. 874.4750  Laryngostroboscope.

874.4760 Nasopharyngoscope	(flexible	or rigid)  and  accessories.

874.4770   Otoscope.

874.4780   Intranasal splint. 874.4800    Bone particle collector.

##### Subpart F—Therapeutic Devices

874.5220   Ear, nose, and throat drug adminis- tration  device.

874.5300    Ear,  nose,  and  throat  examination and  treatment  unit.

874.5350    Suction antichoke device. 874.5370   Tongs antichoke device. 874.5550   Powered nasal irrigator. 874.5800    External nasal splint.

874.5840    Antistammering device.

874.5900   External  upper  esophageal  sphinc- ter  compression  device.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  51  FR  40389,  Nov.  6,  1986,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 874 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 874.1   Scope.

    - This part sets forth the classifica- tion of ear, nose, and throat devices in- tended for human use that are in com- mercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provision of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - To avoid duplicative listings, an ear, nose, and throat device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed in one sub- part only.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  67

FR  67790,  Nov.  7,  2002;  78  FR  18233,  Mar.  26,

2013]

§ 874.3

##### § 874.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

      - Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA declaring completed  a  PDP  for the de- vice. If FDA promulgates a regulation under section 515(b) of the act requir- ing premarket approval for a device, section 501(f)(1)(A) of the act applies to the device.

      - Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before  the  device  is  commercially  dis-

#### 21 CFR Ch. I (4–1–17 Edition)

tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

##### § 874.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic   acid   hybridization   technology

#### Food and Drug Administration, HHS	§ 874.1080

rather   than   culture   or   immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2315, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 874.1050   Audiometer.

- Identification. An audiometer or automated	audiometer	is	an electroacoustic device that produces controlled levels of test tones and sig- nals intended for use in conducting di- agnostic hearing evaluations and as- sisting in the diagnosis of possible otologic disorders.

- Classification. Class II. Except for the otoacoustic emission device, the device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, if it is in com- pliance with American National Stand- ard  Institute  S3.6–1996,  ‘‘Specification

for   Audiometers,’’   and   subject   to   the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  64

FR 14831, Mar. 29, 1999]

##### § 874.1060   Acoustic	chamber	for audiometric testing.

- Identification. An acoustic cham- ber for audiometric testing is a room that is intended for use in conducting diagnostic hearing evaluations and that eliminates sound reflections and provides isolation from outside sounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  61

FR 1121,  Jan. 16,  1996;  66  FR 38800,  July 25,

2001]

##### § 874.1070   Short  increment  sensitivity index (SISI) adapter.

- Identification. A short increment sensitivity index (SISI) adapter is a de- vice used with an audiometer in diag- nostic hearing evaluations. A SISI adapter provides short periodic sound pulses in specific small decibel incre- ments that are intended to be super- imposed on the audiometer’s output tone frequency.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2315, Jan. 14, 2000]

##### § 874.1080   Audiometer calibration set.

- Identification. An audiometer cali- bration set is an electronic reference device that is intended to calibrate an audiometer. It measures the sound fre- quency and intensity characteristics that emanate from an audiometer ear- phone. The device consists of an acous- tic cavity of known volume, a sound level meter, a microphone with calibra- tion traceable to the National Bureau of Standards, oscillators, frequency counters, microphone amplifiers, and a recorder. The device can measure se- lected audiometer test frequencies at a given  intensity  level,  and  selectable

§ 874.1090

audiometer  attenuation  settings  at  a given test frequency.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  61

FR 1121,  Jan. 16,  1996;  66  FR 38800,  July 25,

2001]

##### § 874.1090   Auditory impedance tester.

- Identification. An auditory imped- ance tester is a device that is intended to change the air pressure in the exter- nal auditory canal and measure and graph the  mobility  characteristics  of the tympanic  membrane  to  evaluate the functional condition of the middle ear. The device is used to determine ab- normalities in the mobility of the tym- panic membrane due to stiffness, flac- cidity, or the presence of fluid in the middle ear cavity. The device is also used to measure the acoustic reflex threshold from contractions of the stapedial muscle, to monitor healing of tympanic membrane grafts or stapedectomies, or to monitor followup treatment for inflammation of the middle ear.

- Classification. Class II.

##### § 874.1100   Earphone	cushion	for audiometric testing.

- Identification. An earphone cush- ion for audiometric testing is a device that is used to cover an audiometer earphone during audiometric testing to provide an acoustic coupling (sound connection path) between the audiom- eter earphone and the patient’s ear.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[51 FR 40389, Nov. 9, 1986; 52 FR 18495, May 15,

1987, as amended at 52 FR 32111, Aug. 25, 1987;

65 FR 2315, Jan. 14, 2000]

##### § 874.1120   Electronic   noise   generator for audiometric testing.

- Identification. An electronic noise generator for audiometric testing is a device that consists of a swept fre- quency generator, an amplifier, and an earphone. It is intended to introduce a masking  noise  into  the  non-test  ear

#### 21 CFR Ch. I (4–1–17 Edition)

during an audiometric evaluation. The device minimizes the non-test ear’s sensing of test tones and signals being generated for the ear being tested.

- Classification. Class II.

##### § 874.1325   Electroglottograph.

- Identification.	An electroglottograph is an AC-powered device that employs a pair of elec- trodes that are placed in contact with the skin on both sides of the larynx and held in place by a collar. It is in- tended to measure the electrical im- pedance of the larynx to aid in assess- ing the degree of closure of the vocal cords, confirm larygeal diagnosis, aid behavioral treatment of voice dis- orders, and aid research concerning the laryngeal mechanism.

- Classification. Class II.

##### § 874.1500   Gustometer.

- Identification. A gustometer is a battery-powered device that consists of two electrodes that are intended to be placed on both sides of the tongue at different taste centers and that  pro- vides a galvanic stimulus resulting in taste sensation. It is used for assessing the sense of taste.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9. If the device is not la- beled or otherwise represented as ster- ile, it is exempt from the current good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,  with  respect  to  complaint files.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 2316, Jan. 14, 2000]

##### § 874.1600   Olfactory test device.

- Identification. An olfactory test device is used to determine whether an olfactory loss is present. The device in- cludes one or more odorants that are presented to the patient’s nose to sub- jectively assess the patient’s ability to perceive odors.

- Classification. Class II (special controls). The special control for these

#### Food and Drug Administration, HHS	§ 874.3305

devices is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Olfactory Test Device.’’ For the availability of this guidance document, see § 874.1(e). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9. When indicated for the screening or diagnosis of dis- eases or conditions other than the loss of olfactory function, the device is not exempt from premarket notification procedures.

[71 FR 32835, June 7, 2006]

##### § 874.1800   Air  or  water  caloric  stimu- lator.

  - Identification. An air or water ca- loric stimulator is a device that deliv- ers a stream of air or water to the ear canal at controlled rates of flow and temperature and that is intended for vestibular function testing of a pa- tient’s body balance system. The ves- tibular stimulation of the semicircular canals produce involuntary eye move- ments that are measured and recorded by a nystagmograph.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2316, Jan. 14, 2000]

##### § 874.1820   Surgical nerve stimulator/lo- cator.

- Identification. A surgical nerve stimulator/locator is a device that is intended to provide electrical stimula- tion to the body to locate and identify nerves and to test their excitability.

- Classification. Class II.

##### § 874.1925   Toynbee diagnostic tube.

- Identification. The toynbee diag- nostic tube is a listening device in- tended to determine the degree of openness of the eustachian tube.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 2316, Jan. 14, 2000]

## Subpart C [Reserved]

Subpart D—Prosthetic Devices

##### § 874.3300   Hearing Aid.

- Identification. A hearing aid is wearable sound-amplifying device  that is intended to compensate for impaired hearing. This generic type of device in- cludes the air-conduction hearing aid and the bone-conduction hearing aid, but excludes the group hearing aid or group auditory trainer (§ 874.3320), mas- ter hearing aid (§ 874.3330), and tinnitus masker (§ 874.3400).

- Classification. (1) Class I (general controls) for the air-conduction hear- ing aid. The air-conduction hearing aid is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter subject to § 874.9.

- Class II for the bone-conduction hearing aid.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 2316, Jan. 14, 2000]

##### § 874.3305   Wireless	air-conduction hearing aid.

  - Identification. A wireless air-con- duction hearing aid is a wearable sound-amplifying device, intended to compensate for impaired hearing that incorporates wireless  technology  in its programming or use.

  - Classification: Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing should validate electro magnetic com- patibility (EMC) and safety of exposure to non-ionizing radiation;

- Design, description, and perform- ance data should validate wireless technology functions; and

- Labeling should specify appro- priate instructions, warnings, and in- formation relating to EMC and wireless technology and human exposure to non-ionizing  radiation.

  - Premarket notification. The wire- less air-conduction hearing aid is  ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[76 FR 34846, June 15, 2011]

§ 874.3310

##### § 874.3310   Hearing  aid  calibrator  and analysis system.

    - Identification. A hearing aid cali- brator and analysis system is an elec- tronic reference device intended to calibrate	and	assess	the electroacoustic frequency and sound intensity characteristics emanating from a hearing aid, master hearing aid, group hearing aid or group auditory trainer. The device consists of an acoustic complex of known cavity vol- ume, a sound level meter, a micro- phone, oscillators, frequency counters, microphone amplifiers, a distoration analyzer, a chart recorder, and a hear- ing aid test box.

    - Classification. Class II.

##### § 874.3315   Tympanic   membrane   con- tact hearing aid.

- Identification. A tympanic mem- brane contact hearing aid is a prescrip- tion device that compensates for im- paired hearing. Amplified sound is transmitted by vibrating the tympanic membrane through a transducer that is in direct contact with the tympanic membrane.

- Classification. Class II (special controls). The special controls for this device are:

  - The patient contacting compo- nents must be demonstrated to be bio- compatible.

  - Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use, and must include:

- Mechanical integrity testing;

- Electrical and thermal safety testing;

- Software verification, valida- tion, and hazard analysis;

- Reliability testing consistent with expected device life;

- Electromagnetic compatibility testing; and

- Validation testing of device out- put and mechanical force applied to the tympanic membrane in a clinically appropriate model.

  - Clinical performance testing must characterize any adverse events ob- served during clinical use, and dem- onstrate that the device performs as intended under anticipated conditions of use.

#### 21 CFR Ch. I (4–1–17 Edition)

  - Professional training must in- clude the ear impression procedure, correct placement, fitting, monitoring, care, and maintenance of the device.

  - Labeling must include the fol- lowing:

- A detailed summary of the adverse events and effectiveness outcomes from the clinical performance testing;

- Detailed instructions on how to fit the device to the patient;

- Instructions for  periodic clean- ing of any reusable components;

- Information related to electro- magnetic compatibility; and

- Patient labeling that includes:

  - A patient card that identifies if a patient has been fitted with any non- self- removable components of the de- vice and provides relevant information in cases of emergency;

  - Information on how to correctly use and maintain the device;

  - The potential risks and benefits associated with the use of the device; and

  - Alternative  treatments.

[81 FR 3326, Jan. 21, 2015]

##### § 874.3320   Group hearing aid or group auditory trainer.

    - Identification. A group hearing aid or group auditory trainer is a hearing aid that is intended for use in commu- nicating simultaneously with one or more listeners having hearing impair- ment. The device is used with an asso- ciated transmitter microphone. It may be either monaural or binaural, and it provides coupling to the ear through either earphones or earmolds. The ge- neric type of device includes three types of applications: hardwire sys- tems, inductance loop systems, and wireless systems.

    - Classification. Class II.

##### § 874.3330   Master hearing aid.

- Identification. A master hearing aid is an electronic device intended to simulate a hearing aid during audiometric testing. It has adjustable acoustic output levels, such as those for gain, output, and frequency re- sponse. The device is used to select and adjust a person’s wearable hearing aid.

- Classification. Class II.

#### Food and Drug Administration, HHS	§ 874.3540

##### § 874.3375   Battery-powered	artificial larynx.

  - Identification. A battery-powered artificial larynx is an externally  ap- plied device intended for use in the ab- sence of the larynx to produce sound. When held against the skin in the area of the voicebox, the device generates mechanical vibrations which resonate in the oral and nasal  cavities  and  can be modulated by the tongue and lips in a normal manner, thereby allowing the production of speech.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  59

FR 63009,  Dec. 7,  1994;  66  FR 38800,  July 25,

2001]

##### § 874.3400   Tinnitus masker.

- Identification. A tinnitus  masker is an electronic device intended to gen- erate noise of sufficient intensity and bandwidth to mask ringing in the ears or internal head noises. Because the de- vice is able to mask internal noises, it is also used as an aid in hearing exter- nal noises and speech.

- Classification. Class II. The special control for this device is patient label- ing regarding:

  - Hearing health care professional diagnosis, fitting of the device, and fol- lowup care,

  - Risks,

  - Benefits,

  - Warnings for safe use, and

  - Specifications.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 17145, Mar. 31, 2000]

##### § 874.3430   Middle ear mold.

    - Identification. A middle ear mold is a preformed device that is intended to be implanted to reconstruct the middle ear cavity during repair of the tympanic membrane. The device per- mits an ample air-filled cavity to be maintained in the middle ear and pro- motes regeneration of the mucous membrane lining of the middle ear cav- ity. A middle ear mold is made of ma- terials such as polyamide, polytetra- fluoroethylene,  silicone  elastomer,  or

polyethylene, but does not contain po- rous polyethylene.

    - Classification. Class II.

##### § 874.3450 Partial ossicular replace- ment prosthesis.

- Identification. A partial ossicular replacement prosthesis is a device in- tended to be implanted for the func- tional reconstruction of segments of the ossicular chain and facilitates the conduction of sound wave from the tympanic membrane to the inner ear. The device is made of materials such as stainless steel, tantalum, polytetra- fluoroethylene, polyethylene, poly- tetrafluoroethylene with carbon fibers composite, absorbable gelatin mate- rial, porous polyethylene, or from a combination of these materials.

- Classification. Class II.

##### § 874.3495 Total ossicular replacement prosthesis.

- Identification. A total ossicular re- placement prosthesis is a device in- tended to be implanted for the total functional reconstruction of the ossic- ular chain and facilitates the conduc- tion of sound waves from the tympanic membrance to the inner ear. The de- vice is made of materials such as poly- tetrafluoroethylene, polytetrafluoro- ethylene with vitreous carbon fibers composite, porous polyethylene, or from a combination of these materials.

- Classification. Class II.

##### § 874.3540 Prosthesis modification in- strument for ossicular replacement surgery.

- Identification. A prosthesis modi- fication instrument for ossicular re- placement surgery is a device intended for use by a surgeon to construct ossic- ular replacements. This generic type of device includes the ear, nose, and throat cutting block; wire crimper, wire bending die; wire closure forceps; piston cutting  jib;  gelfoam TM punch; wire cutting scissors; and ossicular fin- ger vise.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9. If the device is not la- beled or otherwise represented as ster- ile, it is exempt from the current good

#### § 874.3620

manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,   with   respect   to   complaint files.

[51  FR  40389,  Nov.  9,  1986,  as  amended  at  52

FR 32111,  Aug. 25,  1987;  65  FR 2316,  Jan. 14,

2000]

##### § 874.3620   Ear,  nose,  and  throat  syn- thetic polymer material.

  - Identification. Ear, nose, and throat synthetic polymer material is a device material that is intended to be implanted for use as a space-occupying substance in the reconstructive sur- gery of the head and neck. The device is used, for example, in augmentation rhinoplasty and in tissue defect clo- sures in the esophagus. The device is shaped and formed by the suregon to conform to the patient’s needs. This generic type of device is made of mate- rial such as polyamide mesh or foil and porous polyethylene.

  - Classification. Class II.

##### § 874.3695   Mandibular   implant   facial prosthesis.

- Identification. A mandibular im- plant facial prosthesis is a device that is intended to be implanted for use in the functional reconstruction of man- dibular deficits. The device is made of materials such as stainless steel, tan- talum, titanium, cobalt-chromium based alloy, polytetrafluoroethylene, silicone elastomer, polyethylene, poly- urethane, or polytetrafluoroethylene with carbon fibers composite.

- Classification. Class II.

##### § 874.3730   Laryngeal  prosthesis  (Taub design).

- Identification. A laryngeal pros- thesis (Taub design) is a device in- tended to direct pulmonary air flow to the pharynx in the absence of the lar- ynx, thereby permitting esophageal speech. The device is interposed be- tween openings in the trachea and the esophagus and may be removed and re- placed each day by the patient. During phonation, air from the lungs is di- rected to flow through the device and over the esophageal mucosa to provide

#### 21 CFR Ch. I (4–1–17 Edition)

a  sound  source  that  is  articulated  as speech.

- Classification. Class II.

##### § 874.3760   Sacculotomy	tack	(Cody tack)

- Identification. A sacculotomy tack (Cody tack) is a device that consists of a pointed stainless steel tack intended to be implanted to relieve the symp- toms of vertigo. The device repet- itively ruptures the utricular mem- brane as the membrane expands under increased endolymphatic pressure.

- Classification. Class II.

##### § 874.3820   Endolymphatic shunt.

- Identification. An endolymphatic shunt is a device that consists of a tube or sheet intended to be implanted to relieve the symptons of vertigo. The device permits the unrestricted flow of excess endolymph from the distended end of the endolymphatic system into the mastoid cavity where resorption occurs. This device is made of poly- tetrafluoroethylene or silicone elas- tomer.

- Classification. Class II.

##### § 874.3850   Endolymphatic   shunt   tube with valve.

- Identification. An endolymphatic shunt tube with valve is a device that consists of a pressure-limiting valve associated with a tube intended to be implanted in the inner ear to relieve symptoms of vertigo and hearing loss due to endolymphatic hydrops (in- crease in endolymphatic fluid) of Meniere’s disease.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document ‘‘Class II Special Controls Guidance Document: Endolymphatic Shunt Tube With Valve; Guidance for Industry and FDA.’’

[67 FR 20894, Apr. 29, 2002]

##### § 874.3880   Tympanostomy tube.

- Identification. A tympanostomy tube is a device that is intended to be implanted for ventilation or  drainage of the middle ear. The device is in- serted through the tympanic mem- brane to permit a free exchange of air between the outer ear and middle ear. A type of tympanostomy tube known

#### Food and Drug Administration, HHS	§ 874.4140

as the malleous clip tube attaches  to the malleous to provide middle ear ventilation. The device is made of ma- terials such as polytetrafluoro- ethylene, polyethylene, silicon elas- tomer, or porous polyethylene.

- Classification. Class II.

##### § 874.3900   Nasal dilator.

- Identification. A nasal dilator is a device intended to provide temporary relief from transient causes of breath- ing difficulties resulting from struc- tural abnormalities and/or transient causes of nasal congestion associated with reduced nasal airflow. The device decreases airway resistance and in- creases nasal airflow. The external nasal dilator is constructed from one or more layers of material upon which a spring material is attached, with a skin adhesive applied to adhere to the skin of the nose; it acts with a pulling action to open the nares. The internal nasal dilator is constructed from metal or plastic and is placed inside the nos- trils; it acts by pushing the nostrils open or by gently pressing on the columella.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[64 FR 10949, Mar. 8, 1999]

##### § 874.3930   Tympanostomy    tube    with semipermeable membrane.

- Identification. A tympanostomy tube with a semipermeable membrane is a device intended to be implanted for ventilation or drainage of the middle ear and for preventing fluids from en- tering the middle ear cavity. The de- vice is inserted through the tympanic membrane to permit a free exchange of air between the outer ear and middle ear. The tube portion of the device is made of silicone elastomer or porous polyethylene, and the membrane por- tion is made of polytetrafluoro- ethylene.

- Classification. Class II. The special control for this device is FDA’s ‘‘Tympanostomy Tubes, Submission Guidance for a 510(k).’’

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 17145, Mar. 31, 2000]

##### § 874.3950   Transcutaneous air conduc- tion hearing aid system.

- Identification. A  transcutaneous air conduction  hearing  aid system is  a wearable sound-amplifying device in- tended to compensate for impaired hearing without occluding  the  ear canal. The device consists of an air conduction hearing aid attached to a surgically fitted tube system, which is placed through soft tissue between the post auricular region and the outer ear canal.

- Classification. Class II (special controls). The special control for this device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document: Transcutaneous Air Conduction Hearing Aid System (TACHAS); Guidance for Industry and FDA.’’ See § 874.1 for the availability of this guidance document.

[67 FR 67790, Nov. 7, 2002]

## Subpart E—Surgical Devices

##### § 874.4100   Epistaxis balloon.

- Identification. An epistaxis balloon is a device consisting of an inflatable balloon intended to control internal nasal bleeding by exerting pressure against the sphenopalatine artery.

- Classification Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 2316, Jan. 14, 2000]

##### § 874.4140   Ear, nose, and throat bur.

- Identification. An ear, nose, and throat bur is a device consisting of an interchangeable drill bit  that  is  in- tended for use in an  ear,  nose,  and throat electric or pneumatic  surgical drill (§ 874.4250) for incising or removing bone in the ear, nose, or throat area. The  bur  consists  of  a   carbide   cutting tip on a metal shank or a coating of di- amond on a metal shank. The device is used in mastoid surgery, frontal sinus surgery, and surgery of  the  facial nerves.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

§ 874.4175

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  61

FR 1122,  Jan. 16,  1996;  66  FR 38800,  July 25,

2001]

##### § 874.4175   Nasopharyngeal catheter.

  - Identification. A nasopharyngeal catheter is a device consisting of a bou- gie or filiform catheter that is in- tended for  use in  probing or  dilating the eustachian tube. This generic type of device includes eustachian cath- eters.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 874.4180   Eustachian tube balloon di- lation system.

- Identification. A Eustachian tube balloon dilation system is a prescrip- tion device that includes a flexible catheter attached to an inflatable bal- loon. The system is intended for use in dilating the cartilaginous portion of the Eustachian tube for treating per- sistent Eustachian tube dysfunction.

- Classification. Class II (special controls). The special controls for this device are:

- Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be eval- uated:

  - Mechanical testing, including ten- sile and flexural testing of catheter joints and materials.

  - Durability testing, including fa- tigue and burst pressure testing of the balloon materials and components.

  - Inflation and deflation charac- terization testing, including time and pressure measurements, and leak test- ing of the balloon.

  - Verification testing of safety fea- tures built into the device must be per- formed, including the characterization of catheter geometries and distal tip insertion limitation mechanisms.

- Simulated use testing in a clini- cally relevant model must demonstrate

#### 21 CFR Ch. I (4–1–17 Edition)

the reliability of the device to remain mechanically functional  throughout the anticipated conditions of use, and validate that the design features limit access to only the cartilaginous por- tion of the Eustachian tube.

- The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

- Performance data must dem- onstrate the sterility of the device.

- Performance data must support shelf life by demonstrating continued sterility of the device, package integ- rity, and device functionality over the identified shelf life.

- Training must include simulated use on cadavers to ensure users can fol- low the instructions for use to allow safe use of the device.

- Labeling must include:

  - Detailed instructions for use.

  - A detailed summary of the device technical parameters, including max- imum allowed inflation pressure, al- lowable catheter geometries, and avail- able balloon sizes.

  - A shelf life.

[81 FR 73041, Oct. 24, 2016]

##### § 874.4250   Ear,  nose,  and  throat  elec- tric or pneumatic surgical drill.

- Identification. An ear, nose, and throat electric or pneumatic  surgical drill is a rotating drilling device, in- cluding the handpiece, that  is  intended to drive various accessories, such as an ear, nose, and throat bur (§ 874.4140), for the controlled incision or removal  of bone  in  the  ear,  nose,  and  throat  area.

- Classification. Class II.

##### § 874.4350   Ear,	nose,	and	throat fiberoptic light source and carrier.

- Identification. An ear, nose,  and throat fiberoptic light source and car- rier is an AC-powered device that gen- erates and  transmits  light  through glass of plastic fibers and that is in- tended to provide illumination at  the tip of an ear, nose, or throat endoscope. Endoscopic devices which utilize fiberoptic light sources and carriers in- clude the bronchoscope, esophagoscope, laryngoscope, mediastinoscope, laryn- geal-bronchial	telescope,	and nasopharyngoscope.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	§ 874.4710

premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 874.4420   Ear,  nose,  and  throat  man- ual surgical instrument.

  - Identification. An ear, nose, and throat manual surgical instrument is one of a variety of devices intended for use in surgical procedures to examine or treat the bronchus, esophagus, tra- chea, larynx, pharynx, nasal and para- nasal sinus, or ear. This generic type of device includes the esophageal dilator; tracheal bistour (a long, narrow sur- gical knife); tracheal dilator; tracheal hook; laryngeal injection set; laryngeal knife; laryngeal saw; laryngeal trocar; laryngectomy tube; adenoid curette; adenotome; metal tongue depressor; mouth gag; oral screw; salpingeal cu- rette; tonsillectome; tonsil guillotine; tonsil screw; tonsil snare; tonsil suc- tion tube; tonsil suturing hook; antom reforator; ethmoid curette; frontal sinus-rasp; nasal curette; nasal rasp; nasal rongeur; nasal saw; nasal scis- sors; nasal snare; sinus irrigator; sinus trephine; ear curette; ear excavator; ear rasp; ear scissor, ear snare; ear spoon; ear suction tube; malleous rip- per; mastoid gauge; microsurgical ear chisel; myringotomy tube inserter; ossici holding clamp; sacculotomy tack inserter; vein press; wire ear loop; microrule; mirror; mobilizer; ear, nose, and throat punch; ear, nose and throat knife; and ear, nose, and throat trocar.

  - Classification Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[51  FR  40389,  Nov.  9,  1986,  as  amended  at  52

FR 32111,  Aug. 25,  1987;  65  FR 2316,  Jan. 14,

2000; 72 FR 17400, Apr. 9, 2007]

##### § 874.4490   Argon laser for otology, rhi- nology, and laryngology.

- Identification. The argon laser de- vice for use in otology, rhinology, and laryngology is an electro-optical device which produces coherent, electro- magnetic radiation with principal wavelength peaks of 488 and 514 nano- meters.  In  otology,  the  device  is  used

for the purpose of coagulating and va- porizing soft and fibrous tissues, in- cluding osseous tissue. In rhinology and laryngology, the device is used to coagulate and vaporize soft and fibrous tissues, but not including osseous tis- sues.

- Classification. Class II.

[58 FR 29534, May 21, 1993]

##### § 874.4500   Ear, nose, and throat micro- surgical carbon dioxide laser.

- Identification. An ear, nose, and throat microsurgical carbon dioxide laser is a device intended for the sur- gical excision of tissue from the ear, nose, and throat area. The device is used, for example, in microsurgical procedures to excise lesions and tumors of the vocal cords and adjacent areas.

- Classification. Class II.

##### § 874.4680   Bronchoscope    (flexible    or rigid) and accessories.

- Identification. A bronchoscope (flexible or rigid) and accessories is a tubular endoscopic device with any of a group of accessory devices which at- tach to the bronchoscope and is in- tended to examine or treat the larynx and tracheobronchial tree. It is typi- cally used with a fiberoptic light source and carrier to provide illumina- tion. The device is made of materials such as stainless steel or flexible plas- tic. This generic type of device in- cludes the rigid ventilating broncho- scope, rigid nonventilating broncho- scope, nonrigid bronchoscope, laryn- geal-bronchial telescope, flexible for- eign body claw, bronchoscope tubing, flexible biopsy forceps, rigid biopsy cu- rette, flexible biopsy brush, rigid bi- opsy forceps, flexible biopsy curette, and rigid bronchoscope aspirating tube, but excludes the fiberoptic light source and carrier.

- Classification. Class II.

##### § 874.4710   Esophagoscope  (flexible  or rigid) and accessories.

- Identification. An esophagoscope (flexible or rigid) and accessories is a tubular endoscopic device with any of a group of accessory devices which at- tach to the esophagoscope and is in- tended to examine or treat esophageal malfunction symptoms, esophageal or

#### § 874.4720

mediastinal disease, or to remove for- eign bodies from the esophagus. When inserted, the device extends from the area of the hypopharynx to the stom- ach. It is typically used with a fiberoptic light source and carrier to provide illumination. The  device  is made of materials such  as  stainless steel or flexible plastic.  This  generic type of device includes the flexible for- eign body claw, flexible biopsy forceps, rigid biopsy curette, flexible biopsy brush, rigid biopsy forceps and flexible biopsy curette, but excludes the fiberoptic light source and carrier.

- Classification. Class II.

##### § 874.4720   Mediastinoscope  and  acces- sories.

- Identification. A mediastinoscope and accessories is a tubular tapered electrical endoscopic device with  any of a group of accessory devices which attach to the mediastinoscope and is intended to examine or treat tissue  in the area separating the lungs. The de- vice is inserted transthoracicly and is used in diagnosis of tumors and lesions and to determine whether excision of certain organs or  tissues  is  indicated. It is typically used with a fiberoptic light source and carrier to provide illu- mination. The device is made of mate- rials such as stainless steel. This ge- neric type of device includes the  flexi- ble foreign body claw, flexible biopsy forceps, rigid biopsy curette, flexible biopsy brush, rigid biopsy forceps, and flexible biopsy  curette,  but  excludes the  fiberoptic  light  source  and  carrier.

- Classification. Class II.

##### § 874.4750   Laryngostroboscope.

- Identification.	A laryngostroboscope is a device that is intended to allow observation of glottic action during phonation. The device operates by focusing a stroboscopic light through a lens for direct  or  mir- ror reflected viewing of glottic action. The light and microphone that ampli- fies acoustic signals from  the  glottic area may or may not contact the pa- tient.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter subject to the limitations in § 874.9.

[55 FR 48440, Nov. 20, 1990, as amended at 59

FR 63009,  Dec. 7,  1994;  66  FR 38801,  July 25,

2001]

##### § 874.4760   Nasopharyngoscope	(flexi- ble or rigid) and accessories.

- Identification.			A nasopharyngoscope   (flexible   or   rigid) and accessories is a tubular endoscopic device with any of a group of accessory devices	which	attach	to	the nasopharyngoscope  and  is  intended  to examine  or  treat  the  nasal  cavity  and nasal pharynx. It is typically used with a fiberoptic light source and carrier to provide    illumination.    The    device    is made   of   materials   such   as   stainless steel  and  flexible  plastic.  This  generic type of device includes the antroscope, nasopharyngolaryngoscope, nasosinuscope,		nasoscope,

postrhinoscope,	rhinoscope, salpingoscope, flexible foreign body claw, flexible biopsy forceps, rigid bi- opsy curette, flexible biospy brush, rigid biopsy forceps and flexible biopsy curette, but excludes the fiberoptic light source and carrier.

- Classification. Class II.

##### § 874.4770   Otoscope.

- Identification. An otoscope is a de- vice intended to allow inspection of the external ear canal and tympanic mem- brane under magnification. The device provides illumination of the ear canal for observation by using an AC- or bat- tery-powered light source and an opti- cal magnifying system.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9 only when used in the external ear canal.

[55 FR 48440, Nov. 20, 1990, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 874.4780   Intranasal splint.

- Identification. An intranasal splint is intended to minimize bleeding and edema and to prevent adhesions be- tween the septum and the nasal cavity. It is placed in the nasal cavity after surgery   or   trauma.   The   intranasal

#### Food and Drug Administration, HHS	§ 874.5370

splint is constructed from plastic, sili- cone, or absorbent material.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[64 FR 10949, Mar. 8, 1999]

##### § 874.4800   Bone particle collector.

- Identification. A bone particle col- lector is a filtering device intended to be inserted into a suction tube during the early stages of otologic surgery to collect bone particles for future use.

- Classification. Class I (general con- trols). The device is exempt from pre- market notification procedures in sub- part E of part 807 of this chapter sub- ject to the limitations in § 874.9.

[64 FR 10949, Mar. 8, 1999]

## Subpart F—Therapeutic Devices

##### § 874.5220   Ear,  nose,  and  throat  drug administration device.

- Identification. An ear, nose, and throat drug administration device is one of a group of ear, nose, and throat devices intended specifically to admin- ister medicinal substances to treat ear, nose, and throat disorders. These in- struments include the powder blower, dropper, ear wick, manual nebulizer pump, and nasal inhaler.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.  If the device is not labeled or otherwise represented as sterile, it  is  exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  59

FR 63009,  Dec. 7,  1994;  66  FR 38801,  July 25,

2001]

##### § 874.5300   Ear, nose, and throat exam- ination and treatment unit.

- Identification. An ear, nose, and throat examination and treatment unit is  an  AC-powered  device  intended  to

support a patient during an otologic examination while providing special- ized features for examination and treatment. The unit consists of a pa- tient chair and table, drawers for equipment, suction and blowing appa- ratus, and receptacles for connection of specialized lights and examining in- struments.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2316, Jan. 14, 2000]

##### § 874.5350   Suction antichoke device.

- Identification. A suction antichoke device is a device intended to be used in an emergency situation to remove, by the application of suction, foreign objects that obstruct a patient’s air- way to prevent asphyxiation to the pa- tient.

- Classification. Class III.

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP for a device is re- quired to be filed with the Food and Drug Administration on or before July 13, 1999 for any suction antichoke de- vice that was  in  commercial  distribu- tion before May 28, 1976, or that has, on or before July 13, 1999, been found to be substantially equivalent to a suction antichoke device that was in commer- cial distribution before May 28, 1976. Any  other  suction  antichoke  device shall have an approved PMA or  de- clared completed PDP in effect before being placed in commercial distribu- tion.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  64

FR 18329,  Apr. 14,  1999;  65  FR 2316,  Jan. 14,

2000]

##### § 874.5370   Tongs antichoke device.

- Identification. A tongs antichoke device is a device that is intended to be used in an emergency situation to grasp and remove foreign objects that obstruct a patient’s airway to prevent asphyxiation of the patient. This ge- neric type of device includes a plastic instrument with serrated ends that is inserted into the airway in a blind manner  to  grasp  and  extract  foreign

§ 874.5550

objects, and a stainless steel forceps with spoon ends that is inserted under tactile guidance to grasp and extract foreign objects from the airway.

- Classification. Class III.

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP for a device is re- quired to be filed with the Food and Drug Administration on or before July 13, 1999 for any tongs antichoke device that was in commercial distribution before May  28, 1976,  or that  has, on  or before July 13, 1999, been found to be substantially equivalent to a tongs antichoke device that was in commer- cial distribution before May 28, 1976. Any other tongs antichoke device shall have an approved PMA or declared completed PDP in effect before being placed  in  commercial  distribution.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  64

FR 18329, Apr. 14, 1999]

##### § 874.5550   Powered nasal irrigator.

- Identification. A powered nasal irrigator is an AC-powered device in- tended to wash the nasal cavity by means of a pressure-controlled pul- sating stream of water. The device con- sists of a control unit and pump con- nected to a spray tube and nozzle.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2316, Jan. 14, 2000]

##### § 874.5800   External nasal splint.

- Identification. An external nasal splint is a rigid or partially rigid de- vice intended for use externally for im- mobilization of parts of the nose.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 874.9.

[51  FR  40389,  Nov.  9,  1986,  as  amended  at  52

FR  32111,  Aug.  25,  1987;  59  FR  63009,  Dec.  7,

1994; 66 FR 38801, July 25, 2001]

##### § 874.5840   Antistammering device.

- Identification. An antistammering device is a device that electronically generates  a  noise  when  activated  or

#### 21 CFR Ch. I (4–1–17 Edition)

when it senses the user’s speech and that is intended to prevent the user from hearing the sounds of his or her own voice. The device is used to mini- mize a user’s involuntary hesitative or repetitive speech.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 874.9.

[51  FR  40389,  Nov.  6,  1986,  as  amended  at  65

FR 2316, Jan. 14, 2000]

##### § 874.5900   External  upper  esophageal sphincter compression device.

- Identification. An external upper esophageal sphincter compression de- vice is a prescription device used to apply external pressure on the cricoid cartilage for the purpose of reducing the symptoms of laryngopharyngeal reflux disease.

- Classification. Class II (special controls). The special controls for this device are:

  - The patient contacting compo- nents must be demonstrated to be bio- compatible.

  - Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be dem- onstrated:

    - Mechanical integrity testing (e.g., tensile strength testing, fatigue test- ing) and

    - Shelf life testing.

  - The technical specifications must include pressure measurement accu- racy to characterize device perform- ance.

  - Clinical performance testing must document any adverse events observed during clinical use, and  demonstrate that the device performs as intended under anticipated conditions of use.

  - Labeling must include the fol- lowing:

    - Appropriate warnings and pre- cautions,

    - A detailed summary of the clin- ical testing pertinent to use of the de- vice including a detailed summary of the device-related complications or ad- verse events,

    - Detailed instructions on how to fit the device to the patient, and

#### Food and Drug Administration, HHS	Pt. 876

    - Instructions  for  reprocessing  of any reusable components.

  - Patient labeling must be provided and must include:

    - Relevant   warnings,   precautions, and adverse effects/complications,

    - Information on how to correctly wear the device,

    - The  potential  risks  and  benefits associated with the use of the device,

    - Alternative treatments, and

    - Reprocessing instructions.

[80 FR 46194, Aug. 4, 2015]

## PART 876—GASTROENTEROLOGY- UROLOGY DEVICES

##### Subpart A—General Provisions

Sec.

876.1    Scope.

876.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

876.1075 Gastroenterology-urology	biopsy instrument.

876.1300 Ingestible	telemetric	gastro- intestinal  capsule  imaging  system.

876.1330   Colon capsule endoscopy system. 876.1400   Stomach pH electrode.

876.1500   Endoscope and accessories.

876.1620    Urodynamics  measurement  system. 876.1725    Gastrointestinal      motility      moni-

toring system.

876.1735    Electrogastrography system. 876.1800   Urine   flow   or   volume   measuring

system.

##### Subpart C—Monitoring Devices

876.2040   Enuresis alarm.

876.2050    Prostate  lesion  documentation  sys- tem.

##### Subpart D—Prosthetic Devices

876.3350    Penile inflatable implant. 876.3630    Penile rigidity implant. 876.3750   Testicular prosthesis.

##### Subpart E—Surgical Devices

876.4020     Fiberoptic light ureteral catheter. 876.4270  Colostomy rod.

876.4300    Endoscopic electrosurgical unit and accessories.

876.4370    Gastroenterology-urology evacuator.

876.4400   Hemorrhoidal ligator.

876.4480     Electrohydraulic lithotriptor.

876.4500    Mechanical lithotriptor.

876.4530 Gastroenterology-urology fiberoptic  retractor.

876.4560   Ribdam.

876.4590    Interlocking urethral sound. 876.4620   Ureteral stent.

876.4650   Water jet renal stone dislodger sys- tem.

876.4680   Ureteral stone dislodger.

876.4730 Manual	gastroenterology-urology surgical  instrument  and  accessories.

876.4770    Urethrotome.

876.4890    Urological table and accessories.

##### Subpart F—Therapeutic Devices

876.5010    Biliary catheter and accessories. 876.5011    Metallic   biliary   stent   system   for

benign strictures.

876.5015    Pancreatic  drainage  stent  and  de- livery   system.

876.5020    External penile rigidity devices. 876.5025   Vibrator for climax control of pre-

mature ejaculation.

876.5030    Continent ileostomy catheter. 876.5090    Suprapubic  urological  catheter  and

accessories.

876.5130    Urological catheter and accessories. 876.5140    Urethral insert with pump for blad-

der drainage.

876.5160    Urological clamp for males. 876.5210  Enema kit.

876.5220    Colonic irrigation system. 876.5250    Urine collector and accessories.

876.5270   Implanted  electrical  urinary  con- tinence  device.

876.5280 Implanted	mechanical/hydraulic urinary   continence   device.

876.5310    Nonimplanted,  peripheral  electrical continence   device.

876.5320    Nonimplanted  electrical  continence device.

876.5365   Esophageal dilator.

876.5450   Rectal dilator.

876.5470    Ureteral dilator.

876.5520    Urethral dilator.

876.5530    Implantable    transprostatic    tissue retractor    system.

876.5540   Blood access device and accessories. 876.5600    Sorbent   regenerated   dialysate   de-

livery system for hemodialysis.

876.5630    Peritoneal  dialysis  system  and  ac- cessories.

876.5665  Water	purification	system	for hemodialysis.

876.5820   Hemodialysis    system    and    acces- sories.

876.5830   Hemodialyzer   with   disposable   in- sert  (Kiil  type).

876.5860  High	permeability	hemodialysis system.

876.5870    Sorbent hemoperfusion system. 876.5880 Isolated	kidney	perfusion	and

transport system and accessories.

876.5885 Tissue culture media for human ex vivo tissue and cell culture processing applications.

§ 876.1

876.5895   Ostomy irrigator.

876.5900   Ostomy pouch and accessories. 876.5920    Protective    garment    for    inconti-

nence.

876.5930    Rectal control system. 876.5955   Peritoneo-venous shunt.

876.5970    Hernia support.

876.5980    Gastrointestinal    tube    and    acces- sories.

876.5990  Extracorporeal	shock	wave lithotripter.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE:  48  FR  53023,  Nov.  23,  1983,  unless

otherwise noted.

## Subpart A—General Provisions

EDITORIAL NOTE: Nomenclature changes  to part 876 appear at 73 FR 35341, June 23, 2008.

##### § 876.1   Scope.

    - This part sets forth the classifica- tion of gastroenterology-urology de- vices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - To avoid duplicative listings, a gastroenterology-urology device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed only in one subpart.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 17737, May 11, 1987; 52 FR 22577, June

12, 1987, as amended at 69 FR 77623, Dec. 28,

2004; 78 FR 18233, Mar. 26, 2013]

21 CFR Ch. I (4–1–17 Edition)

##### § 876.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be commerically  distributed  without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regu- lation under section 515(b) of the act requiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device

#### Food and Drug Administration, HHS	§ 876.1075

before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17737, May 11, 1987]

##### § 876.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies    infectious    agents    by    using

deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2316, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 876.1075   Gastroenterology-urology biopsy instrument.

      - Identification. A gastroenterology- urology biopsy instrument is a device used to remove, by cutting or aspira- tion, a specimen of tissue for micro- scopic examination. This generic type of device includes the biopsy punch, gastrointestinal mechanical biopsy in- strument, suction biopsy instrument, gastro-urology biopsy needle and nee- dle set, and nonelectric biopsy forceps. This section does not apply to biopsy instruments that have specialized uses in  other  medical  specialty  areas  and

§ 876.1300

that are covered by classification regu- lations in other parts of the device classification regulations.

      - Classification. (1)  Class  II  (per- formance standards).

- Class I for the biopsy forceps cover and the non-electric biopsy for- ceps. The devices subject to this para- graph (b)(2) are exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter sub- ject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 876.1300   Ingestible telemetric gastro- intestinal capsule imaging system.

  - Identification. An ingestible tele- metric gastrointestinal capsule  imag- ing system is used for visualization  of the small bowel mucosa as an adjunc- tive tool in the detection of abnormali- ties of the small bowel. The device cap- tures images of the small bowel with a wireless camera contained in a capsule. This device includes an ingestible cap- sule (containing a light source, camera, transmitter, and battery), an antenna array, a receiving/recording unit, a data storage device, computer software to process the images, and accessories.

  - Classification. Class II (special controls). The special control is FDA’s guidance, ‘‘Class II Special Controls Guidance Document: Ingestible Tele- metric Gastrointestinal Capsule Imag- ing Systems; Final Guidance for Indus- try and FDA.’’

[67 FR 3433, Jan. 24, 2002]

##### § 876.1330   Colon	capsule	endoscopy system.

- Identification. A prescription, sin- gle-use ingestible capsule designed to acquire video images during natural propulsion through the digestive sys- tem. It is specifically designed to vis- ualize the colon for the detection of polyps. It is intended for use only in patients who had an incomplete optical colonoscopy with adequate prepara- tion, and a complete evaluation of the colon was not technically possible.

- Classification. Class II (special controls). The special controls for this device are:

- The capsule must be demonstrated to be biocompatible.

#### 21 CFR Ch. I (4–1–17 Edition)

- Non-clinical testing data must demonstrate the mechanical and func- tional integrity of the device under physically stressed conditions. The fol- lowing performance characteristics must be tested and detailed protocols must be provided for each test:

- Bite test to ensure that the cap- sule can withstand extreme cases of biting.

- pH resistance test to evaluate in- tegrity of the capsule when exposed to a range of pH values.

- Battery life test to demonstrate that the capsule’s operating time is not constrained by the battery capacity.

- Shelf-life testing to demonstrate that the device performs as intended at the proposed shelf-life date.

- Optical testing to evaluate funda- mental image quality characteristics such as resolution, field of view, depth of field, distortion, signal-to-noise ratio, uniformity, and  image  artifacts. A test must be performed to evaluate the potential of scratches, caused by travelling through the gastrointestinal tract, on the transparent window of the capsule and their impact on the optical and color performance.

- An optical safety analysis must be performed based on maximum (worst-case) light exposure to internal gastrointestinal mucosa, and covering ultraviolet, visible, and near-infrared ranges, as appropriate. A mitigation analysis must be provided.

- A color performance test must be provided to compare the color dif- ferences between the input scene and output  image.

- The video viewer must clearly present the temporal or spatial rela- tionship between any two frames as a real-time lapse or a travel distance. The video viewer must alert the user when the specific video interval is cap- tured at a frame rate lower than the nominal one due to communication er- rors.

- A performance test evaluating the latency caused by any adaptive al- gorithm such as adjustable frame rate must be provided.

- If the capsule includes a localiza- tion module, a localization perform- ance test must be performed to verify the accuracy and precision of locating the capsule position within the colon.

#### Food and Drug Administration, HHS	§ 876.1500

- A data transmission test must be performed to verify the robustness of the data transmission between the cap- sule and the recorder. Controlled signal attenuation should be included for sim- ulating a non-ideal environment.

- Software	validation, verification, and hazards analysis must be provided.

- Electrical equipment safety, in- cluding thermal and mechanical safety and electromagnetic compatibility (EMC) testing must be performed. If the environments of intended use in- clude locations outside of hospitals and clinics, appropriate higher immunity test levels must be used. Labeling must include appropriate EMC information.

- Information demonstrating im- munity from wireless hazards.

- The clinical performance charac- teristics of the device for the detection of colon polyps must be established. Demonstration of the performance characteristics must include assess- ment of positive percent agreement and negative percent agreement com- pared to a clinically acceptable alter- native structural imaging method.

- Clinician labeling must include:

  - Specific instructions and the clin- ical and technical expertise needed for the safe use of the device.

  - A detailed summary of the clin- ical testing pertinent to use of the de- vice, including the percentage of pa- tients in which a polyp was correctly identified by capsule endoscopy, but also the percent of patients in which the capsule either missed or falsely identified a polyp with respect to the clinically acceptable alternative struc- tural imaging method.

  - The colon cleansing procedure.

  - A detailed summary of the device technical parameters.

  - A detailed summary of the device- and procedure-related complications pertinent to use of the device.

  - An expiration date/shelf life.

- Patient labeling must include:

  - An explanation of the device and the mechanism of operation.

  - Patient preparation procedure.

  - A brief summary of the clinical study. The summary should not only include the percentage of patients in which a polyp was correctly identified by capsule endoscopy, but also the per-

cent of patients in which the capsule either missed or falsely identified a polyp with respect to the clinically ac- ceptable alternative structural imag- ing method.

  - A summary of the device- and procedure-related complications perti- nent to use of the device.

[79 FR 28403, May 16, 2014]

##### § 876.1400   Stomach pH electrode.

- Identification. A stomach pH elec- trode is a device used to measure intragastric and intraesophageal pH (hydrogen ion concentration). The pH electrode is at the end of a flexible lead which may be inserted into the esoph- agus or stomach through the patient’s mouth. The device may include an in- tegral gastrointestinal tube.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122, Jan. 16, 1996]

##### § 876.1500   Endoscope and accessories.

- Identification.  An  endoscope  and accessories is a device used to provide access,  illumination,  and  allow  obser- vation   or   manipulation   of   body   cav- ities,  hollow  organs,   and  canals.  The device consists of various rigid or flexi- ble  instruments  that  are  inserted  into body spaces and may include an optical system  for  conveying  an  image  to  the user’s  eye  and  their  accessories  may assist in gaining access or increase the versatility   and   augment   the   capabili- ties of the devices. Examples of devices that are within this generic type of de- vice   include   cleaning   accessories   for endoscopes,   photographic   accessories for endoscopes, nonpowered anoscopes, binolcular attachments for endoscopes, pocket  battery  boxes,  flexible  or  rigid choledochoscopes,   colonoscopes,   diag- nostic  cystoscopes,  cystourethroscopes, enteroscopes, esophagogastroduodenoscopes,	rigid esophagoscopes, fiberoptic illuminators for endoscopes, incandescent endoscope lamps,	biliary	pancreatoscopes, proctoscopes,		resectoscopes,

nephroscopes,	sigmoidoscopes,

ureteroscopes,		urethroscopes, endomagnetic	retrievers,	cytology

#### § 876.1620

brushes for endoscopes, and lubricating jelly for transurethral surgical instru- ments. This section does not apply to endoscopes that have specialized uses in other medical specialty areas and that are covered by classification regu- lations in other parts of the device classification regulations.

- Classification. (1)  Class  II  (per- formance standards).

(2) Class I for the photographic acces- sories for endoscope,  miscellaneous bulb adapter for endoscope, binocular attachment for endoscope, eyepiece at- tachment for prescription lens, teach- ing attachment, inflation bulb, meas- uring device for panendoscope, photo- graphic equipment for physiologic function monitor, special lens instru- ment for endoscope, smoke removal tube, rechargeable battery box, pocket battery box, bite block for endoscope, and cleaning brush for endoscope. The devices subject to this paragraph (b)(2) are exempt from the premarket notifi- cation procedures in subpart E of part 807of this chapter, subject to the limi- tations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 876.1620   Urodynamics   measurement system.

- Identification. A urodynamics measurement system is a device used to measure volume and pressure in the urinary bladder  when it  is  filled through a catheter with carbon dioxide or water. The device controls the sup- ply of carbon dioxide or water and may also record the electrical activity of the muscles associated with urination. The device system may include trans- ducers, electronic signal conditioning and display equipment, a catheter withdrawal device to enable a urethral pressure profile to be obtained, and special catheters for urethral profilometry and electrodes for electromyography. This generic type of device includes the cystometric gas (carbon	dioxide)	device,	the cystometric hydrualic device, and the electrical recording cystometer, but excludes any device that uses air to fill the bladder.

- Classification. Class II (special controls).  The  device  is  exempt  from

#### 21 CFR Ch. I (4–1–17 Edition)

the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 876.1725   Gastrointestinal	motility monitoring system.

- Identification. A gastrointestinal motility monitoring system is a device used to measure peristalic activity or pressure in the stomach or esophagus by means of a probe with transducers that is introduced through the mouth into the gastrointestinal tract. The de- vice may include signal conditioning, amplifying, and recording equipment. This generic type of device includes the esophageal motility monitor and tube, the gastrointestinal motility (elec- trical) system, and certain accessories, such as a pressure transducer, ampli- fier, and external recorder.

- Classification. Class II (perform- ance standards).

##### § 876.1735   Electrogastrography system.

- Identification.	An electrogastrography system (EGG) is a device used to measure gastric myoelectrical activity as an aid in the diagnosis of gastric motility disorders. The device system includes the exter- nal recorder, amplifier, skin electrodes, strip chart, cables, analytical software, and other accessories.

- Classification. Class  II (Special Controls). The special controls are as follows:

- The sale, distribution and use of this device are  restricted  to  prescrip- tion use in accordance with § 801.109 of this   chapter.

- The labeling must include specific instructions:

  - To describe proper patient set-up prior to the start of the test, including the proper placement of electrodes;

  - To describe how background data should be gathered and used to elimi- nate artifact in the data signal;

  - To describe the test protocol (in- cluding the  measurement of baseline data) that may be followed to  obtain the EGG signal; and

  - To explain how data results may be interpreted.

- The device design should ensure that the EGG signal is distinguishable

#### Food and Drug Administration, HHS	§ 876.3350

from background noise that may inter- fere with the true gastric myoelectric signal.

- Data should be collected to dem- onstrate that the device has adequate precision and the EGG signal is repro- ducible and is interpretable.

[64 FR 51444, Sept. 23, 1999]

##### § 876.1800   Urine flow or volume meas- uring system.

- Identification. A urine flow or vol- ume measuring system is a device that measures directly or indirectly the vol- ume or flow of urine from a patient, ei- ther during the course of normal urina- tion or while the patient is catheter- ized. The device may include a drip chamber to reduce the risk of retro- grade bacterial contamination of the bladder and a transducer and electrical signal conditioning and display equip- ment. This generic type of device in- cludes the electrical urinometer, me- chanical urinometer, nonelectric uri- nometer, disposable nonelectric urine flow rate measuring device, and uroflowmeter.

- Classification. (1) Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR  1122,  Jan.  16,  1996;  63  FR  59228,  Nov.  3,

1998]

## Subpart C—Monitoring Devices

##### § 876.2040   Enuresis alarm.

- Identification. An enuresis alarm is a device intended for use in treat- ment of bedwetting. Through an elec- trical trigger mechanism, the device sounds an alarm when a small quantity of urine is detected on a sensing pad. This generic type of device includes conditioned response enuresis alarms.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 876.2050   Prostate  lesion  documenta- tion system.

- Identification. A prostate lesion documentation system is a prescription device intended for use in producing an image of the prostate as an aid in docu- menting prostate abnormalities pre- viously identified during a digital rec- tal examination. The device uses pres- sure sensors and image reconstruction software to produce a prostate image that highlights regional differences in intraprostatic tissue elasticity or stiff- ness. The device is limited to use as a documentation tool and is not intended for diagnostic purposes or for influ- encing any clinical decisions.

- Classification. Class II (special controls). The special controls for this device are:

- Non-clinical and clinical perform- ance testing must demonstrate the ac- curacy and reproducibility of the con- structed image.

- Appropriate analysis/testing must validate electromagnetic compat- ibility, electrical safety, thermal safe- ty, and mechanical safety.

- Appropriate software verification, validation,  and  hazard  analysis  must be performed.

- All elements of the device that may contact the patient must be dem- onstrated to be biocompatible.

- Methods and instructions for re- processing of any reusable components must be properly validated.

- The labeling must include specific information needed  to  ensure  proper use of the device.

[80 FR 72900, Nov. 23, 2015]

## Subpart D—Prosthetic Devices

##### § 876.3350   Penile inflatable implant.

  - Identification. A penile inflatable implant is a device that consists of two inflatable cylinders implanted in the penis, connected to a reservoir filled with radiopaque fluid implanted in the abdomen, and a subcutaneous manual pump implanted in the scrotum. When the cylinders are inflated, they provide rigidity to the penis. This device is used in the treatment of erectile impo- tence.

  - Classification. Class III (premarket approval).

§ 876.3630

  - Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before July 11, 2000, for any penile inflatable implant that was in commercial distribution before May 28, 1976, or that has, on or before July 11, 2000, been found to be substantially equivalent to a penile inflatable im- plant that was in commercial distribu- tion before May 28, 1976. Any other penile inflatable implant shall have an approved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR  17738,  May  11,  1987;  65  FR  19658,  Apr.  12,

2000]

##### § 876.3630   Penile rigidity implant.

- Identification. A penile rigidity implant is a device that consists of a pair of semi-rigid rods implanted in the corpora cavernosa of the penis to pro- vide rigidity. It is intended to be used in men diagnosed as having erectile dysfunction.

- Classification. Class II. The special control for this device is the FDA guid- ance entitled ‘‘Guidance for the Con- tent of Premarket Notifications for Penile Rigidity Implants.’’

[65 FR 4882, Feb. 2, 2000]

##### § 876.3750   Testicular prosthesis.

- Identification. A testicular pros- thesis is an implanted device that con- sists of a solid or gel-filled silicone rub- ber prosthesis that is implanted sur- gically to resemble a testicle.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of product development protocol (PDP) is required. A PMA or no- tice of completion of a PDP is required to be filed with the Food and Drug Ad- ministration on or before July 5, 1995, for any testicular prosthesis  that  was in commercial distribution before May 28, 1976, or that has on or before July 5, 1995, been found to be substantially equivalent to a testicular  prosthesis that was in commercial distribution before May 28, 1976. Any other testic-

#### 21 CFR Ch. I (4–1–17 Edition)

ular prosthesis shall have an approved PMA or a declared completed PDP in effect before being placed in commer- cial distribution.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR  17738,  May  11,  1987;  60  FR  17216,  Apr.  5,

1995]

## Subpart E—Surgical Devices

##### § 876.4020   Fiberoptic	light	ureteral catheter.

- Identification. A fiberoptic light ureteral catheter is a device that con- sists of a fiberoptic bundle that emits light throughout its length and is shaped so that it can be inserted into the ureter to enable the path of the ureter to be seen during lower abdom- inal or pelvic surgery.

- Classification. Class II (perform- ance standards).

##### § 876.4270   Colostomy rod.

- Identification. A colostomy rod is a device used during the loop colos- tomy procedure. A loop of colon is sur- gically brought out through the ab- dominal wall and the stiff colostomy rod is placed through the loop tempo- rarily to keep the colon from slipping back through the surgical opening.

- Classification. Class II (perform- ance standards).

##### § 876.4300   Endoscopic    electrosurgical unit and accessories.

- Identification. An endoscopic electrosurgical unit and accessories is a	device	used	to	perform electrosurgical procedures through an endoscope. This generic type of device includes the electrosurgical generator, patient plate, electric biopsy forceps, electrode,		flexible		snare, electrosurgical			alarm	system, electrosurgical power supply unit, elec- trical clamp, self-opening rigid snare, flexible suction coagulator electrode, patient return wristlet, contact jelly, adaptor to the cord for transurethral surgical instruments, the electric cord for transurethral surgical instruments, and the transurethral desiccator.

- Classification. Class II (perform- ance standards).

Food and Drug Administration, HHS	§ 876.4590

##### § 876.4370   Gastroenterology-urology evacuator.

  - Identification. A gastroenterology- urology evacuator is a device used to remove debris and fluids during gastro- enterological and urological procedures by drainage, aspiration, or irrigation. This generic type of device includes the fluid evacuator system, manually pow- ered bladder evacuator, and the AC- powered vacuum pump.

  - Classification. (1) Class II (special controls) for the gastroenterology-urol- ogy evacuator when other than manu- ally powered. The device is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 876.9.

(2) Class I for the gastroenterology- urology evacuator when manually pow- ered. The device subject to this para- graph (b)(2) is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR  25049,  June  12,  1989;  63  FR  59228,  Nov.  3,

1998]

##### § 876.4400   Hemorrhoidal ligator.

- Identification. A hemorrhoidal ligator is a device used to cut off the blood flow to hemorrhoidal tissue by means of a ligature or band placed around the hemorrhoid.

- Classification. Class II (perform- ance standards).

##### § 876.4480   Electrohydraulic lithotriptor.

- Identification. An electrohydraulic lithotriptor is an AC-powered device used to fragment urinary bladder stones. It consists of a high voltage source connected by a cable to a bipo- lar electrode that is introduced into the urinary bladder through a cysto- scope. The electrode is held against the stone in a water-filled bladder and re- peated electrical discharges between the two poles of the electrode cause electrohydraulic shock waves which disintegrate the stone.

- Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance for the Content of Premarket No-

tifications	for	Intracorporeal Lithotripters.’’

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR 17738,  May 11,  1987;  65  FR 17145,  Mar. 31,

2000]

##### § 876.4500   Mechanical lithotriptor.

- Identification. A mechanical lithotriptor is a device with steel jaws that is inserted into the urinary blad- der through the urethra to grasp and crush bladder stones.

- Classification. Class II (perform- ance standards).

##### § 876.4530   Gastroenterology-urology fiberoptic retractor.

- Identification. A gastroenterology- urology fiberoptic retractor is a device that consists of a mechanical retractor with a fiberoptic light system that is used to illuminate deep surgical sites.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25049, June 12, 1989; 66 FR 38801, July 25,

2001]

##### § 876.4560   Ribdam.

- Identification. A ribdam is a device that consists of a broad strip of latex with supporting ribs used to drain sur- gical wounds where copious urine drainage is expected.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25049, June 12, 1989; 66 FR 38801, July 25,

2001]

##### § 876.4590   Interlocking urethral sound.

- Identification. An interlocking urethral sound is a device that consists of two metal sounds (elongated instru- ments for exploring or sounding body cavities) with interlocking ends,  such as with male and female threads or a rounded point and mating socket, used in the repair of a ruptured urethra. The device may include a protective cap to fit over the metal threads.

#### § 876.4620

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 876.4620   Ureteral stent.

- Identification. A ureteral stent is a tube-like implanted device that is in- serted into the ureter to provide ureteral rigidity and allow the passage of urine. The device may have finger- like protrusions or hooked ends to keep the tube in place. It is used in the treatment of ureteral injuries and ureteral obstruction.

- Classification. Class II (perform- ance standards).

##### § 876.4650   Water	jet	renal	stone dislodger system.

- Identification. A water jet renal stone dislodger system is a device used to dislodge stones from renal calyces (recesses of the pelvis of the kidney) by means of a pressurized stream of water through a conduit. The device is used in the surgical removal of kidney stones.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 876.4680   Ureteral stone dislodger.

- Identification. A ureteral stone dislodger is a device that consists of a bougie or a catheter with an expand- able wire basket near the tip, a special flexible tip, or other special construc- tion. It is inserted through a cysto- scope and used to entrap and remove stones from the ureter. This generic type of device includes the metal bas- ket and the flexible ureteral stone dislodger.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

#### 21 CFR Ch. I (4–1–17 Edition)

in subpart E  of part  807  of  this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 876.4730 Manual gastroenterology- urology surgical instrument and ac- cessories.

- Identification. A manual gastro- enterology-urology surgical  instru- ment and accessories is a device de- signed to be used for gastro- enterological and urological surgical procedures. The device may be nonpow- ered, hand-held, or hand-manipulated. Manual gastroenterology-urology sur- gical instruments include the biopsy forceps cover, biopsy tray without bi- opsy instruments, line clamp, nonpow- ered rectal probe, nonelectrical clamp, colostomy spur-crushers, locking de- vice for intestinal clamp, needle  hold- er, gastro-urology hook, gastro-urology probe and director, nonself-retaining retractor, laparotomy rings, nonelec- trical snare, rectal specula, bladder neck spreader, self-retaining retractor, and scoop. A manual surgical instru- ment that is intended specifically for use as an aid in the insertion, place- ment, fixation, or anchoring of surgical mesh during urogynecologic procedures are classified under § 884.4910 of this chapter.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25049, June 12, 1989; 66 FR 38801, July 25,

2001; 82 FR 12171, Mar. 1, 2017]

##### § 876.4770   Urethrotome.

- Identification. A urethrotome is a device that is inserted into the urethra and used to cut urethral strictures and enlarge the urethra. It is a metal in- strument equipped with a dorsal-fin cutting blade which can be elevated from its sheath. Some urethrotomes in- corporate an optical channel for visual control.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 876.5011

##### § 876.4890   Urological  table  and  acces- sories.

  - Identification. A urological table and accessories is a device that con- sists of a table, stirrups, and belts used to support a patient in a suitable posi- tion for endoscopic procedures of the lower urinary tract. The table can be adjusted into position manually or electrically.

  - Classification. (1) Class II (special controls) for the electrically powered urological table and accessories. The device is exempt from the premarket notification procedures in subpart E of part   807   of   this   chapter   subject   to

§ 876.9.

- Class I for the manually powered table and accessories, and for stirrups for electrically powered table. The de- vice subject to this paragraph (b)(2) is exempt from the premarket  notifica- tion procedures in subpart E of part 807 of this chapter subject to the limita- tions in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR  1122,  Jan.  16,  1996;  63  FR  59228,  Nov.  3,

1998; 66 FR 38801, July 25, 2001]

## Subpart F—Therapeutic Devices

##### § 876.5010   Biliary  catheter  and  acces- sories.

  - Identification. A biliary catheter and accessories is a tubular flexible de- vice used for temporary or prolonged drainage of the biliary tract, for splinting of the bile duct during heal- ing, or for preventing stricture of the bile duct. This generic type of device may include a bile collecting bag that is attached to the biliary catheter by a connector and fastened to the patient with a strap.

  - Classification. Class II (perform- ance standards).

##### § 876.5011   Metallic biliary stent system for benign strictures.

- Identification. A metallic biliary stent system for benign strictures is a prescription device intended for the treatment of benign biliary strictures. The biliary  stents are  intended  to be left indwelling for a limited amount of time and subsequently removed. The device consists of a metallic stent and a delivery system intended to place the

biliary stent in the bile duct. This de- vice type is not intended for use in the vasculature.

- Classification. Class II (special controls). The special controls for this device are:

- Clinical performance testing must demonstrate or provide the following:

  - The ability to safely place and subsequently remove  the  stent  after the maximum labeled indwell period.

  - All adverse event data including bile duct obstruction and trauma to the bile duct.

  - The stent resolves strictures dur- ing the maximum labeled indwell pe- riod.

  - Stricture resolution is main- tained post-stent removal.

- Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be dem- onstrated:

  - Corrosion testing to demonstrate that the stent maintains its integrity during indwell and does not release po- tentially toxic levels of leachables.

  - Stent dimensional testing sup- ports the intended use.

  - Compression and expansion forces must be characterized.

  - The delivery catheter must de- liver the stent to the intended location and the stent must not be adversely impacted by the delivery catheter dur- ing deployment and catheter with- drawal.

  - The delivery  system  must  with- stand clinically anticipated forces.

  - Compatibility in a magnetic res- onance environment.

- All patient contacting compo- nents of the device must be dem- onstrated to be biocompatible.

- Performance data must dem- onstrate the sterility of the device components intended to be provided sterile.

- Shelf life testing must dem- onstrate that the device maintains its performance characteristics and that packaging maintains sterility for the duration of the labeled shelf life.

- Labeling for the device must in- clude:

  - A detailed summary of the clinical testing including device effectiveness,

§ 876.5015

and  device-  and  procedure-related  ad- verse events.

  - Appropriate warning(s) to accu- rately ensure usage of the device  for the intended patient population.

  - Shelf life.

  - Compatibility information for use in the magnetic resonance environ- ment.

  - Stent foreshortening information supported by dimensional testing.

[81 FR 45231, July 13, 2016]

##### § 876.5015   Pancreatic   drainage   stent and delivery system.

- Identification. A pancreatic drain- age stent is a prescription device that consists of a self-expanding, covered, metallic stent, intended for  placement to facilitate transmural endoscopic drainage of  pancreatic  pseudocysts. This stent is intended to be removed upon confirmation of pseudocyst reso- lution. This device may also include a delivery  system.

- Classification. Class II (special controls). The special controls for this device are:

- The device and elements of the de- livery device that may contact the pa- tient must  be demonstrated  to be  bio- compatible.

- Performance data must dem- onstrate the sterility of patient-con- tacting components of the device.

- Performance data must support the shelf life of the device by dem- onstrating continued sterility, package integrity, and device  functionality over the requested shelf life.

- Non-clinical testing data must demonstrate that the stent and deliv- ery system perform as intended under anticipated conditions of use. The fol- lowing performance characteristics must be tested:

  - Deployment testing of the stent and delivery system must be conducted under simulated use conditions.

  - Removal force testing must be conducted. The removal force testing must demonstrate that the stent  can be safely removed, and that the stent will remain in place when subjected to forces encountered during use.

  - Expansion force testing must be conducted. The expansion force must demonstrate that the forces exerted by

#### 21 CFR Ch. I (4–1–17 Edition)

the  stent  will  not  damage  the  tissue surrounding the stent.

  - Compression force testing must be conducted. The compression force must demonstrate that the stent will withstand the forces encountered dur- ing use.

  - Dimensional verification testing must be conducted.

  - Tensile testing of joints and ma- terials must be conducted. The min- imum acceptance criteria must be ade- quate for its intended use.

  - Fatigue testing must be con- ducted. Material strength must dem- onstrate that the stent will withstand forces encountered during use.

  - Corrosion testing must be con- ducted. Corrosion resistance must dem- onstrate that the stent will withstand conditions encountered during use.

- Non-clinical testing must evalu- ate the compatibility of the stent in a magnetic resonance (MR) environment.

- Well-documented clinical experi- ence must demonstrate safe and effec- tive use, and capture any adverse events observed during clinical use.

- Labeling must include the fol- lowing:

  - Appropriate instructions, warn- ings, cautions, limitations, and infor- mation related to the safe use of the device, including deployment of the de- vice, maintenance of the drainage lumen, and removal of the device.

  - A warning that the safety and pa- tency of the stent has not been estab- lished beyond the duration of the docu- mented clinical experience.

  - Specific instructions and the qualifications and clinical training needed for the safe use of the device, including deployment of the device, maintenance of the drainage lumen, and removal of the device.

  - Information on the patient popu- lation for which the device has been demonstrated to be effective.

  - A detailed summary of the clin- ical experience pertinent to use of the device.

  - A detailed summary of the device technical parameters.

  - A detailed summary of the device- and procedure-related com- plications pertinent to use of the de- vice.

#### Food and Drug Administration, HHS	§ 876.5130

  - An expiration date/shelf life.

[79 FR 30724, May 29, 2014]

##### § 876.5020   External  penile  rigidity  de- vices.

- Identification. External penile ri- gidity devices are devices intended to create or maintain sufficient penile ri- gidity for sexual intercourse. External penile rigidity devices include vacuum pumps, constriction rings, and penile splints which are mechanical, powered, or pneumatic devices.

- Classification. Class II (special controls). The devices are exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9. The special control for these devices is the FDA  guidance  document   entitled ‘‘Class II Special Controls Guidance Document: External Penile Rigidity Devices.’’ See § 876.1(e) for the avail- ability of this guidance document.

[69 FR 77623, Dec. 28, 2004]

##### § 876.5025   Vibrator  for  climax  control of premature ejaculation.

- Identification. A vibrator for cli- max control of premature  ejaculation is used for males who suffer from pre- mature ejaculation. It is designed to increase the time between arousal and ejaculation using the stimulating vi- bratory effects of the device on the penis.

- Classification. Class II (special controls). The special controls for this device are:

- The labeling must include specific instructions regarding the proper placement and use of the device.

- The portions of the device that contact the patient must be dem- onstrated to be biocompatible.

- Appropriate analysis/testing must demonstrate electromagnetic compat- ibility safety, electrical safety, and thermal safety of the device.

- Mechanical safety testing must demonstrate that the device will with- stand forces encountered during use.

[80 FR 30355, May 28, 2015]

##### § 876.5030   Continent   ileostomy   cath- eter.

  - Identification. A continent ileos- tomy catheter is a flexible tubular de-

vice used as a form during surgery for continent ileostomy and it provides drainage after surgery. Additionally, the device may be inserted periodically by the patient for routine care to empty the ileal pouch. This generic type of device includes the rectal cath- eter for continent ileostomy.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25050, June 12, 1989; 66 FR 38801, July 25,

2001]

##### § 876.5090   Suprapubic urological cath- eter and accessories.

- Identification. A suprapubic urological catheter and accessories is a flexible tubular device that is inserted through the abdominal wall into the urinary bladder with the aid of a trocar and cannula. The device is used to pass fluids to and from the urinary tract. This generic type of device includes the suprapubic catheter and tube, Malecot catheter, catheter punch instrument, suprapubic drainage tube, and the suprapubic cannula and trocar.

- Classification. (1)  Class  II  (per- formance standards).

- Class I for the catheter punch in- strument, nondisposable cannula and trocar,  and  gastro-urological  trocar. The devices subject to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 876.5130   Urological  catheter  and  ac- cessories.

  - Identification. A urological cath- eter and accessories is a flexible tubu- lar device that is inserted through the urethra and used to pass fluids to or from the urinary tract. This generic type of device includes radiopaque urological catheters, ureteral cath- eters, urethral catheters, coude´ cath- eters, balloon retention type catheters, straight catheters, upper urinary tract catheters, double lumen female urethrographic   catheters,   disposable

#### § 876.5140

ureteral	catheters,	male urethrographic		catheters,	and urological catheter accessories includ- ing ureteral catheter stylets, ureteral catheter adapters, ureteral catheter holders, ureteral catheter  stylets, ureteral catheterization trays, and the gastro-urological irrigation tray (for urological use).

  - Classification. (1)  Class  II  (per- formance standards).

- Class I for the ureteral stylet (guidewire), stylet for gastrourological catheter, ureteral catheter adapter, ureteral catheter connector, and ureteral catheter holder. The devices subject to this paragraph (b)(2) are ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject  to  the  limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38801,  July 25,

2001]

##### § 876.5140   Urethral  insert  with  pump for bladder drainage.

  - Identification. A urethral insert with pump for bladder drainage is a catheter-like device with internal pump mechanism that is placed in the urethra. Under patient control the in- ternal pump draws urine out of the bladder when voiding is desired, and blocks urine flow when continence is desired. The device is intended for use by women who cannot empty their bladder due to impaired detrusor con- tractility.

  - Classification. Class II (special controls). The special controls for this device are:

- The elements of the device that may contact the urinary tract must be demonstrated to be biocompatible.

- Performance data must dem- onstrate the sterility of the device components that contact the urinary tract.

- Performance data must support shelf life by demonstrating continued sterility of the device (or the sterile components), package integrity, and device functionality over the requested shelf life.

- Non-clinical testing data must demonstrate that the device  performs as  intended  under  anticipated  condi-

#### 21 CFR Ch. I (4–1–17 Edition)

tions of use. The following performance characteristics must be tested:

- Urine flow rate testing.

- Valve integrity testing.

- Bladder neck retention force testing.

- Pump/valve endurance testing.

- Encrustation testing.

- Remote control reliability, me- chanical integrity,  and battery  life testing.

- Clinical testing must demonstrate safe and effective use, document the device acceptance rate and the adverse event profile associated with  clinical use, and demonstrate that the device performs as intended under anticipated conditions of use.

- Labeling must include:

- Specific instructions, contra- indications, warnings, cautions, limita- tions, and the clinical training needed for the safe use of the device.

- Statement of the maximum in- sert indwelling period.

- Information on the patient edu- cation and support program prior to and during initial device use.

- Information on the patient popu- lation for which the device has been demonstrated to be safe and effective.

- Information on how the device op- erates and the recommended treatment regimen.

- A detailed summary of the device- and procedure-related com- plications or adverse events  pertinent to use of the device.

- An expiration date/shelf life.

- Patient labeling must be provided and must include:

  - Relevant contraindications, warn- ings, precautions, and adverse events/ complications.

  - Information on how the device operates and the recommended treat- ment regimen.

  - Information on the patient edu- cation and support program prior to and during initial device use.

  - Information on the patient popu- lation for which there is clinical evi- dence of safety and effectiveness.

  - The potential risks and benefits associated with the use of the device.

  - Post-insertion care instructions.

  - Alternative  treatments.

[80 FR 18309, Apr. 6, 2015]

#### Food and Drug Administration, HHS	§ 876.5250

##### § 876.5160   Urological clamp for males.

    - Identification. A urological clamp for males is a device used to close the urethra of a male to control urinary in- continence or to hold anesthetic or ra- diography contrast media in the ure- thra temporarily. It is an external clamp.

    - Classification. Class I (general con- trols). Except when intended for inter- nal use or use on females, the device is exempt from the premarket  notifica- tion procedures in subpart E of part 807 of this chapter subject to § 876.9.

[48 FR 53023, Nov. 23, 1963, as amended at 65

FR 2317, Jan. 14, 2000]

##### § 876.5210   Enema kit.

- Identification. An enema kit is a device intended to instill water or other fluids into the colon through a nozzle inserted into the rectum to pro- mote evacuation of the contents of the lower colon. The device consists of a container for fluid connected to the nozzle either directly or via tubing. This device does not include the co- lonic irrigation system (§ 876.5220).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 876.9. The device is exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180 of this chapter, with respect to general requirements  concerning  records,  and

§ 820.198 of this chapter, with respect to complaint    files.

[48 FR 53023, Nov. 23, 1963, as amended at 65

FR 2317, Jan. 14, 2000]

##### § 876.5220   Colonic irrigation system.

- Identification. A colonic irrigation system is a device intended to instill water into the colon through a nozzle inserted into the rectum to cleanse (evacuate) the contents of the lower colon. The system is designed to allow evacuation of the contents of the colon during the administration of the co- lonic irrigation. The device consists of a container for fluid connected to the nozzle via tubing and includes a system which enables the pressure, tempera- ture, or flow of water through the noz-

zle to be controlled. The device  may in- clude a console-type toilet and nec- essary fittings to allow the device to be connected to water  and  sewer  pipes. The device may use electrical power to heat the water. The device does not in- clude the enema kit (§ 876.5210).

- Classification. (1)  Class  II  (per- formance standards) when the device is intended for colon cleansing when medically indicated,  such  as  before  ra- diological  or  endoscopic  examinations.

(2) Class III (premarket approval) when the device is intended for other uses, including colon cleansing rou- tinely for general well being.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any colonic irrigation system de- scribed in paragraph (b)(2) of this sec- tion that was  in  commercial  distribu- tion before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equivalent to a co- lonic irrigation system described in paragraph (b)(2) of  this  section  that was in commercial distribution  before May 28, 1976. Any other colonic irriga- tion system shall have  an  approved PMA in effect before being placed in commercial     distribution.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR 17738, May 11, 1987; 61 FR 50707, Sept. 27,

1996]

##### § 876.5250   Urine  collector  and  acces- sories.

- Identification. A urine collector and accessories is a device intended to collect urine. The device and acces- sories consist of tubing, a suitable re- ceptacle, connectors, mechanical sup- ports, and may include a means to pre- vent the backflow of urine or ascent of infection. The two kinds of urine col- lectors are:

- A urine collector and accessories intended to be connected to an indwell- ing catheter, which includes the uri- nary drainage collection kit and the closed urine drainage system and drainage bag; and

- A urine collector and accessories not intended to be connected to an in- dwelling catheter, which includes the corrugated   rubber   sheath,   pediatric

#### § 876.5270

urine collector, leg bag for external use, urosheath type incontinence de- vice, and the paste-on device for incon- tinence.

- Classification—(1) Class II (special controls) for a urine collector and acces- sories intended to be connected to an in- dwelling catheter. The device is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 876.9.

  - Class I (general controls).  For  a urine collector and accessories not in- tended to be connected to an  indwell- ing catheter, the device is exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.  If the device is not labeled or otherwise represented  as  sterile,  it  is   exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[48 FR 53023, Nov. 23, 1983, as amended at 63

FR  59228,  Nov.  3,  1998;  65  FR  2317,  Jan.  14,

2000; 66 FR 38802, July 25, 2001; 73 FR 34860,

June 19, 2008]

##### § 876.5270   Implanted electrical urinary continence device.

    - Identification. An implanted elec- trical urinary device is a device in- tended for treatment of urinary incon- tinence that consists of a receiver im- planted in the abdomen with electrodes for pulsed-stimulation that are im- planted either in the bladder wall or in the pelvic floor, and a battery-powered transmitter outside the body.

    - Classification. Class III (premarket approval).

    - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any implanted electrical urinary continence device that was in commer- cial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to an implanted electrical urinary continence device that was in commer- cial distribution before May 28, 1976. Any   other   implanted   electrical   urinary

#### 21 CFR Ch. I (4–1–17 Edition)

continence device shall have an ap- proved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR 17738, May 11, 1987; 61 FR 50707, Sept. 27,

1996]

##### § 876.5280   Implanted	mechanical/hy- draulic urinary continence device.

- Identification. An implanted me- chanical/hydraulic urinary continence device is a device used to treat urinary incontinence by the application of con- tinuous or intermittent pressure to oc- clude the urethra. The totally  im- planted device may consist of a static pressure pad, or a system with a con- tainer of radiopaque fluid in the abdo- men and a manual pump and valve under the skin surface that is con- nected by tubing to an adjustable pres- sure pad or to a cuff around the ure- thra. The fluid is pumped as needed from the container to inflate the pad or cuff to pass on the urethra.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 2000, for any implanted mechanical/hydrau- lic urinary continence device  that  was in commercial distribution before May 28, 1976, or that has, on or before De- cember 26, 2000, been found to be sub- stantially equivalent to an implanted mechanical/hydraulic urinary con- tinence device that was in commercial distribution before May 28, 1976. Any other implanted mechanical/hydraulic urinary  continence  device  shall  have an approved PMA or a declared com- pleted PDP in  effect  before  being placed in commercial distribution.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR 17738, May 11, 1987; 65 FR 57731, Sept. 26,

2000]

##### § 876.5310   Nonimplanted,	peripheral electrical continence device.

- Identification. A nonimplanted, pe- ripheral electrical continence device is a device that consists of an electrode that is connected by an electrical cable to a battery-powered pulse source. The electrode  is  placed  onto  or  inserted

#### Food and Drug Administration, HHS	§ 876.5520

into the body at a peripheral location and used to stimulate the nerves asso- ciated with pelvic floor function to maintain urinary continence. When necessary, the electrode may be re- moved by the user.

- Classification. Class II, subject to the following special controls:

  - That sale, distribution, and use of this device are  restricted  to  prescrip- tion use in accordance with § 801.109 of this   chapter.

  - That the labeling must bear all in- formation required for the safe and ef- fective use of the device as outlined in

§ 801.109(c) of this chapter, including a detailed summary of the clinical infor- mation upon  which  the  instructions are based.

[65 FR 18237, Apr. 7, 2000]

##### § 876.5320   Nonimplanted	electrical continence device.

- Identification. A nonimplanted electrical continence device is a device that consists  of  a  pair  of  electrodes  on a plug or a  pessary  that  are  connected by an electrical cable to a battery-pow- ered pulse source.  The  plug  or  pessary is inserted into the rectum or into the vagina and used to stimulate the mus- cles of the pelvic floor to maintain uri- nary or fecal continence. When nec- essary, the plug or pessary may be re- moved by the user. This device  ex- cludes an AC-powered nonimplanted electrical continence device and the powered vaginal muscle stimulator for therapeutic use (§ 884.5940).

- Classification. Class II (perform- ance standards).

##### § 876.5365   Esophageal dilator.

- Identification. An esophageal dila- tor is a device that consists of a cylin- drical instrument that may be hollow and weighted with mercury or a metal olive-shaped weight that slides on a guide, such as a string or wire and is used to dilate a stricture of the esoph- agus. This generic type of device in- cludes esophageal or gastrointestinal bougies and the esophageal dilator (metal olive).

- Classification. Class II (perform- ance standards).

##### § 876.5450   Rectal dilator.

- Identification. A rectal dilator is a device designed to dilate the anal sphincter and canal when the size of the anal opening may interfere with its function or the passage of an exam- ining instrument.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38802,  July 25,

2001]

##### § 876.5470   Ureteral dilator.

- Identification. A ureteral dilator is a device that consists of a specially shaped catheter or bougie and is used to dilate the ureter at the place where a stone has become lodged or to dilate a ureteral stricture.

- Classification. Class II (perform- ance standards).

##### § 876.5520   Urethral dilator.

- Identification. A urethral dilator is a device that consists of a slender hol- low or solid instrument made of metal, plastic, or other suitable material in a cylindrical form and in a range of sizes and flexibilities. The device may in- clude a mechanism to expand the por- tion of the device in the urethra and indicate the degree of expansion on a dial. It is used to dilate the urethra. This generic type of device includes the mechanical urethral dilator, urological bougies, metal or plastic urethral sound, urethrometer, filiform, and fili- form follower.

- Classification. (1)  Class  II  (per- formance standards).

- Class I for the urethrometer, urological bougie, filiform and filiform follower, and metal or plastic urethral sound. The devices subject to this para- graph (b)(2) are exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter sub- ject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 61

FR 1122,  Jan. 16,  1996;  66  FR 38802,  July 25,

2001]

#### § 876.5530

##### § 876.5530   Implantable	transprostatic tissue retractor system.

  - Identification. An implantable transprostatic tissue retractor system is a prescription use device that con- sists of a delivery device and implant. The delivery device is inserted transurethrally and deploys the im- plant through the prostate. It is de- signed to increase prostatic urethral patency by providing prostate lobe tis- sue retraction while preserving the po- tential for future prostate procedures and is intended for the treatment of symptoms due to urinary outflow ob- struction secondary to benign prostatic hyperplasia in men.

  - Classification. Class II (special controls). The special controls for this device are:

- The elements of the device that may contact the patient must be dem- onstrated to be biocompatible.

- Performance data must dem- onstrate the sterility of the patient- contacting components of the device.

- Performance data must support shelf life by demonstrating continued sterility of the device (of the patient- contacting components), package in- tegrity, and device functionality over the requested shelf life.

- Non-clinical testing data must demonstrate that the device performs as intended under anticipated condi- tions of use. The following performance characteristics must be tested:

- Deployment testing must be con- ducted.

- Mechanical strength must be con- ducted.

- Resistance-to-degradation test- ing must be conducted.

- Non-clinical testing must evalu- ate the compatibility of the device in a magnetic resonance environment.

- In vivo testing must demonstrate safe and effective use, assess the im- pact of the implants on the ability to perform subsequent treatments, docu- ment the adverse event profile associ- ated with clinical use, and demonstrate that the device performs as intended under anticipated conditions of use. The following performance characteris- tics must be tested:

  - Deployment testing must be con- ducted.

#### 21 CFR Ch. I (4–1–17 Edition)

  - Implant migration must be con- ducted.

- Labeling must bear all informa- tion required for safe and effective use of the device, and must include:

  - Specific instructions, warnings, cautions, limitations, and the clinical training needed for the safe use of the device.

  - Information on the patient popu- lation for which the device has been demonstrated to be effective.

  - A detailed summary of the de- vice  technical  parameters.

  - Information on how the device operates and the typical course of treatment.

  - An expiration date/shelf life.

  - A detailed summary of the device- and procedure-related com- plications or adverse events  pertinent to use of the device.

[79 FR 43249, July 25, 2014]

##### § 876.5540   Blood access device and ac- cessories.

- Identification. A blood access de- vice and accessories is a device in- tended to provide access to a patient’s blood for hemodialysis or other chronic uses. When used in hemodialysis, it is part of an artificial kidney system for the treatment of patients with renal failure or toxemic conditions and pro- vides access to a patient’s blood for hemodialysis. The device includes im- planted blood access devices, non- implanted blood access devices, and ac- cessories for both the implanted and nonimplanted blood access devices.

  - The implanted blood access device is a prescription device and consists of various flexible or rigid tubes, such as catheters, or cannulae, which are sur- gically implanted in appropriate blood vessels, may come through the skin, and are intended to remain in the body for 30 days or more. This generic type of device includes various catheters, shunts, and connectors specifically de- signed to provide access to blood. Ex- amples include single and double lumen catheters with cuff(s), fully sub- cutaneous port-catheter systems, and A–V shunt cannulae (with vessel tips). The implanted blood access device may also contain coatings or additives which may provide additional functionality to the device.

#### Food and Drug Administration, HHS	§ 876.5540

  - The nonimplanted blood access device consists of various  flexible  or rigid tubes, such as  catheters,  cannulae or hollow needles, which are inserted into appropriate blood vessels or a vas- cular graft prosthesis (§§ 870.3450 and 870.3460), and are intended to remain in the body for less than 30 days. This ge- neric type of device includes fistula needles, the single needle dialysis set (coaxial flow needle), and the single needle dialysis set (alternating flow needle).

  - Accessories common to either type include the shunt adaptor, cannula clamp, shunt connector, shunt stabilizer, vessel dilator, disconnect forceps, shunt guard, crimp plier, tube plier, crimp ring, joint ring, fistula adaptor, and declotting tray (including contents).

- Classification. (1) Class II (special controls) for the implanted blood ac- cess device.  The  special  controls  for this device are:

- Components of the device that come into human contact must be dem- onstrated to be biocompatible. Mate- rial names and specific designation numbers must be provided.

- Performance data must dem- onstrate that the device performs as intended under anticipated conditions of use. The following performance characteristics must be tested:

  - Pressure versus flow rates for both arterial and venous lumens, from the minimum flow rate to the max- imum flow rate in 100 milliliter per minute increments,  must  be  estab- lished. The  fluid and  its viscosity  used during testing must be stated.

  - Recirculation rates for both for- ward and reverse flow configurations must be established, along with the protocol used to perform the assay, which must be provided.

  - Priming volumes must be estab- lished.

  - Tensile testing of joints and ma- terials must be conducted. The min- imum acceptance criteria must be ade- quate for its intended use.

  - Air leakage testing and liquid leakage testing must be conducted.

  - Testing of the repeated  clamping of the extensions of the catheter that simulates  use  over  the  life  of  the  de-

vice  must  be  conducted,  and  retested for  leakage.

  - Mechanical hemolysis testing must be conducted for new or altered device designs that affect the blood flow pattern.

  - Chemical tolerance of the device to repeated exposure to commonly used disinfection agents must be estab- lished.

- Performance data must dem- onstrate the sterility of the device.

- Performance data must support the shelf life of the device for  contin- ued sterility, package integrity, and functionality over the requested shelf life that must include tensile, repeated clamping,  and  leakage  testing.

- Labeling of implanted blood ac- cess devices for hemodialysis must in- clude the following:

  - Labeling must provide arterial and venous pressure versus flow rates, either in tabular or graphical format. The fluid and its viscosity used during testing must be stated.

  - Labeling must specify the for- ward and reverse recirculation rates.

  - Labeling must provide the arte- rial and venous priming volumes.

  - Labeling must specify an expira- tion date.

  - Labeling must identify any dis- infecting agents that cannot be used to clean any components of the device.

  - Any contraindicated disinfecting agents due to material incompatibility must be identified by printing a warn- ing on the catheter. Alternatively, con- traindicated disinfecting agents must be identified by a label affixed to the patient’s medical record and with writ- ten instructions provided directly to the patient.

  - Labeling must include a patient implant card.

  - The labeling must contain com- prehensive instructions for the fol- lowing:

    - Preparation and insertion of the device, including recommended site of insertion, method of insertion, and a reference on the proper location for tip placement;

    - Proper care and maintenance of the device and device exit site;

    - Removal of the device;

    - Anticoagulation;

#### § 876.5600

    - Management of obstruction and thrombus formation; and

    - Qualifications for clinical pro- viders performing the insertion, main- tenance, and removal of the devices.

- In addition to Special Controls in paragraphs (b)(1)(i) through (v) of this section, implanted blood access devices that include subcutaneous ports must include the following:

  - Labeling must include the rec- ommended type of needle for access as well as detailed instructions for care and maintenance of the port, subcuta- neous pocket, and skin overlying the port.

  - Performance testing must include results on repeated use of the ports that simulates use over the intended life of the device.

  - Clinical performance testing must demonstrate safe  and  effective use and capture any adverse events ob- served during clinical use.

- In addition to Special  Controls in paragraphs (b)(1)(i) through (v) of this section, implanted blood access de- vices with coatings or additives must include the following:

- A description and material char- acterization of the coating or additive material, the purpose of the coating or additive, duration of effectiveness, and how and where the coating is applied.

- An identification in the  labeling of any coatings or additives and a sum- mary of the results of performance testing for  any  coating  or  material with special characteristics, such as decreased thrombus formation or anti- microbial properties.

- A Warning Statement in the la- beling for potential allergic reactions including anaphylaxis if the coating or additive contains known allergens.

- Performance data must dem- onstrate efficacy of the coating or ad- ditive and the duration of  effective- ness.

- The following must be included for A–V shunt cannulae (with vessel tips):

- The device must comply with Special Controls in paragraphs (b)(1)(i) through (v) of this section with the ex- ception of paragraphs (b)(1)(ii)(B), (b)(1)(ii)(C), (b)(1)(v)(B), and (b)(1)(v)(C), which do not apply.

#### 21 CFR Ch. I (4–1–17 Edition)

- Labeling must include Warning Statements to address the potential for vascular access steal syndrome, arte- rial stenosis, arterial thrombosis, and hemorrhage including exsanguination given that the device accesses the arte- rial  circulation.

- Clinical performance testing must demonstrate safe  and  effective use and capture any adverse events ob- served during clinical use.

- Class II (performance standards) for the nonimplanted blood access de- vice.

- Class II (performance standards) for accessories for both the implanted and the nonimplanted blood access de- vices not listed in paragraph (b)(4) of this  section.

- Class I for the cannula clamp, dis- connect forceps, crimp plier, tube plier, crimp ring, and joint ring, accessories for both the implanted and non- implanted blood access device. The de- vices subject to this paragraph (b)(4) are exempt from the premarket notifi- cation procedures in subpart E of part 807 of this chapter subject to the limi- tations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR  17738,  May  11,  1987;  61  FR  1122,  Jan.  16,

1996; 66 FR 38802, July 25, 2001; 79 FR 43245,

July 25, 2014]

##### § 876.5600   Sorbent	regenerated dialysate delivery system for hemo- dialysis.

- Identification. A sorbent regen- erated dialysate delivery system for hemodialysis is a device that is part of an artificial kidney system for the treatment of patients with renal fail- ure or toxemic conditions, and that consists of a sorbent cartridge and the means to circulate  dialysate  through this cartridge and the dialysate com- partment of the dialyzer. The device is used with the extracorporeal blood sys- tem and the dialyzer of the hemo- dialysis      system      and      accessories (§ 876.5820). The device includes the means to maintain the temperature, conductivity, electrolyte balance, flow rate and pressure of the dialysate, and alarms to indicate abnormal dialysate conditions. The sorbent cartridge may include absorbent, ion exchange and catalytic    materials.

#### Food and Drug Administration, HHS	§ 876.5820

- Classification. Class II (perform- ance standards).

##### § 876.5630   Peritoneal   dialysis   system and accessories.

- Identification. (1) A peritoneal di- alysis system and accessories is a de- vice that is used as an artificial kidney system for the treatment of patients with renal failure or toxemic condi- tions, and that consists of a peritoneal access device, an administration set for peritoneal dialysis, a source of dialysate, and, in some cases, a water purification mechanism. After the dialysate is instilled into the patient’s peritoneal cavity, it is allowed to dwell there so that undesirable substances from the patient’s blood pass through the lining membrane of the peritoneal cavity into this dialysate. These sub- stances are then removed when the dialysate is drained from the patient. The peritoneal dialysis system may regulate and monitor the dialysate temperature, volume, and delivery rate together with the time course of each cycle of filling, dwell time, and drain- ing of the peritoneal cavity or manual controls may be used. This generic de- vice includes the semiautomatic and the automatic peritoneal delivery sys- tem.

  - The peritoneal access device is a flexible tube that is implanted through the abdominal  wall into  the peritoneal cavity and that may  have  attached cuffs to provide anchoring and a skin seal. The device is either a single use peritioneal catheter, intended to re- main in the peritoneal cavity for less than 30 days, or a long term peritoneal catheter. Accessories  include  stylets and trocars to aid in the  insertion  of the catheter and an obturator to main- tain the patency of  the  surgical  fistula in the abdominal wall between treat- ments.

  - The disposable administration set for peritoneal dialysis consists of tub- ing, an optional reservoir bag, and ap- propriate connectors. It may include a peritoneal dialysate filter to trap and remove contaminating particles.

  - The source of dialysate may be sterile prepackaged dialysate (for semi- automatic peritoneal dialysate  deliv- ery systems or ‘‘cycler systems’’) or dialysate  prepared  from  dialysate  con-

centrate and sterile purified water (for automatic peritoneal dialysate deliv- ery systems or ‘‘reverse osmosis’’ sys- tems). Prepackaged dialysate intended for use with either of the peritoneal dialysate delivery systems is regulated by FDA as a drug.

- Classification. Class II (perform- ance standards).

##### § 876.5665   Water   purification   system for hemodialysis.

- Identification. A water purification system for hemodialysis is a device that is intended for use with a hemo- dialysis system and that is intended to remove organic and inorganic sub- stances and microbial contaminants from water used to dilute dialysate concentrate to form dialysate. This ge- neric type of device may include a water softener, sediment filter, carbon filter, and water distillation system.

- Classification. Class II (perform- ance standards).

##### § 876.5820   Hemodialysis system and ac- cessories.

- Identification. A hemodialysis sys- tem and accessories is a device that is used as an artificial kidney system for the treatment of patients with renal failure or toxemic conditions and that consists of an extracorporeal blood sys- tem, a conventional dialyzer, a dialysate delivery system, and acces- sories. Blood from a patient flows through the tubing of the extracorporeal blood system and acces- sories to the blood compartment of the dialyzer, then returns through further tubing of the extracorporeal blood sys- tem to the patient. The dialyzer has two compartments that are separated by a semipermeable membrane. While the blood is in the blood compartment, undesirable substances in the blood pass through the semipermeable mem- brane into the dialysate in the dialysate compartment. The dialysate delivery system controls and monitors the dialysate circulating through the dialysate compartment of the dialyzer.

  - The extracorporeal blood system and accessories consists of tubing, pumps, pressure monitors, air foam or bubble  detectors,  and  alarms  to  keep

#### § 876.5830

blood moving safely from the blood ac- cess device and accessories for hemo- dialysis (§ 876.5540) to the blood com- partment  of  the  dialyzer  and  back   to the   patient.

  - The conventional  dialyzer  allows a transfer of water and solutes between the  blood  and  the  dialysate  through the semipermeable membrane. The semipermeable membrane of the con- ventional dialyzer has  a  sufficiently low permeability to water that an ultrafiltration controller is not re- quired to prevent excessive loss of water from the patient’s blood. This conventional dialyzer does not include hemodialyzers with the disposable in- serts (Kiil type) (§ 876.5830) or dialyzers of high permeability (§ 876.5860).

  - The dialysate delivery system consists of mechanisms  that  monitor and control the temperature, conduc- tivity, flow rate, and pressure of the dialysate  and  circulates  dialysate through the  dialysate  compartment  of the dialyzer. The dialysate delivery system includes the dialysate  con- centrate for hemodialysis (liquid or powder) and alarms to indicate abnor- mal   dialysate   conditions.    This dialysate delivery system does not in- clude the sorbent regenerated dialysate delivery      system      for      hemodialysis (§ 876.5600), the dialysate delivery sys- tem of  the  peritoneal  dialysis  system and accessories (§ 876.5630), or the con- trolled dialysate delivery system of the high   permeability   hemodialysis   system

§ 876.5860).

  - Remote accessories to the hemo- dialysis system include the unpowered dialysis chair without a scale, the pow- ered dialysis chair without a scale, the dialyzer holder set, dialysis tie gun and ties, and hemodialysis start/stop tray.

- Classification. (1)  Class  II  (per- formance standards) for hemodialysis systems and all accessories directly as- sociated with the extracorporeal blood system and the dialysate delivery sys- tem.

(2) Class I for other accessories of the hemodialysis system remote from the extracorporeal blood  system and  the dialysate delivery system, such as the unpowered dialysis chair, hemodialysis start/stop tray, dialyzer holder set, and dialysis tie gun and ties. The devices subject to this paragraph (b)(2) are ex-

#### 21 CFR Ch. I (4–1–17 Edition)

empt from the premarket notification procedures in subpart E of part 807 of this chapter subject  to  the  limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25050, June 12, 1989; 66 FR 38802, July 25,

2001]

##### § 876.5830   Hemodialyzer   with   dispos- able insert (Kiil type).

- Identification. A hemodialyzer with disposable inserts (Kiil type) is a device that is used as a part of an arti- ficial kidney system for  the  treatment of patients with renal failure or tox- emic conditions and that includes dis- posable inserts consisting of layers of semipermeable membranes which are sandwiched between support plates. The device is used with the extracorporeal blood  system and  the dialysate delivery system of the hemo- dialysis      system      and       accessories (§ 876.5820).

- Classification. Class II (perform- ance standards).

[48 FR 53023, Nov. 23, 1983, as amended at 53

FR 11253, Apr. 6, 1988]

##### § 876.5860   High	permeability	hemo- dialysis system.

- Identification. A high permeability hemodialysis system is a device in- tended for use as an artificial kidney system for the treatment of patients with renal failure, fluid overload, or toxemic conditions by performing such therapies	as	hemodialysis, hemofiltration,    hemoconcentration, and hemodiafiltration. Using a hemodialyzer with a semipermeable membrane that is more permeable to water than the semipermeable mem- brane of the conventional hemodialysis system (§ 876.5820), the high perme- ability hemodialysis system removes toxins or excess fluid from the  pa- tient’s blood using the principles of convection (via a high ultrafiltration rate) and/or diffusion (via a concentra- tion gradient in dialysate). During treatment, blood is circulated from the patient through the hemodialyzer’s blood compartment, while  the dialysate solution flows countercurrent through the dialysate compartment. In this process, toxins and/or fluid are transferred across the membrane from

#### Food and Drug Administration, HHS	§ 876.5870

the blood to the dialysate compart- ment. The hemodialysis delivery ma- chine controls and monitors the pa- rameters related to this processing, in- cluding the rate at which blood and dialysate are pumped through the sys- tem, and the rate at which fluid is re- moved from the patient. The high per- meability hemodialysis system con- sists of the following devices:

  - The hemodialyzer consists of a semipermeable membrane with an in vitro ultrafiltration coefficient (Kuf) greater than 8 milliliters per hour per conventional millimeter of mercury, as measured with bovine or expired human blood, and is used  with  either an automated ultrafiltration controller or anther method of ultrafiltration control to prevent fluid imbalance.

  - The hemodialysis delivery ma- chine is similar to the extracorporeal blood system and dialysate delivery system of the hemodialysis system and accessories (§ 876.5820), with the addi- tion of an ultrafiltration controller and mechanisms that monitor and/or con- trol such parameters as fluid balance, dialysate composition, and patient treatment parameters (e.g., blood pres- sure,  hematocrit,  urea,  etc.).

  - The high permeability hemo- dialysis system accessories include, but are not limited to, tubing lines and various treatment related monitors (e.g., dialysate pH, blood pressure, hematocrit, and blood recirculation monitors).

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical Device—Part I: Evaluation and Testing,’  ’’

  - ‘‘Guidance for the Content of 510(k)s for Conventional and High Per- meability Hemodialyzers,’’

  - ‘‘Guidance for Industry and CDRH Reviewers on the Content of Pre- market Notifications for Hemodialysis Delivery  Systems,’’

  - ‘‘Guidance for the Content of Pre- market Notifications for Water Purifi- cation Components and Systems for Hemodialysis,’’ and

  - ‘‘Guidance for Hemodialyzer Reuse   Labeling.’’

[65 FR 17145, Mar. 31, 2000]

##### § 876.5870   Sorbent hemoperfusion sys- tem.

- Identification. A sorbent hemoperfusion system is a prescription device  that  consists  of  an extracorporeal blood system similar to that identified in the hemodialysis sys- tem and accessories (§ 876.5820) and a container filled with  adsorbent  mate- rial that removes a wide range of sub- stances, both toxic and normal, from blood flowing  through  it.  The  adsorb- ent materials are usually activated- carbon or resins which  may  be  coated or immobilized to prevent fine  par- ticles entering the patient’s blood. The generic  type  of  device   may   include lines  and  filters  specifically   designed to connect the device to the extracorporeal blood system. The  de- vice is used in the treatment of  poi- soning, drug overdose, hepatic coma, or metabolic    disturbances.

- Classification. (1) Class II (special controls) when the device is intended for the treatment of  poisoning  and drug overdose. The special controls for this device are:

- The device must be demonstrated to be biocompatible;

- Performance data must dem- onstrate the mechanical  integrity  of the device (e.g., tensile, flexural, and structural strength), including  testing for the possibility of leaks, ruptures, release of particles, and/or disconnec- tions under anticipated conditions of use;

- Performance data must dem- onstrate device sterility and shelf life;

- Bench performance testing must demonstrate device functionality in terms of substances, toxins, and drugs removed by the device, and the extent that these are removed when the device is used according to its labeling, and to validate the device’s safeguards;

- A summary of clinical experience with the device that discusses and ana- lyzes device safety and performance, including a list of adverse events ob- served during the testing, must be pro- vided;

- Labeling must include the fol- lowing:

  - A detailed summary of the de- vice-related and procedure-related complications pertinent to the use of the device;

#### § 876.5880

  - A summary of the performance data provided for the device, including a list of the drugs and/or poisons the device has been demonstrated to re- move, and the extent for removal/de- pletion; and

- For those devices that incor- porate electrical components, appro- priate analysis and testing must be conducted to verify electrical safety and electromagnetic compatibility of the device.

(2) Class III (premarket approval) when the device is intended for the treatment of hepatic coma and meta- bolic  disturbances.

- Date premarket approval application (PMA) or notice of completion of product development protocol (PDP) is required. A PMA or notice of completion of a PDP is required to be filed with FDA by April 17, 2014, for any sorbent hemoperfusion system indicated for treatment of hepatic coma  or  meta- bolic disturbances that was in commer- cial distribution before May 28, 1976, or that has, by April 17, 2014, been found to be substantially equivalent to any sorbent hemoperfusion device indicated for treatment of hepatic coma or meta- bolic disturbances that was in commer- cial distribution before May 28, 1976. Any other sorbent hemoperfusion sys- tem device indicated for treatment of hepatic coma or metabolic disturb- ances shall have an approved PMA or declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[79 FR 3094, Jan. 17, 2014]

##### § 876.5880 Isolated kidney perfusion and transport system and acces- sories.

- Identification. An isolated kidney perfusion and transport system and ac- cessories is a device that is used to sup- port a donated or a cadaver kidney and to maintain the organ in a near-normal physiologic state until it is  trans- planted into a recipient patient. This generic type of device may include tub- ing, catheters, connectors, an ice stor- age or freezing container with or with- out bag or preservatives, pulsatile or nonpulsatile hypothermic  isolated organ perfusion apparatus with or without oxygenator, and disposable perfusion set.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class II (perform- ance standards).

##### § 876.5885 Tissue culture media for human ex vivo tissue and cell cul- ture processing applications.

- Identification. Tissue culture media for human ex vivo tissue and cell culture processing applications consist of cell and tissue culture media and components that are composed of chemically defined components (e.g., amino acids, vitamins, inorganic salts) that are essential for the ex vivo devel- opment, survival, and maintenance of tissues and cells of human origin. The solutions are indicated for use in human ex vivo tissue and cell culture processing  applications.

- Classification. Class II (special controls): FDA guidance document, ‘‘Class II Special Controls Guidance Document: Tissue Culture Media for Human Ex Vivo Processing Applica- tions; Final Guidance for Industry and FDA Reviewers.’’

[66 FR 27025, May 16, 2001]

##### § 876.5895   Ostomy irrigator.

- Identification. An ostomy irrigator is a device that consists of a container for fluid, tubing with a cone-shaped tip or a soft and flexible catheter with a retention shield and that is used to wash out the colon through a colos- tomy, a surgically created opening of the colon on the surface of the body.

- Classification. Class II (perform- ance standards).

##### § 876.5900 Ostomy pouch and acces- sories.

- Identification. An ostomy pouch and accessories is a device that con- sists of a bag that is attached to the patient’s skin by an adhesive material and that is intended for use as a recep- tacle for collection of fecal material or urine following an ileostomy, colos- tomy, or ureterostomy (a surgically created opening of the small intestine, large intestine, or the ureter on the surface of the body). This generic type of device and its accessories includes the ostomy pouch, ostomy adhesive, the disposable colostomy appliance, ostomy collector, colostomy pouch, urinary ileostomy bag, urine collecting ureterostomy   bag,   ostomy   drainage

#### Food and Drug Administration, HHS	§ 876.5955

bag with adhesive,  stomal bag, ostomy protector, and the ostomy size selector, but excludes ostomy pouches which in- corporate arsenic-containing com- pounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25050, June 12, 1989; 66 FR 38802, July 25,

2001]

##### § 876.5920   Protective  garment  for  in- continence.

- Identification. A protective gar- ment for incontinence is a device that consists of absorbent padding and a fluid barrier and that is intended to protect an incontinent patient’s gar- ment from the patient’s excreta. This generic type of device does not include diapers for infants.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[48 FR 53023, Nov. 23, 1983, as amended at 54

FR 25050, June 12, 1989; 66 FR 38802, July 25,

2001]

##### § 876.5930   Rectal control system.

- Identification. A rectal control system is a prescription device in- tended to treat fecal incontinence by controlling the size of the rectal lumen. The device is inserted in the va- gina and includes a portion that ex- pands to reduce the rectal lumen to prevent stool leakage and retracts to allow normal passage of stool. The de- vice includes an external regulator to control the state of expansion.

- Classification. Class II (special controls). The special controls for this device are:

  - Clinical testing must document the device acceptance data and the ad- verse event profile associated with clinical use, and demonstrate that the

device performs as intended under an- ticipated conditions of use.

  - The elements of the device that contact vaginal tissue must be dem- onstrated to be biocompatible.

  - The cleaning and disinfection in- structions for the device must be vali- dated.

  - Non-clinical (bench) testing must demonstrate that the device  performs as intended  under  anticipated condi- tions of use.

  - Non-clinical (bench) testing must demonstrate that the device does not:

    - Enhance the growth of Staphy- lococcus aureus.

    - Increase production of Toxic Shock  Syndrome  Toxin-1  by  S. aureus.

    - Alter the growth of normal vag- inal flora.

  - Labeling must include:

- Specific instructions, contra- indications, warnings, cautions, limita- tions, and the clinical training needed for the safe use of the device.

- The intended patient population and the intended use environment.

- Information on how the device is to be fitted, how the device operates, and recommendations on device main- tenance.

- A detailed summary of the clin- ical testing pertinent to  the use of the device, including a summary of the device- and procedure-related com- plications or adverse events related to use of the device, as well as relevant safety and performance information.

  - Patient labeling must be provided and must include:

    - Relevant contraindications, warn- ings, precautions, and adverse events/ complications.

    - Information on how the device operates and the recommended device maintenance (i.e., care instructions), including cleaning and disinfection.

    - Information on the patient popu- lation for which  there was  a favorable benefit/risk assessment.

    - The potential risks and benefits associated with the use of the device.

[80 FR 30933, June 1, 2015]

##### § 876.5955   Peritoneo-venous shunt.

- Identification. A peritoneo-venous shunt is an implanted device that con- sists of a catheter and a pressure acti- vated  one-way  valve.  The  catheter  is

#### § 876.5970

implanted with one end in the peri- toneal cavity and the other in a large vein. This device enables ascitic fluid in the peritoneal cavity to flow into the venous system for the treatment of intractable ascites.

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

  - ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’ and

  - Backflow specification and testing to prevent reflux of blood into the shunt.

[48 FR 53023, Nov. 23, 1983, as amended at 52

FR 17738,  May 11,  1987;  65  FR 17145,  Mar. 31,

2000]

##### § 876.5970   Hernia support.

    - Identification. A hernia support is a device, usually made of elastic, can- vas, leather, or metal, that is intended to be placed over a hernial opening (a weakness in the abdominal wall) to prevent protrusion of the abdominal contents. This generic type of device includes the umbilical truss.

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 876.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[48 FR 53023, Nov. 23, 1983, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38802,  July 25,

2001]

##### § 876.5980   Gastrointestinal   tube   and accessories.

- Identification. A gastrointestinal tube and accessories is a device that consists of flexible or semi-rigid tubing used for instilling fluids into, with- drawing fluids from, splinting, or sup- pressing bleeding of the alimentary tract. This device may incorporate an integral inflatable balloon for reten- tion or hemostasis. This generic type of device includes the hemostatic bag, irrigation and aspiration catheter (gas-

#### 21 CFR Ch. I (4–1–17 Edition)

tric, colonic, etc.), rectal catheter, sterile infant gavage set, gastro- intestinal string and tubes to locate in- ternal bleeding, double lumen tube for intestinal decompression or intubation, feeding tube, gastroenterostomy tube, Levine tube, nasogastric tube, single lumen tube with mercury weight bal- loon for intestinal intubation or de- compression, and gastro-urological ir- rigation tray (for gastrological use).

- Classification. (1) Class II (special controls). The barium enema retention catheter and tip with or without a bag that is a gastrointestinal tube and ac- cessory is exempt from the premarket notification procedures in subpart E of this part subject to the limitations in

§ 876.9.

- Class I (general controls) for the dissolvable nasogastric feed tube guide for the nasogastric tube. The class I de- vice is exempt from the premarket no- tification procedures in subpart E of part  807  of  this  chapter  subject  to

§ 876.9.

[49 FR 573, Jan. 5, 1984, as amended at 65 FR

2317, Jan. 14, 2000; 65 FR 76932, Dec. 8, 2000]

##### § 876.5990   Extracorporeal  shock  wave lithotripter.

  - Identification. An extracorporeal shock wave lithotripter is a device that focuses ultrasonic shock waves into the body to noninvasively fragment uri- nary calculi within the kidney or ure- ter. The primary components of the de- vice are a shock wave generator, high voltage generator, control console, im- aging/localization system, and patient table. Prior to treatment, the urinary stone is targeted using either an inte- gral or stand-alone localization/imag- ing system. Shock waves are typically generated using electrostatic spark dis- charge (spark gap),  electromagneti- cally repelled membranes, or piezo- electric  crystal  arrays,  and  focused onto the stone with either a specially designed reflector, dish, or acoustic lens. The shock waves are created under water within the shock wave generator, and are transferred to the patient’s body using an appropriate acoustic interface. After the stone has been fragmented by the focused shock waves, the fragments pass out of the body with the patient’s urine.

#### Food and Drug Administration, HHS	Pt. 878

  - Classification. Class II (special controls) (FDA guidance document: ‘‘Guidance for the Content of Pre- market Notifications (510(k)’s) for Extracorporeal	Shock	Wave Lithotripters Indicated for the Frag- mentation of Kidney and Ureteral Calculi.’’)

[65 FR 48612, Aug. 9, 2000]

## PART 878—GENERAL AND PLASTIC SURGERY DEVICES

##### Subpart A—General Provisions

Sec.

878.1    Scope.

878.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

878.1800   Speculum and accessories.

##### Subpart C [Reserved]

Subpart D—Prosthetic Devices

878.3250    External facial fracture fixation ap- pliance.

878.3300   Surgical mesh.

878.3500   Polytetrafluoroethylene  with  car- bon fibers composite implant material.

878.3530    Silicone     inflatable     breast     pros- thesis.

878.3540    Silicone gel-filled breast prosthesis. 878.3550  Chin prosthesis.

878.3590  Ear prosthesis.

878.3610   Esophageal prosthesis.

878.3680  Nose prosthesis.

878.3720   Tracheal prosthesls.

878.3750    External prosthesis adhesive. 878.3800    External aesthetic restoration pros-

thesis.

878.3900     Inflatable extremity splint. 878.3910     Noninflatable extremity splint.

878.3925   Plastic surgery kit and accessories.

##### Subpart E—Surgical Devices

878.4010    Tissue adhesive.

878.4011 Tissue adhesive with adjunct wound closure device  for  topical   approximation of  skin.

878.4014  Nonresorbable gauze/sponge for ex- ternal  use.

878.4015 Wound dressing with poly (diallyl dimethyl	ammonium	chloride) (pDADMAC)    additive.

878.4018   Hydrophilic wound dressing. 878.4020   Occlusive wound dressing.

878.4022 Hydrogel wound dressing and burn dressing.

878.4025    Silicone sheeting.

878.4040   Surgical apparel.

878.4100   Organ bag.

878.4160    Surgical camera and accessories. 878.4200   Introduction/drainage  catheter  and

accessories.

878.4300   Implantable clip. 878.4320   Removable skin clip.

878.4340 Contact cooling system for aes- thetic  use.

878.4350    Cryosurgical unit and accessories. 878.4360    Scalp cooling system to reduce the

likelihood of chemotherapy-induced alo- pecia.

878.4370 Surgical drape and drape acces- sories.

878.4380    Drape adhesive.

878.4400 Electrosurgical cutting and coagu- lation  device  and  accessories.

878.4410 Low energy ultrasound wound cleaner.

878.4420 Electrosurgical device for over-the- counter  aesthetic  use.

878.4440   Eye pad.

878.4450 Nonabsorbable gauze for  internal use.

878.4452 Nonabsorbable expandable hemo- static  sponge  for  temporary  internal  use.

878.4460  Non-powdered surgeon’s glove. 878.4470    Surgeon’s gloving cream.

878.4490 Absorbable hemostatic agent and dressing.

878.4493 Absorbable poly(glycolide/l-lactide) surgical   suture.

878.4494 Absorbable poly(hydroxybutyrate) surgical suture produced by recombinant DNA  technology.

878.4495   Stainless steel suture.

878.4520      Polytetrafluoroethylene  injectable.

878.4580   Surgical lamp.

878.4590 Focused  ultrasound stimulator  sys- tem  for  aesthetic  use.

878.4630 Ultraviolet lamp for dermatologic disorders.

878.4635 Sunlamp products and ultraviolet lamps intended for use in sunlamp prod- ucts.

878.4660    Skin marker.

878.4670   Internal tissue marker.

878.4680 Nonpowered, single patient,  port- able  suction  apparatus.

878.4683 Non-Powered suction apparatus de- vice intended  for  negative  pressure wound  therapy.

878.4700 Surgical microscope and acces- sories.

878.4730 Surgical skin degreaser or adhesive tape  solvent.

878.4750   Implantable staple.

878.4755   Absorbable lung biopsy plug. 878.4760   Removable skin staple.

878.4780    Powered suction pump.

878.4790 Powered surgical instrument for improvement in the appearance of cellulite.

§ 878.1

878.4800    Manual    surgical    instrument    for general  use.

878.4810 Laser surgical instrument for use in general and plastic surgery and in der- matology.

878.4815    Magnetic  surgical  instrument  sys- tem.

878.4820   Surgical instrument motors and ac- cessories/attachments.

878.4830    Absorbable surgical gut suture. 878.4840    Absorbable  polydioxanone  surgical

suture.

878.4930   Suture retention device.

878.4950 Manual operating table and acces- sories and manual operating chair and accessories.

878.4960   Operating   tables   and   accessories and  operating  chairs  and  accessories.

878.5000 Nonabsorbable	poly(ethylene terephthalate)   surgical   suture.

878.5010   Nonabsorbable   polypropylene   sur- gical   suture.

878.5020   Nonabsorbable   polyamide   surgical suture.

878.5030 Natural nonabsorbable silk surgical suture.

878.5035   Nonabsorbable  expanded  polytetra- fluoroethylene   surgical   suture.

878.5040    Suction lipoplasty system.

##### Subpart F—Therapeutic Devices

878.5070    Air-handling   apparatus   for   a   sur- gical  operating  room.

878.5350   Needle-type epilator.

878.5360   Tweezer-type epilator.

878.5400   Low level laser system for aesthetic use

878.5650   Topical   oxygen   chamber   for   ex- tremities.

878.5900    Nonpneumatic tourniquet.

878.5910    Pneumatic tourniquet.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE:  53  FR  23872,  June  24,  1988,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 878 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 878.1   Scope.

    - This part sets forth the classifica- tion of general and plastic surgery de- vices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device   under   part   807   cannot   show

21 CFR Ch. I (4–1–17 Edition)

merely that the device is accurately described by the section title and iden- tification provision of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87 of this  chapter.

    - To avoid duplicative listings, a general and plastic surgery device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed in one subpart only.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[53 FR 23872, June 24, 1988, as amended at 67

FR 77676,  Dec. 19,  2002;  78  FR 18233,  Mar. 26,

2013]

##### § 878.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a device, has an approval under section 515 of the act, FDA must promulgate a regula- tion under section 515(b) of the act re- quiring such approval, except as pro- vided in paragraphs (b) and (c) of this section. Such a regulation under sec- tion 515(b) of the act shall not be effec- tive during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar

#### Food and Drug Administration, HHS	§ 878.9

month after the regulation that classi- fies the device into class III is effec- tive, whichever is later. See section 501(f)(2)(B) of the act. Accordingly, un- less an effective date of the require- ment for premarket approval is shown in the regulation for a device classified into class III in this part, the device may be commercially distributed with- out FDA’s issuance of an order approv- ing a PMA or declaring completed a PDP for the device.  If  FDA  promul- gates a regulation under section 515(b) of the act requiring premarket ap- proval for a device, section 501(f)(1)(A) of the act applies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

- A device identified in a regulation in this part that is classified into class III and that is subject to the transi- tional provisions of section 520(l) of the act is automatically classified by stat- ute into class III and must have an ap- proval under section 515 of the act be- fore being commercially distributed. Accordingly, the regulation for such a class III transitional device states that as of the enactment date of the amend- ments, May 28, 1976, the device must have an approval under section 515 of the act before commercial distribution.

##### § 878.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening,   diagnosis,   or   monitoring

#### § 878.1800

life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2317, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 878.1800   Speculum and accessories.

    - Identification. A speculum is a de- vice intended to be inserted into a body cavity to aid observation. It is either nonilluminated or  illuminated  and may have various accessories.

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13827,  Apr.  5,  1989;  59  FR  63010,  Dec.  7,

1994; 66 FR 38802, July 25, 2001]

## Subpart C [Reserved]

Subpart D—Prosthetic Devices

##### § 878.3250   External facial fracture fixa- tion appliance.

- Identification. An external facial fracture fixation appliance is a metal apparatus intended to be used during surgical reconstruction and repair to immobilize maxillofacial bone frag- ments in their proper facial relation- ship.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13827,  Apr.  5,  1989;  65  FR  2317,  Jan.  14,

2000]

##### § 878.3300   Surgical mesh.

- Identification. Surgical mesh is a metallic or polymeric screen intended to be implanted to reinforce soft tissue or bone where weakness exists. Exam- ples of surgical mesh are metallic and polymeric mesh for hernia repair, and acetabular and cement restrictor mesh used during orthopedic surgery.

- Classification. Class II.

##### § 878.3500   Polytetrafluoroethylene with  carbon  fibers  composite  im- plant material.

- Identification. A polytetrafluoro- ethylene with carbon fibers composite implant material is a porous device material intended to be implanted dur- ing surgery of the chin, jaw, nose, or bones or tissue near the eye or ear. The device material serves as a space-occu- pying substance and is shaped and formed by the surgeon to conform to the patient’s need.

- Classification. Class II.

##### § 878.3530   Silicone	inflatable	breast prosthesis.

- Identification. A silicone inflatable breast prosthesis is a silicone rubber shell made of polysiloxane(s), such as polydimethylsiloxane	and polydiphenylsiloxane, that is inflated to the desired size with sterile isotonic saline before or after  implantation. The device is intended to be implanted to augment or reconstruct the female breast.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before November 17, 1999, for any silicone inflatable breast pros- thesis that was in commercial distribu- tion before May 28, 1976, or that has, on or before November 17, 1999, been found to be  substantially equivalent  to a  sili- cone inflatable breast prosthesis that was  in  commercial  distribution  before

#### Food and Drug Administration, HHS	§ 878.3720

May 28, 1976. Any other silicone inflat- able breast prosthesis shall have an ap- proved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[53 FR 23872, June 24, 1988, as amended at 64

FR 45161, Aug. 19, 1999]

##### § 878.3540   Silicone	gel-filled	breast prosthesis.

  - Identification—(1) Single-lumen sili- cone gel-filled breast prosthesis. A single- lumen silicone gel-filled breast pros- thesis is a silicone rubber shell made of polysiloxane(s),	such	as

polydimethylsiloxane	and polydiphenylsiloxane. The shell either contains a fixed amount cross-linked polymerized silicone gel, filler, and sta- bilizers or is filled to the desired size with injectable silicone gel at time of implantation. The device is intended to be implanted to augment or recon- struct the female breast.

- Double-lumen silicone gel-filled breast prosthesis. A double lumen sili- cone gel-filled breast prosthesis is a sil- icone rubber inner shell and a silicone rubber outer shell, both shells made of polysiloxane(s),	such	as

polydimethylsiloxane	and polydiphenylsiloxane. The inner shell contains fixed amounts of cross-linked polymerized silicone gel, fillers, and stabilizers. The outer shell is inflated to the desired size with sterile isotonic saline before or after  implantation. The device is intended to be implanted to augment or reconstruct the female breast.

- Polyurethane covered silicone gel- filled breast prosthesis. A polyurethane covered silicone gel-filled breast pros- thesis is an inner silicone rubber shell made of polysiloxane(s), such as polydimethylsiloxane	and polydiphenylsiloxane, with an outer silicone adhesive layer and an outer covering of polyurethane; contained within the inner shell is a fixed amount of cross-linked polymerized silicone gel, fillers, and stabilizers and an inert support structure compartmentalizing the silicone gel. The device is intended to be implanted to augment or recon- struct the female breast.

  - Classification. Class III.

  - Date premarket approval application (PMA) is required. A PMA is required to

be filed with the Food and Drug Ad- ministration on or before July 9, 1991 for any silicone gel-filled breast pros- thesis that was in commercial distribu- tion before May 28, 1976, or that has on or before July 9, 1991 been found to be substantially equivalent to a silicone gel-filled breast prosthesis that was in commercial distribution before May 28, 1976. Any other silicone gel-filled breast prosthesis shall have an ap- proved PMA in effect before being placed in commercial distribution.

[53 FR 23872, June 24, 1988, as amended at 56

FR 14627, Apr. 10, 1991]

##### § 878.3550   Chin prosthesis.

- Identification. A chin prosthesis is a silicone rubber solid device intended to be implanted to augment or recon- struct the chin.

- Classification. Class II.

##### § 878.3590   Ear prosthesis.

- Identification. An ear prosthesis is a silicone rubber solid device intended to be implanted to reconstruct the ex- ternal ear.

- Classification. Class II.

##### § 878.3610   Esophageal prosthesis.

- Identification. An esophageal pros- thesis is a rigid, flexible, or expandable tubular device made of a plastic, metal, or polymeric material that is intended to be implanted to restore the structure and/or function of the esoph- agus. The metal esophageal prosthesis may be uncovered or covered with a polymeric material. This device may also include a device delivery system.

- Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance for the Content of Premarket No- tification  Submissions  for  Esophageal and Tracheal Prostheses.’’

[65 FR 17145, Mar. 31, 2000]

##### § 878.3680   Nose prosthesis.

- Identification. A nose prosthesis is a silicone rubber solid device intended to be implanted to augment or recon- struct the nasal dorsum.

- Classification. Class II.

##### § 878.3720   Tracheal prosthesis.

- Identification. The  tracheal pros- thesis is a rigid, flexible, or expandable

#### § 878.3750

tubular device made of a silicone, metal, or polymeric material that is intended to be implanted to restore the structure and/or function of the tra- chea or trachealbronchial tree. It may be unbranched or contain one or two branches. The metal tracheal pros- thesis may be uncovered or covered with a polymeric material. This device may also include a device delivery sys- tem.

- Classification. Class II. The special control for this device is FDA’s ‘‘Guid- ance for the Content of Premarket No- tification  Submissions  for  Esophageal and Tracheal Prostheses.’’

[65 FR 17146, Mar. 31, 2000]

##### § 878.3750   External   prosthesis   adhe- sive.

- Identification. An external pros- thesis adhesive is a silicone-type adhe- sive intended to be used to fasten to the body an external aesthetic restora- tion prosthesis, such as an artificial nose.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38802,  July 25,

2001]

##### § 878.3800   External  aesthetic  restora- tion prosthesis.

- Identification. An external aes- thetic restoration prosthesis is a device intended to be used to construct an ex- ternal artificial body structure, such as an ear, breast, or nose. Usually the de- vice is made of silicone rubber and it may be fastened to the body with an external prosthesis adhesive. The de- vice is not intended to be implanted.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.  If the device is intended for  use  without an external  prosthesis  adhesive to  fas- ten it to the body, the device is exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with  respect  to  general  requirements

#### 21 CFR Ch. I (4–1–17 Edition)

concerning  records,  and  § 820.198,  with respect  to  complaint  files.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38802,  July 25,

2001]

##### § 878.3900   Inflatable extremity splint.

- Identification. An inflatable ex- tremity splint is a device intended to be inflated to immobilize a limb or an extremity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38802,  July 25,

2001]

##### § 878.3910   Noninflatable	extremity splint.

- Identification. A noninflatable ex- tremity splint is a device intended to immobilize a limb or an extremity. It is not inflatable.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9. If the device is not la- beled or otherwise represented as ster- ile, it is exempt from the current good manufacturing  practice   requirements of the quality  system  regulation  in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,  with  respect  to  complaint files.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13827,  Apr.  5,  1989;  65  FR  2317,  Jan.  14,

2000]

##### § 878.3925   Plastic  surgery  kit  and  ac- cessories.

- Identification. A plastic surgery kit and accessories is a device intended to be used to reconstruct maxillofacial deficiencies. The kit contains surgical instruments and materials used to make maxillofacial impressions before molding an external prosthesis.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 878.4015

subpart  E  of  part  807  of  this  chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13827,  Apr.  5,  1989;  65  FR  2317,  Jan.  14,

2000]

## Subpart E—Surgical Devices

##### § 878.4010   Tissue adhesive.

  - Tissue adhesive for the topical ap- proximation of skin—(1) Identification. A tissue adhesive for the topical approxi- mation of skin is a device intended for topical closure of surgical incisions, in- cluding laparoscopic incisions, and simple traumatic lacerations that have easily approximated skin edges. Tissue adhesives for the topical approxima- tion of skin may be used in conjunction with, but not in place of, deep dermal stitches.

(2) Classification. Class II (special con- trols). The special control for this de- vice is FDA’s ‘‘Class II Special Controls Guidance Document: ‘‘Tissue Adhesive for the Topical Approximation  of Skin.’’ See § 878.1(e) of this chapter for the availability of this guidance docu- ment.

  - Tissue adhesive for non-topical use—(1) Identification. A tissue adhesive for non-topical use, including adhesives intended for use in the embolization of brain arteriovenous malformation  or for use in ophthalmic surgery, is a de- vice used for adhesion of internal tis- sues and vessels.

(2) Classification. Class III (premarket approval). As of May 28, 1976, an ap- proval under section 515 of the act is required before this device may be commercially distributed. See § 878.3 of this   chapter.

[73 FR 31033, May 30, 2008]

##### § 878.4011 Tissue adhesive with ad- junct wound closure device for top- ical approximation of skin.

- Identification. A tissue adhesive with adjunct wound closure device in- tended for the topical approximation of skin is a device indicated for topical application only to hold closed easily approximated skin edges of wounds from surgical incisions, including punctures from minimally invasive surgery, and simple, thoroughly cleansed,  trauma-induced  lacerations.

It may be used in conjunction with, but not in place of, deep dermal stitches. Additionally, the adjunct wound clo- sure device component maintains tem- porary skin edge alignment along the length of wound during application of the liquid adhesive.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Guidance for Industry and FDA Staff;  Class II Special  Con- trols Guidance Document: Tissue Adhe- sive with Adjunct Wound Closure De- vice Intended for the Topical Approxi- mation of Skin.’’ See § 878.1(e) for the availability  of  this  guidance  document.

[75 FR 68794, Nov. 10, 2010]

##### § 878.4014 Nonresorbable gauze/sponge for external use.

- Identification. A nonresorbable gauze/sponge for external use is a ster- ile or nonsterile device intended for medical purposes, such as to be placed directly on a patient’s wound to absorb exudate. It consists of a strip, piece, or pad made from open woven or nonwoven mesh cotton cellulose or a simple chemical derivative of cellulose. This classification does not include a nonresorbable gauze/sponge for exter- nal use that contains added drugs such as antimicrobial agents, added bio- logics such as growth factors, or is composed of materials derived from animal sources.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in part 807, subpart E of this chapter sub- ject to the limitations in § 878.9.

[64 FR 53929, Oct. 5, 1999]

##### § 878.4015 Wound dressing with poly (diallyl dimethyl ammonium chlo- ride) (pDADMAC) additive.

- Identification. A wound dressing with pDADMAC  additive  is  intended for use as a primary dressing for exud- ing wounds, 1st and 2d degree burns, and surgical wounds, to secure and pre- vent movement of a primary dressing, and as a wound packing.

- Classification. Class II (special controls). The special control is: the FDA guidance document  entitled ‘‘Class II Special Controls Guidance Document: Wound Dressing With Poly

§ 878.4018

(Diallyl   Dimethyl   Ammonium   Chlo- ride)	(pDADMAC)	Additive.’’	See

§ 878.1(e)  for  availability  of  this  guid- ance document.

[74 FR 53167, Oct. 16, 2009]

##### § 878.4018   Hydrophilic   wound   dress- ing.

  - Identification. A hydrophilic wound dressing is a sterile or non-ster- ile device intended to cover a wound and to absorb exudate. It consists of nonresorbable materials with hydro- philic properties that are capable of ab- sorbing exudate (e.g., cotton, cotton derivatives, alginates, dextran, and rayon). This classification does not in- clude a hydrophilic wound dressing that contains added drugs such as anti- microbial agents, added biologics such as growth factors, or is composed of materials derived from animal sources.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in part 807, subpart E of this chapter sub- ject to the limitations in § 878.9.

[64 FR 53929, Oct. 5, 1999]

##### § 878.4020   Occlusive wound dressing.

- Identification. An occlusive wound dressing is a nonresorbable, sterile or non-sterile device intended to cover a wound, to provide or support a moist wound environment, and to allow the exchange of gases such as oxygen and water vapor through the device. It con- sists of a piece of synthetic polymeric material, such as polyurethane, with or without an adhesive backing. This clas- sification does not include an occlusive wound dressing that contains added drugs such as antimicrobial agents, added biologics such as growth factors, or is composed of materials derived from animal sources.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in part 807, subpart E of this chapter sub- ject to the limitations in § 878.9.

[64 FR 53929, Oct. 5, 1999]

##### § 878.4022   Hydrogel   wound   dressing and burn dressing.

- Identification. A hydrogel wound dressing is a sterile or non-sterile de- vice intended to cover a wound, to ab-

#### 21 CFR Ch. I (4–1–17 Edition)

sorb wound exudate, to control bleed- ing or fluid loss, and to protect against abrasion, friction, desiccation, and con- tamination. It consists of a nonresorbable matrix made of hydro- philic polymers or other material in combination with water (at least 50 percent) and capable of absorbing exudate. This classification does not include a hydrogel wound dressing that contains added drugs such as anti- microbial agents, added biologics  such as growth factors, or is composed of materials derived from animal sources.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in part 807, subpart E of this chapter sub- ject to the limitations in § 878.9.

[64 FR 53929, Oct. 5, 1999]

##### § 878.4025   Silicone sheeting.

- Identification. Silicone sheeting is intended for use in the management of closed hyperproliferative (hypertrophic and keloid) scars.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 878.9.

[69 FR 48148, Aug. 9, 2004]

##### § 878.4040   Surgical apparel.

- Identification. Surgical apparel are devices that are intended to be worn by operating room personnel during sur- gical procedures to protect both the surgical patient and the operating room personnel from transfer of micro- organisms, body fluids, and particulate material. Examples include surgical caps, hoods, masks, gowns, operating room shoes and shoe covers, and isola- tion masks and gowns. Surgical suits and dresses, commonly known as scrub suits, are excluded.

- Classification. (1) Class II (special controls) for surgical gowns and sur- gical  masks.

- Class I (general controls) for sur- gical apparel other than surgical gowns and surgical masks.  The  class  I  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 65

FR 2317, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 878.4350

##### § 878.4100   Organ bag.

  - Identification. An organ bag is a device that is a flexible plastic bag in- tended to be used as a temporary recep- tacle for an organ during surgical pro- cedures to prevent moisture loss.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR  63010,  Dec.  7,  1994;  65  FR  2318,  Jan.  14,

2000]

##### § 878.4160   Surgical  camera  and  acces- sories.

- Identification. A surgical camera and accessories is a device intended to be used to record operative procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR 13827,  Apr. 5,  1989;  66  FR 38802,  July 25,

2001]

##### § 878.4200   Introduction/drainage  cath- eter and accessories.

- Identification. An introduction/ drainage catheter is a device that is a flexible single or multilumen tube in- tended to be used to introduce nondrug fluids into body cavities other than blood vessels, drain fluids from body cavities, or evaluate certain physio- logic conditions. Examples include irri- gation and drainage catheters, pedi- atric catheters, peritoneal catheters (including dialysis), and other general surgical catheters. An introduction/ drainage catheter accessory is intended to aid in the manipulation of or inser- tion of the device into the body. Exam- ples of accessories include adaptors, connectors, and catheter needles.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 878.4300   Implantable clip.

- Identification. An implantable clip is a clip-like device intended to con- nect internal tissues to aid healing. It is not absorbable.

- Classification. Class II.

##### § 878.4320   Removable skin clip.

- Identification. A removable skin clip is a clip-like device intended to connect skin tissues temporarily to aid healing. It is not absorbable.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 878.4340   Contact  cooling  system  for aesthetic use.

- Identification. A contact cooling system for aesthetic use is a device that is a combination of a cooling pad associated with a vacuum or mechan- ical massager intended for the disrup- tion of adipocyte cells intended for non-invasive aesthetic use.

- Classification. Class II (special controls). The special controls for this device is FDA’s ‘‘Guidance for Industry and FDA  Staff; Class  II Special  Con- trols Guidance Document: Contact Cooling  System  for  Aesthetic  Use.’’  See

§ 878.1(e)   for   the   availability   of   this guidance document.

[76 FR 6553, Feb. 7, 2011]

##### § 878.4350   Cryosurgical unit and acces- sories.

- Identification—(1) Cryosurgical unit with a liquid nitrogen cooled cryoprobe and accessories. A cryosurgical  unit with a liquid nitrogen cooled cryoprobe and accessories is a device intended to destroy tissue during surgical proce- dures by applying extreme cold.

- Cryosurgical unit with a nitrous oxide cooled cryoprobe and accessories. A cryosurgical unit with a nitrous oxide cooled cryoprobe and accessories is a device intended to destroy tissue dur- ing surgical procedures, including urological applications, by applying ex- treme cold.

- Cryosurgical unit with a carbon di- oxide cooled cryoprobe or a carbon dioxide

#### § 878.4360

dry ice  applicator and accessories.  A cryosurgical unit with a carbon dioxide cooled cryoprobe or a carbon dioxide dry ice applicator and accessories is a device intended to destroy tissue dur- ing surgical procedures by applying ex- treme cold. The device is intended to treat disease conditions such as tu- mors, skin cancers, acne scars, or hemangiomas (benign tumors con- sisting of newly formed blood vessels) and various benign or malignant gyne- cological conditions affecting vulvar, vaginal, or cervical  tissue.  The  device is not intended for urological applica- tions.

- Classification. Class II.

##### § 878.4360 Scalp cooling system to re- duce the likelihood of chemo- therapy-induced alopecia.

- Identification. A scalp cooling sys- tem to reduce the likelihood of chemo- therapy-induced alopecia is a prescrip- tion device intended to reduce the fre- quency and severity of alopecia during chemotherapy in which alopecia-induc- ing chemotherapeutic agents are used.

- Classification—Class II (special controls). The special controls for this device are:

- Non-clinical performance testing must demonstrate that the device meets all design specifications and per- formance requirements, and that the device performs as intended under an- ticipated conditions of use. This infor- mation must include testing to dem- onstrate accuracy of the temperature control mechanism.

- Performance testing must dem- onstrate the electromagnetic compat- ibility and electrical safety of the de- vice.

- Software verification, validation, and hazard analysis must be performed.

- The patient contacting compo- nents of the device must be dem- onstrated to be biocompatible. Mate- rial names must be provided.

- Labeling must include the fol- lowing:

- A statement describing the poten- tial risk of developing scalp metas- tasis.

- Information on the patient popu- lation and chemotherapeutic agents/ regimen for which the device has been demonstrated to be effective.

#### 21 CFR Ch. I (4–1–17 Edition)

- A summary of the non-clinical and/or clinical testing pertinent to use of the device.

- A summary of the device tech- nical parameters, including tempera- ture cooling range and duration  of cooling.

- A summary of the device- and pro- cedure-related adverse events pertinent to use of the device.

- Information on how the device operates and the typical course of treatment.

- Patient labeling must be provided and must include:

  - Relevant contraindications, warn- ings, precautions, and adverse effects/ complications.

  - Information on how the device operates and the typical course of treatment.

  - Information on the patient popu- lation for which there is clinical evi- dence of effectiveness.

  - The potential risks and benefits associated with use of the device.

  - Postoperative care instructions.

  - A statement describing the po- tential risk of developing scalp metas- tasis.

[81 FR 7453, Feb. 12, 2016]

##### § 878.4370   Surgical   drape   and   drape accessories.

- Identification. A surgical drape and drape accessories is a device made of natural or synthetic materials in- tended to be used as a protective pa- tient covering, such as to isolate a site of surgical incision from microbial and other contamination. The device in- cludes a plastic wound protector that may adhere to the skin around a sur- gical incision or be placed in a wound to cover its exposed edges, and a latex drape with a self-retaining finger cot that is intended to allow repeated in- sertion of the surgeon’s finger into the rectum during performance of a transurethral prostatectomy.

- Classification. Class II.

##### § 878.4380   Drape adhesive.

- Identification. A drape adhesive is a device intended to be placed on the skin to attach a surgical drape.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 878.4440

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38802,  July 25,

2001]

##### § 878.4400   Electrosurgical  cutting  and coagulation device and accessories.

  - Identification. An electrosurgical cutting and coagulation device and ac- cessories is a device intended to re- move tissue and control bleeding by use of high-frequency electrical cur- rent.

  - Classification. Class II.

##### § 878.4410   Low	energy	ultrasound wound cleaner.

- Identification. A low energy ultrasound wound cleaner is a device that uses ultrasound energy to vapor- ize a solution and generate a mist that is used for the cleaning and mainte- nance debridement of wounds. Low lev- els of ultrasound energy may be carried to the wound by the saline mist.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Low Energy Ultrasound Wound Clean- er.’’ See § 878.1(e) for the availability of this guidance document.

[70 FR 67355, Nov. 7, 2005]

##### § 878.4420   Electrosurgical   device   for over-the-counter aesthetic use.

- Identification. An electrosurgical device for over-the-counter aesthetic use is a device using radiofrequency en- ergy to produce localized heating with- in tissues for non-invasive  aesthetic use.

- Classification. Class II (special controls). The special controls for this device are:

  - Non-clinical performance data must demonstrate that the  device meets all design specifications and per- formance requirements. The following performance characteristics must be tested: Over-heating, power accuracy radiofrequency, pulse cycle, waveform, pulse duration, and device character- ization  parameters.

  - Label comprehension and self-se- lection performance evaluation must demonstrate that the intended over- the-counter  users  can  understand  the

package labeling and correctly choose the device for the indicated aesthetic use.

  - Usability performance evaluation must demonstrate that the over-the- counter user can correctly use the de- vice, based solely on reading the direc- tions for use, to treat the indicated aesthetic  use.

  - Clinical performance evaluation must demonstrate that the device per- forms as intended under anticipated conditions of  use  to  achieve  the  in- tended aesthetic results.

  - The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

  - Instructions for cleaning the de- vice must be validated.

  - Performance data must be pro- vided to demonstrate the electro- magnetic compatibility and electrical safety, including the mechanical integ- rity, of the device.

  - Software verification, validation, and hazard analysis must be performed.

  - Labeling must include:

    - Warnings, precautions, and con- traindications to ensure the safe use of the device for the over-the-counter users.

    - A statement that the safety and effectiveness of the device’s  use  for uses other than the indicated aesthetic use are not known.

    - A summary of the clinical infor- mation used to establish effectiveness for each indicated aesthetic usage and observed adverse events.

[81 FR 42244, June 29, 2016]

##### § 878.4440   Eye pad.

- Identification. An eye pad is a de- vice that consists of a pad made of var- ious materials, such as gauze and cot- ton, intended for use as a bandage over the eye for protection or absorption of secretions.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

#### § 878.4450

##### § 878.4450   Nonabsorbable gauze for in- ternal use.

  - Identification. Nonabsorbable gauze for internal use is a device made of an open mesh fabric intended to be used inside the body or a surgical inci- sion or applied to internal organs or structures, to control bleeding, absorb fluid, or protect organs or structures from abrasion, drying, or contamina- tion. The device is woven from mate- rial made of not less than 50 percent by mass cotton, cellulose, or a simple chemical derivative of cellulose, and contains x-ray detectable elements.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38803,  July 25,

2001]

##### § 878.4452 Nonabsorbable expandable hemostatic sponge for temporary internal use.

- Identification. A nonabsorbable ex- pandable hemostatic  sponge  for  tem- porary internal use is a prescription device intended to be placed tempo- rarily into junctional, non-compress- ible wounds, which are not amenable to tourniquet use, to control bleeding until surgical care is acquired. The sponges expand upon contact with blood to fill the wound cavity and pro- vide a physical barrier and pressure that facilitates formation  of  a  clot. The device consists of sterile, non- absorbable radiopaque compressed sponges and may include an applicator to facilitate delivery into a wound.

- Classification. Class II (special controls). The special controls for this device are:

  - Performance data must dem- onstrate the biocompatibility of pa- tient-contacting components.

  - Performance data must dem- onstrate the sterility of patient-con- tacting	components	including endotoxin and pyrogenicity assess- ments.

  - Performance data must support device stability by demonstrating con- tinued sterility of the patient-con- tacting components of the device, package       integrity,       and       device

#### 21 CFR Ch. I (4–1–17 Edition)

functionality over the requested shelf life.

  - Assessment of material character- istics must be sufficient to support safety under anticipated conditions of use. Assessments must include the fol- lowing:

    - Material specifications.

    - Immunogenicity.

    - Viral inactivation for animal-de- rived materials.

  - Non-clinical performance data must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be test- ed:

- Absorption capacity.

- Extent of swelling.

- Mechanical properties.

- Expansion  force/pressure.

- Radiopacity.

- Deployment/applicator functionality.

  - In vivo performance data must demonstrate safe and effective use by verifying that the device performs as intended under anticipated conditions of use. Appropriate analysis/testing must demonstrate that the product: Controls bleeding, does not promote adverse local or systemic effects, and can be completely removed from the wound. The following performance characteristics must be tested:

- Deployment.

- Control of bleeding.

- Radiopacity.

- Retrieval.

- Assessment of local and systemic effects.

  - Human factors testing and anal- ysis must validate that the device de- sign and labeling are sufficient for ap- propriate use by emergency responders deploying the device as well as sur- geons retrieving the device  from wounds.

  - Labeling must include:

    - Specific instructions for deploy- ment by emergency responders and re- trieval by surgeons.

    - Warnings, cautions, and limita- tions needed for safe use of the device.

    - Information on how the device operates and the typical course of treatment.

#### Food and Drug Administration, HHS	§ 878.4495

    - A detailed summary of the in vivo and human factors testing perti- nent to use of the device.

    - Appropriate imaging information to ensure complete retrieval of device.

    - An expiration date/shelf life.

[79 FR 34224, June 16, 2014]

##### § 878.4460   Non-powdered	surgeon’s glove.

- Identification. A non-powdered sur- geon’s glove is a device intended to be worn on the hands of operating room personnel to protect a surgical wound from contamination. A non-powdered surgeon’s glove does not incorporate powder for purposes other than manu- facturing. The final finished glove in- cludes only residual powder from man- ufacturing.

- Classification. Class I (general con- trols).

[53 FR 23872, June 24, 1988, as amended at 66

FR 46952, Sept. 10, 2001; 81 FR 91730, Dec. 19,

2016]

##### § 878.4470   Surgeon’s gloving cream.

- Identification. Surgeon’s gloving cream is an ointment intended to be used to lubricate the user’s hand before putting on a surgeon’s glove.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 878.4490   Absorbable	hemostatic agent and dressing.

- Identification. An absorbable he- mostatic agent or dressing is a device intended to produce hemostasis by ac- celerating the clotting process  of blood. It is absorbable.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 878.3.

##### § 878.4493   Absorbable   poly(glycolide/l- lactide) surgical suture.

- Identification. An absorbable poly(glycolide/l-lactide)  surgical  suture

(PGL suture) is an absorbable sterile, flexible strand as prepared and syn- thesized from homopolymers of glycolide and copolymers made from 90 percent glycolide and 10 percent l- lactide, and is indicated for use in soft tissue approximation. A PGL suture meets United States Pharmacopeia (U.S.P.) requirements as described in the U.S.P. ‘‘Monograph for Absorbable Surgical Sutures;’’ it may be monofilament or multifilament (braid- ed) in form; it may be uncoated or coated; and it may be undyed or dyed with an FDA-approved color additive. Also, the suture may be provided with or without a standard needle attached.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[56 FR 47151, Sept. 18, 1991, as amended at 68

FR 32984, June 3, 2003]

##### § 878.4494   Absorbable poly(hydroxybutyrate)  surgical  su- ture produced by recombinant DNA technology.

- Identification. An absorbable poly(hydroxybutyrate)  surgical  suture is an absorbable surgical  suture  made of material isolated from prokaryotic cells produced by recombinant deoxyribonucleic acid (DNA) tech- nology. The device is  intended  for  use in general soft  tissue  approximation and  ligation.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Absorbable Poly(hydroxybutyrate) Surgical Suture Produced by Recombinant DNA Tech- nology.’’ For the availability of this guidance document see § 878.1(e).

[72 FR 43146, Aug. 3, 2007]

##### § 878.4495   Stainless steel suture.

- Identification. A stainless steel su- ture is a needled or unneedled non- absorbable surgical suture composed of 316L stainless steel, in USP sizes 12–0 through 10, or  a  substantially  equiva- lent  stainless  steel  suture,  intended  for

#### § 878.4520

use in abdominal wound closure, intes- tinal anastomosis, hernia repair, and sternal closure.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[65 FR 19836, Apr. 13, 2000, as amended at 68

FR 32984, June 3, 2003]

##### § 878.4520   Polytetrafluoroethylene injectable.

- Identification. Polytetrafluoro- ethylene injectable is an injectable paste prosthetic device composed of polytetrafluoroethylene intended to be used to augment or reconstruct a vocal cord.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 878.3.

##### § 878.4580   Surgical lamp.

- Identification. A surgical lamp (in- cluding a fixture) is a device intended to be used to provide visible illumina- tion of the surgical field or the patient.

- Classification. Class II.

##### § 878.4590   Focused  ultrasound  stimu- lator system for aesthetic use.

- Identification. A Focused Ultrasound Stimulator System for Aes- thetic Use is a device using focused ultrasound to produce localized, me- chanical motion within tissues and cells for the purpose of producing ei- ther localized heating for tissue coagu- lation or for mechanical cellular mem- brane disruption intended for noninvasive aesthetic use.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Focused Ultrasound Stimulator System for Aes- thetic Use.’’ See § 878.1(e) for the avail- ability of this guidance document.

[76 FR 43121, July 20, 2011]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 878.4630 Ultraviolet lamp for der- matologic disorders.

- Identification. An ultraviolet lamp for dermatologic disorders is a device (including a fixture) intended to pro- vide ultraviolet radiation of the body to photoactivate a drug in the treat- ment of a dermatologic disorder if the labeling of the drug intended for use with the device bears adequate direc- tions for the device’s use with that drug.

- Classification. Class II.

##### § 878.4635 Sunlamp products and ul- traviolet lamps intended for use in sunlamp products.

- Identification. A sunlamp  product is any device designed to incorporate one or more ultraviolet (UV) lamps in- tended for irradiation of any part  of the living human body,  by  UV  radi- ation with wavelengths in air between 200 and 400 nanometers, to induce skin tanning. This definition includes tan- ning beds and tanning booths. A UV lamp intended for use in sunlamp prod- ucts is any lamp that produces UV ra- diation in the wavelength interval of 200 to 400 nanometers in air.

- Classification. Class II (special controls). The special controls for sun- lamp products and  UV lamps  intended for use in sunlamp products are:

- Conduct performance testing that demonstrates the following:

  - Device meets appropriate output performance specifications such as wavelengths, energy density, and lamp life; and

  - Device’s safety features, such as timers to limit UV exposure and alarms, function properly.

- Demonstrate that device is me- chanically safe to prevent user injury.

- Demonstrate	software verification, validation, and hazard analysis.

- Demonstrate that device is bio- compatible.

- Demonstrate that device is elec- trically safe and electromagnetically compatible in its intended use environ- ment.

- Labeling—(i) Sunlamp products. (A) The warning statement below must ap- pear on all sunlamp products and must be placed in a black box. This state- ment must be permanently affixed or

#### Food and Drug Administration, HHS	§ 878.4680

inscribed on the product when fully as- sembled for use so as to be legible and readily accessible to view by the person who will be exposed to UV radiation immediately before the use of the prod- uct. It shall be of sufficient durability to remain legible throughout the ex- pected lifetime of the product. It shall appear on a part or panel displayed prominently under normal conditions of use so that it is readily accessible to view whether the tanning bed canopy (or tanning booth door) is open or closed when the person who will be ex- posed approaches the equipment and the text shall be at least 10 millimeters (height). Labeling on the device must include the following statement:

Attention:  This  sunlamp  product  should  not  be used on persons under the age of 18 years.

(B) Manufacturers shall provide vali- dated instructions on cleaning and dis- infection of sunlamp products between uses in the user instructions.

- Sunlamp products and UV lamps in- tended for use in sunlamp products. Man- ufacturers of sunlamp products and UV lamps intended for use in sunlamp products shall provide or cause to be provided in the user instructions, as well as all consumer-directed catalogs, specification sheets, descriptive bro- chures, and Web pages in which sun- lamp products or UV lamps intended for use in sunlamp products are offered for sale,  the  following contraindication and warning statements:

  - ‘‘Contraindication:  This  product is contraindicated for use on persons under the age of 18 years.’’

  - ‘‘Contraindication: This product must not be used if skin lesions or open wounds are present.’’

  - ‘‘Warning: This product should not be used on individuals who have had skin cancer or have a family his- tory of skin cancer.’’

  - ‘‘Warning: Persons repeatedly ex- posed to UV radiation should be regu- larly evaluated for skin cancer.’’

- Performance standard. Sunlamp products and  UV  lamps  intended  for use in sunlamp products are subject  to the electronic product performance standard at § 1040.20 of this chapter.

[79 FR 31213, June 2, 2014]

##### § 878.4660   Skin marker.

- Identification. A skin marker is a pen-like device intended to be used to write on the patient’s skin, e.g., to out- line surgical incision sites or mark an- atomical sites for accurate blood pres- sure measurement.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 878.4670   Internal tissue marker.

- Identification. An internal tissue marker is a prescription use device that is intended for use prior to or dur- ing general surgical procedures to de- marcate selected sites on internal tis- sues.

- Classification. Class II (special controls). The special controls for this device are:

- The device must be demonstrated to be biocompatible.  Material  names and specific designation numbers must be provided.

- Performance testing must dem- onstrate that the device performs as intended to mark  the  tissue  for  which it  is  indicated.

- Performance data must dem- onstrate the sterility of the device.

- Performance data must support the shelf life of the device by dem- onstrating sterility, package integrity, device functionality, and material sta- bility over the requested shelf life.

- Labeling must include:

  - A warning that the device must not be used on a non-sterile surface prior to use internally.

  - An expiration date/shelf life.

  - Single use only labeling must be labeled directly on the device.

[80 FR 46486, Aug. 5, 2015]

##### § 878.4680   Nonpowered, single patient, portable suction apparatus.

- Identification. A nonpowered, sin- gle patient, portable suction apparatus is a device that consists of a manually operated plastic, disposable evacuation system intended to provide a vacuum

#### § 878.4683

for	suction	drainage	of	surgical wounds.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 878.4683 Non-Powered suction appa- ratus device intended for negative pressure wound therapy.

- Identification. A non-powered suc- tion apparatus device intended for neg- ative pressure wound therapy is a de- vice that is indicated for wound man- agement via application of negative pressure to the wound for removal of fluids, including wound exudate, irriga- tion fluids, and infectious materials. It is further indicated for management of wounds, burns, flaps, and grafts.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Non-powered Suction Apparatus Device Intended for Negative Pressure Wound Therapy (NPWT).’’ See § 878.1(e) for the avail- ability of this guidance document.

[75 FR 70114, Nov. 17, 2010]

##### § 878.4700   Surgical microscope and ac- cessories.

- Identification. A surgical micro- scope and accessories is an AC-powered device intended for use during surgery to provide a magnified view of the sur- gical field.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[55 FR 48440, Nov. 20, 1990, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 878.4730   Surgical  skin  degreaser  or adhesive tape solvent.

- Identification. A surgical skin degreaser or an adhesive tape solvent is a device that consists  of  a  liquid  such as 1,1,2-trichloro-1,2,2-trifluoroethane; 1,1,1-trichloroethane; and 1,1,1-tri- chloroethane  with  mineral  spirits  in-

#### 21 CFR Ch. I (4–1–17 Edition)

tended to be used to dissolve surface skin oil or adhesive tape.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 878.4750   Implantable staple.

- Identification. An implantable sta- ple is a staple-like device intended to connect internal tissues to aid healing. It is not absorbable.

- Classification. Class II.

##### § 878.4755   Absorbable	lung	biopsy plug.

- Identification. A preformed (po- lymerized) absorbable lung biopsy plug is intended to provide accuracy in marking a biopsy location for visual- ization during surgical resection and closure of pleural punctures associated with percutaneous, transthoracic nee- dle lung biopsies. Upon deployment into the biopsy tract, the plug expands to fill the biopsy void and remains in place until resorbed.

- Classification. Class II (special controls). The special controls for this device are:

- The design characteristics of the device must ensure that the geometry and material composition are con- sistent with the intended use.

- Performance testing must dem- onstrate deployment as indicated in the accompanying labeling, including the indicated introducer needles, and demonstrate expansion and resorption characteristics in a clinically relevant environment.

- In vivo evaluation must dem- onstrate performance  characteristics of the device, including the ability  of the plug to not prematurely resorb or migrate and the rate of pneumothorax.

- Sterility testing must dem- onstrate the sterility of the device and the effects of the sterilization process on the physical characteristics of the plug.

- Shelf-life testing must dem- onstrate the shelf-life of the device in- cluding the physical characteristics of the plug.

#### Food and Drug Administration, HHS	§ 878.4800

- The device must be demonstrated to be biocompatible.

- Labeling must include a detailed summary of the device-related and pro- cedure-related complications pertinent to the use of the device and appropriate warnings. Labeling must include iden- tification of compatible introducer needles.

[79 FR 13219, Mar. 10, 2014]

##### § 878.4760   Removable skin staple.

  - Identification. A removable skin staple is a staple-like device intended to connect external tissues temporarily to aid healing. It is not absorbable.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[53 FR 23872, June 24, 1988, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 878.4780   Powered suction pump.

- Identification. A powered suction pump is a portable, AC-powered or compressed air-powered device in- tended to be used to remove infectious materials from wounds or fluids from a patient’s airway or respiratory support system. The device may be used during surgery in the operating room or at the patient’s bedside. The device may in- clude a microbial filter.

- Classification. Class II.

##### § 878.4790 Powered surgical instru- ment for improvement in the ap- pearance of cellulite.

- Identification. A powered surgical instrument for improvement in the ap- pearance of cellulite is a prescription device that is used for the controlled release of subcutaneous tissue for im- provement in the appearance of cellulite. The device consists of a cut- ting tool powered by a motor and a means for instrument guidance to con- trol the areas of subcutaneous tissue cutting underneath the cellulite de- pressions or dimples.

- Classification. Class II (special controls). The special controls for this device are:

- Non-clinical testing must be per- formed to demonstrate that the device meets all design specifications and per-

formance requirements, and to dem- onstrate durability and mechanical in- tegrity of the device.

- In vivo evaluation of the device must demonstrate device performance, including the safety of the release methodology and blood loss at the treatment  sites.

- All elements of the device that may contact the patient must be dem- onstrated to be biocompatible.

- Electrical safety and electro- magnetic compatibility of the device must be demonstrated.

- The labeling must include a sum- mary of in vivo evaluation data and all the device specific warnings, pre- cautions, and/or contraindications.

- Sterility and shelf-life testing for the device must demonstrate the ste- rility of patient  contacting  compo- nents and the shelf life of these compo- nents.

[79 FR 31861, June 3, 2014]

##### § 878.4800   Manual surgical instrument for general use.

  - Identification. A manual surgical instrument for general use is a non- powered, hand-held, or hand-manipu- lated device, either reusable or dispos- able, intended to be used in various general surgical procedures. The device includes the applicator, clip applier, bi- opsy brush, manual  dermabrasion brush, scrub brush, cannula, ligature carrier, chisel, clamp, contractor, cu- rette, cutter, dissector, elevator, skin graft expander, file, forceps, gouge, in- strument guide, needle guide, hammer, hemostat, amputation hook, ligature passing and knot-tying instrument, knife, blood lancet, mallet,  disposable or reusable aspiration and injection needle, disposable or reusable suturing needle, osteotome, pliers, rasp, re- tainer, retractor, saw, scalpel blade, scalpel handle, one-piece scalpel, snare, spatula, stapler, disposable or reusable stripper, stylet, suturing apparatus for the stomach and intestine, measuring tape, and calipers. A surgical instru- ment that has specialized uses in a spe- cific medical specialty is classified in separate regulations in parts 868 through 892.

  - Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 878.4810

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13828,  Apr.  5,  1989;  59  FR  63010,  Dec.  7,

1994; 66 FR 38803, July 25, 2001]

##### § 878.4810 Laser surgical  instrument for use in general and plastic sur- gery and in dermatology.

    - Identification. (1) A carbon dioxide laser for use in general surgery and in dermatology is a laser device intended to cut, destroy, or remove tissue by light energy emitted by carbon dioxide.

(2) An argon laser for use in derma- tology is a laser device intended to de- stroy or coagulate tissue by light en- ergy emitted by argon.

    - Classification. (1) Class II.

- Class I for special laser gas mix- tures used as a lasing medium for this class of lasers. The devices subject to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38803,  July 25,

2001]

##### § 878.4815   Magnetic	surgical	instru- ment system.

  - Identification. A magnetic surgical instrument system is a prescription de- vice used in laparoscopic surgical pro- cedures consisting of several compo- nents, such as surgical instruments, and a magnetic controller. The mag- netic controller is provided separately from the surgical instrument and is used outside the patient. The external magnetic controller is magnetically coupled with the internal surgical in- strument(s) at the surgical site to grasp, hold, retract, mobilize, or ma- nipulate soft tissue and organs.

  - Classification. Class II (special controls). The special controls for this device are:

    - In vivo performance data must demonstrate that the device  performs as intended  under  anticipated condi- tions of use. Testing must demonstrate the ability of the device to grasp, hold, retract, mobilize, or manipulate soft tissue and organs.

    - Non-clinical performance data must demonstrate that the system per- forms   as   intended   under   anticipated

#### 21 CFR Ch. I (4–1–17 Edition)

conditions of use. The following per- formance characteristics must be test- ed:

- Magnetic field strength testing characterization to identify the dis- tances from the magnet that are safe for patients and users with ferromag- netic implants, devices, or objects.

- Ability of the internal surgical instrument(s) to be coupled, de-cou- pled, and re-coupled with the external magnet over the external magnet use life.

    - The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

    - Performance data must dem- onstrate the sterility of the device components that are patient-con- tacting.

    - Methods and instructions for re- processing reusable components must be validated.

    - Performance data must support shelf life by demonstrating continued sterility of the device or the sterile components and device functionality over the labeled shelf life.

    - Training must be developed and validated by human factors testing and analysis to ensure users can follow the instructions for use to allow safe use of the device.

    - Labeling must include:

      - Magnetic field safe zones.

      - Instructions for proper device use.

      - A screening checklist to ensure that  all  patients  and  operating  staff are screened from bringing ferromag- netic implants, devices, or objects near the  external  magnet.

      - Reprocessing instructions for any reusable components.

      - Shelf life.

      - Use life.

[81 FR 64763, Sept. 21, 2016]

##### § 878.4820   Surgical  instrument  motors and accessories/attachments.

- Identification. Surgical instrument motors and accessories are AC-pow- ered, battery-powered, or air-powered devices intended for use during surgical procedures to provide power to operate various accessories or attachments to cut hard tissue or bone and soft tissue. Accessories or attachments may in- clude     a     bur,     chisel     (osteotome),

#### Food and Drug Administration, HHS	§ 878.5000

dermabrasion brush, dermatome, drill bit, hammerhead, pin driver, and saw blade.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2318, 2000]

##### § 878.4830   Absorbable  surgical gut  su- ture.

- Identification. An absorbable sur- gical gut suture, both plain and chro- mic, is an absorbable, sterile, flexible thread prepared from either the serosal connective tissue layer of beef (bovine) or the submucosal fibrous tissue of sheep (ovine) intestine, and is intended for use in soft tissue approximation.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[54 FR 50738, Dec. 11, 1989, as amended at 68

FR 32984, June 3, 2003]

##### § 878.4840   Absorbable	polydioxanone surgical suture.

- Identification. An absorbable polydioxanone surgical suture is an ab- sorbable,	flexible,	sterile, monofilament thread prepared from polyester polymer poly (p-dioxanone) and is intended for use in soft tissue approximation, including pediatric car- diovascular tissue where growth is ex- pected to occur, and ophthalmic sur- gery. It may be coated or uncoated, undyed or dyed, and with or without a standard needle attached.

- Classification. Class II (special controls). The special control for the device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures;  Guidance  for   Industry   and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[67 FR 77676, Dec. 19, 2002]

##### § 878.4930   Suture retention device.

- Identification. A  suture  retention device is a device, such as a retention bridge, a  surgical  button,  or  a  suture

bolster, intended to aid wound healing by distributing suture tension over a larger area in the patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 878.4950 Manual operating table and accessories and manual operating chair and accessories.

- Identification. A manual operating table and accessories and a manual op- erating chair and accessories are non- powered devices, usually with movable components, intended to be used to support a patient during diagnostic ex- aminations or surgical procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13828,  Apr.  5,  1989;  59  FR  63010,  Dec.  7,

1994; 66 FR 38803, July 25, 2001]

##### § 878.4960 Operating tables and acces- sories and operating chairs and ac- cessories.

- Identification. Operating tables and accessories and operating chairs and accessories are AC-powered or air- powered devices, usually with movable components, intended for use during di- agnostic examinations or surgical pro- cedures to support and position a pa- tient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 878.9.

[55 FR 48440, Nov. 20, 1990, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 878.5000   Nonabsorbable poly(ethylene    terephthalate)    sur- gical suture.

- Identification. Nonabsorbable poly(ethylene terephthalate) surgical suture is a multifilament, nonabsorb- able, sterile, flexible thread prepared from fibers of high molecular weight,

#### § 878.5010

long-chain, linear polyesters having re- current aromatic rings as an integral component and is indicated for use in soft tissue approximation. The poly(ethylene terephthalate) surgical suture meets U.S.P. requirements as described in the U.S.P. Monograph for Nonabsorbable Surgical Sutures;  it may be provided uncoated or coated; and it may be undyed or dyed with an appropriate FDA listed color additive. Also, the suture may be provided with or without a standard needle attached.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[56 FR 24685, May 31, 1991, as amended at 68

FR 32984, June 3, 2003]

##### § 878.5010   Nonabsorbable	poly- propylene surgical suture.

- Identification. Nonabsorbable poly- propylene surgical suture is a monofilament, nonabsorbable, sterile, flexible thread prepared from long- chain polyolefin polymer known as polypropylene and is indicated for use in soft tissue approximation. The poly- propylene surgical suture meets United States Pharmacopeia (U.S.P.) require- ments as described in the U.S.P. Mono- graph for Nonabsorbable Surgical Su- tures; it may be undyed or dyed with an FDA approved color additive; and the suture may be provided with or without a standard needle attached.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[56 FR 24685, May 31, 1991, as amended at 68

FR 32984, June 3, 2003]

##### § 878.5020   Nonabsorbable	polyamide surgical suture.

- Identification. Nonabsorbable poly- amide surgical suture is a nonabsorb- able, sterile, flexible thread prepared from long-chain aliphatic polymers Nylon 6 and Nylon 6,6 and is indicated for use in soft tissue approximation. The  polyamide  surgical  suture  meets

#### 21 CFR Ch. I (4–1–17 Edition)

United States Pharmacopeia (U.S.P.) requirements as described in the U.S.P. monograph for nonabsorbable surgical sutures; it may be monofilament or multifilament in form; it may be pro- vided uncoated or coated; and it may be undyed or dyed with an appropriate FDA listed color additive. Also, the su- ture may be provided with or without a standard needle attached.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[56 FR 24685, May 31, 1991, as amended at 68

FR 32985, June 3, 2003]

##### § 878.5030   Natural  nonabsorbable  silk surgical suture.

- Identification. Natural nonabsorb- able silk surgical suture is a non- absorbable, sterile, flexible multifila- ment thread composed of an organic protein called fibroin. This protein is derived from the domesticated species Bombyx mori (B. mori) of the family Bombycidae. Natural nonabsorbable silk surgical suture is indicated for use in soft tissue approximation. Natural nonabsorbable silk surgical suture meets the United States Pharmacopeia (U.S.P.) monograph requirements for Nonabsorbable Surgical Suture (class I). Natural nonabsorbable silk surgical suture may be braided or twisted; it may be provided uncoated or coated; and it may be undyed or dyed with an FDA listed color additive.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[58 FR 57558, Oct. 26, 1993, as amended at 68

FR 32985, June 3, 2003]

##### § 878.5035 Nonabsorbable expanded polytetrafluoroethylene surgical su- ture.

- Identification. Nonabsorbable ex- panded	polytetrafluoroethylene (ePTFE) surgical suture is a monofilament, nonabsorbable, sterile, flexible  thread  prepared  from  ePTFE

#### Food and Drug Administration, HHS	§ 878.5400

and is intended for use in soft tissue approximation and ligation, including cardiovascular surgery. It may be undyed or dyed with an approved color additive and may be provided with or without an attached needle(s).

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance Document: Surgical Su- tures; Guidance for  Industry  and FDA.’’ See § 878.1(e) for the availability of this guidance document.

[65 FR 20735, Apr. 18, 2000, as amended at 68

FR 32985, June 3, 2003]

##### § 878.5040   Suction lipoplasty system.

- Identification. A suction lipoplasty system is a device intended  for  aes- thetic body contouring. The device consists of a powered suction pump (containing a microbial filter on the exhaust  and  a  microbial   in-line   filter in the connecting tubing between the collection bottle and the safety trap), collection bottle, cannula, and con- necting tube.  The  microbial  filters, tubing, collection bottle, and cannula must be capable of being changed be- tween patients. The powered suction pump has a motor with  a  minimum  of 1⁄3 horsepower, a variable vacuum range from 0  to 29.9  inches  of mercury,  vacu- um control valves to regulate the vacu- um   with    accompanying    vacuum gauges, a single or double rotary vane (with or without oil), a single or double diaphragm, a  single  or  double  piston, and  a  safety  trap.

- Classification. Class II (special controls). Consensus standards and la- beling  restrictions.

[63 FR 7705, Feb. 17, 1998]

## Subpart F—Therapeutic Devices

##### § 878.5070   Air-handling  apparatus  for a surgical operating room.

- Identification. Air-handling appa- ratus for a surgical operating room is a device intended to produce a directed, nonturbulent flow of air that has been filtered to remove particulate matter and microorganisms to provide an area free of contaminants to reduce the pos- sibility of infection in the patient.

- Classification. Class II.

##### § 878.5350   Needle-type epilator.

- Identification. A needle-type epila- tor is a device intended to destroy the dermal papilla of a hair by applying electric current at the tip of a fine nee- dle that has been inserted close to the hair shaft, under the skin, and into the dermal papilla. The electric current may be high-frequency AC current, high-frequency AC combined with DC current, or DC current only.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38803,  July 25,

2001]

##### § 878.5360   Tweezer-type epilator.

- Identification. The tweezer-type epilator is an electrical device intended to remove hair. The energy provided at the tip of the tweezer used to remove hair may be radio frequency, galvanic (direct current), or a combination of radio frequency and galvanic energy.

- Classification. Class I (general con- trols). The device is exempt from pre- market notification  procedures in sub- part E of part 807 of this  chapter  sub- ject to § 878.9.

[63 FR 57060, Oct. 26, 1998]

##### § 878.5400   Low  level  laser  system  for aesthetic use

- Identification. A Low Level Laser System for Aesthetic Use is a device using low level laser energy for the dis- ruption of adipocyte cells within  the fat layer for the release of fat and lipids from these cells for noninvasive aesthetic use.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Guidance for Industry and Food and Drug Administration Staff; Class II Special Controls Guidance Document: Low Level Laser System for Aesthetic Use.’’ See § 878.1(e) for the availability  of  this  guidance  document.

[76 FR 20842, Apr. 14, 2011]

§ 878.5650

##### § 878.5650   Topical oxygen chamber for extremities.

  - Identification. A topical oxygen chamber for extremities is a device that is intended to surround a patient’s limb and apply humidified oxygen topi- cally at a pressure  slightly  greater than atmospheric pressure to aid heal- ing of chronic skin ulcers such as bed- sores.

  - Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance: Topical Oxygen Cham- ber for Extremities.’’ See § 878.1(e) for the availability of this guidance docu- ment.

[76 FR 22807, Apr. 25, 2011]

##### § 878.5900   Nonpneumatic tourniquet.

- Identification. A nonpneumatic tourniquet is a device consisting of a strap or tubing intended to be wrapped around a patient’s limb and tightened to reduce circulation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 54

FR  13828,  Apr.  5,  1989;  59  FR  63010,  Dec.  7,

1994; 66 FR 38803, July 25, 2001]

##### § 878.5910   Pneumatic tourniquet.

- Identification. A pneumatic tour- niquet is an air-powered device con- sisting of a pressure-regulating unit, connecting tubing, and an  inflatable cuff. The cuff is intended to be wrapped around a patient’s limb and inflated to reduce or totally occlude circulation during surgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 878.9.

[53 FR 23872, June 24, 1988, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38803,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

PART 880—GENERAL HOSPITAL AND PERSONAL USE DEVICES

##### Subpart A—General Provisions

Sec.

880.1    Scope.

880.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B [Reserved]

Subpart C—General Hospital and Personal Use Monitoring Devices

880.2200   Liquid   crystal   forehead   tempera- ture  strip.

880.2400   Bed-patient monitor.

880.2420    Electronic monitor for gravity flow infusion   systems.

880.2460   Electrically   powered   spinal   fluid pressure monitor.

880.2500    Spinal fluid manometer. 880.2700    Stand-on patient scale. 880.2720   Patient scale.

880.2740   Surgical sponge scale.

880.2800    Sterilization process indicator. 880.2900    Clinical  color  change  thermometer. 880.2910     Clinical electronic thermometer.

880.2920    Clinical mercury thermometer. 880.2930  Apgar timer.

##### Subparts D–E [Reserved]

Subpart F—General Hospital and Personal Use Therapeutic Devices

880.5025   I.V. container.

880.5045     Medical recirculating air cleaner. 880.5075   Elastic bandage.

880.5090    Liquid bandage.

880.5100   AC-powered adjustable hospital bed. 880.5110    Hydraulic adjustable hospital bed. 880.5120    Manual adjustable hospital bed.

880.5130 Infant radiant warmer. 880.5140 Pediatric medical crib. 880.5145   Medical bassinet.

880.5150    Nonpowered flotation therapy mat- tress.

880.5160   Therapeutic medical binder. 880.5180  Burn sheet.

880.5200    Intravascular catheter.

880.5210    Intravascular   catheter   securement device.

880.5240    Medical adhesive tape and adhesive bandage.

880.5270   Neonatal eye pad. 880.5300   Medical absorbent fiber. 880.5400   Neonatal incubator.

880.5410    Neonatal transport incubator. 880.5420    Pressure infusor for an I.V. bag. 880.5430    Nonelectrically powered fluid injec-

tor.

#### Food and Drug Administration, HHS	§ 880.1

880.5440     Intravascular administration set. 880.5445    Intravascular    administration    set,

automated air removal system.

880.5450  Patient	care	reverse	isolation chamber.

880.5475  Jet lavage.

880.5500   AC-powered patient lift. 880.5510   Non-AC-powered patient lift.

880.5550    Alternating   pressure   air   flotation mattress.

880.5560    Temperature  regulated  water  mat- tress.

880.5570   Hypodermic single lumen needle. 880.5580   Acupuncture needle.

880.5630   Nipple shield. 880.5640   Lamb feeding nipple.

880.5680    Pediatric position holder. 880.5700    Neonatal phototherapy unit. 880.5725  Infusion pump.

880.5740  Suction snakebite kit.

880.5760  Chemical cold pack snakebite kit. 880.5780 Medical support stocking.

880.5820    Therapeutic scrotal support. 880.5860   Piston syringe.

880.5950    Umbilical occlusion device. 880.5960   Lice removal kit.

880.5965  Subcutaneous,	implanted, intravascular   infusion   port   and   catheter. 880.5970    Percutaneous, implanted, long-term

intravascular catheter.

##### Subpart G—General Hospital and Personal Use Miscellaneous Devices

880.6025   Absorbent tipped applicator. 880.6050  Ice bag.

880.6060     Medical disposable bedding. 880.6070   Bed board.

880.6080  Cardiopulmonary	resuscitation board.

880.6085   Hot/cold water bottle.

880.6100    Ethylene oxide gas aerator cabinet. 880.6140   Medical chair and table.

880.6150    Ultrasonic  cleaner  for  medical  in- struments.

880.6175  [Reserved]

880.6185   Cast cover.

880.6190    Mattress   cover   for   medical   pur- poses.

880.6200   Ring cutter.

880.6230    Tongue depressor.

880.6250   Non-powdered  patient  examination glove.

880.6260 Filtering facepiece  respirator  for use by the general public in public health medical emergencies.

880.6265   Examination gown.

880.6280    Medical insole.

880.6300 Implantable radiofrequency trans- ponder system for  patient  identification and   health   information.

880.6305   Ingestible event marker. 880.6310    Medical device data system.

880.6315    Remote    medication    management system.

880.6320    AC-powered    medical    examination light.

880.6350    Battery-powered  medical  examina- tion   light.

880.6375    Patient lubricant.

880.6430   Liquid medication dispenser. 880.6450   Skin pressure protectors.

880.6500   Medical ultraviolet air purifier. 880.6600   Ultraviolet (UV) radiation chamber

disinfection device.

880.6710    Medical ultraviolet water purifier. 880.6730   Body waste receptacle.

880.6740   Vacuum-powered body fluid suction apparatus.

880.6760    Protective restraint.

880.6775    Powered patient transfer device. 880.6785    Manual patient transfer device. 880.6800   Washers for body waste receptacles. 880.6820   Medical disposable scissors.

880.6850   Sterilization wrap.

880.6860    Ethylene oxide gas sterilizer. 880.6870   Dry-heat sterilizer.

880.6880   Steam sterilizer.

880.6885 Liquid	chemical	sterilants/high level    disinfectants.

880.6890    General purpose disinfectants. 880.6900   Hand-carried stretcher.

880.6910  Wheeled stretcher.

880.6920   Syringe needle introducer. 880.6960   Irrigating syringe.

880.6970    Liquid crystal vein locator. 880.6980   Vein stabilizer.

880.6990    Infusion stand.

880.6991    Medical washer.

880.6992   Medical washer-disinfector.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  45  FR  69682,  Oct.  21,  1980,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 880 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 880.1   Scope.

    - This part sets forth the classifica- tion of general hospital and personal use devices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

§ 880.3

    - To avoid duplicative listings, a general hospital and personal use de- vice that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed only in one subpart.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 17738, May 11, 1987, as amended at 69

FR  71704,  Dec.  8,  2004;  78  FR  18233,  Mar.  26,

2013]

##### § 880.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com-

#### 21 CFR Ch. I (4–1–17 Edition)

mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ devices defined in this section because of any new intended use or other rea- sons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17738, May 11, 1987]

##### § 880.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction  into  interstate  commerce

#### Food and Drug Administration, HHS	§ 880.2460

for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2318, Jan. 14, 2000]

## Subpart B [Reserved]

Subpart C—General Hospital and Personal Use Monitoring Devices

##### § 880.2200   Liquid crystal forehead tem- perature strip.

- Identification. A liquid crystal forehead temperature strip is a device applied to the forehead that is used to indicate the presence or absence of fever, or to monitor body temperature changes. The device displays the color changes of heat sensitive liquid crys- tals corresponding to the variation in the surface temperature of the skin. The liquid crystals, which are cho- lesteric esters, are sealed in plastic.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 880.2400   Bed-patient monitor.

- Identification. A bed-patient mon- itor is a battery-powered device placed under a mattress and used  to  indicate by an alarm or other signal when a pa- tient attempts to leave the bed.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63010,  Dec. 7,  1994;  66  FR 38803,  July 25,

2001]

##### § 880.2420   Electronic monitor for grav- ity flow infusion systems.

- Identification. An electronic mon- itor for gravity flow infusion systems is a device used to monitor the amount of fluid being infused into a patient. The device consists of an electronic transducer and equipment for signal amplification, conditioning, and dis- play.

- Classification. Class II (perform- ance standards).

##### § 880.2460   Electrically  powered  spinal fluid pressure monitor.

- Identification. An electrically pow- ered spinal fluid pressure monitor is an

§ 880.2500

electrically powered device used to measure spinal fluid pressure by the use of a transducer which converts spi- nal fluid pressure into an electrical sig- nal. The device includes signal amplifi- cation, conditioning, and display equip- ment.

- Classification. Class II (perform- ance standards).

##### § 880.2500   Spinal fluid manometer.

- Identification. A spinal fluid ma- nometer is a device used to measure spinal fluid pressure. The device uses a hollow needle, which is inserted into the spinal column fluid space, to con- nect the spinal fluid to a graduated col- umn so that the pressure can be meas- ured by reading the height of the fluid.

- Classification. Class II (perform- ance standards).

##### § 880.2700   Stand-on patient scale.

- Identification. A stand-on patient scale is a device intended for medical purposes that is used to weigh a pa- tient who is able to stand on the scale platform.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38803, July 25, 2001]

##### § 880.2720   Patient scale.

- Identification. A patient scale is a device intended for medical purposes that is used to measure the weight of a patient who cannot stand on a scale. This generic device includes devices placed under a bed or chair to weigh both the support and the patient, de- vices where the patient is lifted by a sling from a bed to be weighed, and de- vices where the patient is placed on the scale platform to be weighed. The de- vice may be mechanical, battery pow- ered, or AC-powered and may include transducers, electronic signal amplifi-

#### 21 CFR Ch. I (4–1–17 Edition)

cation, conditioning and display equip- ment.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38803,  July 25,

2001]

##### § 880.2740   Surgical sponge scale.

- Identification. A surgical sponge scale is a nonelectrically powered de- vice used to weigh surgical sponges that have been  used to  absorb blood during surgery so that, by comparison with the known dry weight of the sponges, an estimate may be made of the blood lost by the patient during surgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.2800   Sterilization   process   indi- cator.

- Biological sterilization process indi- cator—(1) Identification. A biological sterilization process indicator is a de- vice intended for use by a health care provider to accompany products being sterilized through a sterilization proce- dure and to monitor adequacy of steri- lization. The device consists of a known number of microorganisms, of known resistance to the mode of steri- lization, in or on a carrier and enclosed in a protective package. Subsequent growth or failure of the microorga- nisms to grow under suitable condi- tions indicates the adequacy of steri- lization.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 880.5075

- Physical/chemical sterilization proc- ess indicator—(1) Identification. A phys- ical/chemical sterilization process indi- cator is a device intended for use by a health care provider to accompany products being sterilized through a sterilization procedure and to monitor one or more parameters of the steri- lization process. The adequacy of the sterilization conditions as measured by these parameters is indicated by a visi- ble change in the device.

  - Classification. Class II (perform- ance standards).

##### § 880.2900   Clinical  color  change  ther- mometer.

    - Identification. A clinical color change thermometer is a disposable de- vice used to measure a patient’s oral, rectal, or axillary (armpit) body tem- perature. The device records body tem- perature by use of heat sensitive chemicals which are sealed at the end of a plastic or metal strip. Body heat causes a stable color change in the heat sensitive chemicals.

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38804,  July 25,

2001]

##### § 880.2910   Clinical electronic thermom- eter.

- Identification. A clinical electronic thermometer is a device used to meas- ure the body temperature of a patient by means of a transducer coupled with an electronic signal amplification, con- ditioning, and display unit. The trans- ducer may be in a detachable probe with or without a disposable cover.

- Classification. Class II (perform- ance standards).

##### § 880.2920   Clinical  mercury  thermom- eter.

- Identification. A clinical mercury thermometer is a device used to meas- ure oral, rectal, or axillary (armpit) body temperature using the thermal expansion of mercury.

- Classification. Class II (special controls). The device is exempt from the  premarket  notification  procedures

in subpart E  of part  807  of  this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59228, Nov. 3, 1998]

##### § 880.2930   Apgar timer.

- Identification. The Apgar timer is a device intended to alert a health care provider to take the Apgar score of a newborn infant.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments in part 820 of this chapter, with the exception of § 820.180 of this chap- ter, with respect to general require- ments concerning records, and § 820.198 of this chapter, with respect to com- plaint  files.

[63 FR 59718, Nov. 5, 1998]

## Subparts D–E [Reserved]

Subpart F—General Hospital and Personal Use Therapeutic Devices

##### § 880.5025   I.V. container.

- Identification. An I.V. container is a container made of plastic or glass used to hold a fluid mixture to be ad- ministered to a patient through an intravascular administration set.

- Classification. Class II (perform- ance standards).

##### § 880.5045   Medical	recirculating	air cleaner.

- Identification. A medical recircu- lating air cleaner is a device used to re- move particles from the air for medical purposes. The device may function by electrostatic precipitation or filtra- tion.

- Classification. Class II (perform- ance standards).

##### § 880.5075   Elastic bandage.

- Identification. An elastic bandage is a device consisting of either a long flat strip or a tube of elasticized mate- rial that is used to support and com- press a part of a patient’s body.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

§ 880.5090

premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5090   Liquid bandage.

  - Identification. A liquid bandage is a sterile device that is a liquid, semiliquid, or powder and liquid com- bination used to cover an opening in the skin or as a dressing for burns. The device is also used as a topical skin protectant.

  - Classification. Class I (general con- trols). When used only as a skin pro- tectant, the device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 880.5100   AC-powered adjustable hos- pital bed.

- Identification. An AC-powered ad- justable hospital bed is a device in- tended for medical purposes that con- sists of a bed with a built-in electric motor and remote controls that can be operated by the patient to adjust the height and surface contour of the bed. The device includes movable and latchable side rails.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.5110   Hydraulic   adjustable   hos- pital bed.

- Identification. A hydraulic adjust- able hospital bed is a device intended for medical purposes that consists of a bed with a hydraulic mechanism oper- ated by an attendant to adjust the height and surface contour of the bed.

#### 21 CFR Ch. I (4–1–17 Edition)

The    device    includes    movable    and latchable  side  rails.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5120   Manual  adjustable  hospital bed.

- Identification. A manual adjust- able hospital bed is a device intended for medical purposes that consists of a bed with a manual mechanism operated by an attendant to adjust the height and surface contour of the bed. The de- vice includes movable and latchable side rails.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR 25050, June 12, 1989; 66 FR 38804, July 25,

2001]

##### § 880.5130   Infant radiant warmer.

- Identification. The infant radiant warmer is a device consisting of an in- frared heating element intended to be placed over an infant to maintain the infant’s body temperature by means of radiant heat. The device may also con- tain a temperature monitoring sensor, a heat output control mechanism, and an alarm system (infant temperature, manual mode if present, and failure alarms) to alert operators of a tem- perature condition over or under the set temperature, manual mode time limits, and device component failure, respectively. The device may be placed over a pediatric hospital bed or it may be built into the bed as a complete unit.

- Classification. Class  II (Special Controls):

#### Food and Drug Administration, HHS	§ 880.5145

  - The Association for the Advance- ment of Medical Instrumentation (AAMI) Voluntary Standard for the In- fant Radiant Warmer;

  - A prescription statement in ac- cordance with § 801.109 of this chapter (restricted to use by  or  upon  the  order of  qualified  practitioners  as   deter- mined  by  the  States);  and

  - Labeling for use only in health care facilities and only by persons with specific training and experience in the use of the device.

[62 FR 33350, June 19, 1997]

##### § 880.5140   Pediatric medical crib.

    - Identification. A pediatric medical crib is a prescription device intended for medical purposes for use with a pe- diatric patient that consists of an open crib, fixed end rails, movable and latchable side rail components, and possibly an accompanying mattress. The contour of the crib surface may be adjustable.

    - Classification. Class II (special controls). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 880.9. The special controls for this device are:

- Crib design and performance test- ing shall demonstrate the mechanical and structural stability of the crib under expected conditions of use, in- cluding the security of latches and other locking mechanisms when en- gaged;

- Materials used shall be appro- priate for the conditions of use,  allow for proper sanitation, and be free from surface defects that could result in in- juries;

- The height of the rail and end panel as measured from the top of the rail or panel in its highest  position  to the top of the mattress support in its lowest position shall be at least  26 inches (66 centimeters). Any mattress used in this crib must not exceed a thickness of 6 inches;

- Hardware and fasteners shall be designed and constructed to eliminate mechanical hazards to the patient;

- The distance between components of the side rail (i.e., slats, spindles, and corner posts) shall not be greater than 23⁄8 inches (6 centimeters) apart at any point;

- The mattress must fit tightly around all four sides of the crib base, such that entrapment or  impingement of occupant is prevented;

- The mattress for the crib shall meet the Consumer Product Safety Commission (CPSC) Standard for the flammability of mattresses and mat- tress pads (FF 4–72, amended) and Standard for the flammability (open flame) of mattress sets, 16 CFR parts 1632 and 1633, respectively; and

- Each device must have the fol- lowing label(s) affixed:

- Adequate instructions for users to care for, maintain, and clean the crib; and

- A warning label on at least two sides of the medical crib with the fol- lowing language in text of at least 9 millimeters  in  height:

WARNING: Never leave a child unsu- pervised when the moveable side is open or not secured.

[81 FR 91737, Dec. 19, 2016]

##### § 880.5145   Medical bassinet.

- Identification. A medical  bassinet is a prescription device that is a small bed intended for use with pediatric pa- tients, generally from birth to approxi- mately 5 months of age. It is intended for medical purposes for use in a nurs- ery, labor and delivery unit, or patient room, but may also be used outside of traditional health care settings. A medical bassinet is a non-powered de- vice that consists of two components: The plastic basket or bed component and a durable frame with wheels, which holds the basket  or  bed  component. The basket or bed component is a box- like structure, generally made of a clear, high impact-resistant plastic material, with an open top and four stationary walls to hold the pediatric patient. The frame can include draw- ers, shelving, or cabinetry that  pro- vides space to hold infant care items. The wheels or casters allow the bassi- net to transport the infant throughout the  care  setting.

- Classification. Class II (special controls). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 880.9. The special controls for this device are:

#### § 880.5150

  - The manufacturer must conduct performance testing to determine ma- terial compatibility with cleansing products labeled to clean the device. Testing must demonstrate that the cleaning instructions provided by the manufacturer do not cause crazing, cracking, or deterioration of the de- vice;

  - Manufacturers shall conduct per- formance testing to ensure the me- chanical and structural stability of the bassinet under expected conditions of use, including transport of patients in the bassinet. Testing must dem- onstrate that failures such as wheel or caster breakage do not occur and that the device does not present a tipping hazard due to any mechanical failures under expected conditions of use; and

  - Each device must have the fol- lowing label(s) affixed:

- Adequate instructions for users to care for, maintain, and clean the bassi- net; and

- A warning label on at least two sides of the plastic basket or bed com- ponent with the following language in text of at least 9 millimeters in height:

WARNING: To avoid tipping hazards of this device, make sure that the bas- ket or bed component sits firmly in the base and that all doors, drawers, and casters are secure.

[81 FR 91737, Dec. 19, 2016]

##### § 880.5150   Nonpowered  flotation  ther- apy mattress.

- Identification. A nonpowered flota- tion therapy mattress is a mattress in- tended for medical purposes which con- tains air, fluid, or other materials that have the functionally equivalent effect of supporting a patient and avoiding excess pressure on local body areas. The device is intended to treat or pre- vent decubitus ulcers (bed sores).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### 21 CFR Ch. I (4–1–17 Edition)

and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5160   Therapeutic medical binder.

- Identification. A therapeutic med- ical binder is a device, usually made of cloth, that is intended for medical pur- poses and that can be secured by ties so that it supports the underlying part of the body or holds a dressing in place. This generic type of device includes the abdominal binder, breast binder, and perineal binder.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5180   Burn sheet.

- Identification. A burn sheet is a de- vice made of a porous material that is wrapped aroung a burn victim to retain body heat, to absorb wound exudate, and to serve as a barrier against con- taminants.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38804,  July 25,

2001]

##### § 880.5200   Intravascular catheter.

- Identification. An intravascular catheter is a device that consists of a slender tube and any necessary con- necting fittings and  that  is  inserted into the patient’s vascular system for short term use (less than 30 days) to sample blood, monitor blood  pressure, or administer fluids intravenously. The device  may  be  constructed  of  metal,

#### Food and Drug Administration, HHS	§ 880.5400

rubber,  plastic,  or  a  combination  of these materials.

- Classification. Class II (perform- ance standards).

##### § 880.5210   Intravascular   catheter   se- curement device.

- Identification. An intravascular catheter securement device is a device with an adhesive backing that is placed over a needle or catheter and is used to keep the hub of the needle or the cath- eter flat and securely anchored to the skin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38804,  July 25,

2001]

##### § 880.5240   Medical  adhesive  tape  and adhesive bandage.

- Identification. A medical adhesive tape or adhesive bandage is a device in- tended for medical purposes that con- sists of a strip of fabric material or plastic, coated on one side with an ad- hesive, and may include a pad of sur- gical dressing without a disinfectant. The device is used to cover and protect wounds, to hold together the skin edges of a wound, to support an injured part of the body, or to secure objects to the skin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38804,  July 25,

2001]

##### § 880.5270   Neonatal eye pad.

- Identification. A neonatal eye  pad is an opaque device used to cover and protect the eye of an infant during therapeutic procedures, such as phototherapy.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 880.9. If the device is not la- beled or otherwise represented as ster- ile, it is exempt from the current good

manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,   with   respect   to   complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 880.5300   Medical absorbent fiber.

- Identification. A medical absorbent fiber is a device intended for medical purposes that is made from cotton or synthetic fiber in the shape of a ball or a pad and that is used for applying medication to, or absorbing small amounts of body fluids from, a pa- tient’s body surface. Absorbent fibers intended solely for cosmetic purposes are not included in this generic device category.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5400   Neonatal incubator.

- Identification. A neonatal incu- bator is a device consisting of a rigid boxlike enclosure in which an infant may be kept in a controlled environ- ment for medical care. The device may include an AC-powered heater, a fan to circulate the warmed air, a container for water to add humidity, a control valve through which oxygen may be added, and access ports for nursing care.

- Classification. Class II (perform- ance standards).

#### § 880.5410

##### § 880.5410   Neonatal	transport	incu- bator.

  - Identification. A neonatal trans- port incubator is a device consisting of a portable rigid boxlike enclosure with insulated walls in which an infant may be kept in a controlled environment while being transported for medical care. The device may include straps to secure the infant, a battery-operated heater, an AC-powered battery charger, a fan to circulate the warmed air, a container for water to add humidity, and provision for a portable oxygen bottle.

  - Classification. Class II (perform- ance standards).

##### § 880.5420   Pressure infusor for an I.V. bag.

- Identification. A pressure infusor for an I.V. bag is a device consisting of an inflatable cuff which is placed around an I.V. bag. When the device is inflated, it increases the pressure on the I.V. bag to assist the  infusion  of the fluid.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 65

FR 2318, Jan. 14, 2000]

##### § 880.5430   Nonelectrically	powered fluid injector.

- Identification. A nonelectrically powered fluid injector is a nonelec- trically powered device used by a health care provider to give a hypo- dermic injection by means of a narrow, high velocity jet of fluid which can penetrate the surface of the skin and deliver the fluid to the body. It may be used for mass inoculations.

- Classification. Class II (perform- ance standards).

##### § 880.5440   Intravascular	administra- tion set.

- Identification. An intravascular administration set is a device used to administer fluids from a container to a patient’s vascular system through a needle or catheter inserted into a vein. The device may include the needle or catheter, tubing, a flow regulator, a drip chamber, an infusion line filter, an

#### 21 CFR Ch. I (4–1–17 Edition)

I.V. set stopcock, fluid delivery tubing, connectors between parts of the set, a side tube with a cap to serve as an in- jection site, and a hollow spike to pen- etrate  and  connect  the  tubing  to  an

I.V. bag or other infusion fluid con- tainer.

- Classification. Class II (special controls). The special control for phar- macy compounding systems within this classification is the FDA guidance doc- ument entitled ‘‘Class II Special Con- trols Guidance Document: Pharmacy Compounding Systems; Final Guidance for Industry and FDA Reviewers.’’ Pharmacy compounding systems clas- sified within the intravascular admin- istration set are exempt from the pre- market notification procedures in sub- part E of this part and subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 15798, Mar. 21, 2001]

##### § 880.5445 Intravascular administra- tion set, automated air removal sys- tem.

- Identification. An intravascular administration set, automated air re- moval system, is a prescription device used to detect and automatically re- move air from an intravascular admin- istration set with minimal to no inter- ruption in the flow of the intravascular fluid. The device may include an air identification mechanism, software, an air removal mechanism, tubing, appa- ratus to collect removed air, and safety control mechanisms to address haz- ardous situations.

- Classification. Class II (special controls). The special controls for this device are:

  - Provide an argument dem- onstrating that all reasonably foresee- able hazards have been adequately ad- dressed with respect to the persons for whose use the device is represented or intended and the conditions of use for the device, which includes the fol- lowing:

- Description of the device indica- tions for use, design, and technology, use environments, and users in suffi- cient detail to determine that the de- vice  complies  with  all  special  controls.

- Demonstrate that controls are implemented to address device system hazards and their causes.

#### Food and Drug Administration, HHS	§ 880.5510

- Include a justification sup- porting the acceptability criteria for each hazard control.

- A traceability analysis dem- onstrating that all credible hazards have at least one corresponding control and that all controls have been verified and validated in the final device de- sign.

  - Appropriate software verification, validation,  and  hazard  analysis  must be performed.

  - The device parts that directly or indirectly contact the patient must be demonstrated to be biocompatible.

  - Performance data must dem- onstrate the sterility of fluid path con- tacting components and the  shelf  life of these components.

  - The device must be designed and tested for electrical safety and electro- magnetic compatibility (EMC).

  - Nonclinical performance testing data must demonstrate that the device performs as intended under anticipated conditions of use. The following per- formance characteristics must be test- ed:

- Device system and component re- liability testing must be conducted.

- Fluid ingress protection testing must be conducted.

- Testing of safety controls must be performed to demonstrate adequate mitigation of hazardous situations, in- cluding sensor failure, flow control failure, improper device position, de- vice malfunction,  infusion  delivery error, and release of air to the patient.

  - A human factors validation study must demonstrate that use hazards are adequately addressed.

  - The labeling must include the fol- lowing:

    - The device’s air identification and removal response time.

    - The device’s minimum air volume identification  sensitivity.

    - The minimum and maximum flow rates at which the device is capa- ble of reliably detecting and removing air.

    - Quantification of any fluid loss during device air removal operations as a function of flow rate.

[79 FR 28406, May 16, 2014]

##### § 880.5450   Patient  care  reverse  isola- tion chamber.

- Identification. A patient care re- verse isolation chamber is a device consisting of a roomlike enclosure de- signed to prevent the entry of harmful airborne material. This device protects a patient who is undergoing treatment for burns or is lacking a normal im- munosuppressive defense due to ther- apy or congenital abnormality. The de- vice includes fans and air filters which maintain an atmosphere of clean air at a pressure greater than the air pressure outside the enclosure.

- Classification. Class II (perform- ance standards).

##### § 880.5475   Jet lavage.

- Identification. A jet lavage is a de- vice used to clean a wound by a pulsatile jet of sterile fluid. The device consists of the pulsing head, tubing to connect to a container of sterile fluid, and a means of propelling the fluid through the tubing, such as an electric roller pump.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.5500   AC-powered patient lift.

- Identification. An AC-powered lift is an electrically powered device either fixed or mobile, used to lift and trans- port patients in the horizontal or other required position from one place to an- other, as from a bed to a bath. The de- vice includes straps and slings to sup- port the patient.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.5510   Non-AC-powered	patient lift.

- Identification. A non-AC-powered patient lift is a hydraulic, battery, or mechanically  powered  device,  either

#### § 880.5550

fixed or mobile, used to lift and trans- port a patient in the horizontal or other required position from one place to another, as from a bed to a  bath. The device includes straps and a sling to support the patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR 25050, June 12, 1989; 66 FR 38804, July 25,

2001]

##### § 880.5550   Alternating pressure air flo- tation mattress.

- Identification. An alternating pres- sure air flotation mattress is a device intended for medical purposes that con- sists of a mattress with multiple air cells that can be filled and emptied in an alternating pattern by an associated control unit to provide regular, fre- quent, and automatic changes in the distribution of body pressure. The de- vice is used to prevent and treat decubitus ulcers (bed sores).

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.5560   Temperature	regulated water mattress.

- Identification. A temperature regu- lated water mattress is a device in- tended for medical purposes that con- sists of a mattress of suitable size, filled with water which can be heated or in some cases cooled. The device in- cludes electrical heating and water cir- culating components, and an optional cooling component. The temperature control may be manual or automatic.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38804,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 880.5570   Hypodermic   single   lumen needle.

- Identification. A hypodermic single lumen needle is a device intended to in- ject fluids into, or withdraw  fluids from, parts of the body below the sur- face of the skin. The device consists of a metal tube that is sharpened at one end and at the other end joined to a fe- male connector (hub) designed to mate with a male connector (nozzle) of a pis- ton syringe or an intravascular admin- istration set.

- Classification. Class II (perform- ance standards).

##### § 880.5580   Acupuncture needle.

- Identification. An acupuncture needle is a device intended to pierce the skin in the practice of acupunc- ture. The device consists of a solid, stainless steel needle. The device may have a handle attached to the needle to facilitate the delivery of acupuncture treatment.

- Classification. Class II (special controls). Acupuncture needles must comply with the following special con- trols:

- Labeling for single use only and conformance to the requirements for prescription devices set out in 21 CFR 801.109,

- Device material biocompatibility, and

- Device sterility.

[61 FR 64617, Dec. 6, 1996]

##### § 880.5630   Nipple shield.

  - Identification. A nipple shield is a device consisting of a cover used to protect the nipple of a nursing woman. This generic device does not include nursing pads intended solely to protect the clothing of a nursing woman from milk.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 33804,  July 25,

2001]

##### § 880.5640   Lamb feeding nipple.

- Identification. A lamb feeding nip- ple  is  a  device  intended  for  use  as  a

#### Food and Drug Administration, HHS	§ 880.5760

feeding nipple for infants with oral or facial abnormalities.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38804, July 25, 2001]

##### § 880.5680   Pediatric position holder.

- Identification. A pediatric position holder is a device used to hold an in- fant or a child in a desired position for therapeutic or diagnostic purposes, e.g., in a crib under a radiant warmer, or to restrain a child while an intravascular injection is adminis- tered.

- Classification. Class I (general con- trols). The device is exempt from the good manufacturing   practice  regula- tion in part 820 of this  chapter,  with the exception of § 820.180,  with  respect to general requirements concerning records, and § 820.198, with respect to complaint    files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 46952, Sept. 10, 2001]

##### § 880.5700   Neonatal phototherapy unit.

- Identification. A neonatal phototherapy unit is a device used to treat or prevent hyperbilirubinemia (elevated serum bilirubin level). The device consists of one or more lamps that emit a specific spectral band of light, under which an infant is placed for therapy. This generic type of device may include supports for the patient and equipment and component parts.

- Classification. Class II (perform- ance standards).

##### § 880.5725   Infusion pump.

- Identification. An infusion pump is a device used in a health care facility to pump fluids into a patient in a con- trolled manner. The device may use a

piston pump, a roller pump, or a peri- staltic pump and may be powered elec- trically or mechanically. The device may also operate using a  constant force to propel the fluid through a nar- row tube which determines the flow rate. The device may include means to detect a fault condition, such as air in, or blockage of, the infusion line and to activate an alarm.

- Classification. Class II (perform- ance standards).

##### § 880.5740   Suction snakebite kit.

- Identification. A suction snakebite kit is a device consisting of a knife, suction device, and tourniquet used for first-aid treatment of snakebites by re- moving venom from the wound.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38805,  July 25,

2001]

##### § 880.5760   Chemical  cold  pack  snake- bite kit.

- Identification. A chemical cold pack snakebit kit is a device consisting of a chemical cold pack and tourniquet used for first-aid treatment of snakebites.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for  any  chemical  cold  pack   snakebite kit that was  in  commercial  distribu- tion before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equivalent to a chemical cold  pack  snakebite  kit  that was in commercial distribution  before May 28, 1976. Any other chemical cold pack snakebite kit shall have an ap- proved PMA or  a  declared  completed PDP in effect before being placed in commercial     distribution.

[45 FR 69682, Oct. 21, 1980, as amended at 52

FR 17739, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

#### § 880.5780

##### § 880.5780   Medical support stocking.

  - Medical support stocking to prevent the pooling of blood in the legs—(1) Iden- tification. A medical support  stocking to prevent the pooling of blood in the legs is a device that is constructed of elastic material and designed to apply controlled pressure to the leg and that is intended for use in the prevention of pooling of blood in the leg.

(2) Classification. Class II (perform- ance standards).

  - Medical support stocking for general medical purposes—(1) Identification. A medical support stocking for general medical purposes is a device that is constructed of elastic material and de- signed to apply controlled pressure to the leg and that is intended for medical purposes other than the prevention of pooling of blood in the leg.

- Classification.  Class  I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter, subject  to  the  limita- tions in § 880.9. The device is also ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this    chapter,    with    the    exception    of

§ 820.180,   with   respect   to   general   re- quirements     concerning     records,     and

§ 820.198,    with    respect    to    complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38805,  July 25,

2001]

##### § 880.5820   Therapeutic scrotal support.

  - Identification. A therapeutic scro- tal support is a device intended for medical purposes that consist of a pouch attached to an elastic waistband and that is used to support the scrotum (the sac that contains the testicles).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### 21 CFR Ch. I (4–1–17 Edition)

and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.5860   Piston syringe.

- Identification. A piston syringe is a device intended for medical purposes that consists of a calibrated hollow barrel and a movable plunger. At one end of the barrel there is a male con- nector (nozzle) for fitting the female connector (hub) of a hypodermic single lumen needle. The device is used to in- ject fluids into, or withdraw  fluids from, the body.

- Classification. Class II (perform- ance standards).

##### § 880.5950   Umbilical occlusion device.

- Identification. An umbilical occlu- sion device is a clip, tie, tape, or other article used to close the blood vessels in the umbilical cord of a newborn in- fant.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38805,  July 25,

2001]

##### § 880.5960   Lice removal kit.

- Identification. The lice removal kit is a comb or comb-like device in- tended to remove and/or kill lice and nits from head and body hair. It may or may not be battery operated.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9.

[63 FR 59718, Nov. 5, 1998]

##### § 880.5965 Subcutaneous, implanted, intravascular infusion port and catheter.

- Identification. A subcutaneous, im- planted, intravascular infusion port and catheter is a device that consists of a subcutaneous, implanted reservoir that connects to a long-term intravascular catheter. The device al- lows for repeated access to the vascular system  for  the  infusion  of  fluids  and

#### Food and Drug Administration, HHS	§ 880.6060

medications and the sampling of blood. The device consists of a portal body with a resealable septum and outlet made of metal, plastic, or combination of these materials and a long-term intravascular catheter is either preattached to the port or attached to the port at the time of device place- ment. The device is available in var- ious profiles and sizes and can be of a single or multiple lumen design.

- Classification.   Class    II    (special controls)  Guidance  Document:  ‘‘Guid- ance  on   510(k)  Submissions   for  Im- planted  Infusion  Ports,’’  FDA  October 1990.

[65 FR 37043, June 13, 2000]

##### § 880.5970 Percutaneous, implanted, long-term intravascular catheter.

- Identification. A percutaneous, im- planted, long-term intravascular cath- eter is a device that consists of a slen- der tube and any necessary connecting fittings, such as luer hubs, and acces- sories that facilitate the placement of the device. The device allows for re- peated access to the vascular system for long-term use of 30 days or  more, and it is intended for administration of fluids, medications, and nutrients; the sampling of blood;  and  monitoring blood pressure and temperature. The device may be constructed of metal, rubber, plastic, composite materials, or any combination of  these  materials and may be of single or multiple lumen design.

- Classification.   Class    II    (special controls)  Guidance  Document:  ‘‘Guid- ance	on	Premarket	Notification [510(k)] Submission for Short-Term and Long-Term   Intravascular   Catheters.’’

[65 FR 37043, June 13, 2000]

## Subpart G—General Hospital and Personal Use Miscellaneous Devices

##### § 880.6025 Absorbent tipped appli- cator.

- Identification. An absorbent tipped applicator is a device intended for med- ical purposes that consists of an ab- sorbent swab on a wooden, paper, or plastic stick. The device is used to apply medications to, or to take speci- mens from, a patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.6050   Ice bag.

- Identification. An ice bag is a de- vice intended for medical purposes that is in the form of a container intended to be filled with ice that is used to apply dry cold therapy to an  area  of the body. The device may include a holder that keeps the bag in place against an external area of the patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.6060   Medical disposable bedding.

- Identification. Medical disposable bedding is a device intended for med- ical purposes to be used by one patient for a period of time and then discarded. This generic type of device may in- clude disposable bedsheets, bedpads, pillows and pillowcases, blankets, emergency rescue blankets, or water- proof sheets.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject  to  the  limitations  in  § 880.9.  If

§ 880.6070

the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38805,  July 25,

2001]

##### § 880.6070   Bed board.

  - Identification. A bed board is a de- vice intended for medical purposes that consists of a stiff board used to in- crease the firmness of a bed.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.6080   Cardiopulmonary  resuscita- tion board.

- Identification. A cardiopulmonary resuscitation board is a device con- sisting of a rigid board which is placed under a patient to act as a support dur- ing  cardiopulmonary  resuscitation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 880.6085   Hot/cold water bottle.

- Identification. A hot/cold water bottle is a device intended for medical purposes that is in the form of a con- tainer intended to be filled with hot or cold water to apply heat or cold to an area of the body.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.6100   Ethylene  oxide  gas  aerator cabinet.

- Identification. An ethyene oxide gas aerator cabinet is a device that is intended for use by a health care pro- vider and consists of a cabinet with a ventilation system designed to cir- culate and exchange the air in the cabi- net to shorten the time required to re- move residual ethylene oxide (ETO) from wrapped medical devices that have undergone ETO sterilization. The device may include a heater to warm the circulating air.

- Classification. Class II (perform- ance standards).

##### § 880.6140   Medical chair and table.

- Identification. A medical chair or table is a device intended for medical purposes that consists of a chair or table without wheels and not elec- trically powered which, by reason of special shape or attachments, such as food trays or headrests, or special fea- tures such as a built-in raising and lowering mechanism or removable arms, is intended for use of blood do- nors, geriatric patients, or patients un- dergoing treatment or examination.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The

#### Food and Drug Administration, HHS	§ 880.6230

device is also exempt from the current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38805, July 25, 2001]

##### § 880.6150   Ultrasonic cleaner for med- ical instruments.

  - Identification. An ultrasonic clean- er for medical instruments is a device intended for cleaning medical instru- ments by the emission of high fre- quency soundwaves.

  - Classification. Class I. The device, including any solutions  intended  for use with the device for cleaning and sanitizing the instruments, is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR  25050,  June  12,  1989;  59  FR  63011,  Dec.  7,

1994; 66 FR 38805, July 25, 2001]

##### § 880.6175   [Reserved]

§ 880.6185   Cast cover.

- Identification. A cast cover is a de- vice intended for medical purposes that is made of waterproof material and placed over a cast to protect it from getting wet during a shower or a bath.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6190   Mattress  cover  for  medical purposes.

- Identification. A mattress cover for medical purposes is a device in- tended for medical purposes that is used to protect a mattress. It may be electrically conductive or contain a germicide.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38806,  July 25,

2001]

##### § 880.6200   Ring cutter.

- Identification. A ring cutter is a device intended for medical purposes that is used to cut a ring on a patient’s finger so that the ring can be removed. The device incorporates a guard to pre- vent injury to the patient’s finger.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6230   Tongue depressor.

- Identification. A tongue depressor is a device intended to displace the tongue to facilitate examination of the surrounding organs and tissues.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart  E  of  part  807  of  this  chapter,

#### § 880.6250

subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect  to  complaint  files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6250   Non-powdered   patient   ex- amination glove.

  - Identification. A non-powdered pa- tient examination glove is a disposable device intended for medical purposes that is worn on the examiner’s hand or finger to prevent contamination be- tween patient and examiner. A non- powdered patient examination glove does not incorporate powder for pur- poses other than manufacturing. The final finished glove includes only resid- ual powder from manufacturing.

  - Classification. Class I (general con- trols).

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR 1604,  Jan. 13,  1989;  66  FR 46952,  Sept. 10,

2001; 81 FR 91730, Dec. 19, 2016]

##### § 880.6260 Filtering facepiece res- pirator for use by the general pub- lic in public health medical emer- gencies.

- Identification. A filtering facepiece respirator for use by the general public in public health medical emergencies is a device that is a disposable half-face- piece non-powered air-purifying partic- ulate respirator intended for use to cover the nose and mouth of the wearer to help reduce wearer exposure to path- ogenic biological airborne particulates during a public health medical emer- gency. The device is made of polymeric materials and is intended to fit closely to the face and to function by filtering particulate  material.

- Classification. Class II (special controls). The special controls are:

- Certification by the National In- stitute for Occupational Safety and Health (NIOSH) as a non-powered air- purifying particulate respirator with a minimum filtration efficiency classi- fication of N95, in accordance with 42 CFR part 84.

#### 21 CFR Ch. I (4–1–17 Edition)

- The FDA guidance document enti- tled: ‘‘Guidance for Industry and Food and Drug Administration Staff; Class II Special Controls Guidance Document: Filtering Facepiece Respirator for use by the General Public in Public Health Medical Emergencies.’’ See § 880.1(e) for information on obtaining a copy of this guidance document.

[72 FR 36362, July 3, 2007]

##### § 880.6265   Examination gown.

- Identification. An examination gown is a device intended for medical purposes that is made of cloth, paper, or other material that is draped over or worn by a patient as a body covering during a medical examination.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6280   Medical insole.

- Identification. A medical insole is a device intended for medical purposes that is placed inside a shoe to relieve the symptoms of athlete’s foot infec- tion by absorbing moisture.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR 25050, June 12, 1989; 66 FR 38806, July 25,

2001]

##### § 880.6300 Implantable radiofrequency transponder system for patient identification and health informa- tion.

- Identification. An implantable ra- diofrequency transponder system for patient identification and health infor- mation is a device intended to enable

#### Food and Drug Administration, HHS	§ 880.6315

access to secure patient identification and corresponding health information. This system may include a passive im- planted transponder, inserter, and scanner. The implanted transponder is used only to store a unique electronic identification code that is read by the scanner. The identification code is used to access patient identity and cor- responding health information stored in a database.

- Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Implantable Radiofrequency Trans- ponder System for Patient Identifica- tion   and   Health   Information.’’   See

§ 880.1(e) for the availability of this guidance document. This device is ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject  to  the  limitations in § 880.9.

[69 FR 71704, Dec. 10, 2004]

##### § 880.6305   Ingestible event marker.

- Identification. An ingestible event marker is a prescription device used to record time-stamped, patient-logged events. The ingestible component links wirelessly through intrabody commu- nication to an external recorder which records the date and time of ingestion as well as the unique serial number of the ingestible device.

- Classification. Class II (special controls). The special controls for this device are:

- The device must be demonstrated to be biocompatible and non-toxic;

- Nonclinical, animal, and clinical testing must provide a reasonable as- surance of safety and effectiveness, in- cluding device performance, durability, compatibility, usability (human fac- tors testing), event recording, and proper excretion of the device;

- Appropriate analysis and nonclin- ical testing must validate electro- magnetic compatibility performance, wireless performance, and electrical safety; and

- Labeling must include a detailed summary of the nonclinical and clin- ical testing pertinent to use of the de-

vice and the maximum number of daily device ingestions.

[78 FR 28734, May 16, 2013]

##### § 880.6310   Medical device data system.

- Identification. (1) A medical device data system (MDDS) is a device that is intended to provide one or more of the following uses, without controlling or altering the functions or parameters of any connected medical devices:

- The electronic transfer of medical device data;

- The electronic storage of medical device data;

- The electronic conversion of medical device data from one format to another format in accordance with a preset specification; or

- The electronic display of medical device data.

(2) An MDDS may include software, electronic or electrical hardware such as a physical communications medium (including wireless hardware), modems, interfaces, and a communications pro- tocol. This identification does not in- clude devices intended to be used in connection with active patient moni- toring.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[76 FR 8649, Feb. 15, 2011]

##### § 880.6315   Remote Medication Manage- ment System.

- Identification. A remote medica- tion management system is a device composed of clinical and communica- tions software, a medication delivery unit, and medication packaging. The system is intended to store the pa- tient’s prescribed medications in a de- livery unit, to permit a health care professional to remotely schedule the patient’s prescribed medications, to notify the patient when the prescribed medications are due to be taken, to re- lease the prescribed medications to a tray of the delivery unit accessible to the patient on the patient’s command, and to record a history of the event for the health care professional. The sys- tem is intended for use as an aid to health care professionals in managing

#### § 880.6320

therapeutic  regimens  for  patients  in the home or clinic.

- Classification. Class II (special controls). The special control is: The FDA guidance document entitled ‘‘Guidance for Industry and Food and Drug Administration Staff; Class II Special Controls Guidance Document: Remote Medication Management Sys- tem.’’ See § 880.1(e) for availability of this guidance document.

[72 FR 59177, Oct. 19, 2007]

##### § 880.6320   AC-powered  medical  exam- ination light.

- Identification. An AC-powered medical examination light is an AC- powered device intended for medical purposes that is used to illuminate body surfaces and cavities during a medical examination.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38806,  July 25,

2001]

##### § 880.6350   Battery-powered	medical examination light.

- Identification. A battery-powered medical examination light is a battery- powered device intended for medical purposes that is used to illuminate body surfaces and cavities during a medical examination.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6375   Patient lubricant.

- Identification. A patient lubricant is a device intended for medical pur- poses that is used to lubricate a body

#### 21 CFR Ch. I (4–1–17 Edition)

orifice  to  facilitate  entry  of  a  diag- nostic or therapeutic device.

- Classification. Class I (general con- trols).

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 46952, Sept. 10, 2001]

##### § 880.6430   Liquid	medication	dis- penser.

- Identification. A Liquid medica- tion dispenser is a device intended for medical purposes that is used to issue a measured amount of liquid medication.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6450   Skin pressure protectors.

- Identification. A skin pressure pro- tector is a device intended for medical purposes that is used to reduce pres- sure on the skin over a bony promi- nence to reduce the likelihood of the patient’s developing decubitus ulcers (bedsores).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38806, July 25, 2001]

##### § 880.6500   Medical ultraviolet air puri- fier.

- Identification. A medical ultra- violet air purifier is a device intended

#### Food and Drug Administration, HHS	§ 880.6740

for medical purposes that is used to de- stroy bacteria in the air by exposure to ultraviolet radiation.

- Classification. Class II (perform- ance standards).

##### § 880.6600   Ultraviolet   (UV)   radiation chamber disinfection device.

- Identification. An ultraviolet (UV) radiation chamber disinfection  device is intended for the low-level surface disinfection of non-porous equipment surfaces by dose-controlled UV irradia- tion. This classification does not in- clude self-contained open chamber UV radiation disinfection devices intended for whole room disinfection in a health care  environment.

- Classification—Class II (special controls). The special controls for this device are:

- Performance testing must dem- onstrate the following:

  - The chamber’s ability to control the UV radiation dose during oper- ation.

  - The chamber’s disinfection per- formance through microbial challenge testing.

  - Evidence that  the  equipment in- tended to be processed is UV compat- ible.

  - Validation of the cleaning and disinfection procedures.

  - The ability of the device to con- tinue to perform to all specification after cleaning and disinfection.

  - Whether the device generates ozone (if so, 21 CFR 801.415, Maximum acceptable level of ozone, applies).

- Appropriate software verification, validation,  and  hazard  analysis  must be performed.

- Appropriate analysis and/or test- ing must validate electrical safety, me- chanical safety, and electromagnetic compatibility of the device in its in- tended use environment.

- The labeling must include:

  - UV hazard warning labels.

  - Explanation of all displays and/or labeling on user interface.

  - Explanation of device safety interlocks.

  - Explanation of all disinfection cycle signals, cautions and warnings.

  - Device operating procedures.

  - Identification of the expected UV lamp operational life and instructions

for  procedures  on  replacement  of  the UV lamp when needed.

  - Procedures to follow in case of UV lamp malfunction or failure.

  - Procedures for disposing of mercury-containing UV lamps, if appli- cable.

  - Identification of specific equip- ment that is compatible with the UV radiation dose generated by the device and that can safely undergo UV radi- ation low-level disinfection in the chamber device.

  - Description of the required prepa- ration of equipment for disinfection in the UV radiation chamber device.

  - Identification of the specific mi- crobes used in successful performance testing of the device.

  - Validated instructions for clean- ing and disinfection of the device.

[80 FR 72588, Nov. 20, 2015]

##### § 880.6710   Medical    ultraviolet    water purifier.

- Identification. A medical ultra- violet water purifier is a device in- tended for medical purposes that is used to destroy bacteria in water by exposure to ultraviolet radiation.

- Classification. Class II (perform- ance standards).

##### § 880.6730   Body waste receptacle.

- Identification. A body waste recep- tacle is a device intended for medical purposes that is not attached to the body and that is used to collect the body wastes of a bed patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[66 FR 38806, July 25, 2001]

##### § 880.6740   Vacuum-powered body fluid suction apparatus.

- Identification. A vacuum-powered body fluid suction apparatus is a device

#### § 880.6760

used to aspirate, remove, or  sample body fluids. The device is  powered  by an external source of vacuum. This ge- neric type of device includes vacuum regulators, vacuum collection bottles, suction catheters and tips, connecting flexible aspirating tubes, rigid suction tips, specimen traps, noninvasive tub- ing, and  suction  regulators  (with gauge).

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.6760   Protective restraint.

- Identification. A protective re- straint is a device, including but not limited to a wristlet,  anklet,  vest, mitt, straight jacket, body/limb holder, or other type of strap, that is intended for medical purposes and that limits the patient’s movements to the extent necessary for treatment, examination, or protection of the patient or others.

- Classification. Class I (general con- trols).

[61 FR 8439, Mar. 4, 1996, as amended at 66 FR

46952, Sept. 10, 2001]

##### § 880.6775   Powered    patient    transfer device.

- Identification. A powered patient transfer device is a device consisting of a wheeled stretcher and a powered mechanism that has a broad, flexible band stretched over long rollers that can advance itself under a patient and transfer the patient with minimal dis- turbance in a horizontal position to the stretcher.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.6785   Manual patient transfer de- vice.

- Identification. A manual patient transfer device is a device consisting of a wheeled stretcher and a mechanism on  which  a  patient  can  be  placed  so

#### 21 CFR Ch. I (4–1–17 Edition)

that the patient can  be  transferred with minimal disturbance in a hori- zontal position to the stretcher.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38807, July 25, 2001]

##### § 880.6800   Washers for body waste re- ceptacles.

- Identification. A washer for body waste receptacles is a device intended for medical purposes that is used to clean and sanitize a body waste recep- tacle, such as a bedpan. The device con- sists of a wall-mounted plumbing fix- ture with a door through which a body waste receptacle is inserted. When the door is closed the body waste recep- tacle is cleaned by hot water, steam, or germicide.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device also is exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38807, July 25, 2001]

##### § 880.6820   Medical disposable scissors.

- Identification. Medical disposable scissors are disposable type general cutting devices intended for medical purposes. This generic type of device does not include surgical scissors.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 880.6910

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38807, July 25, 2001]

##### § 880.6850   Sterilization wrap.

  - Identification. A sterilization wrap (pack, sterilization wrapper, bag, or ac- cessories, is a device intended to be used to enclose another medical device that is to be sterilized by a health care provider. It is intended to allow steri- lization of the enclosed medical device and also to maintain sterility of the enclosed device until used.

  - Classification. Class II (perform- ance standards).

##### § 880.6860   Ethylene   oxide   gas   steri- lizer.

- Identification. An ethylene gas sterilizer is a nonportable device in- tended for use by a health care pro- vider that uses ethylene oxide (ETO) to sterilize medical products.

- Classification. Class II (perform- ance standards).

##### § 880.6870   Dry-heat sterilizer.

- Identification. A dry-heat sterilizer is a device that is intended for use by a health care provider to sterilize med- ical products by means of dry heat.

- Classification. Class II (perform- ance standards).

##### § 880.6880   Steam sterilizer.

- Identification. A steam sterilizer (autoclave) is a device that is intended for use by a health care provider to sterilize medical products by means of pressurized steam.

- Classification. Class II (perform- ance standards).

##### § 880.6885   Liquid  chemical  sterilants/ high level disinfectants.

- Identification. A liquid chemical sterilant/high level disinfectant is a germicide that is intended for  use as the terminal step in processing critical and semicritical medical devices prior to patient use. Critical devices make contact with normally sterile tissue or body spaces during use. Semicritical devices make contact during use with mucous membranes or nonintact skin.

- Classification. Class II (special controls). Guidance on the Content and

Format of Premarket Notification (510(k)) Submissions for Liquid Chem- ical Sterilants/High Level Disinfect- ants, and user information and train- ing.

[65 FR 36325, June 8, 2000]

##### § 880.6890   General   purpose   disinfect- ants.

- Identification. A general purpose disinfectant is a germicide intended to process noncritical medical devices and equipment surfaces. A general purpose disinfectant can be used to preclean or decontaminate critical or semicritical medical devices prior to terminal steri- lization or high level disinfection. Non- critical medical devices make only top- ical contact with intact skin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9.

[65 FR 36326, June 8, 2000]

##### § 880.6900   Hand-carried stretcher.

- Identification. A hand-carried stretcher is a device consisting of a lightweight frame, or of two poles with a cloth or metal platform, on which a patient can be carried.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38807,  July 25,

2001]

##### § 880.6910   Wheeled stretcher.

- Identification. A wheeled stretcher is a device consisting of a platform mounted on a wheeled frame that is de- signed to transport patients in a hori- zontal position. The device may have side rails, supports for fluid infusion equipment,   and   patient   securement

#### § 880.6920

straps. The frame may be fixed or col- lapsible for use in an ambulance.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 880.6920   Syringe needle introducer.

- Identification. A syringe needle in- troducer is a device that uses a spring- loaded mechanism to drive a hypo- dermic needle into a patient to a pre- determined depth below the skin sur- face.

- Classification. Class II (perform- ance standards).

##### § 880.6960   Irrigating syringe.

- Identification. An irrigating sy- ringe is a device intended for medical purposes that consists of a bulb or a piston syringe with an integral or a de- tachable tube. The device is used to ir- rigate, withdraw fluid from, or instill fluid into, a body cavity or wound.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38807, July 25, 2001]

##### § 880.6970   Liquid crystal vein locator.

- Identification. A liquid crystal vein locator is a device used to indicate the location of a vein by revealing vari- ations in the surface temperature of the skin by displaying the color changes of heat sensitive liquid crys- tals (cholesteric esters).

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 880.9.

[45 FR 69682, Oct. 21, 1980, as amended at 54

FR 25050, June 12, 1989; 66 FR 38807, July 25,

2001]

##### § 880.6980   Vein stabilizer.

- Identification. A vein stabilizer is a device consisting of a flat piece of plastic with two noninvasive prongs. The device is placed on the skin so that the prongs are on either side of a vein and hold it stable while a hypodermic needle is inserted into the vein.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 880.9.  If the device is not labeled or otherwise represented as sterile, it is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with respect to general requirements concerning records, and § 820.198, with respect to complaint files.

[45 FR 69682, Oct. 21, 1980, as amended at 66

FR 38807, July 25, 2001]

##### § 880.6990   Infusion stand.

- Identification. The infusion stand is a stationary or movable stand in- tended to hold infusion liquids, infu- sion accessories, and other medical de- vices.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 880.9.

[63 FR 59718, Nov. 5, 1998]

##### § 880.6991   Medical washer.

- Identification. A medical washer is a device that is intended for general medical purposes to clean and dry sur- gical instruments, anesthesia equip- ment, hollowware,  and  other  medical devices.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Medical Washers and Medical Washer-Disinfectors.’’ The device is exempt from the premarket

#### Food and Drug Administration, HHS	Pt. 882

notification procedures in subpart E of part  807  of  this  chapter  subject  to

§ 880.9.

[67 FR 69121, Nov. 15, 2002]

##### § 880.6992   Medical washer-disinfector.

  - Identification. A medical washer- disinfector is a device that is intended for general medical purposes to clean, decontaminate, disinfect, and dry sur- gical instruments, anesthesia equip- ment, hollowware,  and  other  medical devices.

  - Classification.  Class   II   (special

882.1400      Electroencephalograph.

882.1410 Electroencephalograph electrode/ lead   tester.

882.1420 Electroencephalogram (EEG) signal spectrum   analyzer.

882.1430 Electroencephalograph test signal generator.

882.1440   Neuropsychiatric	interpretive electroencephalograph   assessment   aid.

882.1450 Brain injury adjunctive interpre- tive electroencephalograph assessment aid.

882.1460     Nystagmograph.

882.1470 Computerized cognitive assessment aid.

| controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Medical Washers | 882.1480 | Neurological endoscope. |
| --- | --- | --- |
|  | 882.1500 | Esthesiometer. |
|  | 882.1525 | Tuning fork. |
|  | 882.1540 | Galvanic skin response measure- |
| and Medical Washer-Disinfectors.’’ | ment | device. |
| (1) Medical washer-disinfectors that are intended to clean, high level dis- | 882.1550 ment | Nerve conduction velocity measure- device. |

882.1471 Computerized cognitive assessment aid  for  concussion.

infect, and dry surgical instruments, anesthesia equipment, hollowware, and other medical devices.

(2) Medical washer-disinfectors that are intended to clean, low or inter- mediate level disinfect, and dry sur- gical instruments, anesthesia equip- ment, hollowware, and other medical devices are exempt from the premarket notification procedures in subpart E of part   807   of   this   chapter   subject   to

§ 880.9.

[67 FR 69121, Nov. 15, 2002]

## PART 882—NEUROLOGICAL DEVICES

##### Subpart A—General Provisions

Sec.

882.1    Scope.

882.3   Effective dates of requirement for pre- market  approval.

882.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Neurological Diagnostic Devices

882.1020    Rigidity analyzer.

882.1030  Ataxiagraph.

882.1200   Two-point discriminator.

882.1240     Echoencephalograph.

882.1275    Electroconductive media.

882.1310   Cortical electrode.

882.1320   Cutaneous electrode.

882.1330  Depth electrode.

882.1340   Nasopharyngeal electrode.

882.1350  Needle electrode.

882.1560    Skin    potential    measurement    de- vice.

882.1561    Evoked  photon  image  capture  de- vice.

882.1570   Powered    direct-contact    tempera- ture  measurement  device.

882.1610   Alpha monitor.

882.1620    Intracranial    pressure    monitoring device.

882.1700   Percussor.

882.1750   Pinwheel.

882.1790   Ocular plethysmograph.

882.1825     Rheoencephalograph.

882.1835     Physiological signal amplifier. 882.1845     Physiological signal conditioner. 882.1855    Electroencephalogram     (EEG)     te-

lemetry system.

882.1870   Evoked  response  electrical  stimu- lator.

882.1880    Evoked response mechanical stimu- lator.

882.1890    Evoked response photic stimulator. 882.1900   Evoked   response   auditory   stimu-

lator.

882.1925    Ultrasonic  scanner  calibration  test block.

882.1935   Near  Infrared  (NIR)  Brain  Hema- toma  Detector.

882.1950   Tremor transducer.

##### Subparts C–D [Reserved]

Subpart E—Neurological Surgical Devices

882.4030    Skull plate anvil. 882.4060    Ventricular cannula.

882.4100    Ventricular catheter.

882.4125   Neurosurgical chair.

882.4150   Scalp clip.

882.4175     Aneurysm clip applier.

882.4190    Clip forming/cutting instrument. 882.4200    Clip removal instrument.

882.4215   Clip rack.

§ 882.1

882.4250    Cryogenic surgical device. 882.4275    Dowel cutting instrument.

882.4300 Manual	cranial	drills,	burrs, trephines,  and  their  accessories.

882.4305   Powered   compound   cranial   drills, burrs,  trephines,  and  their  accessories.

882.4310  Powered	simple	cranial	drills, burrs,  trephines,  and  their  accessories.

882.4325  Cranial drill handpiece (brace). 882.4360   Electric cranial drill motor.

882.4370    Pneumatic cranial drill motor. 882.4400    Radiofrequency lesion generator. 882.4440   Neurosurgical headrests.

882.4460   Neurosurgical   head   holder   (skull clamp).

882.4500    Cranioplasty   material   forming   in- strument.

882.4525    Microsurgical instrument.

882.4535   Nonpowered   neurosurgical   instru- ment.

882.4545    Shunt  system  implantation  instru- ment.

882.4560    Stereotaxic instrument.

882.4600  Leukotome.

882.4650   Neurosurgical suture needle. 882.4700   Neurosurgical paddie.

882.4725   Radiofrequency lesion probe. 882.4750   Skull punch.

882.4800    Self-retaining  retractor  for  neuro- surgery.

882.4840    Manual rongeur.

882.4845    Powered rongeur.

882.4900   Skullplate screwdriver.

##### Subpart F—Neurological Therapeutic Devices

882.5030 Methyl	methacrylate	for aneurysmorrhaphy.

882.5050   Biofeedback device.

882.5070   Bite block.

882.5150     Intravascular occluding catheter. 882.5175   Carotid artery clamp.

882.5200    Aneurysm clip.

882.5225    Implanted malleable clip. 882.5235    Aversive conditioning device. 882.5250  Burr hole cover.

882.5275   Nerve cuff.

882.5300 Methyl	methacrylate	for cranioplasty.

882.5320    Preformed    alterable    cranioplasty plate.

882.5330   Preformed	nonalterable cranioplasty    plate.

882.5360    Cranioplasty plate fastener. 882.5500    Lesion temperature monitor.

882.5550    Central nervous system fluid shunt and  components.

882.5600  Neurovascular	mechanical thrombectomy device for acute ischemic stroke    treatment.

882.5700    Thermal system for insomnia. 882.5800     Cranial electrotherapy stimulator. 882.5805    Repetitive     transcranial     magnetic

stimulation system.

882.5808    Transcranial    magnetic    stimulator for  headache.

#### 21 CFR Ch. I (4–1–17 Edition)

882.5810    External  functional  neuromuscular stimulator.

882.5820    Implanted cerebellar stimulator. 882.5830 Implanted	diaphragmatic/phrenic

nerve stimulator.

882.5840    Implanted  intracerebral/subcortical stimulator  for  pain  relief.

882.5850    Implanted   spinal   cord   stimulator for  bladder  evacuation.

882.5860    Implanted    neuromuscular    stimu- lator.

882.5870   Implanted  peripheral  nerve  stimu- lator  for  pain  relief.

882.5880    Implanted   spinal   cord   stimulator for  pain  relief.

882.5890    Transcutaneous     electrical     nerve stimulator  for  pain  relief.

882.5891    Transcutaneous     electrical     nerve stimulator  to  treat  headache.

882.5900    Preformed craniosynostosis strip. 882.5894    Limited  output  transcutaneous  pi-

ezoelectric stimulator for skin reactions associated with insect bites.

882.5895    Vibratory   counter-stimulation   de- vice.

882.5910   Dura substitute.

882.5940    Electroconvulsive therapy device. 882.5950    Neurovascular  embolization  device. 882.5960    Skull tongs for traction.

882.5970    Cranial orthosis. 882.5975   Human dura mater.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE:  44  FR  51730,  Sept.  4,  1979,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 882 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 882.1   Scope.

- This part sets forth the classifica- tion of neurological devices intended for human use that are in commercial distribution.

- The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

- To avoid duplicative listings, a neurological device that has two or more types of uses (e.g., used both as a

Food and Drug Administration, HHS	§ 882.9

diagnostic device and as a therapeutic device) is listed only in one subpart.

- References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

- Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 17739, May 11, 1987, as amended at 68

FR 70436,  Dec. 18,  2003;  78  FR 18233,  Mar. 26,

2013]

##### § 882.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device.  If  FDA  promulgates  a  regula-

tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section, 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17739, May 11, 1987]

##### § 882.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic

#### § 882.1020

type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2319, Jan. 14, 2000]

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart B—Neurological Diagnostic Devices

##### § 882.1020   Rigidity analyzer.

- Identification. A rigidity analyzer is a device for quantifying the extent of the rigidity of a patient’s limb to de- termine the  effectiveness  of drugs  or other treatments.

- Classification. Class II (perform- ance standards).

##### § 882.1030   Ataxiagraph.

- Identification. An ataxiagraph is a device used to determine the extent of ataxia (failure of muscular coordina- tion) by measuring the amount of swaying of the body when the patient is standing erect and with eyes closed.

- Classification. Class I (general con- trols).

[44 FR 51730, Sept. 4, 1979, as amended at 66

FR 46952, Sept. 10, 2001]

##### § 882.1200   Two-point discriminator.

- Identification. A two-point dis- criminator is a device with points used for testing a patient’s touch discrimi- nation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9. The device is also ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this    chapter,    with    the    exception    of

§ 820.180 of this chapter, with respect to general requirements   concerning records, and § 820.198 of this chapter, with  respect  to  complaint  files.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR 25051,  June 12,  1989;  65  FR 2319,  Jan. 14,

2000]

##### § 882.1240   Echoencephalograph.

- Identification.	An echoencephalograph is an ultrasonic scanning device (including A-scan, B- scan, and doppler systems) that uses noninvasive transducers for measuring intracranial interfaces and blood flow velocity to and in the head.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 882.1440

##### § 882.1275   Electroconductive media.

  - Identification. Electroconductive media are the conductive creams or gels used with external electrodes to reduce the impedance (resistance to al- ternating current) of the contact be- tween the electrode surface and the skin.

  - Classification. Class II (perform- ance standards).

##### § 882.1310   Cortical electrode.

- Identification. A cortical electrode is an electrode which is temporarily placed on the surface of the brain for stimulating the brain or recording the brain’s electrical activity.

- Classification. Class II (perform- ance standards).

##### § 882.1320   Cutaneous electrode.

- Identification. A cutaneous elec- trode is an electrode that is applied di- rectly to a patient’s skin either to record physiological signals (e.g., the electroencephalogram) or to apply electrical  stimulation.

- Classification. Class II (perform- ance standards).

##### § 882.1330   Depth electrode.

- Identification. A depth electrode is an electrode used for temporary stimu- lation of, or recording electrical sig- nals at, subsurface levels of the brain.

- Classification. Class II (perform- ance standards).

##### § 882.1340   Nasopharyngeal electrode.

- Identification. A nasopharyngeal electrode is an electrode which is tem- porarily placed in the nasopharyngeal region for the purpose of recording electrical  activity.

- Classification. Class II (perform- ance standards).

##### § 882.1350   Needle electrode.

- Identification. A needle electrode is a device which is placed subcutaneously to stimulate or to record electrical signals.

- Classification. Class II (perform- ance standards).

##### § 882.1400   Electroencephalograph.

- Identification.	An electroencephalograph is a device used

to measure and record the electrical activity of the patient’s brain  obtained by placing two or more electrodes  on the head.

- Classification. Class II (perform- ance standards).

##### § 882.1410   Electroencephalograph elec- trode/lead tester.

- Identification.	An electroencephalograph electrode/lead tester is a device used for testing the impedance (resistance to alternating current) of the electrode and lead sys- tem of an electroencephalograph to as- sure that an adequate  contact is made between the electrode and the skin.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38807,  July 25,

2001]

##### § 882.1420   Electroencephalogram (EEG) signal spectrum analyzer.

- Identification. An electroencepha- logram (EEG) signal spectrum analyzer is a  device used  to display  the fre- quency content or power spectral den- sity  of  the   electroencephalogram (EEG)  signal.

- Classification. Class I (general con- trols).

[44 FR 51730, Sept. 4, 1979, as amended at 66

FR 46953, Sept. 10, 2001]

##### § 882.1430   Electroencephalograph   test signal generator.

- Identification.	An electroencephalograph test signal gen- erator is a device used to test or cali- brate an electroencephalograph.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38807,  July 25,

2001]

##### § 882.1440 Neuropsychiatric interpre- tive electroencephalograph assess- ment aid.

- Identification.		The neuropsychiatric	interpretive

#### § 882.1450

electroencephalograph  assessment  aid is a prescription device that uses a pa- tient’s electroencephalograph (EEG) to provide an interpretation of the pa- tient’s neuropsychiatric condition. The neuropsychiatric interpretive EEG as- sessment aid is used only as an assess- ment aid for a medical condition for which there exists other valid methods of diagnosis.

- Classification. Class II (special controls). The special controls for this device are:

- The technical parameters of the device, hardware and software, must be fully characterized and must dem- onstrate a reasonable assurance  of safety and effectiveness.

  - Hardware specifications must be provided. Appropriate verification, val- idation, and hazard analysis must be performed.

  - Software, including any propri- etary algorithm(s) used by the device to arrive at its interpretation of the patient’s condition, must be described in detail in the software requirements specification and software design speci- fication.	Appropriate	software verification, validation, and hazard analysis must be performed.

- The device parts that contact the patient must be demonstrated to be biocompatible.

- The device must be designed and tested for electrical safety, electro- magnetic compatibility, thermal, and mechanical safety.

- Clinical performance testing must demonstrate the accuracy, precision, reproducibility, of determining the EEG-based interpretation, including any specified equivocal zones (cutoffs).

- Clinical performance testing must demonstrate the ability of  the  device to function as an assessment aid  for the medical condition for which the de- vice is indicated. Performance meas- ures must demonstrate device perform- ance characteristics per the intended use in the intended use environment. Performance measurements must in- clude sensitivity, specificity, positive predictive value, and negative pre- dictive value per the device intended use. Repeatability of measurements must be demonstrated  using interclass correlation coefficients and illustrated by  qualitative  scatter  plot(s).

#### 21 CFR Ch. I (4–1–17 Edition)

- The device design must include safeguards to prevent use of the device as a stand-alone diagnostic.

- The labeling must include the fol- lowing information:

  - A warning that the  device  is  not to be used as a stand-alone diagnostic.

  - A detailed summary of the clin- ical performance testing, including any adverse events and complications.

  - The qualifications and training requirements for device users including technicians and clinicians.

  - The intended use population and the intended use environment.

  - Any instructions technicians should convey to patients regarding the collection of EEG data.

  - Information allowing clinicians to gauge clinical risk associated with integrating the EEG interpretive as- sessment aid into their diagnostic pathway.

  - Where appropriate, validated methods and instructions for reprocess- ing of any reusable components.

[79 FR 9085, Feb. 18, 2014]

##### § 882.1450 Brain injury adjunctive in- terpretive electroencephalograph assessment aid.

- Identification. A brain injury ad- junctive	interpretive electroencephalograph  assessment  aid is a prescription device that uses a pa- tient’s electroencephalograph (EEG) to provide an interpretation of the struc- tural condition of the patient’s brain in the setting of trauma. A brain injury adjunctive interpretive EEG assess- ment aid is for use as an adjunct to standard clinical practice only as an assessment aid for a medical condition for which there exists other valid methods of diagnosis.

- Classification. Class II (special controls). The special controls for this device are:

- The technical parameters of the device, hardware and software, must be fully characterized and include the fol- lowing  information:

  - Hardware specifications must be provided. Appropriate verification, val- idation, and hazard analysis must be performed.

  - Software, including any propri- etary algorithm(s) used by the device to  arrive  at  its  interpretation  of  the

#### Food and Drug Administration, HHS	§ 882.1470

patient’s condition, must be described in detail in the software requirements specification (SRS) and software design specification (SDS). Appropriate soft- ware verification, validation, and haz- ard analysis must be performed.

- The device parts that contact the patient must be demonstrated to be biocompatible.

- The device must be designed and tested for electrical safety, electro- magnetic compatibility (EMC), ther- mal, and mechanical safety.

- Clinical performance testing must demonstrate the accuracy, precision- repeatability and reproducibility, of determining the EEG-based interpreta- tion, including any specified equivocal zones (cutoffs).

- Clinical performance testing must demonstrate the ability of  the  device to function as an assessment aid  for the medical condition for which the de- vice is indicated. Performance meas- ures must demonstrate device perform- ance characteristics per the intended use in the intended use environment. Performance measurements must in- clude sensitivity, specificity, positive predictive value (PPV), and negative predictive value (NPV) with respect to the study prevalence per the device in- tended use.

- The device design must include safeguards to ensure appropriate clin- ical interpretation of the device output (e.g., use in appropriate patient popu- lation, or for appropriate clinical deci- sion).

- The labeling and training infor- mation must include:

  - A warning that the  device  is  not to be used as a stand-alone diagnostic.

  - A detailed summary of the clin- ical performance testing, including any adverse events and complications.

  - The intended use population and the intended use environment.

  - Any instructions technicians should convey to patients regarding the collection of EEG data.

  - Information allowing clinicians to gauge clinical risk associated with in- tegrating the EEG interpretive assess- ment aid into their diagnostic path- way.

  - Information allowing clinicians to understand how to integrate the de- vice output into their diagnostic path-

way when the device is unable to pro- vide a classification or final result.

[80 FR 16268, Mar. 27, 2015]

##### § 882.1460   Nystagmograph.

- Identification. A nystagmograph is a device used to measure, record, or visually display the involuntary move- ments (nystagmus) of the eyeball.

- Classification. Class II (perform- ance standards).

##### § 882.1470   Computerized  cognitive  as- sessment aid.

- Identification. The computerized cognitive assessment aid is a prescrip- tion device that uses an individual’s score(s) on a battery of cognitive tasks to provide an interpretation of the cur- rent level of cognitive function. The computerized cognitive assessment aid is used only as an assessment aid to de- termine level of cognitive functioning for which there exists other valid methods of cognitive assessment and does not identify the presence or ab- sence of clinical diagnoses. The com- puterized cognitive assessment aid is not intended as a stand-alone or ad- junctive diagnostic device.

- Classification. Class II (special controls). The special  control(s)  for this device are:

- The technical parameters of the device’s hardware and software must be fully characterized and be accompanied by appropriate non-clinical testing:

  - Hardware specifications must be provided. Appropriate verification, val- idation, and hazard analysis must be performed.

  - Software, including any propri- etary algorithm(s) used by the device to arrive at its interpretation of the patient’s cognitive function, must be described in detail in the software re- quirements specification (SRS) and software design specification (SDS). Appropriate software verification, vali- dation, and hazard analysis must be performed.

- The device must be designed and tested  for  electrical  safety.

- The labeling must include:

  - A summary of any testing con- ducted to demonstrate how the device functions as an interpretation of the current level of cognitive function. The summary  of  testing  must  include  the

#### § 882.1471

following, if available: Any expected or observed adverse events and complica- tions; any performance measurements including sensitivity, specificity, posi- tive predictive value (PPV), and nega- tive predictive value (NPV) per the de- vices intended use; a description of the repeatability of measurements; a de- scription of how the cut-off values for categorization of measurements  were determined; and a description of the construct validity of the device.

  - A warning that the device does not identify the presence or absence of clinical  diagnoses.

  - A warning that the device is not a stand-alone diagnostic.

  - The intended use population and the intended use environment.

  - Any instructions technicians must convey to patients regarding the administration of the test and collec- tion of cognitive test data.

[80 FR 49138, Aug. 17, 2015]

##### § 882.1471   Computerized  cognitive  as- sessment aid for concussion.

- Identification. The computerized cognitive assessment aid for concus- sion is a prescription device that uses an individual’s score(s) on a battery of cognitive tasks to provide an indica- tion of the current level of cognitive function in response to concussion. The computerized cognitive assessment aid for concussion is used only as an as- sessment aid in the management of concussion to determine cognitive function for patients after a potential concussive event where other diag- nostic tools are available and does not identify the presence or absence of con- cussion. It is not intended as a stand- alone diagnostic device.

- Classification. Class II (special controls). The special controls for this device are:

  - Software, including any propri- etary algorithm(s) used by the device to arrive at its interpretation of the patient’s cognitive function, must be described in detail in the software re- quirements specification (SRS) and software design specification (SDS). Software verification, validation, and hazard analysis must be performed.

  - Clinical performance data must be provided that demonstrates how the device functions as an interpretation of

#### 21 CFR Ch. I (4–1–17 Edition)

the current level of cognitive function in an individual  that has recently  re- ceived an injury that causes concern about a possible concussion. The test- ing must:

- Evaluate device output and  clin- ical interpretation.

- Evaluate device test-retest reli- ability of the device output.

- Evaluate construct validity of the device cognitive assessments.

- Describe the construction of the normative database, which includes the following:

  - How the clinical workup was completed to establish a ‘‘normal’’ pop- ulation, including the establishment of inclusion and exclusion criteria.

  - Statistical methods and model assumptions used.

  - The labeling must include:

    - A summary of any clinical testing conducted to demonstrate how the de- vice functions as an interpretation of the current level of cognitive function in a patient that has recently received an injury that causes concern about a possible concussion. The summary of testing must include the following:

- Device output and clinical inter- pretation.

- Device test-retest reliability of the device output.

- Construct validity of the device cognitive assessments.

- A description of the normative database, which includes the following:

- How the clinical workup was com- pleted to establish a ‘‘normal’’ popu- lation, including the establishment of inclusion and exclusion criteria.

- How normal values will be re- ported to the user.

- Representative screen shots and reports that will be generated to pro- vide the user results and normative data.

- Statistical methods and model as- sumptions used.

- Whether or not the normative database was adjusted due to dif- ferences in age and gender.

    - A warning that the device should only be used by health care profes- sionals who are trained in concussion management.

    - A warning that the device does not identify the presence or absence of concussion  or  other  clinical  diagnoses.

#### Food and Drug Administration, HHS	§ 882.1561

    - A warning that the device is not a stand-alone diagnostic.

    - Any instructions technicians must convey to patients regarding the administration of the test and collec- tion of cognitive test data.

[81 FR 87811, Dec. 6, 2016]

##### § 882.1480   Neurological endoscope.

- Identification. A neurological en- doscope is an instrument with a light source used to view the inside of the ventricles of the brain.

- Classification. Class II (perform- ance standards).

##### § 882.1500   Esthesiometer.

- Identification. An esthesiometer is a mechanical device which usually con- sists of a single rod or fiber which is held in the fingers of the physician or other examiner and which is used to determine whether a patient has tac- tile  sensitivity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9. The device is also ex- empt from the current good manufac- turing practice requirements of the quality system regulation in part 820 of this    chapter,    with    the    exception    of

§ 820.180 of this chapter, with respect to general requirements   concerning records, and § 820.198 of this chapter, with  respect  to  complaint  files.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR 25051,  June 12,  1989;  65  FR 2319,  Jan. 14,

2000]

##### § 882.1525   Tuning fork.

- Identification. A tuning fork is a mechanical device which resonates at a given frequency and is used to diagnose hearing disorders and to test for vibra- tory sense.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, of this chap- ter, with the exception of § 820.180, with respect  to  general  requirements  con-

cerning records, and § 820.198, with re- spect to complaint files.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR 25051, June 12, 1989; 66 FR 38807, July 25,

2001]

##### § 882.1540   Galvanic	skin	response measurement device.

- Identification. A galvanic skin re- sponse measurement device is a device used to determine autonomic responses as psychological indicators by meas- uring the electrical resistance of the skin and the tissue path between two electrodes applied to the skin.

- Classification. Class II (perform- ance standards).

##### § 882.1550   Nerve   conduction   velocity measurement device.

- Identification. A nerve conduction velocity measurement device is a de- vice which measures nerve conduction time by applying a stimulus, usually to a patient’s peripheral nerve. This de- vice includes the stimulator and the electronic processing equipment for measuring and displaying the nerve conduction time.

- Classification. Class II (perform- ance standards).

##### § 882.1560   Skin potential measurement device.

- Identification. A skin potential measurement device is a general diag- nostic device used to measure skin voltage by means of surface skin elec- trodes.

- Classification. Class II (perform- ance standards).

##### § 882.1561   Evoked  photon  image  cap- ture device.

- Identification. An evoked  photon image capture device is a prescription, electrically powered device intended for use as a noninvasive measurement tool that applies electricity to detect electrophysiological signals emanating from the skin, which are reported nu- merically and as images without clin- ical interpretation. The device is not intended for diagnostic purposes.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 882.1570

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 882.9.

[81 FR 67155, Sept. 30, 2016]

##### § 882.1570   Powered direct-contact tem- perature measurement device.

  - Identification. A powered direct- contact temperature measurement de- vice is a device which contains a power source and is used to measure dif- ferences in temperature between two points on the body.

  - Classification. Class II (perform- ance standards).

##### § 882.1610   Alpha monitor.

- Identification. An alpha monitor is a device with  electrodes  that  are placed on a patient’s scalp to monitor that portion of the electroencepha- logram which is referred to  as  the alpha wave.

- Classification. Class II (perform- ance standards).

##### § 882.1620   Intracranial pressure  moni- toring device.

- Identification. An intracranial pressure monitoring device is a device used for short-term monitoring and re- cording of intracranial pressures and pressure trends. The device includes the transducer, monitor, and inter- connecting hardware.

- Classification. Class II (perform- ance standards).

##### § 882.1700   Percussor.

- Identification. A percussor is a small hammerlike device used by a physician to provide light blows to a body part. A percussor is used as a di- agnostic aid during physical examina- tions.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### 21 CFR Ch. I (4–1–17 Edition)

and § 820.198, with respect to complaint files.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR  25051,  June  12,  1989;  59  FR  63011,  Dec.  7,

1994; 66 FR 38807, July 25, 2001]

##### § 882.1750   Pinwheel.

- Identification. A pinwheel is a de- vice with sharp points on a rotating wheel used for testing pain sensation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR 25051,  June 12,  1989;  65  FR 2319,  Jan. 14,

2000]

##### § 882.1790   Ocular plethysmograph.

- Identification. An ocular plethysmograph is a device used to measure or detect volume changes in the eye produced by pulsations of the artery, to diagnose carotid artery oc- clusive disease (restrictions on blood flow in the carotid artery).

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before September 21, 2004, for any ocular plethysmograph  that was in commercial distribution before May 28, 1976. Any other ocular plethysmograph shall have an approved PMA or declared completed PDP in ef- fect before being placed in commercial distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17739,  May 11,  1987;  69  FR 34920,  June 23,

2004]

##### § 882.1825   Rheoencephalograph.

- Identification.	A rheoencephalograph is a device used to estimate a patient’s cerebral  circula- tion (blood  flow in  the brain)  by elec- trical impedance methods with direct electrical connections to the scalp or neck  area.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be

#### Food and Drug Administration, HHS	§ 882.1935

filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any rheoencephalograph that  was in commercial distribution before May 28, 1976, or that has, on or before De- cember 26, 1996 been found to be sub- stantially	equivalent	to	a rheoencephalograph that was in com- mercial distribution before May 28, 1976. Any other rheoencephalograph shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17740, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 882.1835   Physiological  signal  ampli- fier.

  - Identification. A physiological sig- nal amplifier is a general purpose de- vice used to  electrically  amplify  sig- nals derived from various physiological sources (e.g., the electroencepha- logram).

  - Classification. Class II (perform- ance standards).

##### § 882.1845   Physiological  signal  condi- tioner.

- Identification. A physiological sig- nal conditioner is a device such as an integrator or differentiator used to modify physiological signals for re- cording and processing.

- Classification. Class II (perform- ance standards).

##### § 882.1855   Electroencephalogram (EEG) telemetry system.

- Identification. An electroencepha- logram (EEG) telemetry system con- sists of transmitters, receivers, and other components used for remotely monitoring or measuring EEG signals by means of radio or telephone trans- mission  systems.

- Classification. Class II (perform- ance standards).

##### § 882.1870   Evoked  response  electrical stimulator.

- Identification. An evoked response electrical stimulator is a  device  used to apply an electrical stimulus to a pa- tient by means of skin electrodes for the purpose of measuring the evoked response.

- Classification. Class II (perform- ance standards).

##### § 882.1880   Evoked   response   mechan- ical stimulator.

- Identification. An evoked response mechanical stimulator is a device used to produce a mechanical stimulus or a series of mechanical stimuli for the purpose of measuring a patient’s evoked response.

- Classification. Class II (perform- ance standards).

##### § 882.1890   Evoked	response	photic stimulator.

- Identification. An evoked response photic stimulator is a device used to generate and display a shifting pattern or to apply a brief light stimulus to a patient’s eye for use in evoked response measurements or for electroencepha- logram  (EEG)  activation.

- Classification. Class II (perform- ance standards).

##### § 882.1900   Evoked   response   auditory stimulator.

- Identification. An evoked response auditory stimulator is a device that produces a sound stimulus for use in evoked response measurements or elec- troencephalogram   activation.

- Classification. Class II (perform- ance standards).

##### § 882.1925   Ultrasonic  scanner  calibra- tion test block.

- Identification. An ultrasonic scan- ner calibration  test block  is a block  of material with known properties used to calibrate ultrasonic scanning devices (e.g., the echoencephalograph).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38807,  July 25,

2001]

##### § 882.1935   Near  Infrared  (NIR)  Brain Hematoma Detector.

- Identification. A Near Infrared (NIR) Brain Hematoma Detector is a noninvasive device that employs near- infrared spectroscopy that is intended

#### § 882.1950

to be used to evaluate suspected brain hematomas.

- Classification. Class II (special controls). The special controls for this device are:

- The sale, distribution, and use of this device are  restricted  to  prescrip- tion use in accordance with § 801.109 of this   chapter;

- The labeling must include specific instructions and the clinical training needed for the safe use of this device;

- Appropriate analysis/testing should validate electromagnetic com- patibility (EMC), electrical safety, and battery characteristics;

- Performance data should validate accuracy and precision and safety fea- tures;

- Any elements of the device that may contact the patient should be demonstrated to be biocompatible; and,

- Appropriate software verification, validation, and hazard analysis should be performed.

[77 FR 16927, Mar. 23, 2012]

##### § 882.1950   Tremor transducer.

- Identification. A tremor trans- ducer is a device used to measure the degree of tremor caused by certain dis- eases.

- Classification. Class II (perform- ance standards).

## Subparts C–D [Reserved]

Subpart E—Neurological Surgical Devices

##### § 882.4030   Skull plate anvil.

- Identification. A skull plate  anvil is a device used to form alterable skull plates in the proper shape to fit the curvature of a patient’s skull.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63011,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4060   Ventricular cannula.

- Identification. A ventricular cannula is  a  device  used  to  puncture the ventricles of the brain for aspira-

21 CFR Ch. I (4–1–17 Edition)

tion or for injection. This device is fre- quently referred to as a ventricular needle.

- Classification. Class I (general con- trols).  When  made  only  of  surgical grade stainless steel, the device is ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 65

FR 2319, Jan. 14, 2000]

##### § 882.4100   Ventricular catheter.

- Identification. A ventricular cath- eter is a device used to gain access to the cavities of the  brain for injection of material into,  or removal of  mate- rial from, the brain.

- Classification. Class II (perform- ance standards).

##### § 882.4125   Neurosurgical chair.

- Identification. A neurosurgical chair is an operating room chair used to position and support a patient dur- ing neurosurgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4150   Scalp clip.

- Identification. A scalp clip is a plastic or metal clip used to stop bleed- ing during surgery on the scalp.

- Classification. Class II (perform- ance standards).

##### § 882.4175   Aneurysm clip applier.

- Identification. An aneurysm clip applier is a device used by the surgeon for holding and applying intracranial aneurysm clips.

- Classification. Class II (perform- ance standards).

##### § 882.4190   Clip  forming/cutting  instru- ment.

- Identification. A clip forming/cut- ting instrument is a device used by the physician to make tissue clips from wire stock.

#### Food and Drug Administration, HHS	§ 882.4370

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012, Dec. 7, 1994]

##### § 882.4200   Clip removal instrument.

- Identification. A clip removal in- strument is a device used to remove surgical clips from the patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4215   Clip rack.

- Identification. A clip rack is a de- vice used to hold or store surgical clips during surgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR  25051,  June  12,  1989;  59  FR  63012,  Dec.  7,

1994; 66 FR 38808, July 25, 2001]

##### § 882.4250   Cryogenic surgical device.

- Identification. A cryogenic sur- gical device is a device used to destroy nervous tissue or produce lesions in nervous tissue by the application of ex- treme cold to the selected site.

- Classification. Class II (perform- ance standards).

##### § 882.4275   Dowel cutting instrument.

- Identification. A dowel cutting in- strument is a device used to cut dowels of bone for bone grafting.

- Classification. Class II (perform- ance standards).

##### § 882.4300   Manual cranial drills, burrs, trephines, and their accessories

- Identification. Manual cranial drills, burrs, trephines, and their acces- sories are bone cutting and drilling in- struments that are used without a power source on a patient’s skull.

- Classification. Class II (perform- ance standards).

##### § 882.4305 Powered compound cranial drills, burrs, trephines, and their accessories.

- Identification. Powered compound cranial drills, burrs, trephines, and their accessories are bone cutting and drilling instruments used on a pa- tient’s skull. The  instruments  employ a clutch mechanism to disengage the tip of the instrument after penetrating the skull to prevent plunging of the tip into the brain.

- Classification. Class II (perform- ance standards).

##### § 882.4310 Powered simple cranial drills, burrs, trephines, and their accessories.

- Identification. Powered simple cra- nial drills, burrs,  trephines, and their accessories are bone cutting and drill- ing instruments used on a patient’s skull. The instruments are used with a power source but do not have a clutch mechanism to disengage the tip after penetrating the skull.

- Classification. Class II (perform- ance standards).

##### § 882.4325 Cranial drill handpiece (brace).

- Identification. A cranial drill hand- piece (brace) is a hand holder, which is used without a power source, for drills, burrs, trephines, or other cutting tools that are used on a patient’s skull.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38808,  July 25,

2001]

##### § 882.4360   Electric cranial drill motor.

- Identification. An electric cranial drill motor is an electrically operated power source used with removable ro- tating surgical cutting tools or drill bits on a patient’s skull.

- Classification. Class II (perform- ance standards).

##### § 882.4370 Pneumatic cranial drill motor.

- Identification. A pneumatic cranial drill motor is a pneumatically operated

#### § 882.4400

power source used with removable ro- tating surgical  cutting  tools  or  drill bits on a patient’s skull.

- Classification. Class II (perform- ance standards).

##### § 882.4400   Radiofrequency  lesion  gen- erator.

- Identification. A radiofrequency le- sion generator is a device used to produce lesions in the nervous system or other tissue by the direct applica- tion of radiofrequency currents to se- lected  sites.

- Classification. Class II (perform- ance standards).

##### § 882.4440   Neurosurgical headrests.

- Identification. A neurosurgical headrest is a device used to support the patient’s head during a surgical proce- dure.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4460   Neurosurgical  head  holder (skull clamp).

- Identification. A neurosurgical head holder (skull clamp) is a device used to clamp the patient’s skull to hold head and neck in a particular po- sition during surgical procedures.

- Classification. Class II (perform- ance standards).

##### § 882.4500   Cranioplasty material form- ing instrument.

- Identification. A cranioplasty ma- terial forming instrument is a roller used in the preparation and forming of cranioplasty (skull repair) materials.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4525   Microsurgical instrument.

- Identification. A microsurgical in- strument is a nonpowered surgical in-

#### 21 CFR Ch. I (4–1–17 Edition)

strument  used  in  neurological  micro- surgery procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4535   Nonpowered   neurosurgical instrument.

- Identification. A nonpowered neurosurgical instrument is a hand in- strument or an accessory to a hand in- strument used during neurosurgical procedures to cut, hold, or manipulate tissue. It includes specialized chisels, osteotomes, curettes, dissectors, ele- vators, forceps, gouges, hooks, surgical knives, rasps, scissors, separators, spatulas, spoons, blades, blade holders, blade breakers, probes, etc.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4545   Shunt  system  implantation instrument.

- Identification. A shunt system im- plantation instrument is an instru- ment used in the implantation of cere- brospinal fluid shunts, and includes tunneling instruments for passing shunt components under the skin.

- Classification. Class I (general con- trols).  When  made  only  of  surgical grade stainless steel, the device is ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 65

FR 2319, Jan. 14, 2000]

##### § 882.4560   Stereotaxic instrument.

- Identification. A stereotaxic in- strument is a device consisting of a rigid frame with a calibrated guide mechanism for precisely positioning probes or other devices within a pa- tient’s brain, spinal cord, or other part of the nervous system.

#### Food and Drug Administration, HHS	§ 882.4900

- Classification. Class II (perform- ance standards).

##### § 882.4600   Leukotome.

- Identification. A leukotome is a device used to cut sections out of the brain.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

##### § 882.4650   Neurosurgical   suture   nee- dle.

- Identification. A neurosurgical su- ture needle is a needle used in suturing during neurosurgical procedures or in the repair of nervous tissue.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 54

FR 25051,  June 12,  1989;  65  FR 2319,  Jan. 14,

2000]

##### § 882.4700   Neurosurgical paddie.

- A neurosurgical paddie is a pad used during surgery to protect nervous tissue, absorb fluids, or stop bleeding.

- Classification. Class II (perform- ance standards).

[44 FR 51730, Sept. 4, 1979, as amended at 69

FR 10332, Mar. 5, 2004]

##### § 882.4725   Radiofrequency	lesion probe.

- Identification. A radiofrequency le- sion probe is a device connected to a radiofrequency (RF) lesion generator to deliver the RF energy to the site within the nervous system where a le- sion is desired.

- Classification. Class II (perform- ance standards).

##### § 882.4750   Skull punch.

- Identification. A skull punch is a device used to punch holes through a patient’s skull to allow fixation of cranioplasty plates or bone flaps by wire or other means.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 882.9. This exemption does not apply to powered compound cranial drills, burrs, trephines, and their acces- sories classified under § 882.4305.

[44 FR 51730, Sept. 4, 1979, as amended at 65

FR 2319, Jan. 14, 2000]

##### § 882.4800   Self-retaining  retractor  for neurosurgery.

- Identification. A self-retaining re- tractor for neurosurgery is a self-lock- ing device used to hold the edges of a wound open during neurosurgery.

- Classification. Class II (perform- ance standards).

##### § 882.4840   Manual rongeur.

- Identification. A manual rongeur is a manually operated instrument used for cutting or biting bone during sur- gery involving the skull or spinal col- umn.

- Classification. Class II (perform- ance standards).

##### § 882.4845   Powered rongeur.

- Identification. A powered rongeur is a powered instrument used for cut- ting or biting bone during surgery in- volving the skull or spinal column.

- Classification. Class II (perform- ance standards).

##### § 882.4900   Skullplate screwdriver.

- Identification. A skullplate screw- driver is a tool used by the surgeon to fasten cranioplasty plates  or skullplates to a patient’s skull by screws.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38808,  July 25,

2001]

#### § 882.5030

Subpart F—Neurological Therapeutic Devices

##### § 882.5030   Methyl	methacrylate	for aneurysmorrhaphy.

  - Identification. Methyl methacry- late for aneurysmorrhaphy (repair of aneurysms, which are balloonlike sacs formed on blood vessels) is a self-cur- ing acrylic used to encase and reinforce intracranial aneurysms that are not amenable to conservative manage- ment, removal, or obliteration by an- eurysm  clip.

  - Classification. Class II (perform- ance standards).

##### § 882.5050   Biofeedback device.

- Identification. A biofeedback de- vice is an instrument that provides a visual or auditory signal corresponding to the status of one or more of a pa- tient’s physiological parameters (e.g., brain alpha wave activity, muscle ac- tivity, skin temperature, etc.) so that the patient can control voluntarily these physiological parameters.

- Classification. Class II (special controls). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter when it is a prescription battery pow- ered device that is indicated for relax- ation training and muscle reeducation and prescription use, subject to § 882.9.

[44 FR 51730, Sept. 4, 1979, as amended at 63

FR 59229, Nov. 3, 1998]

##### § 882.5070   Bite block.

- Identification. A bite block is a de- vice inserted into a patient’s mouth to protect the tongue and teeth while the patient is having convulsions.

- Classification. Class II (perform- ance standards).

##### § 882.5150   Intravascular	occluding catheter.

- Identification. An intravascular occluding catheter is a catheter  with an inflatable or detachable balloon tip that is used to block a blood vessel to treat malformations, e.g., aneurysms (balloonlike sacs formed on blood ves- sels) of intracranial blood vessels.

- Classification. Class III (premarket approval).

#### 21 CFR Ch. I (4–1–17 Edition)

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any intravascular occluding  cath- eter that was  in  commercial  distribu- tion before May 28, 1976, or that has, on or before December 26, 1996 been found to be  substantially  equivalent  to  an intravascular occluding  catheter  that was in commercial distribution  before May 28, 1976. Any other intravascular occluding catheter shall have an ap- proved PMA or  a  declared  completed PDP in effect before being placed in commercial     distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17740, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 882.5175   Carotid artery clamp.

- Identification. A carotid artery clamp is a device that is surgically placed around a patient’s carotid ar- tery (the principal artery in the neck that supplies blood to the brain) and has a removable adjusting mechanism that protrudes through the skin of the patient’s neck. The clamp is used to oc- clude the patient’s carotid artery to treat	intracranial	aneurysms (balloonlike sacs formed on blood ves- sels) or other intracranial vascular malformations that are difficult to at- tach directly by reducing the blood pressure and blood flow to the aneu- rysm or malformation.

- Classification. Class II (perform- ance standards).

##### § 882.5200   Aneurysm clip.

- Identification. An aneurysm clip is a device used to occlude an intracranial aneurysm (a balloonlike sac formed on a blood vessel) to pre- vent it from bleeding or bursting.

- Classification. Class II (perform- ance standards).

##### § 882.5225   Implanted malleable clip.

- Identification. An implanted mal- leable clip is a bent wire or staple that is forcibly closed with a special instru- ment to occlude an intracranial blood vessel or aneurysm (a balloonlike sac formed on a blood vessel), stop bleed- ing, or hold tissue or a mechanical de- vice in place in a patient.

#### Food and Drug Administration, HHS	§ 882.5550

- Classification. Class II (perform- ance standards).

##### § 882.5235   Aversive    conditioning   de- vice.

- Identification. An aversive condi- tioning device is an instrument used to administer an electrical shock or other noxious stimulus to a patient to mod- ify undesirable behavioral characteris- tics.

- Classification. Class II (perform- ance standards).

##### § 882.5250   Burr hole cover.

- Identification. A burr hole cover is a plastic or metal device used to cover or plug holes drilled into the skull dur- ing surgery and to reattach cranial bone removed during surgery.

- Classification. Class II (perform- ance standards).

##### § 882.5275   Nerve cuff.

- Identification. A nerve cuff is a tu- bular silicone rubber sheath used to en- case a nerve for aid in repairing the nerve (e.g., to prevent ingrowth of scar tissue) and for capping the end of the nerve to prevent the formation of neuroma (tumors).

- Classification. Class II (perform- ance standards).

##### § 882.5300   Methyl	methacrylate	for cranioplasty.

- Identification. Methyl methacry- late for cranioplasty (skull repair) is a self-curing acrylic that a surgeon uses to repair a skull defect in a patient. At the time of surgery, the surgeon initi- ates polymerization of  the  material and forms it into a plate or other ap- propriate shape to repair the defect.

- Classification. Class II (perform- ance standards).

##### § 882.5320   Preformed	alterable cranioplasty plate.

- Identification. A preformed alterable cranioplasty plate is a device that is implanted into a patient to re- pair a skull defect. It is constructed of a material, e.g., tantalum, that can be altered or reshaped at the time of sur- gery without changing the chemical behavior of the material.

- Classification. Class II (perform- ance standards).

##### § 882.5330   Preformed	nonalterable cranioplasty plate.

- Identification. A preformed nonalterable cranioplasty plate is a de- vice that is implanted in a patient to repair a skull defect and is constructed of a material, e.g., stainless steel or vitallium, that cannot be altered or re- shaped at the time of surgery without changing the chemical behavior of the material.

- Classification. Class II (perform- ance standards).

##### § 882.5360   Cranioplasty plate fastener.

- Identification. A cranioplasty plate fastener is a screw, wire, or other arti- cle made  of  tantalum, vitallium,  or stainless steel used to secure a plate to the patient’s skull to repair a skull de- fect.

- Classification. Class II (perform- ance standards).

##### § 882.5500   Lesion    temperature    mon- itor.

- Identification. A lesion tempera- ture monitor is a device used to mon- itor the tissue temperature at the site where a lesion (tissue destruction) is to be made when a surgeon uses a radio- frequency (RF) lesion generator and probe.

- Classification. Class II (perform- ance standards).

##### § 882.5550   Central	nervous	system fluid shunt and components.

- Identification. A central nervous system fluid shunt is a device or com- bination of devices used to divert fluid from the brain or other part of the cen- tral nervous system to an internal de- livery site or an external receptacle for the purpose of relieving elevated intracranial pressure or fluid volume (e.g., due to hydrocephalus). Compo- nents of a central nervous system shunt include catheters, valved cath- eters, valves, connectors, and other ac- cessory components intended to facili- tate use of the shunt or evaluation of a patient with a shunt.

- Classification. Class II (perform- ance standards).

#### § 882.5600

##### § 882.5600 Neurovascular mechanical thrombectomy device for acute ischemic stroke treatment.

  - Identification. A neurovascular mechanical thrombectomy device for acute ischemic stroke treatment is a prescription device used in the treat- ment of acute ischemic stroke to im- prove clinical outcomes. The device is delivered into the neurovasculature with an endovascular approach, me- chanically removes thrombus from the body, and restores blood flow in the neurovasculature.

  - Classification. Class II (special controls). The special controls for this device are:

    - The patient contacting compo- nents of the device must be dem- onstrated to be biocompatible.

    - Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use, including:

- Mechanical testing to dem- onstrate the device can withstand an- ticipated tensile, torsional, and com- pressive forces.

- Mechanical testing to evaluate the radial forces exerted by the device.

- Non-clinical testing to verify the dimensions of the device.

- Non-clinical testing must dem- onstrate the device can be delivered to the target location in the neurovasculature and retrieve simu- lated thrombus under simulated use conditions.

- Non-clinical testing must dem- onstrate the device is radiopaque and can be visualized.

- Non-clinical testing must evalu- ate the coating integrity and particu- lates under simulated use conditions.

- Animal testing must evaluate the safety of the device, including dam- age to the vessels or tissue under an- ticipated use conditions.

    - Performance data must support the sterility and pyrogenicity of the patient contacting components of the device.

    - Performance data must support the shelf-life of the device by dem- onstrating continued sterility, package integrity, and device  functionality over the specified shelf-life.

    - Clinical performance testing  of the  device  must  demonstrate  the  de-

#### 21 CFR Ch. I (4–1–17 Edition)

vice performs as intended for use in the treatment of acute ischemic stroke and must capture any adverse events asso- ciated with the device and procedure.

    - The labeling must include:

      - Information on the specific pa- tient population for which the device is intended for use in the treatment of acute ischemic stroke,  including  but not limited to, specifying time from symptom onset, vessels or location of the neurovasculature that can be accessed for treatment, and limitations on core infarct size.

      - Detailed instructions on proper device preparation and use for throm- bus	retrieval	from	the neurovasculature.

      - A summary of the clinical test- ing results, including a detailed sum- mary of the device- and procedure-re- lated complications and adverse events.

      - A shelf life.

[81 FR 94253, Dec. 23, 2016]

##### § 882.5700   Thermal  system  for  insom- nia.

- Identification. A thermal system for insomnia is a prescription device for use in patients with insomnia that is used to apply a specified tempera- ture to the skin surface.

- Classification. Class II (special controls). The special controls for this device are:

  - The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

  - Performance testing must dem- onstrate electromagnetic compat- ibility and electrical safety.

  - Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be eval- uated:

    - Thermal performance of the de- vice, including maintenance of the tar- get temperature, must be evaluated under simulated use conditions.

    - Mechanical testing to dem- onstrate the device can withstand forces under anticipated use condi- tions.

    - Mechanical testing to dem- onstrate the device is resistant to leak- age under anticipated use conditions.

#### Food and Drug Administration, HHS	§ 882.5810

  - Software verification, validation, and hazard analysis must be performed.

  - Patient labeling must be provided to convey information regarding safe use of the device, including instruc- tions for assembly.

[81 FR 44772, July 11, 2016]

##### § 882.5800   Cranial electrotherapy stim- ulator.

- Identification. A cranial electrotherapy stimulator is a device that applies electrical current to a pa- tient’s head to treat insomnia, depres- sion, or anxiety.

- Classification. Class III (premarket approval).

- Date a PMA or notice of completion of a PDP is required. No effective date has been established of the require- ment    for    premarket    approval.    See

§ 882.3.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17740,  May 11,  1987;  60  FR 43969,  Aug. 24,

1995;  62  FR 30457,  June 4,  1997;  73  FR  34860,

June 19, 2008]

##### § 882.5805   Repetitive transcranial mag- netic stimulation system.

- Identification. A repetitive transcranial magnetic stimulation sys- tem is an external device that delivers transcranial repetitive pulsed mag- netic fields of sufficient magnitude to induce neural action potentials in the prefrontal cortex to treat the symp- toms of major depressive disorder with- out inducing seizure in patients who have failed at least one antidepressant medication and are currently not on any antidepressant therapy.

- Classification. Class II (special controls). The special control is FDA’s ‘‘Class II Special Controls Guidance Document: Repetitive Transcranial Magnetic    Stimulation    System.’’    See

§ 882.1(e)   for   the   availability   of   this guidance document.

[76 FR 44491, July 26, 2011]

##### § 882.5808   Transcranial magnetic stim- ulator for headache.

- Identification. A transcranial mag- netic stimulator device for headache is a device that delivers brief duration, rapidly alternating, or pulsed, mag- netic fields that are externally directed at   spatially   discrete   regions   of   the

brain to induce electrical currents for the treatment of headache.

- Classification. Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing must demonstrate electromagnetic compat- ibility, electrical safety, and thermal safety.

- Appropriate verification, valida- tion, and hazard analysis must be per- formed on the device software and firmware.

- The elements of the device that contact the patient must be assessed to be biocompatible.

- Non-clinical testing data must demonstrate that the device performs as intended under anticipated condi- tions of use. This includes full charac- terization of the magnetic pulse output and resulting magnetic field map. This also includes characterization of the sound level of the device during use.

- Clinical testing must demonstrate that the device is safe and effective for treating headache in the indicated pa- tient population.

- The physician and patient label- ing must include the following:

  - A summary of the clinical per- formance testing, including any ad- verse events and complications.

  - The intended use population in terms of the types of headaches appro- priate for use with the device.

  - Information on how to report ad- verse events and device malfunctions.

  - A diagram or picture depicting the proper placement of the device on the user.

[78 FR 38458, July 8, 2014]

##### § 882.5810   External   functional  neuro- muscular stimulator.

- Identification. An external func- tional neuromuscular stimulator is an electrical stimulator that uses external electrodes for stimulating muscles in the leg and ankle of partially paralyzed patients (e.g., after stroke) to provide flexion of the foot and thus  improve the patient’s gait.

- Classification. Class II (perform- ance standards).

#### § 882.5820

##### § 882.5820   Implanted cerebellar stimu- lator.

  - Identification. An implanted cere- bellar stimulator is a device used to stimulate electrically a patient’s cere- bellar cortex  for  the  treatment  of  in- tractable epilepsy, spasticity, and some movement disorders. The stimulator consists of an implanted receiver with electrodes that are placed on the pa- tient’s cerebellum and an external transmitter for transmitting the stim- ulating pulses across the patient’s skin to the implanted receiver.

  - Classification. Class III (premarket approval).

  - Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or  notice  of  completion of a PDP is required to  be  filed  with the  Food  and  Drug   Administration   on or before September 26, 1984. Any im- planted cerebellar  stimulator  that  was not in commercial distribution  before May 28, 1976, or that has not on or be- fore September  26,  1984  been  found  by FDA  to  be  substantially  equivalent   to an   implanted    cerebellar    stimulator that was in commercial  distribution before May 28, 1976 shall have an ap- proved  PMA  or  declared   completed PDP in effect before beginning com- mercial    distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 49

FR 26574, June 28, 1984]

##### § 882.5830   Implanted	diaphragmatic/ phrenic nerve stimulator.

- Identification. An implanted dia- phragmatic/phrenic nerve stimulator is a device that provides electrical stimu- lation of a patient’s phrenic nerve to contract the diaphragm rhythmically and produce breathing in patients who have hypoventilation (a state in  which an abnormally low amount of air en- ters the lungs) caused by brain stem disease, high cervical spinal cord in- jury, or chronic lung disease. The stim- ulator consists of an implanted re- ceiver with electrodes that are placed around the patient’s phrenic nerve and an external transmitter for  transmit- ting the stimulating pulses across the patient’s skin to the implanted re- ceiver.

- Classification. Class III (premarket approval).

#### 21 CFR Ch. I (4–1–17 Edition)

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before July 7, 1986 for any implanted diaphragmatic/phrenic nerve stimulator that was in commercial dis- tribution before May 28, 1976, or that has on or before July 7, 1986 been found to be substantially equivalent to an implanted diaphragmatic/phrenic nerve stimulator that was in commercial dis- tribution before May 28, 1976. Any other implanted diaphragmatic/phrenic nerve stimulator shall have an ap- proved PMA or a declared completed PDP in effect before being placed in commercial   distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 51

FR 12101, Apr. 8, 1986]

##### § 882.5840 Implanted intracerebral/ subcortical stimulator for pain re- lief.

- Identification. An implanted intracerebral/subcortical    stimulator for pain relief is a device that applies electrical current to subsurface  areas of a patient’s brain to treat severe in- tractable pain. The stimulator  consists of an implanted receiver with elec- trodes that are placed within a  pa- tient’s brain and an external trans- mitter for transmitting  the  stimu- lating pulses across the  patient’s  skin to the implanted receiver.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before March 1, 1989, for any implanted	intracerebral/subcortical stimulator for pain relief that was in commercial distribution before May 28, 1976, or that has on or before March 1, 1989, been found to be substantially equivalent	to	an	implanted intracerebral/subcortical    stimulator for pain relief that was in commercial distribution before May 28, 1976. Any other implanted intracerebral/subcor- tical   stimulator   for   pain   relief   shall

#### Food and Drug Administration, HHS	§ 882.5870

have an approved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 53

FR 48621, Dec. 1, 1988]

##### § 882.5850   Implanted spinal cord stim- ulator for bladder evacuation.

  - Identification. An implanted spinal cord stimulator for bladder evacuation is an electrical stimulator used to empty the bladder of a paraplegic pa- tient who has a complete transection of the spinal cord and who is unable to empty his or her bladder by reflex means or by the intermittent use of catheters. The stimulator consists of an implanted receiver with electrodes that are placed on the conus medullaris portion of the patient’s spinal cord and an external transmitter for transmit- ting the stimulating pulses across the patient’s skin to the implanted re- ceiver.

  - Classification. Class III (premarket approval).

  - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any implanted spinal  cord  stimu- lator  for  bladder  evacuation   that   was in commercial distribution before May 28, 1976, or that has, on or before De- cember 26, 1996 been found to be sub- stantially equivalent to an implanted spinal cord stimulator for bladder evacuation that was in commercial dis- tribution before May 28, 1976. Any other implanted spinal cord stimulator for bladder evacuation shall have an approved PMA or a declared completed PDP in effect before being placed in commercial     distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17740, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 882.5860   Implanted	neuromuscular stimulator.

- Identification. An implanted neu- romuscular stimulator is a device that provides electrical stimulation to a pa- tient’s peroneal or femoral nerve to cause muscles in the leg to contract, thus improving the gait in a patient with a paralyzed leg. The stimulator consists of an implanted receiver with

electrodes that are placed around a pa- tient’s nerve and an external trans- mitter for transmitting the stimu- lating pulses across the patient’s skin to the implanted receiver. The external transmitter is activated by a switch in the heel in the patient’s shoe.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP for a device de- scribed in paragraph (b) of this section is required to be filed with the Food and Drug Administration on or before July 13, 1999 for any implanted neuro- muscular stimulator that was in com- mercial distribution before May 28, 1976, or that has, on or before July 13, 1999, been found to be substantially equivalent to an implanted neuro- muscular stimulator that was in com- mercial distribution before May 28, 1976. Any other implanted neuro- muscular stimulator shall have an ap- proved PMA or declared completed PDP in effect before being placed in commercial distribution.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR  17740,  May  11,  1987;  64  FR  18329,  Apr.  14,

1999]

##### § 882.5870   Implanted  peripheral nerve stimulator for pain relief.

- Identification. An implanted pe- ripheral nerve stimulator for pain re- lief is a device that is used to stimulate electrically a peripheral nerve in a pa- tient to relieve severe intractable pain. The stimulator consists of an im- planted receiver with electrodes that are placed around a peripheral nerve and an external transmitter for trans- mitting the stimulating pulses across the patient’s skin to the implanted re- ceiver.

- Classification. Class II (perform- ance standards).

[44 FR 51730, Sept. 4, 1979, as amended at 78

FR 18234, Mar. 26, 2013]

#### § 882.5880

##### § 882.5880   Implanted spinal cord stim- ulator for pain relief.

  - Identification. An implanted spinal cord stimulator for pain relief is a de- vice that is used to stimulate elec- trically a patient’s spinal cord to re- lieve severe intractable pain. The stim- ulator consists of an implanted re- ceiver with electrodes that are placed on the patient’s spinal cord and an ex- ternal transmitter for transmitting the stimulating pulses across the patient’s skin to the implanted receiver.

  - Classification. Class II (perform- ance standards).

##### § 882.5890   Transcutaneous	electrical nerve stimulator for pain relief.

- Identification. A transcutaneous electrical nerve stimulator for pain re- lief is a device used to apply an elec- trical current to electrodes on a pa- tient’s skin to treat pain.

- Classification. Class II (perform- ance standards).

##### § 882.5891   Transcutaneous	electrical nerve stimulator to treat headache.

- Identification. A transcutaneous electrical nerve stimulator to treat headache is a device used to apply an electrical current to a patient’s cra- nium through electrodes placed on the skin.

- Classification. Class II (special controls). The special controls for this device are:

  - The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

  - Appropriate analysis/testing must validate electromagnetic compatibility and electrical, mechanical, and ther- mal safety.

  - The technical parameters of the device, including waveform, output modes, maximum output voltage and current (with 500, 2,000, and 10,000 ohm loads), pulse duration, frequency, net charge (C) per pulse, maximum phase charge at 500 ohms, maximum current density (mA/cm2, r.m.s.), maximum av- erage current (mA), maximum average power density (W/cm2), and  the type of impedance monitoring system must be fully    characterized.

  - Electrical performance, adhesive integrity,   shelf   life,   reusability,   and

#### 21 CFR Ch. I (4–1–17 Edition)

current distribution testing of the elec- trodes must be conducted.

  - Appropriate software verification, validation,  and  hazard  analysis  must be performed.

  - Clinical performance data must demonstrate that the device is safe and effective as a treatment for headache in the indicated patient population.

  - Labeling must include the fol- lowing:

- Appropriate contraindications such as not for use in subjects with an implanted metallic or electronic device in the head, a cardiac pacemaker, or an implanted or wearable defibrillator.

- Appropriate warnings such as not to apply the device  on  the  neck  or chest, not to use the device in the pres- ence of electronic monitoring equip- ment, not to use in the bath or shower, not to use while sleeping, not to use while driving, not to use while oper- ating  machinery.

- Appropriate precautions such as the long-term  effects  of chronic  use  of the device are unknown.

- A summary of the expected risks and benefits of using the device.

- A summary of the clinical per- formance data, including information on the patient population for which the device has and has not been dem- onstrated to be effective, and any ad- verse events and complications.

- Information on how the device operates and the typical sensations ex- perienced during treatment.

- A detailed summary of the de- vice  technical  parameters.

- An expiration date/shelf life for the electrodes and the number of times they can be reused.

- Disposal instructions.

[79 FR 37948, July 3, 2014]

##### § 882.5894 Limited output trans- cutaneous piezoelectric stimulator for skin reactions associated with insect bites.

- Identification. A limited output transcutaneous piezoelectric stimu- lator for skin reactions associated with insect bites is a device intended to al- leviate skin reactions associated with insect bites via cutaneous, piezo- electric stimulation at the local site of the bite.

#### Food and Drug Administration, HHS	§ 882.5940

- Classification. Class II (special controls). The special controls for this device are:

  - Appropriate testing to charac- terize the electrical output specifica- tions of the device (i.e., total charge de- livered, maximum instantaneous out- put current, maximum instantaneous output voltage, pulse duration, charge density) must be conducted.

  - Mechanical bench testing must demonstrate that the device will with- stand the labeled number duration of uses.

  - All elements of the device that may contact the patient must be as- sessed to be biocompatible.

  - Labeling must include:

    - Validated instructions which ad- dresses the following:

      - Identification of areas of the body which are appropriate and not appro- priate for contact with the device.

      - Whether use of the device in con- junction  with  flammable  materials (e.g., insect repellent) is appropriate.

      - Use of the device on or near im- planted devices.

      - How to identify the correct type of skin condition.

    - Technical parameters of the de- vice (maximum output voltage (instan- taneous), maximum output current (in- stantaneous), and pulse duration).

    - Language to direct end users to contact the device manufacturer and MedWatch if they experience any ad- verse events with this device.

    - The anticipated number of device uses prior to failure.

[80 FR 15165, Mar. 23, 2015]

##### § 882.5895   Vibratory	counter-stimula- tion device.

- Identification. A vibratory counter-stimulation device is a pre- scription device that provides elec- trically powered mechanical vibration to improve the quality of sleep in pa- tients with primary Restless Legs Syn- drome.

- Classification. Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing must demonstrate electromagnetic compat- ibility (EMC), electrical safety, and thermal safety.

- If the device contains software or firmware, appropriate verification, val- idation, and hazard analysis must be performed.

- The elements of the device that contact the patient must be assessed to be biocompatible.

- Non-clinical testing data (includ- ing vibration frequency, amplitude, and acceleration) must demonstrate that the device performs as intended under anticipated conditions of use.

- Labeling must include:

  - Specific information pertinent to use of the device by the intended pa- tient population and the treatment regimen;

  - Warning to only use the device on normal, intact, clean, healthy skin;

  - Warning to not use the device if the user has leg skin disorders, such as eczema, psoriasis, cellulitis, non-heal- ing wounds;

  - Warning to discontinue use if Restless Leg Syndrome symptoms worsen; and

  - Instructions for end users to con- tact the  device  manufacturer  and MedWatch in case they experience any adverse events when using this device.

[82 FR 13554, Mar. 14, 2017]

##### § 882.5900   Preformed  craniosynostosis strip.

- Identification. A preformed craniosynostosis strip is a plastic strip used to cover bone edges of craniec- tomy sites (sites where the skull has been cut) to prevent the bone from re- growing in patients whose skull  su- tures are abnormally fused together.

- Classification. Class II (perform- ance standards).

##### § 882.5910   Dura substitute.

- Identification. A dura substitute is a sheet or material that is used to re- pair the dura mater (the membrane surrounding the brain).

- Classification. Class II (perform- ance standards).

##### § 882.5940   Electroconvulsive	therapy device.

- Identification.	An electroconvulsive therapy device is a device used for treating severe psy- chiatric disturbances (e.g., severe de- pression) by inducing in the patient a

#### § 882.5950

major motor seizure by applying a brief intense electrical current to the pa- tient’s head.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. No effective date has been established of the requirement for premarket approval. See § 882.3.

[44 FR 51730, Sept. 4, 1979, as amended at 52

FR 17740, May 11, 1987]

##### § 882.5950   Neurovascular  embolization device.

- Identification. A neurovascular embolization device is an intravascular implant intended to permanently oc- clude blood flow to cerebral aneurysms and cerebral ateriovenous malforma- tions. This does not include cyanoacrylates and other embolic agents, which act by polymerization or precipitation.  Embolization  devices used in other vascular applications are also not included in this classification, see § 870.3300.

- Classification. Class II (special controls.) The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Vascular and Neurovascular Embolization Devices.’’ For availability of this guidance docu- ment, see § 882.1(e).

[69 FR 77900, Dec. 29, 2004]

##### § 882.5960   Skull tongs for traction.

- Identification. Skull tongs for traction is an instrument used to im- mobilize a patient with a cervical spine injury (e.g., fracture or dislocation). The device is caliper shaped with tips that penetrate the skin. It is anchored to the skull and has a heavy weight at- tached to it that maintains, by trac- tion, the patient’s position.

- Classification. Class II (perform- ance standards).

##### § 882.5970   Cranial orthosis.

- Identification. A cranial orthosis is a device that is intended for medical purposes to apply pressure to promi- nent regions of an infant’s cranium in order to improve cranial symmetry and/or shape in infants from 3 to 18 months of age, with moderate to severe nonsynostotic	positional

#### 21 CFR Ch. I (4–1–17 Edition)

plagiocephaly, including infants with plagiocephalic-, brachycephalic-, and scaphocephalic-shaped heads.

- Classification. Class II (special controls) (prescription use in accord- ance with § 801.109 of this chapter, bio- compatibility testing, and labeling (contraindications, warnings, pre- cautions,  adverse  events,   instructions for  physicians  and  parents)).

[63 FR 40651, July 30, 1998]

##### § 882.5975   Human dura mater.

- Identification. Human dura mater is human pachymeninx tissue intended to repair defects in human dura mater.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Human Dura Mater.’’ See § 882.1(e) for the avail- ability  of  this  guidance.

- Scope. The classification set forth in this section is only applicable to human dura mater recovered prior to May 25, 2005.

[68 FR 70436, Dec. 18, 2003, as amended at 76

FR 36993, June 24, 2011]

## PART 884—OBSTETRICAL AND GYNECOLOGICAL DEVICES

##### Subpart A—General Provisions

Sec.

884.1    Scope.

884.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Obstetrical and Gynecological Diagnostic Devices

884.1040    Viscometer for cervical mucus. 884.1050   Endocervical aspirator.

884.1060   Endometrial aspirator.

884.1100   Endometrial brush.

884.1175    Endometrial    suction    curette    and accessories.

884.1185   Endometrial washer.

884.1300  Uterotubal	carbon	dioxide insufflator   and   accessories.

884.1425    Perineometer.

884.1550  Amniotic	fluid	sampler (amniocentesis    tray).

884.1560    Fetal blood sampler.

884.1600  Transabdominal	amnioscope (fetoscope)  and  accessories.

884.1630   Colposcope.

Food and Drug Administration, HHS	Pt. 884

884.1640   Culdoscope and accessories.

884.1660  Transcervical	endoscope (amnioscope)  and  accessories.

884.1690   Hysteroscope and accessories. 884.1700    Hysteroscopic insufflator.

884.1720    Gynecologic laparoscope and acces- sories.

884.1730    Laparoscopic insufflator.

##### Subpart C—Obstetrical and Gynecological Monitoring Devices

884.2050   Obstetric data analyzer.

884.2225   Obstetric-gynecologic	ultrasonic imager.

884.2600   Fetal cardiac monitor.

884.2620  Fetal electroencephalographic mon- itor.

884.2640    Fetal   phonocardiographic   monitor and  accessories.

884.2660    Fetal ultrasonic monitor and acces- sories.

884.2675    Fetal   scalp   circular   (spiral)   elec- trode  and  applicator.

884.2685   Fetal scalp clip electrode and appli- cator.

884.2700   Intrauterine  pressure  monitor  and accessories.

884.2720    External  uterine  contraction  mon- itor  and  accessories.

884.2730   Home uterine activity monitor. 884.2740   Perinatal  monitoring  system  and

accessories.

884.2800    Computerized labor monitoring sys- tem.

884.2900   Fetal stethoscope.

884.2960    Obstetric ultrasonic transducer and accessories.

884.2980    Telethermographic system.

884.2982    Liquid   crystal   thermographic   sys- tem.

884.2990    Breast   lesion   documentation   sys- tem.

##### Subpart D—Obstetrical and Gynecological Prosthetic Devices

884.3200   Cervical drain.

884.3575    Vaginal pessary.

884.3650    Fallopian tube prosthesis. 884.3900   Vaginal stent.

##### Subpart E—Obstetrical and Gynecological Surgical Devices

884.4050  Gynecologic	laparoscopic	power morcellation    containment    system.

884.4100    Endoscopic  electrocautery  and  ac- cessories.

884.4120    Gynecologic  electrocautery  and  ac- cessories.

884.4150    Bipolar  endoscopic  coagulator-cut- ter  and  accessories.

884.4160   Unipolar     endoscopic     coagulator- cutter  and  accessories.

884.4250    Expandable cervical dilator.

884.4260    Hygroscopic  Laminaria  cervical  di- lator.

884.4270    Vibratory cervical dilators. 884.4340    Fetal vacuum extractor.

884.4400   Obstetric forceps.

884.4500    Obstetric   fetal   destructive   instru- ment.

884.4520  Obstetric-gynecologic general man- ual  instrument.

884.4530   Obstetric-gynecologic	specialized manual  instrument.

884.4550    Gynecologic surgical laser. 884.4900    Obstetric table and accessories.

884.4910 Specialized  surgical  instrumenta- tion for use with urogynecologic surgical mesh.

##### Subpart F—Obstetrical and Gynecological Therapeutic Devices

884.5050    Metreurynter-balloon  abortion  sys- tem.

884.5070    Vacuum abortion system. 884.5100    Obstetric anesthesia set. 884.5150  Nonpowered breast pump. 884.5160  Powered breast pump.

884.5200   Hemorrhoid    prevention    pressure wedge.

884.5225   Abdominal decompression chamber. 884.5250  Cervical cap.

884.5300    Condom.

884.5310    Condom with spermicidal lubricant. 884.5320   Glans sheath.

884.5330   Female condom.

884.5350    Contraceptive diaphragm and acces- sories.

884.5360    Contraceptive    intrauterine    device (IUD)  and  introducer.

884.5380    Contraceptive   tubal   occlusion   de- vice  (TOD)  and  introducer.

884.5390    Perineal heater.

884.5400  Menstrual cup.

884.5425   Scented or scented deodorized men- strual  pad.

884.5435   Unscented menstrual pad.

884.5460   Scented or scented deodorized men- strual   tampon.

884.5470    Unscented menstrual tampon. 884.5900    Therapeutic   vaginal   douche   appa-

ratus.

884.5920    Vaginal insufflator.

884.5940    Powered  vaginal  muscle  stimulator for  therapeutic  use.

884.5960    Genital    vibrator    for    therapeutic use.

884.5970    Clitoral engorgement device. 884.5980    Surgical  mesh  for  transvaginal  pel-

vic organ prolapse repair.

##### Subpart G—Assisted Reproduction Devices

884.6100   Assisted reproduction needles. 884.6110    Assisted reproduction catheters. 884.6120    Assisted reproduction accessories. 884.6130    Assisted reproduction microtools. 884.6140    Assisted  reproduction  micropipette

fabrication instruments.

#### § 884.1

884.6150 Assisted	reproduction	micro- manipulators   and   microinjectors.

884.6160   Assisted reproduction labware. 884.6165    Intravaginal culture system. 884.6170   Assisted   reproduction   water   and

water purification systems.

884.6180   Reproductive    media    and    supple- ments.

884.6190   Assisted   reproductive   microscopes and  microscope  accessories.

884.6195 Assisted	reproduction	embryo image  assessment  system.

884.6200    Assisted reproduction laser system.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  45  FR  12684,  Feb.  26,  1980,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 884.1   Scope.

    - This part sets forth the classifica- tion of obstetrical and gynecological devices intended for human use that are in commercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - To avoid duplicative listings, an obstetrical and gynecological device that has two or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed only in one subpart.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 17740, May 11, 1987, as amended at 68

FR 44415, Aug. 27,  2003; 78  FR 18233, Mar. 26,

2013]

21 CFR Ch. I (4–1–17 Edition)

##### § 884.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device

#### Food and Drug Administration, HHS	§ 884.1040

before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17740, May 11, 1987]

##### § 884.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies    infectious    agents    by    using

deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2319, Jan. 14, 2000]

## Subpart B—Obstetrical and Gyne- cological Diagnostic Devices

##### § 884.1040   Viscometer	for	cervical mucus.

- Identification. A viscometer for cervical mucus  is a  device  that is  in- tended to measure the relative viscoelasticity of cervical mucus col- lected from a female patient. Measure- ments of relative viscoelasticity are in- tended for use as an adjunct in the clinical evaluation of a female with chronic infertility, to determine the time of ovulation and the penetrability of cervical mucus to motile sperm.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

§ 884.1050

premarket notification procedures in subpart E of part 807 of this chapter subject to § 884.9.

[47  FR  14706,  Apr.  6,  1982,  as  amended  at  65

FR 2320, Jan. 14, 2000]

##### § 884.1050   Endocervical aspirator.

  - Identification. An endocervical as- pirator is a device designed to remove tissue from the endocervix (mucous membrane lining the canal of the cer- vix of the uterus) by suction with a sy- ringe, bulb and pipette, or catheter. This device is used to evaluate endocervical tissue to detect malig- nant and premalignant lesions.

  - Classification. Class II (perform- ance standards).

##### § 884.1060   Endometrial aspirator.

- Identification. An endometrial as- pirator is a device designed to remove materials from the endometrium (the mucosal lining of the uterus) by suc- tion with a syringe, bulb and pipette, or catheter. This device is used to study endometrial cytology (cells).

- Classification. Class II. The special controls for this device are:

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and Testing,’ ’’ and

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

  - Labeling:

    - Indication: Only to evaluate the endometrium, and

    - Contraindications: Pregnancy, history of uterine perforation, or a re- cent cesarean section, and

  - The sampling component is cov- ered within vagina.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741,  May 11,  1987;  65  FR 17146,  Mar. 31,

2000]

##### § 884.1100   Endometrial brush.

- Identification. An endometrial brush is a device designed to remove samples of the endometrium (the mucosal lining of the uterus) by brush- ing its surface. This device is used to study endometrial cytology (cells).

- Classification. Class II. The special controls for this device are:

  - FDA’s:

#### 21 CFR Ch. I (4–1–17 Edition)

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and Testing,’ ’’ and

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

  - Labeling:

    - Indication: Only to evaluate the endometrium, and

    - Contraindications: Pregnancy, history of uterine perforation, or a re- cent cesarean section, and

  - Design and testing:

    - The sampling component is cov- ered within the vagina, and

    - For adherence of the bristles and brush head.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741,  May 11,  1987;  65  FR 17146,  Mar. 31,

2000]

##### § 884.1175   Endometrial suction curette and accessories.

- Identification. An endometrial suc- tion curette is a device used to remove material from the uterus and from the mucosal lining of the uterus by scrap- ing and vacuum suction. This device is used to obtain tissue for biopsy or for menstrual extraction. This  generic type of device may include catheters, syringes, and tissue filters or traps.

- Classification. Class II (perform- ance standards).

##### § 884.1185   Endometrial washer.

- Identification. An endometrial washer is a device used to remove ma- terials from the endometrium (the mucosal lining of the uterus) by wash- ing with water or saline solution and then aspirating with negative pressure. This device is used to study endometrial cytology (cells).

- Classification. Class II. The special controls for this device are:

  - FDA’s:

    - ‘‘Use of International Organiza- tion for Standardization’s ISO 10993 ‘Biological Evaluation of Medical De- vices—Part I: Evaluation and  Test- ing,’ ’’ and

    - ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

  - Labeling:

    - Indication: Only to evaluate the endometrium,

#### Food and Drug Administration, HHS	§ 884.1600

    - Contraindications: Pregnancy, history of uterine perforation, or a re- cent cesarean section, and

    - Warning: Do not attach to a wall or any external suction, and

  - Design and Testing:

    - The sampling component is cov- ered within the vagina, and

    - Intrauterine pressure should not exceed 50 millimeters of mercury.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741,  May 11,  1987;  65  FR 17146,  Mar. 31,

2000]

##### § 884.1300   Uterotubal   carbon   dioxide insufflator and accessories.

- Identification. A uterotubal carbon dioxide insufflator and accessories is a device used to test the patency (lack of obstruction) of the fallopian tubes by pressurizing the uterus and fallopian tubes and filling them with carbon di- oxide gas.

- Classification. Class II (perform- ance standards).

##### § 884.1425   Perineometer.

- Identification. A perineometer is a device consisting of a fluid-filled sack for intravaginal use that is attached to an external manometer. The devices measure the strength of the perineal muscles by offering resistence to a pa- tient’s voluntary contractions of these muscles and is used to diagnose and to correct, through exercise, uninary in- continence or sexual dysfunction.

- Classification. Class II (perform- ance standards).

##### § 884.1550   Amniotic	fluid	sampler (amniocentesis tray).

- Identification. The amniotic fluid sampler (amniocentesis tray) is a col- lection of devices used to aspirate amniotic fluid from the amniotic sac via a transabdominal approach. Com- ponents of the amniocentesis tray in- clude a disposable 3 inch 20 gauge nee- dle with stylet and a 30 cc. syringe, as well as the various sample collection accessories, such as vials, specimen containers, medium, drapes, etc. The device is used at 16–18 weeks gestation for antepartum diagnosis of certain congenital abnormalities or anytime after 24 weeks gestation when used to assess  fetal  maturity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 884.9.

[61  FR 1123,  Jan. 16,  1996,  as  amended  at 66

FR 33808, July 25, 2001]

##### § 884.1560   Fetal blood sampler.

- Identification. A fetal blood sam- pler is a device used to obtain fetal blood transcervically through an endo- scope by puncturing the fetal skin with a short blade and drawing blood into a heparinized tube. The fetal blood pH is determined and used in the diagnosis of fetal distress and fetal hypoxia.

- Classification. Class II (perform- ance standards).

##### § 884.1600   Transabdominal amnioscope (fetoscope) and accessories.

- Identification. A transabdominal amnioscope is a device designed to per- mit direct visual examination of  the fetus by a telescopic system via ab- dominal entry. The device is used to ascertain fetal abnormalities, to obtain fetal blood samples, or to obtain fetal tissue. This generic type of device may include the following accessories: tro- car and cannula, instruments used through an operating channel  or through a separate cannula associated with the amnioscope, light source and cables, and component parts.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food  and  Drug  Administra- tion on or before January 29, 1987  for any transabdominal amnioscope (fetoscope) and accessories that was in commercial distribution before May 28, 1976, or that has on or  before  January 29, 1987 been found to be substantially equivalent to a transabdominal amnioscope (fetoscope) and accessories that was in commercial  distribution before May 28, 1976. Any other transabdominal amnioscope (fetoscope) and accessories shall have an approved PMA  or  a  declared  completed  PDP  in

#### § 884.1630

effect before being placed in commer- cial distribution.

[45 FR 12684, Feb. 26, 1980, as amended at 51

FR 39845, Oct. 31, 1986]

##### § 884.1630   Colposcope.

  - Identification. A colposcope is a device designed to permit direct view- ing of the tissues of the vagina and cer- vix by a telescopic system located out- side the vagina. It is used to diagnose abnormalities and select areas for bi- opsy. This generic type of device may include a light source, cables, and com- ponent parts.

  - Classification. Class II (perform- ance standards).

##### § 884.1640   Culdoscope and accessories.

- Identification. A culdoscope is a device designed to permit direct view- ing of the organs within the peri- toneum by a telescopic system intro- duced into the pelvic cavity through the posterior vaginal fornix. It is used to perform diagnostic and surgical pro- cedures on the female genital organs. This generic type of device may in- clude trocar and cannula, instruments used through an operating channel, scope preheaters, light source and ca- bles, and component parts.

- Classification. (1)  Class  II  (per- formance standards).

(2) Class I for culdoscope accessories that are not part of a specialized in- strument or device delivery system; do not have adapters, connectors, chan- nels, or do not have portals for electrosurgical, laser, or other power sources. Such culdoscope accessory in- struments include: lens cleaning brush, biopsy brush, clip applier  (without clips), applicator, cannula (without trocar or valves), ligature carrier/nee- dle holder, clamp/hemostat/grasper, cu- rette, instrument guide, ligature pass- ing and knotting instrument, suture needle (without  suture), retractor,  me- chanical (noninflatable), snare, stylet, forceps, dissector, mechanical (non- inflatable) scissors, and suction/irriga- tion probe.  The  devices  subject to  this paragraph (b)(2) are exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38808,  July 25,

2001]

##### § 884.1660   Transcervical	endoscope (amnioscope) and accessories.

- Identification. A transcervical en- doscope is a device designed to permit direct viewing of the fetus and amniotic sac by means of an open tube introduced into the uterus through the cervix. The device may be used to vis- ualize the fetus or amniotic fluid and to sample fetal blood or amniotic fluid. This generic type of device may in- clude obturators, instruments used through an operating channel, light sources and cables, and component parts.

- Classification. Class II (perform- ance standards).

##### § 884.1690   Hysteroscope	and	acces- sories.

- Identification. A hysteroscope is a device used to permit direct viewing of the cervical canal and the uterine cav- ity by a telescopic system introduced into the uterus through the cervix. It is used to perform diagnostic and surgical procedures other than sterilization. This generic type of device may in- clude obturators and sheaths, instru- ments used through an operating chan- nel, scope preheaters, light sources and cables, and component parts.

- Classification. (1)  Class  II  (per- formance standards).

- Class I for hysteroscope acces- sories that are not part of a specialized instrument or device delivery system; do not have adapters, connectors, chan- nels, or do not have portals for electrosurgical, laser, or other power sources. Such hysteroscope accessory instruments include: lens cleaning brush, cannula (without trocar or valves), clamp/hemostat/grasper, cu- rette, instrument guide, forceps, dis- sector, mechanical (noninflatable), and scissors. The devices subject to this paragraph (b)(2) are exempt from the premarket   notification   procedures   in

#### Food and Drug Administration, HHS	§ 884.2050

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1123,  Jan. 16,  1996;  66  FR 38808,  July 25,

2001]

##### § 884.1700   Hysteroscopic insufflator.

  - Identification. A hysteroscopic insufflator is a device designed to dis- tend the  uterus  by filling  the  uterine cavity with a liquid or gas to facilitate viewing with a hysteroscope.

  - Classification. (1)  Class  II  (per- formance standards).

(2) Class I for tubing and tubing/filter fits which only include accessory in- struments that are not used to effect intrauterine access, e.g., hysteroscopic introducer sheaths, etc.; and single-use tubing kits used for only intrauterine insufflation. The devices subject to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38808,  July 25,

2001]

##### § 884.1720   Gynecologic	laparoscope and accessories.

- Identification. A gynecologic laparoscope is a device used to permit direct viewing of the organs within the peritoneum by a telescopic system in- troduced through the abdominal wall. It is used to perform diagnostic and surgical procedures on the female gen- ital organs. This generic type of device may include: Trocar and cannula, in- struments used through an operating channel, scope preheater, light source and cables, and component parts.

- Classification. (1)  Class  II  (per- formance standards).

(2) Class I for gynecologic laparoscope accessories that are not part of a specialized instrument or de- vice delivery system, do not have adapters, connector channels, or do not have portals for electrosurgical, lasers, or other power sources. Such gynecologic laparosope accessory in- struments include: the lens cleaning brush, biopsy brush, clip applier (with- out clips), applicator, cannula (without trocar or valves), ligature carrier/nee- dle holder, clamp/hemostat/grasper, cu- rette, instrument guide, ligature pass-

ing and knotting instrument, suture needle (without suture), retractor, me- chanical (noninflatable), snare, stylet, forceps, dissector, mechanical (non- inflatable), scissors, and suction/irriga- tion probe. The devices subject to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38808,  July 25,

2001]

##### § 884.1730   Laparoscopic insufflator.

- Identification. A laparoscopic insufflator is a device used to facilitate the use of the laparoscope by filling the peritoneal cavity with gas  to  distend it.

- Classification. (1)  Class  II  (per- formance standards).

(2) Class I for tubing and tubing/filter kits which include accessory  instru- ments that are not used to effect intra- abdominal	insufflation (pneumoperitoneum). The devices sub- ject to this paragraph (b)(2) are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,   subject   to   the   limitations   in

§ 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38809,  July 25,

2001]

## Subpart C—Obstetrical and Gyne- cological Monitoring Devices

##### § 884.2050   Obstetric data analyzer.

- Identification. An obstetric data analyzer (fetal status data analyzer) is a device used during labor to analyze electronic signal data obtained from fetal and maternal monitors. The ob- stetric data analyzer provides clinical diagnosis of fetal status and rec- ommendations for labor management and clinical interventions. This generic type of device may include signal anal- ysis and display equipment, electronic interfaces for other equipment, and power supplies and component parts.

- Classification: Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP is required to be

§ 884.2225

filed with the Food and Drug Adminis- tration on or before October 3, 2000, for any obstetric data analyzer  described in paragraph (a) of this section that was in commercial distribution before May 28, 1976, or that has been found, on or before October 3, 2000, to be substan- tially equivalent to an obstetric data analyzer described in paragraph (a) of this section that was in commercial distribution before May 28, 1976. Any other obstetric data analyzer described in paragraph (a) of this section shall have an approved PMA or declared completed PDP in effect before being placed in commercial distribution.

[65 FR 41332, July 5, 2000]

##### § 884.2225   Obstetric-gynecologic  ultra- sonic imager.

  - Identification. An obstetric- gynecologic ultrasonic imager is a de- vice designed to transmit and receive ultrasonic energy into and from a fe- male patient by pulsed echoscopy. This device is used to provide a visual rep- resentation of some physiological or artificial structure, or of a fetus, for di- agnostic purposes during a limited pe- riod of time. This generic type of de- vice may include the following: signal analysis and display equipment, elec- tronic interfaces for other equipment, patient and equipment supports, cou- pling gel, and component parts. This generic type of device does not include devices used to monitor the changes in some physiological condition over long periods of time.

  - Classification. Class II (perform- ance standards).

##### § 884.2600   Fetal cardiac monitor.

- Identification. A fetal cardiac monitor is a device used to ascertain fetal heart activity during pregnancy and labor. The device is designed to separate fetal heart signals from ma- ternal heart signals by analyzing elec- trocardiographic signals (electrical po- tentials generated during contraction and relaxation of heart muscle) ob- tained from the maternal abdomen with external electrodes. This generic type of device may include an alarm that signals when the heart rate crosses a preset threshold. This generic type of device includes the ‘‘fetal cardiotachometer (with sensors)’’ and

#### 21 CFR Ch. I (4–1–17 Edition)

the  ‘‘fetal  electrocardiographic  mon- itor.’’

- Classification. Class II (perform- ance standards).

##### § 884.2620   Fetal electroencephalographic monitor.

- Identification.	A	fetal electroencephalographic monitor is a device used to detect, measure, and record in graphic form (by means of one or more electrodes placed transcervically on the fetal scalp dur- ing labor) the rhythmically varying electrical skin potentials produced by the fetal brain.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any fetal electroencephalographic monitor that was in commercial dis- tribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a fetal electroencephalographic monitor in commercial distribution be- fore May 28, 1976. Any other fetal electroencephalographic monitor shall have an approved PMA or a declared completed PDP in effect before being placed in commercial distribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 884.2640   Fetal	phonocardiographic monitor and accessories.

- Identification.	A	fetal phonocardiographic monitor is a device designed to detect, measure, and record fetal heart sounds electronically, in graphic form, and noninvasively, to as- certain fetal condition during  labor. This generic type of device includes the following accessories: signal analysis and display equipment, patient and equipment supports, and other compo- nent parts.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 884.2720

##### § 884.2660   Fetal	ultrasonic	monitor and accessories.

  - Identification. A fetal ultrasonic monitor is a device designed to trans- mit and receive ultrasonic energy into and from the pregnant woman, usually by means of continuous wave (doppler) echoscopy. The device is used to rep- resent some physiological condition or characteristic in a measured value over a period of time (e.g., perinatal moni- toring during labor) or in an imme- diately perceptible form (e.g., use of the ultrasonic stethoscope). This ge- neric type of device may include the following accessories: signal analysis and display equipment, electronic interfaces for other equipment, patient and equipment supports, and compo- nent parts. This generic type of device does not include devices used to image some relatively unchanging physio- logical structure or interpret a physio- logical condition, but does include de- vices which may be set to alarm auto- matically at a predetermined threshold value.

  - Classification. Class II (perform- ance standards).

##### § 884.2675   Fetal scalp circular (spiral) electrode and applicator.

- Identification. A fetal scalp cir- cular (spiral) electrode and applicator is a device used to obtain a fetal elec- trocardiogram during labor and deliv- ery. It establishes electrical contact between fetal skin and an external monitoring device by a shallow sub- cutaneous puncture of  fetal scalp tis- sue with a curved  needle or needles. This generic type of device includes nonreusable spiral electrodes and reus- able circular electrodes.

- Classification. Class II (perform- ance standards).

##### § 884.2685   Fetal   scalp   clip   electrode and applicator.

- Identification. A fetal scalp clip electrode and applicator is a device de- signed to establish electrical contact between fetal skin and an external monitoring device by means of pinch- ing skin tissue with a nonreusable clip. This device is used to obtain a fetal electrocardiogram. This generic type of device may include a clip electrode ap- plicator.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any fetal scalp clip electrode and applicator that was in commercial dis- tribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a fetal scalp clip electrode and applicator that was in commercial dis- tribution before May 28, 1976. Any other fetal scalp clip electrode and ap- plicator shall  have  an  approved  PMA  or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 884.2700   Intrauterine  pressure  mon- itor and accessories.

- Identification. An intrauterine pressure monitor is a device designed to detect and measure intrauterine and amniotic fluid pressure with a catheter placed transcervically into the uterine cavity. The device is used to monitor intensity, duration, and frequency of uterine contractions during labor. This generic type of device may include the following accessories: signal analysis and display equipment, patient and equipment supports, and component parts.

- Classification. Class II (perform- ance standards).

##### § 884.2720   External    uterine    contrac- tion monitor and accessories.

- Identification. An external uterine contraction monitor (i.e., the tokodynamometer) is a device used to monitor the progress of labor. It meas- ures the duration, frequency, and rel- ative pressure of uterine contractions with a transducer strapped to the ma- ternal abdomen. This generic type of device may include an external pres- sure transducer, support straps, and other  patient  and  equipment  supports.

- Classification. Class II (perform- ance standards).

#### § 884.2730

##### § 884.2730   Home uterine activity mon- itor.

  - Identification. A home uterine ac- tivity monitor (HUAM) is an electronic system for at home antepartum meas- urement of uterine contractions, data transmission by telephone to a clinical setting, and for receipt and display of the uterine contraction data at the clinic. The HUAM system comprises a tocotransducer, an at-home recorder, a modem, and a computer and monitor that receive, process, and display data. This device is intended for use in women with a previous preterm deliv- ery to aid in the detection of preterm labor.

  - Classification. Class II (special controls); guidance document (Class II Special Controls Guidance for Home Uterine  Activity  Monitors).

[66 FR 14076, Mar. 9, 2001]

##### § 884.2740   Perinatal monitoring system and accessories.

- Identification. A perinatal moni- toring system is a device used to show graphically the relationship between maternal labor and the fetal heart rate by means of combining and coordi- nating uterine contraction and fetal heart monitors with appropriate dis- plays of the well-being of the fetus dur- ing pregnancy, labor, and delivery. This generic type of device may in- clude  any  of  the  devices  subject  to

§§ 884.2600,    884.2640,    884.2660,    884.2675,

884.2700, and 884.2720. This generic type of device may include the following ac- cessories:  Central  monitoring  system and remote repeaters,  signal  analysis and display equipment, patient and equipment supports, and component parts.

- Classification. Class II (perform- ance standards).

##### § 884.2800   Computerized  Labor  Moni- toring System.

- Identification. A computerized labor monitoring system is a system intended to continuously measure cer- vical dilation and fetal head descent and provide a display that indicates the progress of labor. The computerized labor monitoring system includes a monitor and ultrasound transducers. Ultrasound transducers are placed on

#### 21 CFR Ch. I (4–1–17 Edition)

the maternal abdomen and cervix  and on the fetal scalp to provide the matrix of measurements used to produce the display.

- Classification. Class II (special controls). The special controls are the FDA guidance document entitled: ‘‘Guidance for Industry and Food and Drug Administration Staff; Class II Special Controls Guidance Document: Computerized Labor Monitoring Sys- tems.’’ See § 884.1(e) for availability of this guidance document.

[72 FR 20227, Apr. 24, 2007]

##### § 884.2900   Fetal stethoscope.

- Identification. A fetal stethoscope is a device used for  listening to  fetal heart sounds. It is designed to transmit the fetal heart sounds not only through sound channels by air conduction, but also through the user’s head by tissue conduction into the user’s ears. It does not use ultrasonic energy. This device is designed to eliminate noise inter- ference commonly caused by handling conventional  stethoscopes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 66

FR 38809, July 25, 2001]

##### § 884.2960   Obstetric   ultrasonic   trans- ducer and accessories.

- Identification. An obstetric ultra- sonic transducer is a device used to apply ultrasonic energy to, and to re- ceive ultrasonic energy from, the body in conjunction with an obstetric mon- itor or imager. The device converts electrical signals into ultrasonic en- ergy, and vice versa, by means of an as- sembly distinct from an ultrasonic gen- erator. This generic type of device may include the following accessories: cou- pling gel, preamplifiers, amplifiers, sig- nal conditioners with their power sup- ply, connecting cables, and component parts. This generic type of device does not include devices used to generate the ultrasonic frequency electrical sig- nals for application.

- Classification. Class II (perform- ance standards).

#### Food and Drug Administration, HHS	§ 884.2990

##### § 884.2980   Telethermographic system.

  - Telethermographic system intended for adjunctive diagnostic screening for de- tection of breast cancer or other uses—(1) Identification. A telethermographic sys- tem for adjunctive diagnostic screen- ing for detection of breast cancer or other uses is an electrically powered device with a detector that is intended to measure, without touching the pa- tient’s skin, the self-emanating infra- red radiation that reveals the tempera- ture variations of the surface of the body. This generic type of device may include signal analysis and display equipment, patient and equipment sup- ports, component parts, and acces- sories.

(2) Classification. Class I (general con- trols).

  - Telethermographic system intended for use alone in diagnostic screening for detection of breast cancer or other uses—

- Identification. A telethermographic system for use as the sole diagnostic screening tool for detection of breast cancer or other uses is an electrically powered device with a detector that is intended to measure, without touching the patient’s skin, the self-emanating infrared radiation that reveals the temperature variations of the surface of the body. This generic type of device may include signal analysis and dis- play equipment, patient and equipment supports, component parts, and acces- sories.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of the enactment date of the amendments, May  28,  1976, an approval under section 515 of the act is required before the  device  described in paragraph (b)(1) may be  commer- cially distributed. See § 884.3.

[53  FR 1566,  Jan. 20,  1988,  as  amended  at 55

FR 48440, Nov. 20, 1990; 66 FR 46953, Sept. 10,

2001]

##### § 884.2982   Liquid	crystal	thermo- graphic system.

  - A nonelectrically powered or an AC- powered liquid crystal thermographic sys- tem intended for adjunctive use in diag- nostic screening for detection of breast cancer or other uses—(1) Identification. A nonelectrically powered or an AC-pow- ered liquid crystal thermographic sys- tem  intended  for  use  as  an  adjunct  to

physical palpation or mammography in diagnostic screening for detection of breast cancer or other uses is a non- electrically powered or an AC-powered device applied to the skin that displays the color patterns of heat sensitive cholesteric liquid crystals that respond to temperature variations of the sur- face of the body. This generic type of device may include patient and equip- ment supports, a means to ensure ther- mal contact between the patient’s skin and the liquid  crystals,  component parts, and accessories.

(2) Classification. Class I (general con- trols).

  - A nonelectrically powered or an AC-powered liquid crystal thermo- graphic system intended for use alone in diagnostic screening for detection of breast cancer or other uses—

    - Identification. A nonelectrically powered or an AC-powered liquid crys- tal thermographic system intended for use as the sole diagnostic screening tool for detection of breast cancer or other uses is a nonelectrically powered or an AC-powered device applied to the skin that displays the color patterns of heat sensitive cholesteric liquid crys- tals that respond to temperature vari- ations of the surface of the body. This generic type of device may include image display and recording equip- ment, patient and equipment supports, a means to ensure thermal contact be- tween the patient’s skin and the liquid crystals, component parts, and acces- sories.

    - Classification. Class III.

    - Date PMA or notice of completion of a PDP is required. As of the enactment date of the amendments, May  28,  1976, an approval under section 515 of the act is required before the  device  described in paragraph (b)(1) may be  commer- cially distributed. See § 884.3.

[53  FR 1566,  Jan. 20,  1988,  as  amended  at 55

FR 48441, Nov. 20, 1990; 66 FR 46953, Sept. 10,

2001]

##### § 884.2990   Breast lesion documentation system.

      - Identification. A breast lesion doc- umentation system is a device for use in producing a surface map of the breast as an aid to document palpable breast lesions identified during a clin- ical breast examination.

#### § 884.3200

      - Classification. Class II (special controls). The special control is FDA’s guidance entitled ‘‘Class II Special Controls Guidance Document: Breast Lesion   Documentation   System.’’   See

§ 884.1(e)   for   the   availability   of   this guidance document.

[68 FR 44415, Aug. 27, 2003]

## Subpart D—Obstetrical and Gynecological Prosthetic Devices

##### § 884.3200   Cervical drain.

- Identification. A cervical drain is a device designed  to  provide  an  exit channel for draining discharge from the cervix after pelvic surgery.

- Classification. Class II (perform- ance standards).

##### § 884.3575   Vaginal pessary.

- Identification. A vaginal pessary is a removable structure placed in the va- gina to support the pelvic organs and is used to treat conditions such as uter- ine prolapse (falling down of uterus), uterine retroposition (backward dis- placement), or gynecologic hernia.

- Classification. Class II (perform- ance standards).

##### § 884.3650   Fallopian tube prosthesis.

- Identification. A fallopian tube prosthesis is a device designed to main- tain the patency (openness) of the fal- lopian tube and is used after recon- structive surgery.

- Classification. Class II (perform- ance standards).

##### § 884.3900   Vaginal stent.

- Identification. A vaginal stent is a device used  to  enlarge  the  vagina  by stretching, or to support the vagina and to hold a skin graft after recon- structive surgery.

- Classification. Class II (perform- ance standards).

## Subpart E—Obstetrical and Gynecological Surgical Devices

##### § 884.4050 Gynecologic laparoscopic power morcellation containment system.

- Identification. A gynecologic laparoscopic power morcellation con- tainment system is a prescription de-

21 CFR Ch. I (4–1–17 Edition)

vice consisting of an instrument port and tissue containment method that creates a working space allowing for direct visualization during a power morcellation procedure following a laparoscopic procedure for the excision of benign gynecologic tissue that is not suspected to contain malignancy.

- Classification. Class II (special controls). The special controls for this device are:

  - The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible;

  - Device components that are la- beled sterile must be validated to a ste- rility assurance level of 10¥6;

  - Performance data must support shelf life by demonstrating continued sterility of the device or the sterile components, package integrity, and de- vice functionality over the intended shelf life;

  - Non-clinical performance data must demonstrate that the  device meets all design specifications and per- formance requirements. The following performance characteristics must be tested:

- Demonstration of the device im- permeability to tissue, cells,  and fluids;

- Demonstration that the device al- lows for the insertion  and  withdrawal of laparoscopic instruments while maintaining  pneumoperitoneum;

- Demonstration that the contain- ment system provides adequate space to perform morcellation and adequate visualization of the laparoscopic in- struments and tissue specimen relative to the external viscera;

- Demonstration that intended laparoscopic	instruments	and morcellators do not compromise the integrity of the containment system; and

- Demonstration that intended users can adequately deploy the device, morcellate a specimen without com- promising the integrity of the device, and remove the device without spillage of contents;

  - Training must be developed and validated to ensure users can follow the instructions for use; and

  - Labeling must include the fol- lowing:

#### Food and Drug Administration, HHS	§ 884.4150

    - A contraindication for use in gynecologic surgery in which the tissue to be morcellated is known or sus- pected to contain malignancy;

    - Unless clinical performance data demonstrates that it can be removed or modified, a contraindication for re- moval of uterine tissue containing sus- pected fibroids in patients who are: Peri- or postmenopausal, or candidates for en bloc tissue removal, for example, through the vagina or via a mini-lapa- rotomy incision;

    - The following boxed warning: ‘‘Warning: Information regarding the potential risks of a procedure with this device should be shared with patients. Uterine	tissue	may	contain unsuspected cancer. The use of laparoscopic power morcellators during fibroid surgery may spread cancer. The use of this containment system has not been clinically demonstrated to reduce this  risk.’’

    - A statement limiting use of de- vice to physicians who have completed the training program; and

    - An expiration date or shelf life.

[81 FR 40183, June 21, 2016]

##### § 884.4100   Endoscopic	electrocautery and accessories.

- Identification. An endoscopic electrocautery is a device used to per- form female sterilization under endoscopic observation. It is designed to coagulate fallopian tube tissue with a probe heated by low-voltage energy. This generic type of device may in- clude the following accessories: elec- trical generators, probes, and electrical cables.

- Classification. Class II. The special controls for this device are:

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review Guidance 2/12/90 (K–90),’’ and

- ‘‘Guidance (‘Guidelines’) for Evaluation of Laproscopic Bipolar and Thermal Coagulators (and Acces- sories),’’

  - International Electrotechnical Commission’s IEC 60601–1–AM2 (1995– 03),   Amendment   2,   ‘‘Medical   Electrical

Equipment—Part   1:   General   Require- ments  for  Safety,’’

  - American National Standards In- stitute/American Association for Med- ical Instrumentation’s HF–18, 1993, ‘‘Electrosurgical  Devices,’’

  - Labeling:

    - Indication: For female tubal steri- lization, and

    - Instructions for use:

      - Destroy at  least 2  centimeters of the fallopian tubes,

      - Use a cut or undampened sinus- oidal waveform,

      - Use a minimum power of 25 watts, and

      - For devices with ammeters: con- tinue electrode activation for 5 seconds after the visual endpoint (tissue blanching) is reached or current flow ceases indicating adequate tissue de- struction.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741,  May 11,  1987;  65  FR 17146,  Mar. 31,

2000]

##### § 884.4120   Gynecologic   electrocautery and accessories.

- Identification. A gynecologic electrocautery is a device designed to destroy tissue with high temperatures by tissue contact with an electrically heated probe. It is used to excise cer- vical lesions, perform biopsies, or treat chronic cervicitis under direct visual observation. This generic type of de- vice may include the following acces- sories: an electrical generator, a probe, and electrical cables.

- Classification. Class II (perform- ance standards).

##### § 884.4150   Bipolar   endoscopic   coagu- lator-cutter and accessories.

- Identification. A bipolar endoscopic coagulator-cutter is a de- vice used to perform female steriliza- tion and other operative procedures under endoscopic observation. It de- stroys tissue with high temperatures by directing a high frequency electrical current through tissue between two electrical contacts of a probe. This ge- neric type of device may include the following accessories: an electrical generator, probes, and electrical ca- bles.

- Classification. Class II. The special controls for this device are:

#### § 884.4160

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review Guidance 2/12/90 (K–90),’’ and

- ‘‘Guidance (‘Guidelines’) for Evaluation of Laproscopic Bipolar and Thermal Coagulators (and Acces- sories),’’

  - International Electrotechnical Commission’s IEC 60601–1–AM2 (1995– 03), Amendment 2, ‘‘Medical Electrical Equipment—Part 1: General Require- ments  for  Safety,’’

  - American National Standards In- stitute/American Association for Med- ical Instrumentation’s HF–18, 1993, ‘‘Electrosurgical  Devices,’’

  - Labeling:

    - Indication: For female tubal steri- lization, and

    - Instructions for use:

      - Destroy at  least 2  centimeters of the fallopian tubes,

      - Use a cut or undampened sinus- oidal waveform,

      - Use a minimum power of 25 watts, and

      - For devices with ammeters: con- tinue electrode activation for 5 seconds after the visual endpoint (tissue blanching) is reached or current flow ceases indicating adequate tissue de- struction.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741,  May 11,  1987;  65  FR 17146,  Mar. 31,

2000]

##### § 884.4160   Unipolar  endoscopic  coagu- lator-cutter and accessories.

- Identification. A unipolar endoscopic coagulator-cutter is a de- vice designed to destroy tissue with high temperatures by directing a high frequency electrical current through the tissue between an energized probe and a grounding plate. It is used in fe- male sterilization and in other opera- tive procedures under endoscopic obser- vation. This generic type of device may include the following accessories: an electrical generator, probes and elec- trical cables, and a patient grounding plate. This generic type of device does not include devices used to perform fe- male sterilization under hysteroscopic observation.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class II (perform- ance standards).

##### § 884.4250   Expandable cervical dilator.

- Identification. An expandable cer- vical dilator is an instrument with two handles and two opposing blades used manually to  dilate  (stretch open) the cervical os.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for  any  expandable  cervical   dilator that was in commercial distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to an ex- pandable cervical dilator that was in commercial distribution before May 28, 1976. Any other expandable cervical di- lator shall have an approved PMA or a declared completed PDP  in  effect  be- fore being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 884.4260   Hygroscopic Laminaria cer- vical dilator.

- Identification. A hygroscopic Lam- inaria cervical dilator is a device de- signed to dilate (stretch open) the cer- vical os by cervical insertion of a con- ical and expansible material made from the root of a seaweed (Laminaria digitata or Laminaria japonica). The de- vice is used to induce abortion.

- Classification. Class II (perform- ance standards).

##### § 884.4270   Vibratory cervical dilators.

- Identification. A vibratory cervical dilator is a device designed to  dilate the cervical os by stretching it with a power-driven vibrating probe head. The device is used to gain access to the uterus or to induce abortion, but is not to be used during labor when a viable fetus is desired or anticipated.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be

#### Food and Drug Administration, HHS	§ 884.4530

filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any vibratory cervical dilator that was in commercial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be sub- stantially equivalent to a vibratory cervical dilator that was in commer- cial distribution before May 28, 1976. Any other vibratory cervical dilator shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50708, Sept. 27,

1996]

##### § 884.4340   Fetal vacuum extractor.

  - Identification. A fetal vacuum ex- tractor is a device used to facilitate de- livery. The device enables traction to be applied to the fetal head (in the birth canal) by means of a suction cup attached to the scalp and is powered by an external vacuum source. This ge- neric type of device may include the cup, hosing, vacuum source, and vacu- um control.

  - Classification. Class II (perform- ance standards).

##### § 884.4400   Obstetric forceps.

- Identification. An obstetric forceps is a device consisting of two blades, with handles, designed to grasp and apply traction to the fetal head in the birth passage and facilitate delivery.

- Classification. Class II (perform- ance standards).

##### § 884.4500   Obstetric   fetal   destructive instrument.

- Identification. An obstetric fetal destructive instrument is a device de- signed to crush or pull the fetal body to facilitate the delivery of a dead or anomalous (abnormal) fetus. This ge- neric type of device includes the cleidoclast, cranioclast, craniotribe, and destructive hook.

- Classification. Class II (perform- ance standards).

##### § 884.4520   Obstetric-gynecologic	gen- eral manual instrument.

- Identification. An obstetric- gynecologic general manual instru- ment is one of a group of devices used

to perform simple obstetric and gynecologic manipulative functions. This generic type of device consists of the following:

- An episiotomy scissors is a cut- ting instrument, with two opposed shearing blades, used for surgical inci- sion of the vulvar orifice for obstet- rical purposes.

- A fiberoptic metal vaginal spec- ulum is a metal instrument, with fiberoptic light, used to expose and il- luminate the interior of the vagina.

- A metal vaginal speculum is a metal instrument used to expose the interior of the vagina.

- An umbilical scissors is a cutting instrument, with two opposed shearing blades, used to cut the umbilical cord.

- A uterine clamp is an instrument used to hold the uterus by compres- sion.

- A uterine packer is an instrument used to introduce dressing into the uterus or vagina.

- A vaginal applicator is an instru- ment used to insert medication into the vagina.

- A vaginal retractor is an instru- ment used to maintain vaginal expo- sure by separating the edges of the va- gina and holding back the tissue.

- A gynecological fibroid hook is an instrument used to exert traction upon a fibroid.

- A pelvimeter  (external)  is  an  in- strument used to measure the external diameters of the pelvis.

- Classification. Class I (general con- trols). The devices are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 54

FR 25052, June 12, 1989; 66 FR 38809, July 25,

2001]

##### § 884.4530   Obstetric-gynecologic	spe- cialized manual instrument.

- Identification. An obstetric- gynecologic specialized manual instru- ment is one of a group of devices used during obstetric-gynecologic proce- dures to perform manipulative diag- nostic and surgical functions (e.g., di- lating, grasping, measuring, and scrap- ing), where structural integrity is the chief criterion of device performance.

#### § 884.4550

This type of device consists of the fol- lowing:

  - An amniotome is an instrument used to rupture the fetal membranes.

  - A circumcision clamp is an instru- ment used to compress the foreskin of the penis during circumcision of a male infant.

  - An umbilical clamp is an instru- ment used to compress the umbilical cord.

  - A uterine curette is an instrument used to scrape and remove material from the uterus.

  - A fixed-size cervical dilator is any of a series of bougies of various sizes used to dilate the cervical os by stretching the cervix.

  - A uterine elevator is an instru- ment inserted into the uterus used to lift and manipulate the uterus.

  - A gynecological surgical forceps is an instrument with two blades and handles used to pull, grasp, or com- press during gynecological examina- tion.

  - A cervical cone knife is a cutting instrument used to excise and remove tissue from the cervix.

  - A gynecological cerclage needle is a looplike instrument used to suture the cervix.

  - A hook-type contraceptive intra- uterine device (IUD) remover is an in- strument used to remove an IUD from the uterus.

  - A gynecological fibroid screw is an instrument used to hold onto a fi- broid.

  - A uterine sound is an instrument used to determine the depth of the uterus by inserting it into the uterine cavity.

  - A cytological  cervical  spatula  is a blunt instrument used to scrape and remove cytological material from the surface of the cervix or vagina.

  - A gynecological biopsy forceps is an instrument with two blades and handles used for gynecological biopsy procedures.

  - A uterine tenaculum is a hook- like instrument used to seize and hold the cervix or fundus.

  - An internal pelvimeter is an in- strument used within the vagina to measure the diameter and capacity of the  pelvis.

#### 21 CFR Ch. I (4–1–17 Edition)

  - A nonmetal vaginal speculum is a nonmetal instrument used to expose the interior of the vagina.

  - A fiberoptic nonmetal vaginal speculum is a nonmetal instrument, with fiberoptic light, used to expose and illuminate the interior of the va- gina.

- Classification. (1)  Class  II  (per- formance standards).

- Class I for the amniotome, uterine curette, cervical dilator (fixed-size bou- gies), cerclage needle, IUD remover, uterine sound, and gynecological bi- opsy forceps.  The  devices  subject  to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38809,  July 25,

2001]

##### § 884.4550   Gynecologic surgical laser.

  - Identification. A  gynecologic sur- gical laser is a continuous wave carbon dioxide laser designed to destroy tissue thermally or to remove tissue by radi- ant light energy. The device is used only in conjunction with a colposcope as part of a gynecological surgical sys- tem. A colposcope is a magnifying lens system used to examine the vagina and cervix.

  - Classification. Class II (perform- ance standards).

##### § 884.4900   Obstetric  table  and  acces- sories.

- Identification. An obstetric table is a device with adjustable sections de- signed to support a patient in the var- ious positions required during obstetric and gynecologic procedures. This ge- neric type of device may include the following accessories: patient equip- ment, support attachments, and cabi- nets for warming instruments and dis- posing of wastes.

- Classification. Class II (perform- ance standards).

##### § 884.4910 Specialized surgical instru- mentation for use with urogynecologic surgical mesh.

- Identification. Specialized surgical instrumentation for use with urogynecologic surgical mesh is a pre- scription  device  specifically  intended

#### Food and Drug Administration, HHS	§ 884.5100

for use as an aid in the insertion, place- ment, fixation, or anchoring of surgical mesh during urogynecologic  proce- dures. These procedures include transvaginal pelvic organ prolapse re- pair, sacrocolpopexy (transabdominal pelvic organ prolapse repair), and treatment of female stress urinary in- continence. Examples of specialized surgical  instrumentation  include   nee- dle passers and trocars, needle guides, fixation tools, and tissue anchors. This device is  not a manual gastro- enterology-urology   surgical    instru- ment and accessories (§ 876.4730) or a manual surgical instrument for general use (§ 878.4800).

- Classification. Class II (special controls). The special controls for spe- cialized surgical instrumentation for use with urogynecologic surgical mesh are:

- The device must be demonstrated to be biocompatible;

- The device must be demonstrated to be sterile and, if reusable, it must be demonstrated that the device can be adequately reprocessed;

- Performance data must support the shelf life of the device by dem- onstrating package integrity and de- vice functionality over the requested shelf life;

- Non-clinical performance testing must demonstrate that the  device meets all design specifications and per- formance requirements, and that the device performs as intended under an- ticipated conditions of use; and

- Labeling must include:

  - Information regarding the mesh design that may be used with the de- vice;

  - Detailed summary of the clinical evaluations pertinent to use of the de- vice;

  - Expiration date; and

  - Where components are intended to be sterilized by the user prior to ini- tial use and/or are reusable, validated methods and instructions for steriliza- tion and/or reprocessing of any reus- able components.

[82 FR 1603, Jan. 6, 2017]

## Subpart F—Obstetrical and Gyne- cological Therapeutic De- vices

##### § 884.5050   Metreurynter-balloon   abor- tion system.

- Identification. A metreurynter-bal- loon abortion system is a device used to induce abortion. The device is in- serted into the uterine cavity, inflated, and slowly extracted. The extraction of the balloon from the uterus causes di- lation of the cervical os. This generic type of device may include pressure sources and pressure controls.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any metreurynter-balloon abortion system that was in commercial dis- tribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a metreurynter-balloon  abor- tion system that was in commercial distribution before May 28, 1976. Any other metreurynter-balloon abortion system shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50709, Sept. 27,

1996]

##### § 884.5070   Vacuum abortion system.

- Identification. A vacuum abortion system is a device designed to aspirate transcervically the products of concep- tion or menstruation from the uterus by using a cannula connected to a suc- tion source. This device is used for pregnancy termination or menstrual regulation. This type of device may in- clude aspiration cannula, vacuum source, and vacuum controller.

- Classification. Class II (perform- ance standards).

##### § 884.5100   Obstetric anesthesia set.

- Identification. An obstetric anes- thesia set is an assembly of antiseptic solution, needles, needle guides, sy- ringes, and other accessories, intended

§ 884.5150

for use with an anesthetic drug. This device is used to administer regional blocks (e.g., paracervical, uterosacral, and pudendal) that may be used during labor, delivery, or both.

- Classification. Class II (perform- ance standards).

##### § 884.5150   Nonpowered breast pump.

- Identification. A nonpowered breast pump is a manual suction device used to express milk from the breast.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter, subject to the limita- tions in § 884.9, if the device is using ei- ther a bulb or telescoping mechanism which does not develop more than 250 mm Hg suction, and the device mate- rials that contact breast or breast milk do not produce cytotoxicity,  irritation, or  sensitization  effects.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38809,  July 25,

2001]

##### § 884.5160   Powered breast pump.

- Identification. A powered breast pump in an electrically powered suc- tion device used to express milk from the breast.

- Classification. Class II (perform- ance standards).

##### § 884.5200   Hemorrhoid	prevention pressure wedge.

- Identification. A hemorrhoid pre- vention pressure wedge provides me- chanical support to the perianal region during the labor and delivery process. External mechanical support of the perianal region is intended to help pre- vent the occurrence of external hemor- rhoids associated with vaginal child- birth.

- Classification. Class II (special controls). The special controls for this device are:

- The sale, distribution, and use of this device are  restricted  to  prescrip- tion use in accordance with § 801.109 of this   chapter.

- The labeling must include specific instructions regarding the proper placement and use of the device.

- The device must be demonstrated to be biocompatible.

#### 21 CFR Ch. I (4–1–17 Edition)

- Mechanical bench testing of mate- rial strength  must  demonstrate  that the device will withstand forces en- countered during use.

- Safety and effectiveness data must demonstrate that the device pre- vents hemorrhoids  in women  under- going spontaneous vaginal delivery, in addition to general controls.

[76 FR 21238, Apr. 15, 2011]

##### § 884.5225   Abdominal	decompression chamber.

  - Identification. An abdominal de- compression chamber is a hoodlike de- vice used to reduce pressure on the pregnant patient’s abdomen for the re- lief of abdominal pain during preg- nancy or labor.

  - Classification. Class III (premarket approval).

  - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any abdominal decompression chamber that was in commercial dis- tribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to an abdominal decompression chamber that was in commercial dis- tribution before May 28, 1976. Any other abdominal decompression  cham- ber shall have an approved PMA or a declared completed PDP  in  effect  be- fore being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 17741, May 11, 1987; 61 FR 50709, Sept. 27,

1996]

##### § 884.5250   Cervical cap.

- Identification. A cervical cap is a flexible cuplike receptacle that fits over the cervix to collect menstrual flow or to aid artificial insemination. This generic type of device is not for contraceptive use.

- Classification. Class II (perform- ance standards).

##### § 884.5300   Condom.

- Identification. A condom is a sheath which completely covers the penis with a closely fitting membrane. The condom is used for contraceptive

#### Food and Drug Administration, HHS	§ 884.5350

and for prophylactic purposes (pre- venting transmission of sexually trans- mitted infections). The device may also be used to collect semen to aid in the diagnosis of infertility.

- Classification. (1) Class II (special controls) for condoms made of mate- rials other than natural rubber latex, including natural membrane (skin) or synthetic.

(2) Class II (special controls) for nat- ural rubber latex condoms. The guid- ance document entitled ‘‘Class II Spe- cial Controls Guidance Document: La- beling for Natural Rubber Latex Condoms Classified Under 21 CFR 884.5300’’ will serve as the special con- trol. See § 884.1(e) for the availability of this guidance document.

[73 FR 66538, Nov. 10, 2008]

##### § 884.5310   Condom   with   spermicidal lubricant.

- Identification. A condom with spermicidal lubricant is a sheath which completely covers the penis with a closely fitting membrane with a lubri- cant that contains a spermicidal agent, nonoxynol–9. This condom is used for contraceptive and prophylactic pur- poses (preventing transmission of vene- real disease).

- Classification. Class II (perform- ance standards).

[47 FR 49022, Oct. 29, 1982]

##### § 884.5320   Glans sheath.

- Identification. A glans sheath de- vice is a sheath which covers only the glans penis or part thereof and may also cover the area in the immediate proximity thereof, the corona and frenulum, but not the entire shaft  of the penis. It is indicated only for the prevention of pregnancy and not for the prevention of sexually-transmitted diseases.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before September 12, 2002, for any glans sheath that was in commer- cial distribution before May 28, 1976, or

that has, on or before September 12, 2002, been found to be substantially equivalent to a glans  sheath  that  was in commercial distribution before May 28, 1976. Any other glans sheath shall have an approved PMA or declared completed PDP in effect before being placed in commercial distribution.

[59 FR 67187, Dec. 29, 1994, as amended at 67

FR 40849, June 14, 2002]

##### § 884.5330   Female condom.

- Identification. A female condom is a sheath-like device that lines the vag- inal wall and is inserted into the va- gina prior to the initiation of coitus. It is indicated for contraceptive and pro- phylactic (preventing the transmission of sexually transmitted diseases) pur- poses.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before November 21, 2011, for any female condom that was in commercial distribution before May 28, 1976, or that has, on or before Novem- ber 21, 2011, been found to be substan- tially equivalent to any female condom that was in commercial distribution before May 28, 1976. Any other female condom shall have an approved PMA or declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[65 FR 31455, May 18, 2000, as amended at 76

FR 50667, Aug. 16, 2011]

##### § 884.5350   Contraceptive	diaphragm and accessories.

- Identification. A contraceptive dia- phragm is a closely fitting membrane placed between the posterior aspect of the pubic bone and the posterior vag- inal fornix. The device covers the cer- vix completely and is used with a spermicide to prevent pregnancy. This generic type of device may include an introducer.

- Classification. Class II (perform- ance standards).

#### § 884.5360

##### § 884.5360   Contraceptive    intrauterine device (IUD) and introducer.

  - Identification. A contraceptive intrauterine device (IUD) is a device used to prevent pregnancy. The device is placed high in the uterine fundus with a string extending from the device through the cervical os into the va- gina. This generic type of device in- cludes the introducer, but does not in- clude contraceptive IUD’s that  func- tion by drug activity, which are sub- ject to the new drug provisions of the Federal Food, Drug, and Cosmetic Act (see § 310.502).

  - Classification. Class III (premarket approval).

  - Labeling. Labeling requirements for contraceptive IUD’s are set forth in

§ 801.427.

  - Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with the Food and Drug Administra- tion on or before August 4, 1986, for any IUD and introducer that was in com- mercial distribution before May 28, 1976, or that has on or before August 4, 1986, been found to be substantially equivalent to an IUD and introducer that was in commercial distribution before May 28, 1976. Any other IUD and introducer shall have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 51

FR 16649, May 5, 1986]

##### § 884.5380   Contraceptive   tubal   occlu- sion device (TOD) and introducer.

- Identification. A contraceptive tubal occlusion device (TOD) and intro- ducer is a device designed to close a fallopian tube with a mechanical struc- ture, e.g., a band or clip on the outside of the fallopian tube or a plug or valve on the inside. The devices are used to prevent pregnancy.

- Classification. Class III (premarket approval).

- Date premarket approval application (PMA) or notice of completion of a prod- uct development protocol (PDP) is re- quired. A PMA or a notice of comple- tion of a PDP is required to be filed with  the  Food  and  Drug  Administra-

#### 21 CFR Ch. I (4–1–17 Edition)

tion on or before December 30, 1987, for any TOD and introducer that was in commercial distribution before May 28, 1976, or that has on or before December 30, 1987, been found to be substantially equivalent to a  TOD  and  introducer that was in commercial distribution before May 28, 1976. Any other TOD and introducer shall have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR 36883, Oct. 1, 1987]

##### § 884.5390   Perineal heater.

- Identification. A perineal heater is a device designed to apply heat di- rectly by contact, or indirectly from a radiant source, to the surface of the perineum (the area between the vulva and the anus) and is used to soothe or to help heal the perineum after an episiotomy (incision of the vulvar ori- fice for obstetrical purposes).

- Classification. Class II (perform- ance standards).

##### § 884.5400   Menstrual cup.

- Identification. A menstrual cup is a receptacle placed in the vagina to collect menstrual flow.

- Classification. Class II (perform- ance standards).

##### § 884.5425   Scented  or  scented  deodor- ized menstrual pad.

- Identification. A scented or scent- ed deodorized menstrual pad is a device that is a pad made of cellulosic or syn- thetic material which is used to absorb menstrual or other vaginal discharge. It has scent (i.e., fragrance materials) added for aesthetic purposes (scented menstrual pad) or for deodorizing pur- poses (scented deodorized menstrual pad). This generic type of device in- cludes sterile scented menstrual pads used for medically indicated condi- tions, but does not include menstrual pads treated with added antimicrobial agents or other drugs.

- Classification. (1) Class I (general controls) for menstrual pads made of common cellulosic and synthetic mate- rial with an established safety profile. The devices subject to this paragraph (b)(1) are exempt from the premarket notification procedures in subpart E of

#### Food and Drug Administration, HHS	§ 884.5920

part 807 of this chapter, subject to the limitations in § 884.9. This exemption does not include the intralabial pads and reusable menstrual pads.

- Class II (special controls) for scented or scented deodorized men- strual pads made of materials not de- scribed in paragraph (b)(1).

[45 FR 12684, Feb. 26, 1980, as amended at 45

FR  51185,  Aug.  1,  1980;  61  FR  67714,  Dec.  24,

1996; 66 FR 38809, July 25, 2001]

##### § 884.5435   Unscented menstrual pad.

  - Identification. An unscented men- strual pad is a device that is a pad made of cellulosic or synthetic mate- rial which is used to absorb menstrual or other vaginal discharge. This ge- neric type of device includes sterile unscented menstrual pads used for medically indicated conditions, but does not include menstrual pads treat- ed with scent (i.e., fragrance materials) or those with added antimicrobial agents or other drugs.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter only when the device is made of com- mon cellulosic and synthetic material with an established safety profile. This exemption does not include the intralabial pads  and reusable men- strual pads.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 67714,  Dec. 24,  1996;  65  FR 2320,  Jan. 14,

2000; 73 FR 34860, June 19, 2008]

##### § 884.5460   Scented  or  scented  deodor- ized menstrual tampon.

- Identification. A scented or scent- ed deodorized menstrual tampon is a device that is a plug made of cellulosic or synthetic material that is inserted into the vagina and used to absorb menstrual or other vaginal discharge. It has scent (i.e., fragrance materials) added for aesthetic purposes (scented menstrual tampon) or for deodorizing purposes (scented deodorized menstrual tampon). This generic type of device does not include menstrual tampons treated with added antimicrobial agents or other drugs.

- Classification. Class II (perform- ance standards).

[45 FR 12684, Feb. 26, 1980, as amended at 45

FR 51186, Aug. 1, 1980]

##### § 884.5470   Unscented   menstrual   tam- pon.

- Identification. An unscented men- strual tampon is a device that is a plug made of cellulosic or synthetic mate- rial that is inserted into the vagina and used to absorb menstrual or other vaginal discharge. This generic type of device does not include menstrual tam- pons treated with scent (i.e., fragrance materials) or those with added anti- microbial agents or other drugs.

- Classification. Class II (perform- ance standards).

##### § 884.5900   Therapeutic vaginal douche apparatus.

- Identification. A therapeutic vag- inal douche apparatus is a device that is a bag or bottle with tubing and a nozzle. The apparatus does not include douche solutions. The apparatus is in- tended and labeled for use in the treat- ment of medical conditions except it is not for contraceptive use. After filling the therapeutic vaginal douche appa- ratus with a solution, the patient uses the device to direct a stream of solu- tion into the vaginal cavity.

- Classification. (1)  Class  II  (per- formance standards).

(2) Class I if the device is operated by gravity feed. Devices subject to this paragraph (b)(2) are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38809,  July 25,

2001]

##### § 884.5920   Vaginal insufflator.

- Identification. A vaginal insufflator is a device used to treat vaginitis by introducing medicated powder from a hand-held bulb into the vagina through an open speculum.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807

#### § 884.5940

of  this  chapter,  subject  to  the  limita- tions in § 884.9.

[45 FR 12684, Feb. 26, 1980, as amended at 54

FR 25052, June 12, 1989; 66 FR 38809, July 25,

2001]

##### § 884.5940   Powered	vaginal	muscle stimulator for therapeutic use.

  - Identification. A powered vaginal muscle stimulator is an electrically powered device designed to stimulate directly the muscles of the vagina with pulsating electrical current. This de- vice is intended and labeled for thera- peutic use in increasing muscular tone and strength in the treatment of sexual dysfunction. This generic type of de- vice does not include devices used to treat urinary incontinence.

  - Classification. Class III (premarket approval).

  - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP for a device is re- quired to be filed with the Food and Drug Administration on or before July 12, 2000, for any powered vaginal mus- cle stimulator for therapeutic use that was in commercial distribution before May 28, 1976, or that has, on or before July 12, 2000, been found to be substan- tially equivalent to a powered vaginal muscle stimulator that was in commer- cial distribution before May 28, 1976. Any other powered vaginal muscle stimulator for therapeutic use shall have an approved PMA or declared completed PDP in effect before being placed in commercial distribution.

[45 FR 12684, Feb. 26, 1980, as amended at 52

FR  17741,  May  11,  1987;  65  FR  19834,  Apr.  13,

2000]

##### § 884.5960   Genital  vibrator  for  thera- peutic use.

- Identification. A genital vibrator for therapeutic use is an electrically operated device intended and labeled for therapeutic use in the treatment of sexual dysfunction or as an adjunct to Kegel’s exercise (tightening of the muscles of the pelvic floor to increase muscle tone).

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 884.5970 Clitoral engorgement de- vice.

- Identification. A clitoral engorgement device is designed to apply a vacuum to the clitoris. It is in- tended for use in the treatment of fe- male sexual arousal disorder.

- Classification. Class II (special controls). The special control is a guid- ance document entitled: ‘‘Guidance for Industry and FDA Reviewers: Class II Special Controls Guidance  Document for  Clitoral  Engorgement  Devices.’’

[65 FR 47306, Aug. 2, 2000]

##### § 884.5980   Surgical	mesh	for transvaginal pelvic organ prolapse repair.

- Identification. Surgical mesh for transvaginal pelvic organ prolapse  re- pair  is  a  prescription   device   intended to reinforce soft tissue  in  the  pelvic floor. This device is a porous implant that is made of synthetic material, non-synthetic material, or  a  combina- tion of synthetic and non-synthetic materials. This device does not include surgical mesh for other  intended  uses (§ 878.3300 of this chapter).

- Classification. Class III (premarket approval).

- Date premarket application approval or notice of completion of a product devel- opment protocol is required. A premarket application approval or notice of com- pletion of a product development pro- tocol for a device is required to be filed with the Food and Drug Administra- tion on or before July 5, 2018, for any surgical  mesh  described  in  paragraph

(a) of this section that was in commer- cial distribution before May 28, 1976, or that has, on or before July 5, 2018, been found substantially equivalent to a surgical  mesh  described  in  paragraph

- of this section that was in commer- cial distribution before May 28, 1976. Any other surgical mesh  for transvaginal pelvic organ prolapse re- pair shall have an approved premarket application or declared completed product development protocol in effect before being placed in commercial dis- tribution.

[81 FR 361, Jan. 5, 2016, as amended at 81 FR

369, Jan. 5, 2016]

#### Food and Drug Administration, HHS	§ 884.6140

Subpart G—Assisted Reproduction Devices

SOURCE: 63 FR 48436, Sept. 10, 1998, unless

otherwise noted.

##### § 884.6100   Assisted  reproduction  nee- dles.

  - Identification. Assisted reproduc- tion needles are devices used in in vitro fertilization	(IVF),	gamete intrafallopian transfer (GIFT), or other assisted reproduction procedures to ob- tain gametes from the body or intro- duce gametes, zygote(s), preembryo(s) and/or embryo(s) into the body. This generic type of device may include a single or double lumen needle and com- ponent parts, including needle guides, such as those used with ultrasound.

  - Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, design specifications, label- ing requirements, biocompatibility testing, and clinical testing).

##### § 884.6110   Assisted  reproduction  cath- eters.

- Identification. Assisted reproduc- tion catheters are devices used in in vitro fertilization (IVF), gamete intrafallopian transfer (GIFT), or other assisted reproduction procedures to in- troduce or remove gametes, zygote(s), preembryo(s), and/or embryo(s) into or from the body. This generic type of de- vice may include catheters, cannulae, introducers, dilators, sheaths, stylets, and component parts.

- Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, design specifications, label- ing requirements, biocompatibility testing, and clinical testing).

##### § 884.6120   Assisted reproduction acces- sories.

- Identification. Assisted reproduc- tion accessories are a group of devices used during assisted reproduction pro- cedures, in conjunction with assisted reproduction needles and/or assisted re- production catheters, to aspirate, incu- bate, infuse, and/or maintain tempera- ture. This generic type of device may include:

- Powered aspiration pumps used to provide low flow, intermittent vacuum for the aspiration of eggs (ova).

- Syringe pumps (powered or man- ual) used to activate a syringe to in- fuse or aspirate small volumes of fluid during assisted reproduction proce- dures.

- Collection tube warmers, used to maintain the temperature of egg (oo- cyte) collection tubes at or near body temperature. A dish/plate/microscope stage warmer is a device used to main- tain the temperature of the egg (oo- cyte)  during  manipulation.

- Embryo incubators, used to store and preserve gametes and/or embryos at or near body temperature.

- Cryopreservation instrumentation and devices,  used  to  contain,  freeze, and maintain gametes and/or embryos at an appropriate freezing temperature.

- Classification. Class II (special controls) (design specifications, label- ing requirements, and clinical testing).

##### § 884.6130   Assisted	reproduction microtools.

- Identification. Assisted reproduc- tion microtools are pipettes or other devices used in the laboratory to denude, micromanipulate, hold, or transfer human gametes or embryos for assisted hatching, intracytoplasmic sperm injection (ICSI), or other as- sisted reproduction methods.

- Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, design specifications, label- ing requirements, and clinical testing).

##### § 884.6140   Assisted	reproduction micropipette fabrication instru- ments.

- Identification. Assisted reproduc- tion micropipette fabrication devices are instruments intended to pull, bevel, or forge a micropipette or needle for intracytoplasmic sperm  injection (ICSI), in vitro fertilization (IVF) or other similar assisted reproduction procedures.

- Classification. Class II (special controls) (design specifications, label- ing requirements, and clinical testing).

#### § 884.6150

##### § 884.6150   Assisted	reproduction micromanipulators		and microinjectors.

  - Identification. Assisted reproduc- tion micromanipulators are devices in- tended to control the position of an as- sisted reproduction microtool. Assisted reproduction microinjectors are any device intended to control aspiration or expulsion of the contents of an as- sisted reproduction microtool.

  - Classification. Class II (special controls) (design specifications, label- ing requirements, and clinical testing).

##### § 884.6160   Assisted	reproduction labware.

- Identification. Assisted reproduc- tion labware consists of laboratory equipment or supplies intended to pre- pare, store, manipulate, or transfer human gametes or embryos for in vitro fertilization	(IVF),	gamete intrafallopian transfer (GIFT), or other assisted reproduction procedures. These include syringes, IVF tissue cul- ture dishes, IVF tissue culture plates, pipette tips, dishes, plates, and other vessels that come into physical contact with gametes, embryos or tissue cul- ture media.

- Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, design specifications, label- ing requirements, and clinical testing).

##### § 884.6165   Intravaginal culture system.

- Identification. An intravaginal cul- ture system is a prescription device in- tended for preparing, holding, and transferring human gametes or em- bryos during intravaginal in vitro fer- tilization or intravaginal culture pro- cedures.

- Classification. Class II (special controls). The special controls for this device are:

- Clinical performance testing must demonstrate the following:

- Comfort and retention of the intravaginal culture device;

- Adverse vaginal tissue reactions associated with intravaginal culture;

- Maximum number of gametes and/or embryos that can be placed in a device; and

- Rates of embryo development to the    designated    stage,    implantation

#### 21 CFR Ch. I (4–1–17 Edition)

rates, clinical pregnancy rates, live birth rates, and any adverse events or outcomes.

- Nonclinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. The following per- formance characteristics must be dem- onstrated:

- Mouse embryo assay testing to as- sess embryotoxicity by evaluating the gamete and embryo-contacting device components effect on the growth and development of mouse embryos to the blastocyst stage;

- Endotoxin testing on gamete and embryo-contacting components of the device;

- Cleaning and disinfection valida- tion of reusable device components;

- Sterility maintenance of the cul- ture media within the device through- out the vaginal incubation period and subsequent embryo extraction; and

- Ability of the device to permit ox- ygen and carbon dioxide exchange be- tween the media contained within the device and the external environment throughout the vaginal incubation pe- riod.

- The patient-contacting compo- nents of the device must be dem- onstrated to be biocompatible.

- Performance data must dem- onstrate the sterility of the device components intended to be provided sterile.

- Shelf life testing must dem- onstrate that the device maintains its performance characteristics and  the packaging of device components la- beled as sterile maintain integrity and sterility for the duration of the shelf life.

- Labeling for the device must in- clude:

  - A detailed summary of the clinical testing, including device effectiveness, device-related complications, and ad- verse events;

  - Validated methods and instruc- tions for reprocessing of reusable com- ponents;

  - The maximum number of gametes or embryos that can be loaded into the device;

  - A warning that informs  users that  the  embryo  development  is  first

#### Food and Drug Administration, HHS	§ 884.6195

evaluated  following  intravaginal  cul- ture; and

  - A statement that instructs the user to use legally marketed assisted reproductive technology media that contain elements to mitigate the con- tamination risk (e.g., antibiotics) and to support continued embryonic devel- opment over the intravaginal culture period.

- Patient labeling must be provided and must include:

  - Relevant warnings, precautions, and  adverse  effects  and  complications;

  - Information on how to use the de- vice;

  - The risks and benefits associated with the use of the device; and

  - A summary of the principal clin- ical device effectiveness results.

[81 FR 379, Jan. 6, 2016]

##### § 884.6170   Assisted reproduction water and water purification systems.

- Identification. Assisted reproduc- tion water purification systems are de- vices specifically intended to generate high quality, sterile, pyrogen-free water for reconstitution of media used for aspiration, incubation, transfer or storage of gametes or embryos for in vitro fertilization (IVF) or other as- sisted reproduction procedures. These devices may also be intended as the final rinse for labware or other assisted reproduction devices that will contact the gametes or embryos. These devices also include bottled water ready for re- constitution available from a vendor that is specifically intended for recon- stitution of media used for aspiration, incubation, transfer, or storage of gametes or embryos for IVF or other assisted reproduction procedures.

- Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, water quality testing, de- sign specifications, labeling require- ments, biocompatibility testing, and clinical  testing).

##### § 884.6180   Reproductive	media	and supplements.

- Identification. Reproductive media and supplement are products that are used for assisted reproduction proce- dures. Media include liquid and powder versions   of   various   substances   that

come in direct physical contact with human gametes or embryos (including water, acid solutions used to treat gametes or embryos, rinsing solutions, sperm separation media, supplements, or oil used to cover the media) for the purposes of preparation, maintenance, transfer or storage. Supplements are specific reagents added to media to en- hance specific properties of the media (e.g.,  proteins,  sera,  antibiotics,  etc.).

- Classification. Class II (special controls) (mouse embryo assay infor- mation, endotoxin testing, sterilization validation, design specifications, label- ing requirements, biocompatibility testing, and clinical testing).

##### § 884.6190 Assisted reproductive mi- croscopes and microscope acces- sories.

- Identification. Assisted reproduc- tion microscopes and microscope acces- sories (excluding microscope stage warmers, which are classified under as- sisted reproduction accessories) are op- tical instruments used to enlarge im- ages of gametes or embryos. Variations of microscopes and accessories used for these purposes would include phase contrast microscopes, dissecting mi- croscopes and inverted stage micro- scopes.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807 of this chapter, subject to the limita- tions in § 884.9.

[63 FR 48436, Sept. 10, 1998, as amended at 64

FR 62977, Nov. 18, 1999; 66 FR 38809, July 25,

2001]

##### § 884.6195 Assisted Reproduction Em- bryo Image Assessment System.

- Identification. An Assisted Repro- duction Embryo Image Assessment System is a prescription device that is designed to obtain and analyze light microscopy images of developing em- bryos. This device provides information to aid in the selection of embryo(s) for transfer when there are multiple em- bryos deemed suitable for transfer or freezing.

- Classification. Class II (special controls). The special  control(s)  for this device are:

- Clinical performance testing must demonstrate a reasonable assurance of

#### § 884.6200

safety and effectiveness of  the  device to predict embryo development. Classi- fication performance (sensitivity and specificity) and predictive accuracy (Positive Predictive Value  and  Nega- tive Predictive Value) must be assessed at the subject and embryo levels.

- Software validation, verification, and hazard analysis must be provided.

- Non-clinical performance testing data must demonstrate the perform- ance characteristics of the device. Testing must include the following:

  - Total light exposure and output testing;

  - A safety analysis must be per- formed based on maximum (worst-case) light exposure to embryos, which also includes the safety of the light wave- length(s) emitted by the device;

  - Simulated-use testing;

  - Mouse Embryo Assay testing to assess whether device operation im- pacts growth and development of mouse embryos to the blastocyst stage;

  - Cleaning and disinfection valida- tion of reusable components;

  - Package integrity and transit testing;

  - Hardware fail-safe validation;

  - Electrical equipment safety and electromagnetic compatibility testing; and

  - Prediction algorithm reproduc- ibility.

- Labeling must include the fol- lowing:

  - A detailed summary of clinical performance testing, including any ad- verse events;

#### 21 CFR Ch. I (4–1–17 Edition)

to ablate a small tangential hole in, or to thin, the zona pellucida of an em- bryo for assisted hatching or other as- sisted reproduction procedures.

(b) Classification. Class II (special controls). The special control is FDA’s guidance document entitled ‘‘Class II Special Controls Guidance Document: Assisted Reproduction Laser Systems.’’ See § 884.1(e) for the availability of this guidance document.

[69 FR 77624, Dec. 28, 2004]

## PART 886—OPHTHALMIC DEVICES

##### Subpart A—General Provisions

Sec.

886.1    Scope.

886.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

886.1040   Ocular esthesiometer.

886.1050   Adaptometer (biophotometer).

886.1070    Anomaloscope.

886.1090  Haidinger brush.

886.1120   Ophthalmic camera.

886.1140   Ophthalmic chair. 886.1150    Visual acuity chart.

886.1160    Color vision plate illuminator. 886.1170   Color vision tester.

886.1190  Distometer.

886.1200   Optokinetic drum.

886.1220   Corneal electrode.

886.1250    Euthyscope.

886.1270      Exophthalmometer.

886.1290   Fixation device.

886.1300   Afterimage flasher.

| (ii) Specific instructions, warnings, | 886.1320 | Fornixscope. |
| --- | --- | --- |
| precautions, and training needed for | 886.1330 | Amsler grid. |

safe use of the device

- Appropriate electromagnetic compatibility information;

- Validated methods and instruc- tions for cleaning and disinfection of reusable components; and

- Information identifying compat- ible cultureware and explain how they are used with the device.

[80 FR 10332, Feb. 26, 2015]

##### § 884.6200   Assisted  reproduction  laser system.

(a) Identification. The assisted repro- duction laser system is a device that images, targets, and controls the power and pulse duration of a laser beam used

886.1340   Haploscope.

886.1342    Strabismus detection device. 886.1350  Keratoscope.

886.1360    Visual field laser instrument. 886.1375   Bagolini lens.

886.1380   Diagnostic condensing lens.

886.1385   Polymethylmethacrylate    (PMMA) diagnostic  contact  lens.

886.1390    Flexible diagnostic Fresnel lens. 886.1395    Diagnostic Hruby fundus lens.

886.1400   Maddox lens.

886.1405    Ophthalmic trial lens set. 886.1410    Ophthalmic trial lens clip. 886.1415    Ophthalmic trial lens frame. 886.1420   Ophthalmic lens gauge.

886.1425    Lens measuring instrument.

886.1430    Ophthalmic    contact    lens    radius measuring   device.

886.1435    Maxwell spot.

886.1450    Corneal radius measuring device.

Food and Drug Administration, HHS	§ 886.1

886.1460    Stereopsis measuring instrument. 886.1500  Headband mirror.

886.1510   Eye movement monitor. 886.1570  Ophthalmoscope.

886.1605  Perimeter.

886.1630   AC-powered photostimulator.

886.1640    Ophthalmic preamplifier. 886.1650   Ophthalmic bar prism.

886.1655    Ophthalmic Fresnel prism. 886.1660   Gonioscopic prism.

886.1665    Ophthalmic rotary prism.

886.1670    Ophthalmic isotope uptake probe. 886.1680   Ophthalmic projector.

886.1690    Pupillograph.

886.1700  Pupillometer.

886.1750    Skiascopic rack.

886.1760    Ophthalmic refractometer.

886.1770   Manual refractor.

886.1780    Retinoscope.

886.1790  Nearpoint ruler.

886.1800    Schirmer strip.

886.1810   Tangent screen (campimeter). 886.1840    Simulatan   (including   crossed   cyl-

inder).

886.1850  AC-powered	slitlamp	biomicro- scope.

886.1860    Ophthalmic instrument stand. 886.1870  Stereoscope.

886.1880    Fusion and stereoscopic target. 886.1905   Nystagmus tape.

886.1910     Spectacle dissociation test system. 886.1925    Diurnal pattern recorder system. 886.1930   Tonometer and accessories.

886.1940   Tonometer sterilizer.

886.1945   Transilluminator.

##### Subpart C [Reserved]

Subpart D—Prosthetic Devices

886.3100    Ophthalmic tantalum clip. 886.3130   Ophthalmic conformer.

886.3200   Artificial eye.

886.3300    Absorbable   implant   (scleral   buck- ling   method).

886.3320    Eye sphere implant.

886.3340     Extraocular orbital implant. 886.3400  Keratoprosthesis.

886.3600   Intraocular lens.

886.3800   Scleral shell.

886.3920    Aqueous shunt.

##### Subpart E—Surgical Devices

886.4070   Powered corneal burr.

886.4100    Radiofrequency electrosurgical cau- tery   apparatus.

886.4115   Thermal cautery unit.

886.4150    Vitreous  aspiration  and  cutting  in- strument.

886.4155   Scleral plug.

886.4170    Cryophthalmic unit.

886.4230    Ophthalmic knife test drum. 886.4250     Ophthalmic electrolysis unit. 886.4270   Intraocular gas.

886.4275   Intraocular fluid.

886.4280   Intraocular  pressure  measuring  de- vice.

886.4300   Intraocular lens guide. 886.4335   Operating headlamp.

886.4350    Manual  ophthalmic  surgical  instru- ment.

886.4360    Ocular surgery irrigation device. 886.4370  Keratome.

886.4390   Ophthalmic laser.

886.4392  Nd:YAG	laser	for	posterior capsulotomy  and  peripheral  iridotomy.

886.4400    Electronic metal locator. 886.4440  AC-powered magnet.

886.4445   Permanent magnet.

886.4570    Ophthalmic surgical marker. 886.4610   Ocular pressure applicator. 886.4670    Phacofragmentation system.

886.4690    Ophthalmic photocoagulator. 886.4750   Ophthalmic eye shield.

886.4770    Ophthalmic     operating     spectacles (loupes).

886.4790   Ophthalmic sponge.

886.4855    Ophthalmic instrument table.

##### Subpart F—Therapeutic Devices

886.5100    Ophthalmic beta radiation source. 886.5120   Low-power binocular loupe.

886.5200    Eyelid thermal pulsation system. 886.5420   Contact lens inserter/remover.

886.5540   Low-vision magnifier.

886.5600    Ptosis crutch.

886.5700   Eyelid weight.

886.5800   Ophthalmic bar reader. 886.5810   Ophthalmic prism reader.

886.5820    Closed-circuit     television     reading system.

886.5838    Nasolacrimal compression device. 886.5840    Magnifying spectacles.

886.5842   Spectacle frame.

886.5844    Prescription spectacle lens. 886.5850   Sunglasses (nonprescription).

886.5870   Low-vision telescope. 886.5900    Electronic vision aid. 886.5905    Oral electronic vision aid.

886.5910    Image intensification vision aid. 886.5915   Optical vision aid.

886.5916    Rigid gas permeable contact lens. 886.5918    Rigid   gas   permeable   contact   lens

care products.

886.5925    Soft (hydrophilic) contact lens. 886.5928    Soft (hydrophilic) contact lens care

products.

886.5933  [Reserved]

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE:  52  FR  33355,  Sept.  2,  1987,  unless

otherwise noted.

EDITORIAL    NOTE:   Nomenclature   changes   to part 886 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 886.1   Scope.

    - This part sets forth the classifica- tion of ophthalmic devices intended for

§ 886.3

human use that are in commercial dis- tribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provision of a regulation in this part but shall state why the device is substantially equivalent to other de- vices, as required by § 807.87.

    - To avoid duplicative listings, an ophthalmic device that has two  or more types of uses (e.g., used both as a diagnostic device and as a therapeutic device) is listed in one subpart only.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 33355, Sept. 2, 1987, as amended at 73

FR 34860, June 19, 2008; 78 FR 18233, Mar. 26,

2013]

##### § 886.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a device, has an approval under section 515 of the act, FDA must promulgate a regula- tion under section 515(b) of the act re- quiring such approval, except as pro-

#### 21 CFR Ch. I (4–1–17 Edition)

vided in paragraphs (b) and (c) of this section. Such a regulation under sec- tion 515(b) of the act shall not be effec- tive during the grace period ending on the 90th  day  after  its promulgation or on the last day of the 30th full calendar month after the regulation that classi- fies the device into class III is effec- tive, whichever is later. See section 501(f)(2)(B) of the act. Accordingly, un- less an effective date of the require- ment for premarket approval is shown in the regulation for a device classified into class III in this part, the device may be commercially distributed with- out FDA’s issuance of an order approv- ing a PMA or declaring completed a PDP for the device.  If  FDA  promul- gates a regulation under section 515(b) of the act requiring premarket ap- proval for a device, section 501(f)(1)(A) of the act applies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

- A device identified in a regulation in this part that is classified into class III and that is subject to the transi- tional provisions of section 520(1) of the act is automatically classified by stat- ute into class III and must have an ap- proval under section 515 of the act be- fore being commercially distributed. Accordingly, the regulation for such a class III transitional device states that as of the enactment date of the amend- ments, May 28, 1976, the device must

#### Food and Drug Administration, HHS	§ 886.1050

have an approval under section 515 of the act before commercial distribution.

##### § 886.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2320, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 886.1040   Ocular esthesiometer.

- Identification. An ocular esthesiometer is a device, such as a sin- gle-hair brush, intended to touch the cornea to assess corneal sensitivity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35603,  Sept.  14,  1988;  59  FR  63012,  Dec.  7,

1994; 66 FR 38809, July 25, 2001]

##### § 886.1050   Adaptometer	(biophotom- eter).

- Identification. An adaptometer (biophotometer) is an AC-powered de- vice that provides a stimulating light source which has various controlled in- tensities intended to measure the time required for retinal adaptation (regen- eration of the visual purple) and the minimum light threshold.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

§ 886.1070

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38809,  July 25,

2001]

##### § 886.1070   Anomaloscope.

  - Identification. An anomaloscope is an AC-powered device intended to test for anomalies of color vision by dis- playing mixed spectral lines to be matched by the patient.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

##### § 886.1090   Haidinger brush.

- Identification. A  Haidinger brush is an AC-powered device that provides two conical brushlike images with apexes touching which are viewed by the patient through a Nicol prism and intended to evaluate visual function. It may include a component for meas- uring macular integrity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001; 72 FR 17400, Apr. 9, 2007]

##### § 886.1120   Ophthalmic camera.

- Identification. An ophthalmic cam- era is an AC-powered device intended to take photographs of the eye and the surrounding area.

- Classification. Class II.

[55 FR 48441, Nov. 20, 1990]

##### § 886.1140   Ophthalmic chair.

- Identification. An ophthalmic chair is an AC-powered or manual de- vice with adjustable positioning in which a patient is to sit or recline dur- ing ophthalmological examination or treatment.

- Classification. Class I. The AC- powered device and the manual device are exempt from the premarket notifi- cation procedures in subpart E of part

#### 21 CFR Ch. I (4–1–17 Edition)

807 of this chapter, subject to the limi- tations in § 886.9. The manual device is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,    with    respect    to    complaint files.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

##### § 886.1150   Visual acuity chart.

- Identification. A visual acuity chart is a device that is a chart, such as a Snellen chart with block letters or other symbols in graduated sizes, in- tended to test visual acuity.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35603,  Sept. 14,  1988;  53  FR 40825,  Oct. 18,

1988; 66 FR 38810, July 25, 2001]

##### § 886.1160   Color	vision	plate	illu- minator.

- Identification. A color vision plate illuminator is an AC-powered device that is a lamp intended to properly il- luminate color vision testing plates. It may include a filter.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

##### § 886.1170   Color vision tester.

- Identification. A color vision tester is a device that consists of various col- ored materials, such as colored yarns or  color  vision  plates  (multicolored

#### Food and Drug Administration, HHS	§ 886.1270

plates which patients with color vision deficiency would perceive as being of one color), intended to evaluate color vision.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35603, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1190   Distometer.

- Identification. A distometer is a device intended to measure the dis- tance between the cornea and a correc- tive lens during refraction to help measure the change of the visual image when a lens is in place.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35603, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1200   Optokinetic drum.

- Identification. An optokinetic drum is a drum-like device covered with alternating white and dark stripes or pictures that can be rotated on its handle. The device is intended to elicit and evaluate nystagmus (involuntary rapid movement of the eyeball) in pa- tients.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The

device is also exempt from the current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1220   Corneal electrode.

- Identification. A corneal electrode is an AC-powered device, usually part of a special contact lens, intended to be applied directly to the cornea to pro- vide data showing the changes in elec- trical potential in the retina after electroretinography (stimulation by light).

- Classification. Class II.

##### § 886.1250   Euthyscope.

- Identification. A euthyscope is a device that is a  modified  AC-powered or battery-powered ophthalmoscope (a perforated mirror device intended to inspect the interior of the eye) that projects a bright light encompassing an arc of about 30 degrees onto the fundus of the eye. The center of the light bun- dle is blocked by a black disk covering the fovea (the central depression of the macular retinae where only cones are present and blood vessels are lacking). The device is intended for use in the treatment of amblyopia (dimness of vi- sion without apparent disease of the eye).

- Classification. Class I for the bat- tery powered device. The battery pow- ered device is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter, sub- ject to the  limitations  in  § 886.9.  Class II for the AC-powered device.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

##### § 886.1270   Exophthalmometer.

- Identification.	An exophthalmometer is a device, such as a ruler, gauge, or caliper, intended to measure the degree of exophthalmos (abnormal protrusion of the eyeball).

#### § 886.1290

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1290   Fixation device.

- Identification. A fixation device is an AC-powered device intended for use as a fixation target for the patient dur- ing ophthalmological examination. The patient directs his or her gaze so that the visual image of the object falls on the fovea  centralis (the  center  of the macular retina of the eye.)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

##### § 886.1300   Afterimage flasher.

- Identification. An afterimage flasher is an AC-powered light that automatically switches on and off to allow performance of  an  afterimage test in which the patient indicates the positions of afterimages  after  the  light is off. The device is intended to deter- mine harmonious/anomalous retinal correspondence (the condition in which corresponding points on the retina have the same directional value).

- Classification. Class II.

[55 FR 48441, Nov. 20, 1990]

##### § 886.1320   Fornixscope.

- Identification. A fornixscope is a device intended to pull back and hold open the eyelid to aid examination of the conjunctiva.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### 21 CFR Ch. I (4–1–17 Edition)

and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1330   Amsler grid.

- Identification. An Amsler grid is a device that is a series of charts with grids of different sizes that are held at

30 centimeters distance from the pa- tient and intended to rapidly detect central  and  paracentral  irregularities in the visual field.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1340   Haploscope.

- Identification. A haploscope is an AC-powered device that consists of two movable viewing tubes, each con- taining a slide carrier, a low-intensity light source for the illumination of the slides, and a high-intensity light source for creating afterimages. The device is intended to measure stra- bismus (eye muscle imbalance), to as- sess binocular vision (use of both eyes to see), and to treat suppression and amblyopia (dimness of vision without any apparent disease of the eye).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR 63012,  Dec. 7,  1994;  66  FR 38810,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 886.1375

##### § 886.1342   Strabismus	detection	de- vice.

  - Identification. A strabismus detec- tion device is a prescription device de- signed to simultaneously illuminate both eyes with polarized light for auto- mated detection of strabismus by ana- lyzing foveal birefringence properties.

  - Classification. Class II (special controls). The special controls for this device are:

- Clinical performance testing must demonstrate the device performs as in- tended under anticipated conditions of use. Testing must be conducted in a representative patient population and clinical setting for the indicated use. Demonstration of clinical performance must include assessment of sensitivity and specificity compared to a clearly defined reference standard (e.g., com- prehensive ophthalmological examina- tion comprises age-appropriate visual acuity testing, examination of the ex- ternal ocular adnexae and orbit, ante- rior segment evaluation, extraocular motility evaluation, assessment of stereopsis, cycloplegic refraction, and dilated fundus examination).

- Non-clinical performance testing must demonstrate the device performs as intended under anticipated condi- tions of use. The following technical characteristics must be evaluated:

  - Verification of lowest detectable amount of deviation; and

  - Validation of the accuracy and precision at the lowest detectable amount of deviation.

- Software verification, validation, and hazard analysis must be performed.

- Optical radiation safety testing must demonstrate the device is safe per the directions for use.

- Performance testing must dem- onstrate the electromagnetic compat- ibility of the device.

- Performance testing must dem- onstrate the electrical safety of the de- vice.

- Labeling must include the fol- lowing:

  - Summaries of non-clinical and clinical performance testing;

  - Instructions on how to correctly use and maintain the device;

  - Instructions and explanation of all user-interface components; and

  - Information related to electro- magnetic compatibility and optical ra- diation  classification.

[81 FR 65280, Sept. 22, 2016]

##### § 886.1350   Keratoscope.

- Identification. A keratoscope is an AC-powered or battery-powered device intended to measure and evaluate the corneal curvature of the eye. Lines and circles within the keratoscope are used to observe the corneal reflex. This ge- neric type of device includes the photokeratoscope which records cor- neal curvature by taking  photographs of the cornea.

- The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 886.9. The battery-powered device is exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,  with  respect  to  complaint  files

[55 FR 48441, Nov. 20, 1990, as amended at 59

FR  63012,  Dec.  7,  1994;  65  FR  2320,  Jan.  14,

2000]

##### § 886.1360   Visual   field   laser   instru- ment.

- Identification. A visual field laser instrument is an AC-powered device in- tended to provide visible laser radi- ation that produces an interference pattern on the retina to evaluate ret- inal function.

- Classification. Class II.

##### § 886.1375   Bagolini lens.

- Identification. A Bagolini lens is a device that consists of a plane lens con- taining almost imperceptible striations that do not obscure visualization of ob- jects. The device is placed in a trial frame and intended to determine har- monious/anomalous retinal correspond- ence (a condition in which cor- responding points on the retina have the same directional values).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The

#### § 886.1380

device is also exempt from the current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1380   Diagnostic condensing lens.

  - Identification. A diagnostic con- densing lens is a device used in bin- ocular indirect ophthalmoscopy (a pro- cedure that produces an inverted or re- versed direct magnified image of the eye) intended to focus reflected light from the fundus of the eye.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38810, July 25,

2001]

##### § 886.1385   Polymethylmethacrylate (PMMA) diagnostic contact lens.

- Identification.	A polymethylmethacrylate (PMMA) diag- nostic contact lens is a device that is a curved shell of PMMA intended to be applied for a short period of time di- rectly on the globe or cornea of the eye for diagnosis or therapy of intraocular abnormalities.

- Classification. Class II.

##### § 886.1390   Flexible  diagnostic  Fresnel lens.

- Identification. A flexible diag- nostic Fresnel lens is a device that is a very thin lens which has its surface a concentric series of increasingly re- fractive zones. The device is intended to be applied to the back of the spec- tacle lenses of patients with aphakia (absence of the lens of the eye).

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1395   Diagnostic	Hruby	fundus lens.

- Identification. A diagnostic Hruby fundus lens is a device that is a 55 diop- ter lens  intended  for use in the exam- ination of the vitreous body and the fundus of the eye under slitlamp illu- mination  and  magnification.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1400   Maddox lens.

- Identification. A Maddox lens is a device that is a series of red cylinders that change the size, shape, and color of an image. The device is intended to be handheld or placed in a trial frame to evaluate eye muscle dysfunction.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation

#### Food and Drug Administration, HHS	§ 886.1435

in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1405   Ophthalmic trial lens set.

  - Identification. An ophthalmic trial lens set is a device that is a set of lenses of various dioptric powers in- tended to be handheld or inserted in a trial frame for vision testing to deter- mine refraction.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38811,  July 25,

2001]

##### § 886.1410   Ophthalmic trial lens clip.

- Identification. An ophthalmic trial lens clip is a device intended to hold prisms, spheres, cylinders, or occluders on a trial frame or spectacles for vision testing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1415   Ophthalmic trial lens frame.

- Identification. An opthalmic trial lens frame is a mechanical device in- tended to hold trial lenses for vision testing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1420   Ophthalmic lens gauge.

- Identification. An ophthalmic lens gauge is a calibrated device intended to manually measure the curvature of a spectacle lens.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35604, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1425   Lens measuring instrument.

- Identification. A lens measuring instrument is an AC-powered device in- tended to measure the power of lenses, prisms, and their centers (e.g., lensometer).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38811,  July 25,

2001]

##### § 886.1430   Ophthalmic contact lens ra- dius measuring device.

- Identification. An ophthalmic con- tact lens radius measuring device is an AC-powered device that is a microscope and dial gauge intended to measure the radius of a contact lens.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38811,  July 25,

2001]

##### § 886.1435   Maxwell spot.

- Identification. A Maxwell spot is an AC-powered device that is a light source with a red and blue filter in- tended to test macular function.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### § 886.1450

premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38811,  July 25,

2001]

##### § 886.1450   Corneal   radius   measuring device.

  - Identification. A corneal radius measuring device is an AC-powered de- vice intended to measure corneal size by superimposing the image of the cor- nea on a scale at the focal length of the lens of a small, hand held, single tube penscope or eye gauge magnifier.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations  in  § 886.9, only when the device does not include computer software in the unit or topographers.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38811,  July 25,

2001]

##### § 886.1460   Stereopsis	measuring	in- strument.

- Identification. A stereopsis meas- uring instrument is a device intended to measure depth perception by illu- mination of objects placed on different planes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1500   Headband mirror.

- Identification. A headband mirror is a device intended to be strapped to the head of the user to reflect light for use in examination of the eye.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38811, July 25,

2001]

##### § 886.1510   Eye movement monitor.

- Identification. An eye movement monitor is an AC-powered device with an electrode intended to measure and record ocular movements.

- Classification. Class II.

##### § 886.1570   Ophthalmoscope.

- Identification. An ophthalmoscope is an AC-powered or battery-powered device containing illumination and viewing optics intended to examine the media (cornea, aqueous, lens, and vit- reous) and the retina of the eye.

- Classification. Class II.

##### § 886.1605   Perimeter.

- Identification. A perimeter is an AC-powered or manual device intended to determine the extent of the periph- eral visual  field  of  a  patient. The  de- vice projects light on various points of a curved surface, and the patient indi- cates whether he or she sees the light.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[55 FR 48442, Nov. 20, 1990, as amended at 66

FR 38811, July 25, 2001]

#### Food and Drug Administration, HHS	§ 886.1665

##### § 886.1630   AC-powered photostimulator.

  - Identification. An AC-powered photostimulator is an AC-powered de- vice intended to provide light stimulus which allows measurement of retinal or visual function by perceptual or electrical methods (e.g., stroboscope).

  - Classification. Class II.

##### § 886.1640   Ophthalmic preamplifier.

- Identification. An ophthalmic pre- amplifier is an AC-powered or battery- powered device intended to amplify electrical signals from the eye in electroretinography (recording retinal action currents from the surface of the eyeball after stimulation by light), electrooculography (testing for retinal dysfunction by comparing the standing potential in the front and the back of the eyeball), and electromyography (recording electrical currents gen- erated in active muscle).

- Classification. Class II.

##### § 886.1650   Ophthalmic bar prism.

- Identification. An ophthalmic bar prism is a device that is a bar com- posed of fused prisms of gradually in- creasing strengths intended to measure latent and manifest strabismus (eye muscle deviation) or the power of fu- sion of a patient’s eyes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1655   Ophthalmic Fresnel prism.

- Identification. An ophthalmic Fresnel prism is a device that is a thin plastic sheet with embossed rulings which provides the optical effect of a prism. The device is intended to be ap-

plied to spectacle lenses to give a pris- matic effect.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1660   Gonioscopic prism.

- Identification. A gonioscopic prism is a device that is a prism intended to be placed on the eye to study the ante- rior chamber. The device may have an- gled mirrors to facilitate visualization of anatomical features.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35605,  Sept.  14,  1988;  59  FR  63013,  Dec.  7,

1994; 66 FR 38812, July 25, 2001]

##### § 886.1665   Ophthalmic rotary prism.

- Identification. An ophthalmic ro- tary prism is a device with various prismatic powers intended to be handheld and used to measure ocular deviation in patients with latent or manifest strabismus (eye muscle devi- ation).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### § 886.1670

and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1670   Ophthalmic  isotope  uptake probe.

  - Identification. An ophthalmic iso- tope uptake probe is an AC-powered de- vice intended to measure, by a probe which is placed in close proximity to the eye, the uptake of a radioisotope (phosphorus 32) by tumors to detect tumor masses on, around, or within the eye.

  - Classification. Class II.

##### § 886.1680   Ophthalmic projector.

- Identification. An ophthalmic pro- jector is an AC-powered device in- tended to project an image on a screen for vision testing.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38812,  July 25,

2001]

##### § 886.1690   Pupillograph.

- Identification. A pupillograph is an AC-powered device intended to meas- ure the pupil of the eye by reflected light and record the responses of the pupil.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38812,  July 25,

2001]

##### § 886.1700   Pupillometer.

- Identification. A pupillometer is an AC-powered or manual device intended to measure by reflected light the width or diameter of the pupil of the eye.

- Classification. Class I (general con- trols). The  AC-powered device  and the manual device are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The

#### 21 CFR Ch. I (4–1–17 Edition)

manual device is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with re- spect to general requirements con- cerning records, and § 820.198, with re- spect  to  complaint  files.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38812,  July 25,

2001]

##### § 886.1750   Skiascopic rack.

- Identification. A skiascopic rack is a device that is a rack and a set of at- tached ophthalmic lenses of various dioptric strengths intended as an aid in refraction.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38812,  July 25,

2001]

##### § 886.1760   Ophthalmic refractometer.

- Identification. An ophthalmic re- fractometer is an automatic AC-pow- ered device that consists of a fixation system, a measurement and recording system, and an alignment system in- tended to measure the refractive power of the eye by measuring light reflexes from the retina.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38812,  July 25,

2001]

##### § 886.1770   Manual refractor.

- Identification. A manual refractor is a device that is a set of lenses of varous dioptric powers intended to measure the refractive error of the eye.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation

#### Food and Drug Administration, HHS	§ 886.1810

in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1780   Retinoscope.

  - Identification. A retinoscope is an AC-powered or battery-powered device intended to measure the refraction of the eye by illuminating the retina and noting the direction of movement of the light on the retinal surface and of the refraction by the eye of the emer- gent rays.

  - Classification. (1) Class II (special controls) for the AC-powered device.

- Class I (general controls) for the battery-powered device. The class I battery-powered device is exempt from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 886.9. The battery-powered device is exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180 of this chapter, with re- spect to general requirements con- cerning records, and § 820.198 of this chapter,   with   respect   to   complaint files.

[55 FR 48442, Nov. 20, 1990; 55 FR 51799, Dec.

17, 1990, as amended at 65 FR 2320, Jan. 14,

2000]

##### § 886.1790   Nearpoint ruler.

  - Identification. A nearpoint ruler is a device calibrated in centimeters in- tended to measure the nearpoint of convergence (the point to which the visual lines are directed when conver- gence is at its maximum).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605,  Sept. 14,  1988;  53  FR 40825,  Oct. 18,

1988; 66 FR 38812, July 25, 2001]

##### § 886.1800   Schirmer strip.

- Identification. A Schirmer strip is a device made of filter paper or similar material intended to be inserted under a patient’s lower eyelid to stimulate and evaluate formation of tears.

- Classification. Class I (general con- trols). If the device is made of the same materials that were used in the device before May 28, 1976, the device is ex- empt from the premarket notification procedures in subpart E of part 807 of this chapter, subject  to  the  limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1810   Tangent    screen    (campim- eter).

- Identification. A tangent screen (campimeter) is an AC-powered or bat- tery-powered device that is a large square cloth chart with a central mark of fixation intended to map on a flat surface the central 30 degrees of a pa- tient’s visual field. This generic type of device includes projection tangent screens, target tangent screens and targets,  felt  tangent  screens,   and stereo campimeters.

- Classification. Class I (general con- trols). The  AC-powered device  and the battery-powered device are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 886.9. The battery-powered device is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,    with    respect    to    complaint files.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38812,  July 25,

2001]

#### § 886.1840

##### § 886.1840   Simulatan	(including crossed cylinder).

  - Identification. A simulatan (in- cluding crossed cylinder) is a device that is a set of pairs of cylinder lenses that provides various equal plus and minus refractive strengths. The lenses are arranged so that the user can ex- change the positions of plus and minus cylinder lenses of equal strengths. The device is intended for subjective refrac- tion (refraction in which the patient judges whether a given object is clearly in focus, as the examiner uses different lenses).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35605, Sept. 14, 1988; 66 FR 38812, July 25,

2001]

##### § 886.1850   AC-powered    slitlamp    bio- microscope.

- Identification. An AC-powered slitlamp biomicroscope is an AC-pow- ered device that is a microscope in- tended for use in eye examination that projects into a patient’s eye through a control diaphragm a thin, intense beam of light.

- Classification. Class II.

##### § 886.1860   Ophthalmic	instrument stand.

- Identification. An ophthalmic in- strument stand is an AC-powered or nonpowered device intended to store ophthalmic instruments in a readily accessible position.

- Classification. Class I (general con- trols). The  AC-powered device  and the battery-powered device are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 886.9. The battery-powered device is also exempt from the current good manufacturing     practice     requirements

#### 21 CFR Ch. I (4–1–17 Edition)

of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,  with  respect  to  complaint files.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38812,  July 25,

2001]

##### § 886.1870   Stereoscope.

- Identification. A stereoscope is an AC-powered or battery-powered device that combines the images of two simi- lar objects to produce a three-dimen- sional appearance of solidity and relief. It is intended to measure the angle of strabismus (eye muscle deviation), evaluate binocular vision (usage of both eyes to see), and guide a patient’s corrective exercises of eye muscles.

- Classification. Class I (general con- trols). The  AC-powered device  and the battery-powered device are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 886.9. The battery-powered device is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,  with  respect  to  complaint files.

[55 FR 48442, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38813,  July 25,

2001]

##### § 886.1880   Fusion and stereoscopic tar- get.

- Identification. A fusion and stereo- scopic target is a device intended for use as a viewing object with a stereo- scope (§ 886.1870).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### Food and Drug Administration, HHS	§ 886.1925

and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35606, Sept. 14, 1988; 66 FR 38813, July 25,

2001]

##### § 886.1905   Nystagmus tape.

  - Identification. Nystagmus tape is a device that is a long, narrow strip of fabric or other flexible material on which a series of objects are printed. The device is intended to be moved across a patient’s field of vision to elic- it optokinetic nystagmus (abnormal and irregular eye movements) and to test for blindness.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35606, Sept. 14, 1988; 66 FR 38813, July 25,

2001]

##### § 886.1910   Spectacle   dissociation   test system.

- Identification. A spectacle disso- ciation test system is an AC-powered or battery-powered device, such as a Lancaster test system, that consists of a light source and various filters, usu- ally red or green filters, intended to subjectively measure imbalance of ocu- lar muscles.

- Classification. Class I (general con- trols). The  AC-powered device  and the battery-powered device are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 886.9. The battery-powered device is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,    with    respect    to    complaint files.

[55 FR 48442, Nov. 20, 1990; 55 FR 51799, Dec.

17, 1990, as amended at 59 FR 63013, Dec. 7,

1994; 66 FR 38813, July 25, 2001]

##### § 886.1925   Diurnal	pattern	recorder system.

- Identification. A diurnal pattern recorder system is a nonimplantable, prescription device incorporating a tel- emetric sensor to detect changes in oc- ular dimension for monitoring diurnal patterns of intraocular pressure (IOP) fluctuations.

- Classification. Class II (special controls). The special controls for this device are:

  - Clinical performance data must demonstrate that the device and  all  of its components perform as intended under anticipated conditions of  use. The following performance characteris- tics must be demonstrated:

- Ability of the device to detect di- urnal changes.

- Tolerability of the system at the corneoscleral interface in the intended use population.

  - Nonclinical testing must validate measurements in an appropriate non- clinical testing model to  ensure  ability to detect changes in intraocular pres- sure.

  - Patient-contacting components must be demonstrated to be biocompat- ible.

  - Any component that is intended to contact the eye must be dem- onstrated to be sterile throughout its intended shelf life.

  - Software verification, validation, and hazard analysis must be performed.

  - Performance testing must dem- onstrate the electromagnetic compat- ibility and electromagnetic inter- ference of the device.

  - Performance testing must dem- onstrate electrical safety of the device.

  - Labeling must include the fol- lowing:

- Warning against activities and en- vironments that may put the user at greater risk.

- Specific instructions for the safe use of the device, which includes:

  - Description of all device compo- nents and instructions for assembling the device;

#### § 886.1930

  - Explanations of all available pro- grams and instructions for their use;

  - Instructions and  explanation  of all user-interface components;

  - Instructions on all safety fea- tures of the device; and

  - Instructions for properly main- taining the device.

- A summary of nonclinical test- ing information to describe EMC safety considerations.

- A summary of safety information obtained from clinical testing.

- Patient labeling to convey infor- mation regarding appropriate use of de- vice.

[81 FR 34270, May 31, 2016]

##### § 886.1930   Tonometer and accessories.

- Identification. A tonometer and ac- cessories is a manual device intended to measure intraocular pressure by ap- plying a known force on the globe of the eye and measuring the amount of indentation produced (Schiotz type) or to measure intraocular tension by applanation (applying a small flat disk to the cornea). Accessories for the de- vice may include a tonometer cali- brator or a tonograph recording sys- tem. The device is intended for use in the diagnosis of glaucoma.

- Classification. Class II.

##### § 886.1940   Tonometer sterilizer.

- Identification. A tonometer steri- lizer is an AC-powered device intended to heat sterilize a tonometer (a device used to measure intraocular pressure).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 886.9.

[55 FR 48443, Nov. 20, 1990, as amended at 65

FR 2321, Jan. 14, 2000]

##### § 886.1945   Transilluminator.

- Identification. A transilluminator is an AC-powered or battery-powered device that is a light source intended to transmit light through tissues to aid examination of patients.

- Classification. Class I for the bat- tery-powered device. The battery-pow- ered device is also exempt from the premarket notification procedures in subpart  E  of  part  807  of  this  chapter,

#### 21 CFR Ch. I (4–1–17 Edition)

subject   to   the   limitations   in   § 886.9. Class II for the AC-powered device.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38813,  July 25,

2001]

## Subpart C [Reserved]

Subpart D—Prosthetic Devices

##### § 886.3100   Ophthalmic tantalum clip.

- Identification. An ophthalmic tan- talum clip is a malleable metallic de- vice intended to be implanted perma- nently or temporarily to bring to- gether the edges of a wound to aid healing or prevent bleeding from small blood vessels in the eye.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 63

FR 59230, Nov. 3, 1998]

##### § 886.3130   Ophthalmic conformer.

- Identification. An ophthalmic con- former is a device usually made of molded plastic intended to be inserted temporarily between the eyeball and eyelid to maintain space in the orbital cavity and prevent closure or adhesions during the healing process following surgery. ]

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 63

FR 59230, Nov. 3, 1998]

##### § 886.3200   Artificial eye.

- Identification. An artificial  eye  is a device resembling the anterior por- tion of the eye, usually made of glass or plastic, intended to be inserted in a patient’s eye socket anterior to an or- bital implant, or the eviscerated eye- ball, for cosmetic purposes. The device is not intended to be implanted.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject  to  the  limitations  in  § 886.9,  if

Food and Drug Administration, HHS	§ 886.4070

the device is made from the same ma- terials, has the same chemical com- position, and uses the same manufac- turing processes as currently legally marketed devices.

[61  FR 1124,  Jan. 16,  1996,  as  amended  at 66

FR 38813, July 25, 2001]

##### § 886.3300   Absorbable  implant  (scleral buckling method).

  - Identification. An absorbable im- plant (scleral buckling method) is a de- vice intended to be implanted on the sclera to aid retinal reattachment.

  - Classification. Class II.

##### § 886.3320   Eye sphere implant.

- Identification. An eye sphere im- plant is a device intended to be im- planted in the eyeball to occupy space following the removal of the contents of the eyeball with the sclera left in- tact.

- Classification. Class II.

##### § 886.3340   Extraocular orbital implant.

- Identification. An extraocular or- bital implant is a nonabsorbable device intended to be implanted during scleral surgery for buckling or building up the floor of the eye, usually in conjunction with retinal reattachment. Injectable substances are excluded.

- Classification. Class II.

##### § 886.3400   Keratoprosthesis.

- Identification. A keratoprosthesis is a device intended to provide a trans- parent optical pathway through an opacified	cornea,	either intraoperatively or permanently, in an eye that is not a reasonable candidate for a corneal transplant.

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

  - ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’ and

  - ‘‘Guidance on 510(k) Submissions for  Keratoprostheses.’’

[65 FR 17147, Mar. 31, 2000]

##### § 886.3600   Intraocular lens.

    - Identification. An intraocular lens is a device made of materials such as

glass or plastic intended to be im- planted to replace the natural  lens  of an eye.

    - Classification. Class III.

    - Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 886.3.

##### § 886.3800   Scleral shell.

- Identification. A scleral shell is a device made of glass or plastic that is intended to be inserted for short time periods over the cornea and proximal- cornea sclera for cosmetic or recon- structive purposes. An artificial eye is usually painted on the device. The de- vice is not intended to be implanted.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 63

FR 59230, Nov. 3, 1998]

##### § 886.3920   Aqueous shunt.

- Identification. An aqueous shunt is an implantable device intended to re- duce intraocular pressure in the ante- rior chamber of the eye in patients with neovascular glaucoma or with glaucoma when medical and conven- tional surgical treatments have failed.

- Classification. Class II. The special controls for this device are FDA’s:

  - ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

  - ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’ and

  - ‘‘Aqueous Shunts—510(k) Submis- sions.’’

[65 FR 17147, Mar. 31, 2000, as amended at 66

FR 18542, Apr. 10, 2001]

## Subpart E—Surgical Devices

##### § 886.4070   Powered corneal burr.

    - Identification. A powered corneal burr is an AC-powered or battery-pow- ered device that is a motor and drilling tool intended to remove rust rings from the cornea of the eye.

    - Classification. Class I (general con- trols). When intended only for rust ring

§ 886.4100

removal, the device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 886.9.

[55 FR 48443, Nov. 20, 1990; 55 FR 51799, Dec.

17, 1990, as amended at 65 FR 2321, Jan. 14,

2000]

##### § 886.4100   Radiofrequency electrosurgical cautery apparatus.

      - Identification. A radiofrequency electrosurgical cautery apparatus is an AC-powered or battery-powered device intended for use during ocular surgery to coagulate tissue or arrest  bleeding by a high frequency electric current.

      - Classification. Class II.

##### § 886.4115   Thermal cautery unit.

- Identification. A thermal cautery unit is an AC-powered or battery-pow- ered device intended for use during oc- ular surgery to coagulate tissue or ar- rest bleeding by heat conducted through a wire tip.

- Classification. Class II.

##### § 886.4150   Vitreous aspiration and cut- ting instrument.

- Identification. A vitreous aspira- tion and cutting instrument is an elec- trically powered device, which may use ultrasound, intended to remove vit- reous matter from the vitreous cavity or remove a crystalline lens.

- Classification. Class II.

##### § 886.4155   Scleral plug.

- Identification. A scleral plug is a prescription device intended to provide temporary closure of a scleral incision during an ophthalmic surgical proce- dure. These plugs prevent intraocular fluid and pressure loss when instru- ments are withdrawn from the eye. Scleral plugs include a head portion re- maining above the sclera, which can be gripped for insertion and removal, and a shaft that fits inside the scleral inci- sion. Scleral plugs are removed before completing the surgery.

- Classification. Class II (special controls). The special controls for the scleral plug are as follows:

  - The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to the limitations in § 886.9  if the  material  is  a  surgical  grade  stain-

#### 21 CFR Ch. I (4–1–17 Edition)

less steel with or without a gold, silver, or titanium coating. The special con- trols for the surgical grade stainless steel scleral plug (with or without a gold, silver, or titanium coating) are:

- The device must be demonstrated to be sterile during the labeled shelf life;

- The device must be demonstrated to be biocompatible; and

- Labeling must include all infor- mation required for the safe and effec- tive use of the device, including spe- cific instructions regarding the proper sizing, placement, and removal of the device.

  - The device is not exempt from premarket notification procedures if it is composed of a material other than surgical grade stainless steel (with or without a gold, silver, or titanium coating). The special controls for scleral plugs made of other materials are:

    - The device must be demonstrated to be sterile during the labeled shelf life;

    - The device must be demonstrated to be biocompatible;

    - Characterization of the device materials must be performed;

    - Performance data must dem- onstrate acceptable mechanical prop- erties under simulated clinical use con- ditions including insertion and removal of the device;

    - Performance data must dem- onstrate adequately low levels of the extractables or residues from manufac- turing (or processing) of the device; and

    - Labeling must include all infor- mation required for the safe and effec- tive use of the device, including spe- cific instructions regarding the proper sizing, placement, and removal of the device.

[78 FR 68715, Nov. 15, 2013]

##### § 886.4170   Cryophthalmic unit.

- Identification. A cryophthalmic unit is a device that is a probe with a small tip that becomes extremely cold through the controlled use of a refrig- erant or gas. The device may be AC- powered. The device is intended to re- move cataracts by the formation of an adherent ice ball in the lens, to freeze

#### Food and Drug Administration, HHS	§ 886.4335

the  eye  and  adjunct  parts  for  surgical removal of scars, and to freeze tumors.

- Classification. Class II.

##### § 886.4230   Ophthalmic knife test drum.

- Identification. An ophthalmic knife test drum is a device intended to test the keenness of ophthalmic sur- gical knives to determine whether re- sharpening is needed.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35606, Sept. 14, 1988; 66 FR 38813, July 25,

2001]

##### § 886.4250   Ophthalmic	electrolysis unit.

- Identification. An ophthalmic elec- trolysis unit is an AC-powered or bat- tery-powered device intended to de- stroy ocular hair follicles by applying a galvanic electrical current.

- Classification. Class I for the bat- tery-powered device. Class II for the AC-powered device. The battery-pow- ered device is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter, sub- ject to the limitations in § 886.9.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63013,  Dec. 7,  1994;  66  FR 38813,  July 25,

2001]

##### § 886.4270   Intraocular gas.

- Identification. An intraocular gas is a device consisting of a gaseous fluid intended to be introduced into the eye to place pressure on a detached retina.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 886.3.

##### § 886.4275   Intraocular fluid.

- Identification. An intraocular fluid is a device consisting of a nongaseous fluid intended to be introduced into the eye to aid performance of surgery, such as to maintain anterior chamber depth, preserve tissue integrity, protect tissue from surgical trauma, or function as a tamponade during retinal reattach- ment.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 886.3.

##### § 886.4280   Intraocular  pressure  meas- uring device.

- Identification. An intraocular pres- sure measuring device is a manual or AC-powered device intended to meas- ure intraocular pressure. Also included are any devices found by FDA to be substantially equivalent to such de- vices. Accessories for the device may include calibrators or recorders. The device is intended for use in the diag- nosis of glaucoma.

- Classification. Class III.

- Date PMA or notice of completion of PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 886.3.

##### § 886.4300   Intraocular lens guide.

- Identification. An intraocular lens guide is a device intended to be in- serted into the eye during surgery to direct the insertion of an intraocular lens and be removed after insertion is completed.

- Classification. Class I (general con- trols). Except when used as folders or injectors for  soft  or  foldable  intra- ocular  lenses,  the  device  is   exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter subject to § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 65

FR 2321, 2000]

##### § 886.4335   Operating headlamp.

- Identification. An operating headlamp is an AC-powered or battery- powered device intended to be worn on the   user’s   head   to   provide   a   light

#### § 886.4350

source to aid visualization during sur- gical, diagnostic, or therapeutic proce- dures.

- Classification. Class I for the bat- tery-powered device. Class II for the AC-powered device. The battery-pow- ered device is exempt from the pre- market notification procedures in sub- part E of part 807 of this chapter, sub- ject to the limitations in § 886.9.

[55 FR 48443, Nov. 20, 1990, as amended at 66

FR 38813, July 25, 2001]

##### § 886.4350   Manual ophthalmic surgical instrument.

- Identification. A manual oph- thalmic surgical instrument is a non- powered, handheld device intended to aid or perform ophthalmic surgical procedures. This generic type of device includes the manual corneal burr, oph- thalmic caliper, ophthalmic cannula, eyelid clamp, ophthalmic muscle clamp, iris retractor clip, orbital com- pressor,	ophthalmic	curette, cystotome, orbital depressor, lach- rymal dilator, erisophake, expressor, ophthalmic forcep, ophthalmic hook, sphere introducer, ophthalmic knife, ophthalmic suturing needle, lachrymal probe, trabeculotomy probe, cornea- sclera punch, ophthalmic retractor, ophthalmic ring (Flieringa), lachrymal sac rongeur, ophthalmic scissors, enucleating snare, ophthalmic spatula, ophthalmic specula, ophthalmic spoon, ophthalmic spud, trabeculotome or ophthalmic manual trephine.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35606,  Sept.  14,  1988;  59  FR  63013,  Dec.  7,

1994; 60 FR 15872, Mar. 28, 1995; 66 FR 38813,

July 25, 2001]

##### § 886.4360   Ocular    surgery    irrigation device.

- Identification. An ocular surgery irrigation device is a device intended to be suspended over the ocular area during ophthalmic surgery to deliver continuous, controlled irrigation to the surgical field.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35606,  Sept.  14,  1988;  59  FR  63013,  Dec.  7,

1994; 66 FR 38813, July 25, 2001]

##### § 886.4370   Keratome.

- Identification. A keratome is an AC-powered or battery-powered device intended to shave tissue from sections of the cornea for a lamellar (partial thickness) transplant.

- Classification. Class I.

[55 FR 48443, Nov. 20, 1990]

##### § 886.4390   Ophthalmic laser.

- Identification. An ophthalmic laser is an AC-powered device intended to coagulate or cut tissue of the eye, orbit, or surrounding skin by a laser beam.

- Classification. Class II.

##### § 886.4392 Nd:YAG laser for posterior capsulotomy and peripheral iridotomy.

- Identification. The Nd:YAG laser for posterior capsulotomy and periph- eral iridotomy consists of a mode- locked or Q-switched solid state Nd:YAG laser intended for  disruption of the posterior capsule or the iris via optical breakdown. The Nd:YAG laser generates short pulse, low energy, high power, coherent optical radiation. When the laser output is combined with focusing optics, the high irradi- ance at the target causes tissue disrup- tion via optical breakdown. A visible aiming system is utilized to target the invisible Nd:YAG laser radiation on or in close proximity to the target tissue.

- Classification. Class II (special controls). Design Parameters: Device must emit a laser beam with the fol- lowing parameters: wavelength = 1064 nanometers; spot size = 50 to 100 micros; pulse width = 3 to 30 nano- seconds; output energy per  pulse  =  0.5 to 15 millijoules  (mJ);  repetition  rate  = 1 to 10 pulses; and total energy = 20 to 120 mJ.

[65 FR 6894, Feb. 11, 2000]

##### § 886.4400   Electronic metal locator.

- Identification. An electronic metal locator is an AC-powered device with

#### Food and Drug Administration, HHS	§ 886.4770

probes  intended  to  locate  metallic  for- eign bodies in the eye or eye socket.

- Classification. Class II.

##### § 886.4440   AC-powered magnet.

- Identification. An AC-powered magnet is an AC-powered device that generates a magnetic field intended to find and remove metallic foreign bodies from eye tissue.

- Classification. Class II.

##### § 886.4445   Permanent magnet.

- Identification. A permanent mag- net is a nonelectric device that gen- erates a magnetic field intended to find and remove metallic foreign bodies from eye tissue.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35606, Sept. 14, 1988; 66 FR 38813, July 25,

2001]

##### § 886.4570   Ophthalmic  surgical  mark- er.

- Identification. An ophthalmic sur- gical marker is a device intended to mark by use of ink, dye, or indentation the location of ocular or scleral sur- gical manipulation.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35606,  Sept.  14,  1988;  59  FR  63013,  Dec.  7,

1994; 66 FR 38813, July 25, 2001]

##### § 886.4610   Ocular pressure applicator.

- Identification. An ocular pressure applicator is a manual device that con- sists of a sphygmomanometer-type squeeze bulb, a dial indicator, a band, and bellows, intended to apply pressure

on the eye in preparation for oph- thalmic  surgery.

- Classification. Class II.

##### § 886.4670   Phacofragmentation system.

- Identification.	A phacofragmentation system is  an AC- powered device with a fragmenting needle intended for use in cataract sur- gery to disrupt a cataract with ultrasound and extract the cataract.

- Classification. Class II.

##### § 886.4690   Ophthalmic photocoagulator.

- Identification. An ophthalmic photocoagulator is an AC-powered de- vice intended to use the energy from an extended noncoherent light source to occlude blood vessels of the retina, choroid, or iris.

- Classification. Class II.

##### § 886.4750   Ophthalmic eye shield.

- Identification. An ophthalmic eye shield is a device that consists of a plastic or aluminum eye covering in- tended to protect the eye or retain dressing materials in place.

- Classification. Class I (general con- trols). When made only of plastic or aluminum,  the  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 886.9. When made only of plastic or aluminum, the devices are exempt from the current good manu- facturing practice requirements of the quality system regulation in part 820 of this    chapter,    with    the    exception    of

§ 820.180 of this chapter, with respect to general requirements   concerning records, and § 820.198 of this chapter, with  respect  to  complaint  files.

[52 FR 33355, Sept. 2, 1987, as amended at 59

FR  63014,  Dec.  7,  1994;  65  FR  2321,  Jan.  14,

2000]

##### § 886.4770   Ophthalmic  operating  spec- tacles (loupes).

- Identification. Ophthalmic oper- ating spectacles (loupes) are devices that consist of convex lenses or lens systems intended to be worn by a sur- geon to magnify the surgical site dur- ing ophthalmic surgery.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### § 886.4790

subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35606, Sept. 14, 1988; 66 FR 38813, July 25,

2001]

##### § 886.4790   Ophthalmic sponge.

  - Identification. An ophthalmic sponge is a device that is an absorbant sponge, pad, or spear made of folded gauze, cotton, cellulose, or other mate- rial intended to absorb fluids from the operative field in ophthalmic surgery.

  - Classification. Class II.

##### § 886.4855   Ophthalmic	instrument table.

- Identification. An ophthalmic in- strument table is an AC-powered or manual device on which ophthalmic in- struments are intended to be placed.

- Classification. Class I (general con- trols). The  AC-powered device  and the manual device are exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The manual device is also exempt from the current good manufacturing practice requirements of the quality system regulation in part 820 of this chapter, with the exception of § 820.180, with re- spect to general requirements con- cerning records, and § 820.198, with re- spect to complaint files.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38814,  July 25,

2001]

## Subpart F—Therapeutic Devices

##### § 886.5100   Ophthalmic   beta   radiation source.

- Identification. An ophthalmic beta radiation source is a device intended to apply superficial radiation to benign and malignant ocular growths.

- Classification. Class II.

21 CFR Ch. I (4–1–17 Edition)

##### § 886.5120   Low-power binocular loupe.

- Identification. A low-power bin- ocular loupe is a device that consists of two eyepieces, each with a lens or lens system, intended for medical purposes to magnify the appearance of objects.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5200   Eyelid	thermal	pulsation system.

- Identification. An eyelid thermal pulsation system is an electrically- powered device intended for use in the application of localized heat and pres- sure therapy to the eyelids. The device is used in adult patients with chronic cystic conditions of the eyelids, includ- ing meibomian gland dysfunction (MGD), also known as evaporative dry eye or lipid deficiency dry eye. The sys- tem consists of a component that is in- serted around the eyelids and a compo- nent to control the application of heat and pressure to the eyelids.

- Classification. Class II (special controls). The special controls for this device are:

- Appropriate analysis/testing should validate electromagnetic com- patibility (EMC) and safety of exposure to non-ionizing radiation;

- Design, description, and perform- ance data should validate safeguards related to the temperature and pres- sure aspects of the device, including during fault conditions;

- Performance data should dem- onstrate the sterility of patient-con- tacting components and the  shelf-life of these components;

- The device should be dem- onstrated to be biocompatible; and

- Performance data should dem- onstrate      that      any      technological

#### Food and Drug Administration, HHS	§ 886.5700

changes  do  not  adversely  effect  safety and effectiveness.

[76 FR 51878, Aug. 19, 2011]

##### § 886.5420   Contact	lens	inserter/re- mover.

  - Identification. A contact lens inserter/remover is a handheld device intended to insert or remove contact lenses by surface adhesion or suction.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5540   Low-vision magnifier.

- Identification. A low-vision mag- nifier is a device that consists of a magnifying lens intended for use by a patient who has impaired vision. The device may be held in the hand or at- tached to spectacles.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5600   Ptosis crutch.

- Identification. A ptosis crutch is a device intended to be mounted on the spectacles of a patient who has ptosis (drooping of the upper eyelid as a re- sult of faulty development or paralysis) to hold the upper eyelid open.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation

in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5700   Eyelid weight.

- Identification. An eyelid weight is a prescription device made of gold, tan- talum, platinum, iridium, or surgical grade stainless steel that is rectan- gular in shape and contoured to the shape of the eye. The device is intended for the gravity assisted treatment of lagophthalmos (incomplete eyelid clo- sure).

  - The external eyelid weight is ad- hered to the outer skin of the upper eyelid.

  - The implantable eyelid weight is implanted into the upper eyelid.

- Classification. (1) Class II (special controls) for  the  external  eyelid weight. The external eyelid weight is exempt from the premarket  notifica- tion procedures in subpart E of part 807 of this chapter subject to the limita- tions in § 886.9. The special controls for the external eyelid weight are:

- Testing demonstrating the bio- compatibility of the device; and

- Labeling must include the fol- lowing information:

  - Specific instructions regarding the proper placement, sizing, and re- moval of the device; and

  - A warning stating that the pa- tient should be instructed to remove the device prior to entering a magnetic resonance environment.

- Class II (special controls) for the implantable eyelid weight. The special controls for the implantable eyelid weight are:

  - Testing demonstrating the bio- compatibility of the device;

  - Testing demonstrating the ste- rility and shelf life of the device;

  - Nonclinical testing evaluating the compatibility of the device in a magnetic resonance environment.

  - Patient labeling to convey infor- mation regarding the safety and com- patibility of the device in a magnetic resonance environment, the conditions under which a patient with the device

#### § 886.5800

can be safely scanned, and a mecha- nism for a healthcare provider to ob- tain detailed information about mag- netic resonance safety and compat- ibility if needed.

[79 FR 22015, Apr. 21, 2014]

##### § 886.5800   Ophthalmic bar reader.

    - Identification. An ophthalmic bar reader is a device that consists of a magnifying lens intended for use by a patient who has impaired vision. The device is placed directly onto reading material to magnify print.

    - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5810   Ophthalmic prism reader.

- Identification. An ophthalmic prism reader is a device intended for use by a patient who is in a supine po- sition to change the angle of print to aid reading.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5820   Closed-circuit	television reading system.

- Identification. A closed-circuit tel- evision reading system is a device that

#### 21 CFR Ch. I (4–1–17 Edition)

consists of a lens, video camera, and video monitor  that  is intended  for  use by a patient who has subnormal vision to  magnify  reading  material.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38814,  July 25,

2001]

##### § 886.5838   Nasolacrimal	compression device.

- Identification. A nasolacrimal compression device is a prescription device that is fitted to apply mechan- ical pressure to the nasal aspect of the orbital rim to reduce outflow through the nasolacrimal ducts.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[81 FR 37500, June 10, 2016]

##### § 886.5840   Magnifying spectacles.

- Identification. Magnifying spec- tacles are devices that consist of spec- tacle frames with convex lenses in- tended to be worn by a patient who has impaired vision to enlarge images.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 866.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35607,  Sept.  14,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38814, July 25, 2001]

##### § 886.5842   Spectacle frame.

- Identification. A spectacle frame is a device made of metal or plastic in- tended to  hold  prescription  spectacle lenses worn by a patient to correct re- fractive errors.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38814,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 886.5905

##### § 886.5844   Prescription spectacle lens.

  - Identification. A prescription spec- tacle lens is a glass or plastic device that is a lens intended to be worn by a patient in a spectacle frame to provide refractive corrections in  accordance with a prescription for the patient. The device may be modified to protect the eyes from bright sunlight (i.e., pre- scription sunglasses). Prescription sun- glass lenses may be reflective, tinted, polarizing, or photosensitized.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR  35607,  Sept.  14,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38814, July 25, 2001]

##### § 886.5850   Sunglasses	(nonprescrip- tion).

- Identification. Sunglasses (non- prescription) are devices that consist of spectacle frames or clips with absorb- ing, reflective, tinted, polarizing, or photosensitized lenses intended to be worn by a person to protect the eyes from bright sunlight but not to provide refractive corrections. This device is usually  available  over-the-counter.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 886.9.

[52 FR 33355, Sept. 2, 1987, as amended at 65

FR 2321, 2000]

##### § 886.5870   Low-vision telescope.

- Identification. A low-vision tele- scope is a device that consists of an ar- rangement of lenses or mirrors in- tended for use by a patient who has im- paired vision to increase the apparent size of objects. This generic type of de- vice includes handheld or spectacle telescopes.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex-

ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5900   Electronic vision aid.

- Identification. An electronic vision aid is an AC-powered or battery-pow- ered device that consists of an elec- tronic sensor/transducer intended for use by a patient who has impaired vi- sion or blindness to translate visual images of objects into tactile or audi- tory signals.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38814,  July 25,

2001]

##### § 886.5905   Oral electronic vision aid.

- Identification. An oral electronic vision aid  is a  battery-powered  pre- scription device that contains an elec- trode stimulation array to generate electrotactile   stimulation    patterns that are derived from digital object im- ages captured by a camera. It is in- tended to aid profoundly blind patients in orientation, mobility, and object recognition as an adjunctive device to other assistive methods such as a white cane or a guide dog.

- Classification. Class II (special controls). The special controls for this device are:

- Clinical performance testing must demonstrate an acceptable adverse event profile, including adverse events involving the mouth, tongue, and gums and demonstrate the effect of the stim- ulation to provide clinically meaning- ful outcomes. The clinical performance testing must also investigate the an- ticipated conditions of use, including potential use error, intended environ- ment of use, and duration of use.

- Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use, including simulated moisture  ingress,  device  durability, and  battery  reliability.

#### § 886.5910

- Software verification, validation, and hazard analysis must be performed.

- Analysis/testing must validate electromagnetic  compatibility.

- Analysis/testing must validate electrical  safety.

- Analysis/testing must assess and validate  wireless  coexistence  concerns.

- Any elements  of  the  device  that contact the patient must be dem- onstrated to be biocompatible.

- Training must include elements to ensure that the  healthcare  provider and user can identify the safe environ- ments for device use, use all safety fea- tures of the device, and operate the de- vice in  the  intended  environment  of use.

- Labeling for the trainer and user must include a summary of the clinical testing including adverse events en- countered under use conditions, sum- mary of study outcomes and endpoints, and information pertinent to use of the device including the conditions under which the device was studied (e.g., level of supervision or assistance, and envi- ronment of use).

[80 FR 57092, Sept. 22, 2015]

##### § 886.5910   Image intensification vision aid.

- Identification. An image inten- sification vision aid is a battery-pow- ered device intended for use by a pa- tient who has limited dark adaptation or impaired vision to amplify ambient light.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 886.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[52 FR 33355, Sept. 2, 1987, as amended at 53

FR 35607, Sept. 14, 1988; 66 FR 38814, July 25,

2001]

##### § 886.5915   Optical vision aid.

- Identification. An optical vision aid is a device that consists of a magni-

#### 21 CFR Ch. I (4–1–17 Edition)

fying lens with an accompanying AC- powered or battery-powered light source intended for use by a patient who has impaired vision to increase the apparent size of object detail.

- Classification. Class I (general con- trols). The  AC-powered device  and the battery-powered device are exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 886.9. The battery-powered device is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, with respect to general requirements   concerning   records,   and

§ 820.198,  with  respect  to  complaint files.

[55 FR 48443, Nov. 20, 1990, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38815,  July 25,

2001]

##### § 886.5916   Rigid gas permeable contact lens.

- Identification. A rigid  gas per- meable contact lens is a device in- tended to be worn directly against the cornea of the eye to correct vision con- ditions. The device is made of various materials, such as cellulose acetate bu- tyrate, polyacrylate-silicone, or sili- cone elastomers, whose main polymer molecules generally do not absorb or attract water.

- Classification. (1) Class II if the de- vice is intended for daily wear only.

- Class III if the device is intended for extended wear.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before a device described in paragraph (b)(2) of this section may be commercially  distributed.  See  § 886.3.

[52 FR 33355, Sept. 2, 1987, as amended at 59

FR 10284, Mar. 4, 1994]

##### § 886.5918   Rigid gas permeable contact lens care products.

- Identification. A rigid  gas per- meable contact lens care product is a device intended for use in the cleaning, conditioning, rinsing, lubricating/re- wetting, or storing of a rigid gas per- meable contact lens. This includes all solutions  and   tablets   used   together

#### Food and Drug Administration, HHS	Pt. 888

with    rigid    gas    permeable    contact lenses.

- Classification.   Class    II    (Special Controls)  Guidance  Document:  ‘‘Guid- ance  for  Industry  Premarket  Notifica- tion  (510(k))  Guidance  Document  for Contact  Lens  Care  Products.’’

[62 FR 30987, June 6, 1997]

##### § 886.5925   Soft   (hydrophilic)   contact lens.

- Identification. A soft (hydrophilic) contact lens is a device intended to be worn directly  against  the  cornea  and adjacent limbal and scleral areas of the

##### § 886.5933   [Reserved]

PART 888—ORTHOPEDIC DEVICES

Subpart A—General Provisions

Sec.

888.1    Scope.

888.3   Effective dates of requirement for pre- market  approval.

  - Resurfacing technique.

  - Degree of constraint.

888.9 Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

| eye to correct vision conditions or act | 888.1100 | Arthroscope. |
| --- | --- | --- |
| as a therapeutic bandage. The device is | 888.1240 | AC-powered dynamometer. |
| made of various polymer materials the | 888.1250 | Nonpowered dynamometer. |
| main polymer molecules of which ab- | 888.1500 | Goniometer. |

sorb  or  attract  a  certain  volume  (per- centage) of water.

- Classification. (1) Class II if the de- vice is intended for daily wear only.

(2) Class III if the device is intended for extended wear.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before a device described in paragraph (b)(2) of this section may be commercially  distributed.  See  § 886.3.

[52 FR 33355, Sept. 2, 1987, as amended at 59

FR 10284, Mar. 4, 1994]

##### § 886.5928   Soft   (hydrophilic)   contact lens care products.

- Identification. A soft (hydrophilic) contact lens care product is a device intended for use in the cleaning, rins- ing, disinfecting, lubricating/rewetting, or storing of a soft (hydrophilic) con- tact lens.  This  includes  all  solutions and tablets used together with soft (hy- drophilic) contact lenses and heat dis- infecting units intended to disinfect a soft (hydrophilic) contact lens by means of heat.

- Classification.   Class    II    (Special Controls)  Guidance  Document:  ‘‘Guid- ance  for  Industry  Premarket  Notifica- tion  (510(k))  Guidance  Document  for Contact  Lens  Care  Products.’’

[62 FR 30988, June 6, 1997]

888.1520    Nonpowered goniometer.

##### Subpart C [Reserved]

Subpart D—Prosthetic Devices

888.3000   Bone cap.

888.3010    Bone fixation cerclage. 888.3015   Bone heterograft.

888.3020    Intramedullary fixation rod. 888.3025   Passive tendon prosthesis.

888.3027   Polymethylmethacrylate    (PMMA) bone cement.

888.3030    Single/multiple component metallic bone  fixation  appliances  and  accessories. 888.3040    Smooth  or  threaded  metallic  bone

fixation fastener.

888.3045    Resorbable  calcium  salt  bone  void filler   device.

888.3050     Spinal  interlaminal  fixation  ortho- sis.

888.3060    Spinal  intervertebral  body  fixation orthosis.

888.3070   Thoracolumbosacral   pedicle   screw system.

888.3080    Intervertebral body fusion device. 888.3100    Ankle  joint  metal/composite  semi-

constrained cemented prosthesis.

888.3110    Ankle   joint   metal/polymer   semi- constrained   cemented   prosthesis.

888.3120    Ankle joint metal/polymer non-con- strained  cemented  prosthesis.

888.3150   Elbow    joint    metal/polymer    con- strained  cemented  prosthesis.

888.3160   Elbow   joint   metal/polymer   semi- constrained   cemented   prosthesis.

888.3170   Elbow    joint    radial    (hemi-elbow) polymer   prosthesis.

888.3180   Elbow  joint  humeral  (hemi-elbow) metallic   uncemented   prosthesis.

888.3200 Finger	joint	metal/metal	con- strained  uncemented  prosthesis.

888.3210 Finger	joint	metal/metal	con- strained  cemented  prosthesis.

#### Pt. 888

888.3220 Finger joint metal/polymer con- strained  cemented  prosthesis.

888.3230 Finger joint polymer constrained prosthesis.

888.3300 Hip joint metal constrained  ce- mented  or  uncemented  prosthesis.

888.3310 Hip joint  metal/polymer  con- strained cemented or uncemented pros- thesis.

888.3320 Hip joint metal/metal semi-con- strained, with a cemented acetabular component,   prosthesis.

888.3330 Hip joint metal/metal semi-con- strained, with an uncemented acetabular component,   prosthesis.

888.3340 Hip joint metal/composite semi- constrained   cemented   prosthesis.

888.3350 Hip joint metal/polymer semi-con- strained  cemented  prosthesis.

888.3353 Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented   prosthesis.

888.3358 Hip joint metal/polymer/metal semi-constrained	porous-coated uncemented   prosthesis.

888.3360 Hip joint femoral (hemi-hip) metal- lic  cemented  or  uncemented  prosthesis.

888.3370 Hip joint  (hemi-hip)  acetabular metal  cemented  prosthesis.

888.3380 Hip joint femoral (hemi-hip) trun- nion-bearing metal/polyacetal cemented prosthesis.

888.3390 Hip joint femoral (hemi-hip) metal/ polymer cemented or uncemented pros- thesis.

888.3400 Hip joint femoral (hemi-hip) metal- lic   resurfacing   prosthesis.

888.3410 Hip joint metal/polymer or ceramic/ polymer   semiconstrained   resurfacing

888.3480 Knee joint femorotibial metallic constrained   cemented   prosthesis.

888.3490 Knee joint femorotibial metal/com- posite non-constrained cemented pros- thesis.

888.3500    Knee joint femorotibial metal/com-

#### 21 CFR Ch. I (4–1–17 Edition)

888.3565 Knee joint patellofemorotibial metal/polymer	porous-coated uncemented   prosthesis.

888.3570 Knee joint femoral (hemi-knee) me- tallic   uncemented   prosthesis.

888.3580    Knee joint patellar (hemi-knee) me- tallic   resurfacing   uncemented   prosthesis. 888.3590    Knee  joint  tibial  (hemi-knee)  me- tallic   resurfacing   uncemented   prosthesis.

888.3640 Shoulder joint metal/metal or metal/polymer constrained cemented prosthesis.

888.3650 Shoulder joint metal/polymer non- constrained   cemented   prosthesis.

888.3660 Shoulder joint metal/polymer semi- constrained   cemented   prosthesis.

888.3670 Shoulder joint metal/polymer/metal nonconstrained or semi-constrained po- rous-coated  uncemented  prosthesis.

888.3680 Shoulder joint glenoid (hemi-shoul- der)  metallic  cemented  prosthesis.

888.3690 Shoulder joint humeral (hemi- shoulder) metallic uncemented  pros- thesis.

888.3720 Toe joint polymer constrained pros- thesis.

888.3730 Toe joint phalangeal (hemi-toe) polymer   prosthesis.

888.3750 Wrist joint carpal lunate polymer prosthesis.

888.3760 Wrist joint carpal scaphoid polymer prosthesis.

888.3770 Wrist joint carpal trapezium poly- mer  prosthesis.

888.3780 Wrist joint polymer constrained prosthesis.

888.3790 Wrist joint metal constrained ce- mented   prosthesis.

888.3800 Wrist joint metal/polymer semi- constrained   cemented   prosthesis.

888.3810 Wrist joint ulnar (hemi-wrist) poly- mer  prosthesis.

##### Subpart E—Surgical Devices

| posite semi-constrained cemented pros- | 888.4150 | Calipers for clinical use. |
| --- | --- | --- |
| thesis. | 888.4200 | Cement dispenser. |

888.3510    Knee joint femorotibial metal/poly-

mer constrained cemented prosthesis.

888.3520 Knee joint femorotibial metal/poly- mer non-constrained cemented pros- thesis.

888.3530 Knee joint femorotibial metal/poly- mer semi-constrained cemented  pros- thesis.

888.3535 Knee joint femorotibial (uni-com- partmental)  metal/polymer   porous-coat- ed  uncemented  prosthesis.

888.3540 Knee joint patellofemoral polymer/ metal semi-constrained cemented pros- thesis.

888.3550 Knee joint patellofemorotibial polymer/metal/metal constrained ce- mented   prosthesis.

888.3560 Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis.

888.4210    Cement mixer for clinical use. 888.4220   Cement monomer vapor evacuator. 888.4230    Cement ventilation tube.

888.4300   Depth gauge for clinical use.  888.4540   Orthopedic manual surgical instru-

ment.

888.4580    Sonic  surgical  instrument  and  ac- cessories/attachments.

888.4600   Protractor for clinical use. 888.4800   Template for clinical use.

888.5850   Nonpowered orthopedic traction ap- paratus  and  accessories.

888.5890    Noninvasive traction component. 888.5940   Cast component.

888.5960    Cast removal instrument.

888.5980    Manual   cast   application   and   re- moval    instrument.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

#### Food and Drug Administration, HHS	§ 888.3

SOURCE:  52  FR  33702,  Sept.  4,  1987,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 888 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 888.1   Scope.

- This part sets forth the classifica- tion of orthopedic devices intended for human use that are in commercial dis- tribution.

- The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provision of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

- To avoid duplicative listings, an orthopedic device that has two or more types of uses (e.g., used both as a diag- nostic device and as a surgical device) is listed in one subpart only.

- References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21 unless otherwise noted.

- Guidance documents referenced in this  part  are  available  on  the  Internet.

[52 FR 33702, Sept. 4, 1987, as amended at 68

FR 14137, Mar. 24, 2003; 78 FR 18233, Mar. 26,

2013]

##### § 888.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com-

pleted a product development protocol (PDP) for the device.

  - Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act, FDA must promulgate a  regula- tion under section 515(b) of the act re- quiring such approval, except as pro- vided in paragraphs (b) and (c) of this section. Such a regulation under sec- tion 515(b) of the act shall not be effec- tive during the grace period ending on the 90th  day  after  its promulgation or on the last day of the 30th full calendar month after the regulation that classi- fies the device into class III is effec- tive, whichever is later. See section 501(f)(2)(B) of the act. Accordingly, un- less an effective date of the require- ment for premarket approval is shown in the regulation for a device classified into class III in this part, the device may be commercially distributed with- out FDA’s issuance of an order approv- ing a PMA or declaring completed a PDP for the device. If  FDA  promul- gates a regulation under section 515(b) of the act requiring premarket ap- proval for a device, section 501(f)(1)(A) of the act applies to the device.

  - Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

§ 888.5

  - A device identified in a regulation in this part that is classified into class III and that is subject to the transi- tional provisions of section 520(1) of the act is automatically classified by stat- ute into class III and must have an ap- proval under section 515 of the act be- fore being commercially distributed. Accordingly, the regulation for such a class III transitional device states that as of the enactment date of the amend- ments, May 28, 1976, the device must have an approval under section 515 of the act before commercial distribution.

##### § 888.5   Resurfacing technique.

Because of resurfacing techniques, certain joint prostheses require far less bone resection than other devices in- tended to repair or replace the same joint. The amount  of  bone  resection may or may not affect the safety and effectiveness of the implantation of the prosthesis. When a resurfacing tech- nique is used, the name of the pros- thesis  includes  this  information.

##### § 888.6   Degree of constraint.

Certain joint prostheses provide more constraint of joint movement than oth- ers. FDA believes that the degree of constraint is an important factor af- fecting the safety and effectiveness of orthopedic prostheses. FDA is defining the following standard terms for cat- egorizing the degree of constraint.

- A ‘‘constrained’’ joint  prosthesis is used for joint replacement and pre- vents dislocation of the prosthesis in more than one anatomic plane and con- sists of either a single, flexible, across- the-joint component or more than one component linked together or affined.

- A ‘‘semi-constrained’’ joint pros- thesis is used for partial or total joint replacement and limits translation and rotation of the prosthesis in one or more planes via the geometry of its ar- ticulating surfaces. It has no across- the-joint  linkage.

- A ‘‘non-constrained’’ joint pros- thesis is used for partial or total joint replacement and restricts minimally prosthesis movement in one or more planes. Its components have no across- the-joint  linkage.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 888.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening,   diagnosis,   or   monitoring

#### Food and Drug Administration, HHS	§ 888.1520

life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2321, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 888.1100   Arthroscope.

- Identification. An arthroscope is an electrically powered endoscope in- tended to make visible the interior of a joint. The arthroscope and accessories also is intended to perform surgery within a joint.

- Classification. (1)  Class  II  (per- formance standards).

- Class I for the following manual arthroscopic instruments: cannulas, currettes, drill guides, forceps, gouges, graspers,	knives,	obturators, osteotomes, probes, punches, rasps, re- tractors, rongeurs, suture passers, su- ture knotpushers, suture punches, switching rods, and trocars. The de- vices subject to this paragraph (b)(2) are exempt from the premarket notifi- cation procedures in subpart E of part 807 of this chapter, subject to the limi- tations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38815,  July 25,

2001]

##### § 888.1240   AC-powered dynamometer.

  - Identification. An AC-powered dy- namometer is an AC-powered device in- tended for medical purposes to assess neuromuscular function or degree of neuromuscular blockage  by  measuring,

with a force transducer (a device that translates force into electrical im- pulses), the grip-strength of a patient’s hand.

  - Classification. Class II.

##### § 888.1250   Nonpowered dynamometer.

- Identification. A nonpowered dyna- mometer is a mechanical device in- tended for medical purposes to measure the pinch and grip muscle strength of a patient’s hand.

- Classification. Class I.  The  device is exempt from the premarket notifica- tion procedures in subpart E of part 807.

##### § 888.1500   Goniometer.

- Identification. A goniometer is an AC-powered or battery powered device intended to evaluate joint function by measuring and recording ranges of mo- tion, acceleration, or forces exerted by a joint.

- Classification. (1) Class I (general controls) for  a  goniometer  that  does not use electrode  lead  wires  and  pa- tient  cables.  This  device   is   exempt from the premarket notification proce- dures of subpart E of part 807 of this chapter subject to § 888.9.

- Class II (special controls) for a go- niometer that uses electrode lead wires and patient cables. The special controls consist of:

  - The performance standard under part 898 of this chapter, and

  - The guidance entitled ‘‘Guidance on the Performance Standard for Elec- trode Lead Wires and Patient Cables.’’ This device is exempt from the pre- market notification procedures of sub- part E of part 807 of this chapter sub- ject to § 888.9.

[65 FR 19319, Apr. 11, 2000]

##### § 888.1520   Nonpowered goniometer.

- Identification. A nonpowered goni- ometer is a mechanical device intended for medical purposes to measure the range of motion of joints.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

§ 888.3000

Subpart C [Reserved]

Subpart D—Prosthetic Devices

##### § 888.3000   Bone cap.

  - Identification. A bone cap is a mushroom-shaped device intended to be implanted made of either silicone elastomer or ultra-high molecular weight polyethylene. It is used to cover the severed end of a long bone, such as the humerus or tibia, to control bone overgrowth in juvenile amputees.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 1124,  Jan. 16,  1996;  66  FR 38815,  July 25,

2001]

##### § 888.3010   Bone fixation cerclage.

- Identification. A bone fixation cerclage is a device intended to be im- planted that is made of alloys, such as cobalt-chromium-molybdenum,	and that consists of a metallic ribbon  or flat sheet or a wire. The device is wrapped around the shaft of a long bone, anchored to the  bone  with  wire or screws, and used in the fixation of fractures.

- Classification. Class II.

##### § 888.3015   Bone heterograft.

- Identification. Bone heterograft is a device intended to be implanted that is made from mature (adult) bovine bones and used to replace human bone following surgery in the cervical region of the spinal column.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. As of May 28, 1976, an approval under section 515 of the act is required before this device may be commercially  distributed.  See  § 888.3.

##### § 888.3020   Intramedullary fixation rod.

- Identification. An intramedullary fixation rod is a device intended to be implanted that consists of  a  rod  made of alloys such  as cobalt-chromium-mo- lybdenum and stainless steel. It is in- serted into the medullary (bone mar- row) canal of long bones for the fixa- tion of fractures.

#### 21 CFR Ch. I (4–1–17 Edition)

- Classification. Class II.

##### § 888.3025   Passive tendon prosthesis.

- Identification. A passive tendon prosthesis is a device intended to be implanted made of silicon elastomer or a polyester reinforced medical grade silicone elastomer intended for use in the surgical reconstruction of a flexor tendon of the hand. The device is im- planted for a period of 2 to 6 months to aid growth of a new tendon sheath. The device is not intended as a permanent implant nor to function as a replace- ment for the ligament or tendon nor to function as a scaffold for soft tissue ingrowth.

- Classification. Class II.

##### § 888.3027   Polymethylmethacrylate (PMMA) bone cement.

- Identification.

Polymethylmethacrylate (PMMA) bone cement is a device intended to be im- planted	that	is	made	from methylmethacrylate, polymethylmethacrylate,	esters		of methacrylic  acid,  or  copolymers  con- taining   polymethylmethacrylate   and polystyrene. The device is intended for use in arthroplastic procedures of the hip, knee, and other joints for the fixa- tion of polymer or metallic prosthetic implants to living bone.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance	Document: Polymethylmethacrylate		(PMMA) Bone Cement.’’

[67 FR 46855, July 17, 2002]

##### § 888.3030 Single/multiple component metallic bone fixation appliances and accessories.

- Identification. Single/multiple component metallic bone fixation ap- pliances and accessories are devices in- tended to be implanted consisting of one or more metallic components and their metallic fasteners. The devices contain a plate, a nail/plate combina- tion, or a blade/plate combination that are made of alloys, such as cobalt-chro- mium-molybdenum,   stainless   steel, and titanium, that are intended to be held in position with fasteners, such as screws  and  nails,  or  bolts,  nuts,  and

#### Food and Drug Administration, HHS	§ 888.3070

washers. These devices are used for fix- ation of fractures of the proximal or distal end of long bones, such as intracapsular, intertrochanteric, inter- cervical, supracondylar, or condylar fractures of the femur; for fusion of a joint; or for surgical procedures that involve cutting a bone. The devices may be implanted or attached through the skin so that a pulling force (trac- tion) may be applied to the skeletal system.

- Classification. Class II.

##### § 888.3040   Smooth or threaded metallic bone fixation fastener.

- Identification. A smooth or thread- ed metallic bone fixation fastener is a device intended to be implanted that consists of a stiff wire segment or rod made of alloys, such as cobalt-chro- mium-molybdenum and stainless steel, and that may be smooth on the out- side, fully or  partially  threaded, straight or U-shaped; and may be ei- ther blunt pointed, sharp pointed, or have a formed, slotted head on the end. It may be used for fixation of bone fractures, for  bone  reconstructions,  as a guide pin for insertion of other im- plants, or it may be implanted through the skin so that a pulling force (trac- tion) may be applied to the skeletal system.

- Classification. Class II.

##### § 888.3045   Resorbable	calcium	salt bone void filler device.

- Identification. A resorbable cal- cium salt bone void filler device is a resorbable implant intended to fill bony voids or gaps of the extremities, spine, and pelvis that are caused by trauma or surgery and are not intrinsic to the stability of the bony structure.

- Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance: Resorbable Calcium Salt Bone Void Filler Device; Guidance for Industry and FDA.’’ See § 888.1(e) of this chapter for the availability of this guidance.

[68 FR 32636, June 2, 2003]

##### § 888.3050   Spinal interlaminal fixation orthosis.

- Identification. A spinal interlam- inal fixation orthosis is a device in- tended to  be implanted  made of  an alloy, such as stainless steel, that con- sists of various hooks and a posteriorly placed compression or distraction rod. The device is implanted, usually across three adjacent vertebrae, to straighten and immobilize the spine to allow bone grafts to unite and fuse the vertebrae together. The device is  used  primarily in the treatment of scoliosis (a lateral curvature of the spine), but it also may be used in the treatment of fracture or dislocation of the spine, grades 3 and 4 of spondylolisthesis  (a  dislocation  of the spinal column), and lower back syndrome.

- Classification. Class II.

##### § 888.3060   Spinal   intervertebral   body fixation orthosis.

- Identification.        A	spinal intervertebral body fixation orthosis is a device intended to be implanted made of titanium. It consists of various vertebral plates that are punched into each of a series of vertebral bodies. An eye-type screw is inserted in a hole in the center of each of the plates. A braided cable is threaded through each eye-type screw. The cable is tightened with a tension device and it is fastened or crimped at each eye-type screw. The device is used to apply force to a series of vertebrae to correct ‘‘sway back,’’ scoliosis (lateral curvature of the spine), or other conditions.

- Classification. Class II.

##### § 888.3070   Thoracolumbosacral pedicle screw system.

- Identification. (1) Rigid pedicle screw systems are comprised of mul- tiple components, made from a variety of materials that allow the surgeon to build an implant system to fit the pa- tient’s anatomical and physiological requirements. Such a spinal implant assembly consists of a combination of screws, longitudinal members (e.g., plates, rods including dual diameter rods, plate/rod combinations), trans- verse or cross connectors, and inter- connection mechanisms (e.g., rod-to- rod connectors, offset connectors).

#### § 888.3070

(2) Semi-rigid systems are defined as systems that contain one or more  of the following features  (including  but not limited to): Non-uniform longitu- dinal elements, or features that allow more motion or flexibility compared to rigid  systems.

- Classification. (1) Class II (special controls), when intended to provide im- mobilization and stabilization of spinal segments in skeletally mature  patients as an adjunct to fusion in the treat- ment of the following acute and chron- ic instabilities or deformities of the thoracic, lumbar, and sacral spine: se- vere spondylolisthesis (grades 3 and 4) of the L5–S1 vertebra; degenerative spondylolisthesis with objective evi- dence of neurologic impairment; frac- ture; dislocation; scoliosis; kyphosis; spinal tumor; and failed previous fu- sion (pseudarthrosis). These pedicle screw spinal systems must comply with the following special controls:

- Compliance with material stand- ards;

- Compliance with mechanical test- ing standards;

- Compliance with biocompat- ibility standards; and

- Labeling that contains these two statements in addition to other appro- priate labeling information:

‘‘Warning: The safety and effectiveness of pedicle screw spinal systems have been es- tablished only for spinal conditions with sig- nificant mechanical instability or deformity requiring  fusion  with   instrumentation. These conditions are significant mechanical instability or deformity of the thoracic, lum- bar, and sacral spine secondary to severe spondylolisthesis (grades 3 and 4) of the L5– S1 vertebra, degenerative spondylolisthesis with objective evidence of neurologic impair- ment, fracture, dislocation, scoliosis, kypho- sis, spinal tumor, and failed previous fusion (pseudarthrosis). The safety and effective- ness of these devices for any other conditions are unknown.’’

‘‘Precaution: The implantation of pedicle screw spinal systems should be performed only by experienced spinal surgeons with specific training in the use of this pedicle screw spinal system because this is a tech- nically demanding procedure presenting a risk of serious injury to the patient.’’

- Class II (special controls), when a rigid pedicle screw system is intended to provide immobilization and sta- bilization of spinal segments in the thoracic,  lumbar,  and  sacral  spine  as

#### 21 CFR Ch. I (4–1–17 Edition)

an adjunct to fusion  in  the  treatment of degenerative disc disease and spondylolisthesis other than either se- vere spondylolisthesis (grades 3 and 4) at	L5–S1	or	degenerative spondylolisthesis with objective evi- dence of neurologic impairment. These pedicle screw systems must comply with the following special controls:

- The design characteristics of the device, including engineering sche- matics, must ensure that the geometry and material composition are con- sistent with the intended use.

- Non-clinical performance testing must demonstrate the mechanical function and durability of the implant.

- Device components must be dem- onstrated to be biocompatible.

- Validation testing must dem- onstrate the cleanliness and sterility of, or the ability to clean and sterilize, the device components and device-spe- cific instruments.

- Labeling must include the fol- lowing:

  - A clear description of the techno- logical features of the device including identification of device materials and the principles of device operation;

  - Intended use and indications for use, including levels of fixation;

  - Identification of magnetic reso- nance (MR) compatibility status;

  - Cleaning and sterilization in- structions for devices and instruments that are provided non-sterile to the end user; and

  - Detailed instructions of each sur- gical step, including device removal.

- Class II (special controls), when a semi-rigid system is intended to pro- vide immobilization and stabilization of spinal segments in the thoracic, lumbar, and sacral spine as an adjunct to fusion for any indication. In addi- tion to complying with the special con- trols in paragraphs (b)(2)(i) through (v) of this section, these pedicle screw sys- tems must comply with the following special controls:

  - Demonstration that clinical per- formance characteristics of the device support the intended use of the prod- uct, including assessment of fusion compared to a clinically acceptable fu- sion rate.

#### Food and Drug Administration, HHS	§ 888.3120

  - Semi-rigid systems marketed prior to the effective date of this re- classification must submit an amend- ment to their previously cleared pre- market notification (510(k)) dem- onstrating compliance with the special controls in paragraphs (b)(2)(i) through

- and paragraph (b)(3)(i) of this sec- tion.

[66 FR 28053, May 22, 2001, as amended at 81

FR 96373, Dec. 30, 2016]

##### § 888.3080 Intervertebral body fusion device.

  - Identification. An intervertebral body fusion device is an implanted sin- gle or multiple component spinal de- vice made from a variety of materials, including titanium and polymers. The device is inserted into the intervertebral body space of the cer- vical or lumbosacral spine, and is in- tended for intervertebral body fusion.

  - Classification. (1) Class II (special controls) for intervertebral body fusion devices that contain bone grafting ma- terial. The special control is the FDA guidance document entitled ‘‘Class II Special Controls Guidance Document: Intervertebral  Body  Fusion  Device.’’ See § 888.1(e) for the availability of this guidance document.

(2) Class III (premarket approval) for intervertebral body fusion devices that include any  therapeutic  biologic  (e.g., bone	morphogenic	protein). Intervertebral body fusion devices that contain any therapeutic biologic re- quire premarket approval.

  - Date premarket approval application (PMA) or notice of product development protocol (PDP) is required. Devices de- scribed in paragraph (b)(2) of this sec- tion shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[72 FR 32172, June 12, 2007]

##### § 888.3100 Ankle joint metal/composite semi-constrained cemented pros- thesis.

- Identification. An ankle joint metal/composite semi-constrained ce- mented prosthesis is a device intended to be implanted to replace an ankle joint. The device limits translation and rotation: in one or more planes via the geometry   of   its   articulating   surfaces.

It has no linkage across-the-joint. This generic type of device includes pros- theses that consist of a talar  resur- facing component made of alloys, such as cobalt-chromium-molybdenum,  and a tibial resurfacing component fab- ricated from ultra-high molecular weight polyethylene with carbon fibers composite, and is limited to those pros- theses intended for use with bone ce- ment (§ 888.3027).

- Classification. Class II.

##### § 888.3110 Ankle joint metal/polymer semi-constrained cemented pros- thesis.

- Identification. An ankle joint metal/polymer semi-constrained ce- mented prosthesis is a device intended to be implanted to replace an ankle joint. The device limits translation and rotation in one or more planes via the geometry of its  articulating  surfaces and has no linkage across-the-joint. This generic type of device includes prostheses that have a talar resur- facing component made of alloys, such as cobalt-chromium-molybdenum,  and a tibial resurfacing component made of ultra-high molecular weight poly- ethylene and is limited to those pros- theses intended for use with bone ce- ment (§ 888.3027).

- Classification. Class II.

##### § 888.3120 Ankle joint metal/polymer non-constrained cemented pros- thesis.

- Identification. An ankle joint metal/polymer non-constrained ce- mented prosthesis is a device intended to be implanted to replace an ankle joint. The device limits minimally (less than normal anatomic constraints) translation in one or more planes. It has no linkage across-the-joint. This generic type of device includes pros- theses that have a tibial component made of alloys, such as cobalt-chro- mium-molybdenum, and a talar compo- nent made of ultra-high molecular weight polyethylene, and is limited to those prostheses intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be

#### § 888.3150

filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any ankle joint metal/polymer non- constrained cemented  prosthesis  that was in commercial distribution  before May 28,  1976, or  that has,  on or  before December 26, 1996, been found to be substantially  equivalent  to  an  ankle joint metal/polymer  non-constrained cemented prosthesis that was in com- mercial distribution before  May  28, 1976. Any other ankle joint metal/poly- mer non-constrained cemented pros- thesis shall have an approved PMA or a declared completed PDP  in  effect  be- fore being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3150   Elbow   joint   metal/polymer constrained cemented prosthesis.

  - Identification. An elbow joint metal/polymer constrained cemented prosthesis is a device intended to be implanted to replace an elbow joint. It is made of alloys, such as cobalt-chro- mium-molybdenum, or of these alloys and of an ultra-high molecular weight polyethylene bushing. The device pre- vents dislocation in more than one anatomic plane and consists of two components that are linked together. This generic type of  device  is  limited to those prostheses intended for use with bone cement (§ 888.3027).

  - Classification. Class II. The special controls for this device are:

    - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

- ‘‘Guidance Document for Testing Orthopedic Implants with Modified Me- tallic Surfaces Apposing Bone or Bone Cement,’’

- ‘‘Guidance Document for the Preparation of Premarket Notification (510(k)) Application for Orthopedic De- vices,’’

- ‘‘Guidance Document for Testing Non-articulating,	‘Mechanically Locked’ Modular Implant Compo- nents,’’

    - International Organization for Standardization’s (ISO):

#### 21 CFR Ch. I (4–1–17 Edition)

- ISO 5832–3:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 3: Wrought Titanium 6-Aluminum 4- Vandium   Alloy,’’

- ISO 5832–4:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 4: Co- balt-Chromium-Molybdenum Casting Alloy,’’

- ISO 5832–12:1996 ‘‘Implants for Surgery—Metallic Materials—Part 12: Wrought	Cobalt-Chromium-Molyb- denum    Alloy,’’

- ISO 5833:1992 ‘‘Implants for Sur- gery—Acrylic   Resin   Cements,’’

- ISO 5834–2:1998 ‘‘Implants for Sur- gery—Ultra High Molecular Weight Polyethylene—Part	2:	Moulded Forms,’’

- ISO 6018:1987 ‘‘Orthopaedic Im- plants—General Requirements for Marking,  Packaging,  and  Labeling,’’

- ISO 9001:1994 ‘‘Quality Systems— Model for Quality Assurance in Design/ Development, Production, Installation, and  Servicing,’’  and

- ISO 14630:1997 ‘‘Non-active Sur- gical  Implants—General  Require- ments,’’

    - American Society for Testing and Materials’:

      - F 75–92 ‘‘Specification for Cast Co- balt-28 Chromium-6 Molybdenum Alloy for Surgical Implant Material,’’

      - F 648–98 ‘‘Specification for Ultra- High-Molecular-Weight Polyethylene Powder and Fabricated Form for Sur- gical   Implants,’’

      - F 799–96 ‘‘Specification for Co- balt-28 Chromium-6 Molybdenum Alloy Forgings for Surgical Implants,’’

      - F 981–93 ‘‘Practice for Assess- ment of Compatibility of Biomaterials (Nonporous) for Surgical Implant with Respect to Effect of Material on Mus- cle and Bone,’’

      - F 1044–95 ‘‘Test Method for Shear Testing of Porous Metal Coatings,’’

      - F 1108–97 ‘‘Specification for Tita- nium-6 Aluminum-4 Vanadium Alloy Castings for Surgical Implants,’’

      - F 1147–95 ‘‘Test Method for Ten- sion Testing of Porous Metal Coatings, ’’ and

      - F 1537–94 ‘‘Specification for Wrought Cobalt-28 Chromium-6 Molyb- denum Alloy for Surgical Implants.’’

[65 FR 17147, Mar. 31, 2000]

#### Food and Drug Administration, HHS	§ 888.3210

##### § 888.3160 Elbow joint metal/polymer semi-constrained cemented pros- thesis.

        - Identification. An elbow joint metal/polymer semi-constrained ce- mented prosthesis is a device intended to be implanted to replace an elbow joint. The device limits translation and rotation in one or more planes via the geometry  of  its   articulating   surfaces. It has no linkage across-the-joint. This generic type of device includes pros- theses that consist of a humeral resur- facing component made of alloys, such as cobalt-chromium-molybdenum,  and a radial resurfacing component  made of ultra-high molecular weight poly- ethylene. This generic type of device is limited to  those  prostheses  intended for use with bone cement (§ 888.3027).

        - Classification. Class II.

##### § 888.3170 Elbow joint radial (hemi- elbow) polymer prosthesis.

- Identification. An elbow joint ra- dial (hemi-elbow) polymer prosthesis is a device intended to be implanted made of medical grade silicone elastomer used to replace the proximal end of the radius.

- Classification. Class II.

##### § 888.3180 Elbow joint humeral (hemi- elbow) metallic uncemented pros- thesis.

- Identification. An elbow joint hu- meral	(hemi-elbow)	metallic uncemented prosthesis is a device in- tended to be implanted made of alloys, such as cobalt-chromium-molybdenum, that is used to replace the distal end of the humerus formed by the trochlea humeri and the capitulum humeri. The generic type of device is limited to prostheses intended for use without bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any elbow joint humeral (hemi- elbow) metallic uncemented prosthesis that was in commercial distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to an elbow joint     humeral     (hemi-elbow)     metallic

uncemented prosthesis that was in commercial distribution before May 28, 1976. Any other elbow joint humeral (hemi-elbow) metallic uncemented prosthesis shall have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3200   Finger	joint	metal/metal constrained uncemented prosthesis.

- Identification. A finger joint metal/metal constrained uncemented prosthesis is a device intended to be implanted	to	replace		a metacarpophalangeal or proximal interphalangeal (finger) joint. The de- vice prevents dislocation in more than one anatomic plane and consists of two components which are linked together. This generic type of device includes prostheses made of alloys, such as co- balt-chromium-molybdenum,	or protheses made from alloys and ultra- high molecular weight polyethylene. This generic type of  device  is  limited to prostheses intended for use without bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any finger joint metal/metal con- strained uncemented  prosthesis  that was in commercial distribution  before May 28,  1976, or  that has,  on or  before December 26, 1996 been found to be sub- stantially equivalent to a finger joint metal/metal constrained uncemented prosthesis that was in commercial dis- tribution before May 28, 1976. Any other finger joint metal/metal con- strained uncemented prosthesis  shall have an approved PMA or a declared completed PDP in effect before being placed  in  commercial  distribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3210   Finger	joint	metal/metal constrained cemented prosthesis.

- Identification. A finger joint metal/metal constrained cemented prosthesis is  a  device  intended  to  be implanted	to	replace	a

#### § 888.3220

metacarpophalangeal  (finger)  joint. This device  prevents  dislocation  in more than one anatomic plane and has components which are linked together. This generic type of device includes prostheses that are made  of  alloys, such as cobalt-chromium-molybdenum, and is limited to those prostheses in- tended   for   use    with    bone    cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any finger joint metal/metal con- strained cemented prosthesis that  was in commercial distribution before May 28, 1976, or that has, on or before De- cember 26, 1996 been found to be sub- stantially equivalent to a finger joint metal/metal constrained cemented prosthesis that was in commercial dis- tribution before May 28, 1976. Any other finger joint metal/metal con- strained   cemented   prosthesis    shall have an approved PMA or a declared completed PDP in effect before being placed  in  commercial  distribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3220   Finger  joint  metal/polymer constrained cemented prosthesis.

- Identification. A finger joint metal/polymer constrained cemented prosthesis is a device intended to be implanted	to	replace	a metacarpophalangeal or proximal interphalangeal (finger) joint. The de- vice prevents dislocation in more than one anatomic plane,  and  consists  of two components which are linked to- gether. This generic type of device in- cludes prostheses that are made of al- loys, such as cobalt-chromium-molyb- denum,  and  ultra-high  molecular weight polyethylene, and is limited to those prostheses intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any finger joint metal/polymer con- strained  cemented  prosthesis  that  was

#### 21 CFR Ch. I (4–1–17 Edition)

in commercial distribution before May 28, 1976, or that has, on or before De- cember 26, 1996 been found to be sub- stantially equivalent to a finger joint metal/polymer constrained cemented prosthesis that was in commercial dis- tribution before May 28, 1976. Any other finger joint metal/polymer con- strained   cemented   prosthesis    shall have an approved PMA or a declared completed PDP in effect before being placed  in  commercial  distribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3230 Finger joint polymer con- strained prosthesis.

- Identification. A finger joint poly- mer constrained prosthesis is a device intended to be implanted to replace a metacarpophalangeal or proximal interphalangeal (finger) joint. This ge- neric type of device includes prostheses that consist of a single flexible across- the-joint component made from either a silicone elastomer or a combination pf polypropylene and polyester mate- rial. The flexible across-the-joint com- ponent may be covered with a silicone rubber sleeve.

- Classification. Class II.

##### § 888.3300 Hip joint metal constrained cemented or uncemented pros- thesis.

- Identification. A hip joint metal constrained cemented or uncemented prosthesis is a device intended to be implanted to replace a hip joint. The device prevents dislocation in more than one anatomic plane and has com- ponents that are linked together. This generic type of device includes pros- theses that have components made of alloys, such as cobalt-chromium-mo- lybdenum, and is intended for use with or without bone cement (§ 888.3027). This device is not intended for biologi- cal   fixation.

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any hip joint metal constrained ce- mented or uncemented prosthesis that was in commercial distribution  before May 28,  1976, or that  has, on or  before

#### Food and Drug Administration, HHS	§ 888.3330

December 26, 1996 been found to be sub- stantially equivalent to a hip joint metal constrained cemented or uncemented prosthesis that was in commercial distribution before May 28, 1976. Any other hip joint metal con- strained cemented or uncemented pros- thesis shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50709, Sept. 27, 1996]

##### § 888.3310 Hip joint metal/polymer con- strained cemented or uncemented prosthesis.

  - Identification. A hip joint metal/ polymer constrained cemented or uncemented prosthesis is a device in- tended to be implanted to replace a hip joint. The device prevents  dislocation in more than one anatomic plane and has components that are linked to- gether. This generic type of device in- cludes prostheses that have a femoral component made of alloys, such as co- balt-chromium-molybdenum, and an acetabular component made of ultra- high-molecular-weight polyethylene with or without a metal shell, made of alloys, such as cobalt-chromium-mo- lybdenum and titanium alloys. This ge- neric type of device is intended for use with     or     without      bone      cement (§ 888.3027).

  - Classification. Class II (special controls). The special control for this device is the FDA guidance document entitled ‘‘Class II Special Controls Guidance: Hip Joint Metal/Polymer Constrained Cemented or Uncemented Prosthesis.’’

[67 FR 21173, Apr. 30, 2002]

##### § 888.3320 Hip joint metal/metal semi- constrained, with a cemented ace- tabular component, prosthesis.

- Identification. A hip joint metal/ metal semi-constrained, with a ce- mented acetabular component, pros- thesis is a two-part device intended to be implanted to replace a hip joint. The device limits translation and  rotation in one or more planes via the geometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that consist  of  a  femoral  and  an  acetabular

component, both made of  alloys,  such as cobalt-chromium-molybdenum. This generic type of device is  limited  to those prostheses intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before May 18, 2016, for any hip joint metal/metal semi-con- strained prosthesis with a cemented acetabular component that was in com- mercial distribution before May 28, 1976, or that has, on or before May 18, 2016, been found to be substantially equivalent to a hip joint metal/metal semi-constrained prosthesis with a ce- mented acetabular component that was in commercial distribution before May 28, 1976. Any other hip joint metal/ metal semi-constrained prosthesis with a  cemented  acetabular   component shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[52 FR 33702, Sept. 4, 1987, as amended at 81

FR 8149, Feb. 18, 2016]

##### § 888.3330 Hip joint metal/metal semi- constrained, with an uncemented acetabular component, prosthesis.

- Identification. A hip joint metal/ metal semi-constrained, with an uncemented acetabular component, prosthesis is a two-part device intended to be implanted to replace a hip  joint. The device limits translation and rota- tion in one or more planes via the ge- ometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes pros- theses that consist of a femoral and an acetabular component, both made of alloys, such as cobalt-chromium-mo- lybdenum. The femoral component is intended to be fixed with bone cement. The acetabular component is intended for use without bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before May 18, 2016, for any hip joint metal/metal semi-con- strained	prosthesis	with	an

#### § 888.3340

uncemented  acetabular   component that was in commercial distribution before May 28, 1976, or that has, on or before May 18, 2016, been found to be substantially equivalent to a hip joint metal/metal semi-constrained pros- thesis with an uncemented acetabular component that was in commercial dis- tribution before May 28, 1976. Any other hip joint metal/metal semi-con- strained	prosthesis	with	an uncemented  acetabular  component shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[52 FR 33702, Sept. 4, 1987, as amended at 81

FR 8149, Feb. 18, 2016]

##### § 888.3340 Hip joint metal/composite semi-constrained cemented pros- thesis.

  - Identification. A hip joint metal/ composite semi-constrained cemented prosthesis is a two-part device intended to be implanted to replace a hip  joint. The device limits translation and rota- tion in one or more planes via the ge- ometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes pros- theses that consist of a femoral compo- nent made of alloys, such as cobalt- chromium-molybdenum, and an ace- tabular component made of ultra-high molecular weight polyethylene with carbon fibers composite. Both compo- nents are intended for use with bone cement (§ 888.3027).

  - Classification. Class II.

##### § 888.3350 Hip joint metal/polymer semi-constrained cemented pros- thesis.

- Identification. A hip joint metal/ polymer semi-constrained cemented prosthesis is  a  device  intended  to  be implanted to replace a hip joint. The device limits translation and  rotation in one or more planes via the geometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that have a femoral component made of al- loys, such as cobalt-chromium-molyb- denum, and an acetabular resurfacing component made of ultra-high molec- ular weight polyethylene and is limited

#### 21 CFR Ch. I (4–1–17 Edition)

to   those   prostheses   intended   for   use with bone cement (§ 888.3027).

- Classification. Class II.

##### § 888.3353 Hip joint metal/ceramic/ polymer semi-constrained cemented or nonporous uncemented pros- thesis.

- Identification. A hip joint metal/ ceramic/polymer semi-constrained ce- mented or nonporous uncemented pros- thesis is a device intended to be im- planted to replace a hip joint. This de- vice limits translation and rotation in one or more planes via the geometry of its articulating surfaces. It  has  no linkage across-the-joint.  The  two-part femoral component consists of a fem- oral stem made of alloys to be fixed in the intramedullary canal of the  femur by impaction with or without use of bone cement. The proximal end of the femoral stem is tapered with a surface that ensures positive locking with the spherical ceramic (aluminium oxide, A1203) head of the femoral component. The acetabular component is made of ultra-high molecular weight poly- ethylene  or  ultra-high  molecular weight polyethylene reinforced with nonporous metal alloys, and used with or without bone cement.

- Classification. Class II.

[54 FR 48239, Nov. 22, 1989; 54 FR 51342, Dec.

14, 1989]

##### § 888.3358 Hip joint metal/polymer/ metal semi-constrained porous-coat- ed uncemented prosthesis.

- Identification. A hip joint metal/ polymer/metal semi-constrained po- rous-coated uncemented prosthesis is a device intended to be implanted to re- place a hip joint. The device limits translation and rotation in one or more planes via the geometry of its articu- lating surfaces. It  has  no  linkage across the joint. This generic type of device has a femoral  component  made of a cobalt-chromium-molybdenum (Co-Cr-Mo) alloy or a titanium-alu- minum-vanadium (Ti-6Al-4V) alloy and an acetabular component composed of an ultra-high molecular weight poly- ethylene articulating bearing surface fixed in a metal shell made of Co-Cr-Mo or Ti-6Al-4V. The femoral stem and acetabular  shell  have  a  porous  coating

#### Food and Drug Administration, HHS	§ 888.3380

made of, in the case of Co-Cr-Mo sub- strates, beads of the same alloy,  and  in the case of Ti-6Al-4V  substrates,  fibers of commercially pure titanium or Ti- 6Al-4V alloy. The porous coating has a volume porosity between 30 and 70 per- cent, an average pore size between 100 and 1,000 microns, interconnecting po- rosity, and a porous coating thickness between 500 and 1,500 microns. The ge- neric type of device has a design to achieve biological fixation  to  bone without the use of bone cement.

- Classification. Class II.

[58 FR 3228, Jan. 8, 1993]

##### § 888.3360 Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis.

- Identification. A hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis is a device in- tended to be implanted to replace a portion of the hip joint. This generic type of device includes prostheses that have a femoral component made of al- loys, such as cobalt-chromium-molyb- denum. This generic type of device in- cludes designs which are intended to be fixed to the bone  with  bone  cement (§ 888.3027) as well as designs which have large window-like holes in  the stem of the device and which are in- tended for use without bone cement. However, in these latter designs, fixa- tion of the device is not achieved by means of bone ingrowth.

- Classification. Class II.

##### § 888.3370 Hip joint (hemi-hip) ace- tabular metal cemented prosthesis.

- Identification. A hip joint (hemi- hip) acetabular metal cemented pros- thesis is a device intended to be im- planted to replace a portion of the hip joint. This generic type of device in- cludes prostheses that have an ace- tabular  component  made  of  alloys, such as cobalt-chromium-molybdenum. This generic type of  device  is  limited to those prostheses intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for  any  hip  joint  (hemi-hip)  acetabular

metal cemented prosthesis that was in commercial distribution before May 28, 1976, or that has, on or before Decem- ber 26, 1996 been found to be substan- tially equivalent to a hip joint (hemi- hip) acetabular metal cemented pros- thesis that was in commercial distribu- tion before May 28, 1976. Any other hip joint metal  (hemi-hip)  acetabular metal cemented prosthesis shall have an approved PMA or a declared com- pleted PDP in  effect  before  being placed in commercial distribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996]

##### § 888.3380 Hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal cemented prosthesis.

- Identification. A hip joint femoral (hemi-hip) trunnion-bearing metal/ polyacetal cemented prosthesis is a two-part device intended to be im- planted to replace the head and neck of the femur. This generic type of device includes prostheses that consist of a metallic stem made of alloys, such as cobalt-chromium-molybdenum,    with an integrated cylindrical trunnion bearing at the upper end of the stem that fits into a recess in  the  head  of the device. The head of the device is made of polyacetal (polyoxymethylene) and it is covered by a metallic alloy, such as cobalt-chromium-molybdenum. The trunnion bearing allows  the  head of the device to rotate on its stem. The prosthesis is intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal ce- mented prosthesis that was in commer- cial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal ce- mented prosthesis that was in commer- cial distribution before May 28, 1976. Any other hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal ce- mented prosthesis shall have an ap- proved   PMA   or   a   declared   completed

#### § 888.3390

PDP  in  effect  before  being  placed  in commercial distribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996]

##### § 888.3390 Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis.

  - Identification. A hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis is a two-part device intended to be implanted to re- place the head and neck of the femur. This generic type of device includes prostheses that have a femoral compo- nent made of alloys, such as cobalt- chromium-molybdenum, and a snap-fit acetabular component made  of  an alloy, such as cobalt-chromium-molyb- denum,  and  ultra-high  molecular weight polyethylene. This generic  type of device may be fixed to the bone with bone cement (§ 888.3027) or implanted by impaction.

  - Classification. Class II.

##### § 888.3400 Hip joint femoral (hemi-hip) metallic resurfacing prosthesis.

- Identification. A hip joint femoral (hemi-hip) metallic resurfacing pros- thesis is a device intended to be im- planted to replace a portion of the hip joint. This generic type of device in- cludes prostheses that have a femoral resurfacing component made of alloys, such  as  cobalt-chromium-molybdenum.

- Classification. Class II.

##### § 888.3410 Hip joint metal/polymer or ceramic/polymer semiconstrained resurfacing cemented prosthesis.

- Identification. A hip joint metal/ polymer or ceramic/polymer semi-con- strained resurfacing cemented pros- thesis is a two-part device intended to be implanted to replace the articu- lating surfaces of the hip while pre- serving the femoral head and neck. The device limits translation and  rotation in one or more planes via the geometry of its articulating surfaces. It has no linkage across the joint. This generic type of device includes prostheses that consist of a femoral cap component made of a metal alloy, such as cobalt- chromium-molybdenum, or a ceramic material, that is placed over a sur- gically prepared femoral head, and an acetabular resurfacing polymer compo-

#### 21 CFR Ch. I (4–1–17 Edition)

nent.    Both    components    are    intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before January 3, 2005, for any hip joint metal/polymer or ce- ramic/polymer semiconstrained resur- facing cemented prosthesis that was in commercial distribution before May 28, 1976, or that has, on or before January 3, 2005, been found to be substantially equivalent to a hip joint metal/polymer or ceramic/polymer semiconstrained resurfacing cemented prosthesis that was in commercial distribution before May 28, 1976. Any other hip joint metal/ polymer	or	ceramic/polymer semiconstrained resurfacing cemented prosthesis must have an approved PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[69 FR 59134, Oct. 4, 2004]

##### § 888.3480 Knee joint femorotibial me- tallic constrained cemented pros- thesis.

- Identification. A knee joint femorotibial metallic constrained ce- mented prosthesis is a device intended to be implanted to replace part of a knee joint. The device prevents dis- location in more than one anatomic plane and has components that are linked together. The only knee joint movement allowed by the device is  in the sagittal  plane.  This generic type of device includes prostheses that have an intramedullary stem at both the proxi- mal and distal locations. The upper and lower components may be joined either by a solid bolt or pin, an internally threaded bolt  with  locking  screw, or  a bolt retained by circlip. The compo- nents of the device are made of alloys, such as cobalt-chromium-molybdenum. The stems of the device may be per- forated, but are intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996

#### Food and Drug Administration, HHS	§ 888.3510

for any knee joint  femorotibial  metal- lic  constrained   cemented   prosthesis that was in commercial distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to  a  knee joint femorotibial metallic constrained cemented prosthesis that was in com- mercial distribution before  May  28, 1976. Any other knee joint femorotibial metallic constrained cemented pros- thesis shall have an approved PMA or a declared completed PDP  in  effect  be- fore being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996]

##### § 888.3490 Knee joint femorotibial metal/composite non-constrained cemented prosthesis.

  - Identification. A knee joint femorotibial metal/composite non-con- strained cemented prosthesis is a de- vice intended to be implanted to re- place part of a knee joint. The device limits minimally (less than normal anatomic  constraints)  translation  in one or more planes. It has no linkage across-the-joint. This generic type of device includes prostheses that have a femoral condylar resurfacing compo- nent or components made  of  alloys, such as cobalt-chromium-molybdenum, and a tibial condylar component or components made of ultra-high molec- ular weight polyethylene with carbon fibers composite and are intended for use with bone cement (§ 888.3027).

  - Classification. Class II.

##### § 888.3500 Knee joint femorotibial metal/composite semi-constrained cemented prosthesis.

- Identification. A knee joint femorotibial metal/composite semi- constrained cemented prosthesis is a two-part device intended to be im- planted to replace part of a knee joint. The device limits translation and rota- tion in one or more planes via the ge- ometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes pros- theses that have a femoral component made of alloys, such as cobalt-chro- mium-molybdenum, and a tibial com- ponent with the articulating surfaces made   of   ultra-high   molecular   weight

polyethylene with carbon-fibers com- posite and is limited to those pros- theses intended for use with bone ce- ment (§ 888.3027).

- Classification. Class II.

##### § 888.3510 Knee joint femorotibial metal/polymer constrained ce- mented prosthesis.

- Identification. A knee joint femorotibial metal/polymer con- strained cemented prosthesis is a de- vice intended to be implanted to re- place part of a knee joint. The device limits translation or rotation in one or more planes and has components that are linked together or affined. This ge- neric type of device includes prostheses composed of a ball-and-socket joint lo- cated between a stemmed femoral and a stemmed tibial component and a run- ner and track joint  between  each  pair of femoral and tibial condyles. The ball-and-socket joint is composed of a ball at the  head  of  a  column  rising from the stemmed tibial  component. The ball, the  column,  the  tibial  pla- teau, and the stem for fixation of the tibial component are made of an alloy, such as cobalt-chromium-molybdenum. The ball of the tibial component is held within the socket of the femoral com- ponent by the femoral component’s flat outer surface. The flat outer surface of the tibial component abuts both a re- ciprocal flat surface within  the  cavity of the femoral component and  flanges on the femoral component designed to prevent distal displacement. The  stem of the femoral component is made of an alloy, such as cobalt-chromium-molyb- denum, but the socket of the compo- nent is made of ultra-high molecular weight polyethylene. The femoral com- ponent has  metallic  runners  which align with the ultra-high molecular weight polyethylene tracks that  press- fit into the metallic tibial component. The generic class also includes devices whose upper and lower components are linked with a solid bolt passing through a journal bearing of greater radius, permitting some rotation in the transverse plane, a minimal arc of ab- duction/adduction. This generic type of device is limited to those prostheses intended  for  use   with   bone   cement (§ 888.3027).

- Classification. Class II.

#### § 888.3520

##### § 888.3520 Knee joint femorotibial metal/polymer non-constrained ce- mented prosthesis.

  - Identification. A knee joint femorotibial metal/polymer non-con- strained cemented prosthesis is a de- vice intended to be implanted to re- place part of a knee joint. The device limits minimally (less than normal anatomic  constraints)  translation  in one or more planes. It has no linkage across-the-joint. This generic type of device includes prostheses that have a femoral condylar resurfacing compo- nent or components made  of  alloys, such as cobalt-chromium-molybdenum, and a tibial component or components made of ultra-high molecular weight polyethylene and are intended for use with bone cement (§ 888.3027).

  - Classification. Class II.

##### § 888.3530 Knee joint femorotibial metal/polymer semi-constrained ce- mented prosthesis.

- Identification. A knee joint femorotibial metal/polymer semi-con- strained cemented prosthesis is a de- vice intended to be implanted to re- place part of a knee joint. The device limits translation and rotation  in  one or more planes via the geometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that consist of a femoral component made of alloys, such as cobalt-chromium-molybdenum, and a tibial component made of ultra- high  molecular  weight  polyethylene and is limited to those prostheses in- tended   for   use    with    bone    cement (§ 888.3027).

- Classification. Class II.

##### § 888.3535 Knee joint femorotibial (uni- compartmental) metal/polymer po- rous-coated uncemented prosthesis.

- Identification. A knee joint femorotibial	(uni-compartmental) metal/polymer		porous-coated uncemented prosthesis is a device in- tended to be implanted to replace part of a knee joint. The device limits translation and rotation in one or more planes via the geometry of its articu- lating surface. It has no linkage across- the-joint. This generic type of device is designed to achieve biological  fixation to  bone  without  the  use  of  bone  ce-

#### 21 CFR Ch. I (4–1–17 Edition)

ment. This identification includes fixed-bearing knee prostheses where the ultra-high molecular weight poly- ethylene tibial bearing is rigidly se- cured to the metal tibial baseplate.

- Classification. Class II (special controls). The special control is FDA’s guidance: ‘‘Class II Special Controls Guidance Document: Knee Joint Patellofemorotibial and Femorotibial Metal/Polymer	Porous-Coated Uncemented Prostheses; Guidance for Industry and FDA.’’ See § 888.1 for the availability of this guidance.

[68 FR 14137, Mar. 24, 2003]

##### § 888.3540 Knee joint patellofemoral polymer/metal semi-constrained ce- mented prosthesis.

- Identification. A knee joint patellofemoral polymer/metal semi- constrained cemented prosthesis is a two-part device intended to be im- planted to replace part of a knee joint in the treatment of primary patellofemoral	arthritis	or chondromalacia. The device limits translation and rotation in one or more planes via the geometry of its articu- lating surfaces. It has no linkage across-the-joint. This generic type of device includes a component made of alloys, such as cobalt-chromium-mo- lybdenum or austenitic steel, for resur- facing the intercondylar groove (fem- oral sulcus) on the anterior aspect  of the distal femur, and a patellar compo- nent made of ultra-high molecular weight polyethylene. This generic  type of device is limited to those devices in- tended   for   use   with    bone    cement (§ 888.3027). The patellar component is designed to be implanted only with its femoral  component.

- Classification. Class II. The special controls for this device are:

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

- ‘‘Guidance Document for Testing Orthopedic Implants with Modified Me- tallic Surfaces Apposing Bone or Bone Cement,’’

- ‘‘Guidance Document for the Preparation of Premarket Notification

#### Food and Drug Administration, HHS	§ 888.3550

(510(k)) Applications for Orthopedic De- vices,’’  and

- ‘‘Guidance Document for Testing Non-articulating,	‘Mechanically Locked’ Modular Implant Compo- nents,’’ and

  - International Organization for Standardization’s (ISO):

- ISO 5832–3:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 3: Wrought Titanium 6-Aluminum 4- Vandium   Alloy,’’

- ISO 5832–4:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 4: Co- balt-Chromium-Molybdenum Casting Alloy,’’

- ISO 5832–12:1996 ‘‘Implants for Surgery—Metallic Materials—Part 12: Wrought	Cobalt-Chromium-Molyb- denum    Alloy,’’

- ISO 5833:1992 ‘‘Implants for Sur- gery—Acrylic   Resin   Cements,’’

- ISO 5834–2:1998 ‘‘Implants for Sur- gery—Ultra-high Molecular Weight Polyethylene—Part	2:	Moulded Forms,’’

- ISO 6018:1987 ‘‘Orthopaedic Im- plants—General Requirements for Marking,  Packaging,  and  Labeling,’’

- ISO 7207–2:1998 ‘‘Implants for Surgery—Components for Partial and Total Knee Joint Prostheses—Part 2: Articulating   Surfaces   Made   of   Metal,

Ceramic and Plastic Materials,’’ and

- ISO 9001:1994 ‘‘Quality Sys- tems—Model for Quality Assurance in Design/Development, Production, In- stallation,   and   Servicing,’’   and

  - American Society for Testing and Materials’:

    - F 75–92 ‘‘Specification for Cast Co- balt-28 Chromium-6 Molybdenum Alloy for Surgical Implant Material,’’

    - F 648–98 ‘‘Specification for Ultra- High-Molecular-Weight Polyethylene Powder and Fabricated Form for Sur- gical   Implants,’’

    - F 799–96 ‘‘Specification for Co- balt-28 Chromium-6 Molybdenum Alloy Forgings for Surgical Implants,’’

    - F 1044–95 ‘‘Test Method for Shear Testing of Porous Metal Coatings,’’

    - F 1108–97 ‘‘Titanium-6 Aluminum- 4 Vanadium Alloy Castings for Surgical Implants,’’

    - F 1147–95 ‘‘Test Method for Ten- sion Testing of Porous Metal Coat- ings,’’

    - F 1537–94 ‘‘Specification for Wrought Cobalt-28 Chromium-6 Molyb- denum Alloy for Surgical Implants,’’ and

    - F 1672–95 ‘‘Specification for Re- surfacing Patellar Prosthesis.’’

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996; 65 FR 17147, Mar. 31,

2000]

##### § 888.3550   Knee	joint patellofemorotibial polymer/metal/ metal constrained cemented pros- thesis.

- Identification. A knee joint patellofemorotibial	polymer/metal/ metal constrained  cemented  prosthesis is a device intended to be implanted to replace a knee joint. The device pre- vents dislocation in more than one anatomic plane and has  components that are linked together. This generic type of device includes prostheses that have a femoral component, a tibial component, a cylindrical bolt and ac- companying locking hardware that are all made of alloys, such as cobalt-chro- mium-molybdenum,		and	a retropatellar resurfacing component made of ultra-high molecular weight polyethylene. The retropatellar sur- facing component may be attached to the resected patella either with a me- tallic screw or bone cement. All stemmed metallic components within this generic type are intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any knee joint patellofemorotibial polymer/metal/metal constrained ce- mented prosthesis that was in commer- cial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a knee joint patellofemorotibial polymer/metal/metal constrained ce- mented prosthesis that was in commer- cial distribution before May 28, 1976. Any	other	knee	joint patellofemorotibial		polymer/metal/ metal constrained  cemented  prosthesis shall have an approved PMA or a de- clared  completed  PDP  in  effect  before

#### § 888.3560

being  placed  in  commercial  distribu- tion.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996]

##### § 888.3560   Knee	joint patellofemorotibial polymer/metal/ polymer semi-constrained cemented prosthesis.

  - Identification. A knee joint patellofemorotibial	polymer/metal/ polymer semi-constrained cemented prosthesis is a device intended to be implanted to replace a knee joint. The device limits translation and  rotation in one or more planes via the geometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that have a femoral component made of al- loys, such as cobalt-chromium-molyb- denum, and a tibial component or com- ponents and a retropatellar resurfacing component made of ultra-high molec- ular weight polyethylene. This generic type of device is limited to those pros- theses intended for use with bone ce- ment (§ 888.3027).

  - Classification. Class II.

##### § 888.3565   Knee	joint patellofemorotibial metal/polymer porous-coated uncemented pros- thesis.

- Identification. A knee joint patellofemorotibial metal/polymer po- rous-coated uncemented prosthesis is a device intended to be implanted to re- place a knee joint. The device limits translation and rotation in one or more planes via the geometry of its articu- lating surfaces. It has no linkage across-the-joint. This generic type of device is designed to achieve biological fixation to bone without the use of bone cement. This identification in- cludes fixed-bearing knee prostheses where the ultra high molecular weight polyethylene tibial bearing is rigidly secured to the metal tibial base plate.

- Classification. Class II (special controls). The special control is FDA’s guidance: ‘‘Class II Special Controls Guidance Document: Knee Joint Patellofemorotibial and Femorotibial Metal/Polymer	Porous-Coated Uncemented Prostheses; Guidance for

#### 21 CFR Ch. I (4–1–17 Edition)

Industry  and  FDA.’’  See  § 888.1  for  the availability  of  this  guidance.

[68 FR 14137, Mar. 24, 2003]

##### § 888.3570 Knee joint femoral (hemi- knee) metallic uncemented pros- thesis.

- Identification. A knee joint fem- oral (hemi-knee) metallic uncemented prosthesis is a device made of alloys, such as cobalt-chromium-molybdenum, intended to be  implanted  to  replace part of a knee joint. The device limits translation and rotation in one or more planes via the geometry of its articu- lating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that consist of a femoral component with or with- out protuberance(s) for the enhance- ment of fixation and is limited to those prostheses intended for use without bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any knee joint femoral (hemi-knee) metallic uncemented prosthesis that was in commercial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be sub- stantially equivalent to a knee joint femoral	(hemi-knee)	metallic uncemented prosthesis that was in commercial distribution before May 28, 1976. Any other knee joint femoral (hemi-knee) metallic uncemented pros- thesis shall have an approved PMA or a declared completed PDP in effect be- fore being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50710, Sept. 27, 1996]

##### § 888.3580 Knee joint patellar (hemi- knee) metallic resurfacing uncemented prosthesis.

- Identification. A knee joint patellar (hemi-knee) metallic resur- facing uncemented prosthesis is a de- vice made of alloys, such as cobalt- chromium-molybdenum, intended to be implanted to replace the retropatellar articular surface of the patellofemoral joint. The device limits minimally (less than    normal    anatomic    constraints)

#### Food and Drug Administration, HHS	§ 888.3650

translation in one or more planes. It has no linkage across-the-joint. This generic type of device includes pros- theses that have a retropatellar resur- facing component and an orthopedic screw to transfix the patellar remnant. This generic type of  device  is  limited to those prostheses intended for use without bone cement (§ 888.3027).

- Classification. (1) Class II when in- tended for treatment of degenerative and  posttraumatic  patellar  arthritis.

(2) Class III when intended for uses other than treatment of degenerative and posttraumatic patellar arthritis.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any knee joint patellar (hemi-knee) metallic resurfacing uncemented pros- thesis described in paragraph (b)(2) of this section that was in commercial distribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a knee joint  patellar  (hemi- knee) metallic resurfacing uncemented prosthesis that was in commercial dis- tribution before May 28, 1976. Any other knee joint patellar (hemi-knee) metallic resurfacing uncemented pros- thesis shall have an approved PMA or a declared completed PDP  in  effect  be- fore being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50711, Sept. 27, 1996]

##### § 888.3590 Knee joint tibial  (hemi- knee) metallic resurfacing uncemented prosthesis.

- Identification. A knee joint tibial (hemi-knee) metallic resurfacing uncemented prosthesis is a device in- tended to be implanted to replace part of a knee joint. The device limits mini- mally (less than normal anatomic con- straints) translation in one or more planes. It has no linkage  across-the- joint. This prosthesis is made of alloys, such as cobalt-chromium-molybdenum, and is intended to resurface one tibial condyle. The generic type of device is limited to  those  prostheses  intended for use without bone cement (§ 888.3027).

- Classification. Class II.

##### § 888.3640 Shoulder joint  metal/metal or metal/polymer constrained ce- mented prosthesis.

- Identification. A shoulder joint metal/metal or metal/polymer con- strained cemented prosthesis is a de- vice intended to be implanted to re- place a shoulder joint. The device pre- vents dislocation in more than one anatomic plane and has  components that are linked together. This generic type of device includes prostheses that have a humeral component made of al- loys, such as cobalt-chromium-molyb- denum, and a glenoid component made of this alloy or a combination of this alloy and ultra-high molecular weight polyethylene. This generic type of de- vice is limited to those prostheses in- tended   for   use    with    bone    cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any shoulder joint metal/metal or metal/polymer constrained cemented prosthesis that was in commercial dis- tribution before May 28, 1976, or that has, on or before December 26, 1996 been found to be substantially equiva- lent to a shoulder joint metal/metal or metal/polymer constrained cemented prosthesis that was in commercial dis- tribution before May 28, 1976. Any other shoulder joint metal/metal or metal/polymer constrained cemented prosthesis shall have an  approved  PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50711, Sept. 27, 1996]

##### § 888.3650 Shoulder joint metal/poly- mer non-constrained cemented prosthesis.

- Identification. A shoulder joint metal/polymer non-constrained ce- mented prosthesis is a device intended to be implanted to replace a shoulder joint. The device limits minimally (less than normal anatomic constraints) translation in one or more planes. It has no linkage across-the-joint. This generic type of device includes pros- theses that have a humeral component

#### § 888.3660

made of alloys, such as cobalt-chro- mium-molybdenum, and a glenoid re- surfacing component made of ultra- high molecular weight polyethylene, and is limited to those prostheses in- tended   for   use    with    bone    cement (§ 888.3027).

- Classification. Class II. The special controls for this device are:

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

- ‘‘Guidance Document for Testing Orthopedic Implants with Modified Me- tallic Surfaces Apposing Bone or Bone Cement,’’

- ‘‘Guidance Document for the Preparation of Premarket Notification (510(k)) Application for Orthopedic De- vices,’’ and

- ‘‘Guidance Document for Testing Non-articulating,	‘Mechanically Locked’ Modular Implant Compo- nents,’’

  - International Organization for Standardization’s (ISO):

- ISO 5832–3:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 3: Wrought Titanium 6-Aluminum 4- Vandium   Alloy,’’

- ISO 5832–4:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 4: Co- balt-Chromium-Molybdenum Casting Alloy,’’

- ISO 5832–12:1996 ‘‘Implants for Surgery—Metallic Materials—Part 12: Wrought	Cobalt-Chromium-Molyb- denum    Alloy,’’

- ISO 5833:1992 ‘‘Implants for Sur- gery—Acrylic   Resin   Cements,’’

- ISO 5834–2:1998 ‘‘Implants for Sur- gery—Ultra-high Molecular Weight Polyethylene—Part	2:	Moulded Forms,’’

- ISO 6018:1987 ‘‘Orthopaedic Im- plants—General Requirements for Marking,  Packaging,  and  Labeling,’’ and

- ISO 9001:1994 ‘‘Quality Systems— Model for Quality Assurance in Design/ Development, Production, Installation, and  Servicing,’’  and

  - American Society for Testing and Materials’:

#### 21 CFR Ch. I (4–1–17 Edition)

- F 75–92 ‘‘Specification for Cast Co- balt-28 Chromium-6 Molybdenum Alloy for Surgical Implant Material,’’

- F 648–98 ‘‘Specification for Ultra- High-Molecular-Weight Polyethylene Powder and Fabricated Form for Sur- gical   Implants,’’

- F 799–96 ‘‘Specification for Co- balt-28 Chromium-6 Molybdenum Alloy Forgings for Surgical Implants,’’

- F 1044–95 ‘‘Test Method for Shear Testing of Porous Metal Coatings,’’

- F 1108–97 ‘‘Titanium-6 Aluminum- 4 Vanadium Alloy Castings for Surgical Implants,’’

- F 1147–95 ‘‘Test Method for Ten- sion Testing of Porous Metal Coat- ings,’’

- F 1378–97 ‘‘Specification for Shoulder  Prosthesis,’’  and

- F 1537–94 ‘‘Specification for Wrought Cobalt-28 Chromium-6 Molyb- denum Alloy for Surgical Implants.’’

[52 FR 33702, Sept. 4, 1987, as amended at 65

FR 17148, Mar. 31, 2000]

##### § 888.3660 Shoulder joint metal/poly- mer semi-constrained cemented prosthesis.

- Identification. A shoulder joint metal/polymer semi-constrained ce- mented prosthesis is a device intended to be implanted to replace a shoulder joint. The device limits translation and rotation in one or more planes via the geometry  of  its   articulating   surfaces. It has no linkage across-the-joint. This generic type of device includes pros- theses that have a humeral resurfacing component made of alloys, such as co- balt-chromium-molybdenum, and a glenoid resurfacing component made of ultra-high molecular weight poly- ethylene, and is limited to those pros- theses intended for use with bone ce- ment (§ 888.3027).

- Classification. Class II. The special controls for this device are:

  - FDA’s:

- ‘‘Use of International Standard ISO 10993 ‘Biological Evaluation of Medical  Devices—Part  I:  Evaluation and  Testing,’ ’’

- ‘‘510(k) Sterility Review  Guidance of 2/12/90 (K90–1),’’

- ‘‘Guidance Document for Testing Orthopedic Implants with Modified Me- tallic Surfaces Apposing Bone or Bone Cement,’’

#### Food and Drug Administration, HHS	§ 888.3670

- ‘‘Guidance Document for the Preparation of Premarket Notification (510(k)) Application for Orthopedic De- vices,’’ and

- ‘‘Guidance Document for Testing Non-articulating,	‘Mechanically Locked’ Modular Implant Compo- nents,’’

  - International Organization for Standardization’s (ISO):

- ISO 5832–3:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 3: Wrought Titanium 6-aluminum 4- vandium   Alloy,’’

- ISO 5832–4:1996 ‘‘Implants for Sur- gery—Metallic Materials—Part 4: Co- balt-chromium-molybdenum casting alloy,’’

- ISO 5832–12:1996 ‘‘Implants for Surgery—Metallic Materials—Part 12: Wrought	Cobalt-chromium-molyb- denum   alloy,’’

- ISO 5833:1992 ‘‘Implants for Sur- gery—Acrylic   Resin   Cements,’’

- ISO 5834–2:1998 ‘‘Implants for Sur- gery—Ultra-high Molecular Weight Polyethylene—Part	2:	Moulded Forms,’’

- ISO 6018:1987 ‘‘Orthopaedic Im- plants—General Requirements for Marking,  Packaging,  and  Labeling,’’ and

- ISO 9001:1994 ‘‘Quality Systems— Model for Quality Assurance in Design/ Development, Production, Installation, and  Servicing,’’  and

  - American Society for Testing and Materials’:

    - F 75–92 ‘‘Specification for Cast Co- balt-28 Chromium-6 Molybdenum Alloy for Surgical Implant Material,’’

    - F 648–98 ‘‘Specification for Ultra- High-Molecular-Weight Polyethylene Powder and Fabricated Form for Sur- gical   Implants,’’

    - F 799–96 ‘‘Specification for Co- balt-28 Chromium-6 Molybdenum Alloy Forgings for Surgical Implants,’’

    - F 1044–95 ‘‘Test Method for Shear Testing of Porous Metal Coatings,’’

    - F 1108–97 ‘‘Specification for Tita- nium-6 Aluminum-4 Vanadium Alloy Castings for Surgical Implants,’’

    - F 1147–95 ‘‘Test Method for Ten- sion Testing of Porous Metal,’’

    - F 1378–97 ‘‘Standard Specifica- tion for Shoulder Prosthesis,’’ and

    - F 1537–94 ‘‘Specification for Wrought Cobalt-28 Chromium-6 Molyb- denum Alloy for Surgical Implants.’’

[52 FR 33702, Sept. 4, 1987, as amended at 65

FR 17148, Mar. 31, 2000]

##### § 888.3670 Shoulder joint metal/poly- mer/metal nonconstrained or semi- constrained	porous-coated uncemented prosthesis.

- Identification. A shoulder joint metal/polymer/metal  nonconstrained or semi-constrained porous-coated uncemented prosthesis is a device in- tended to be implanted to replace a shoulder joint. The device limits move- ment in one or more planes. It has no linkage across-the-joint. This generic type of device includes prostheses that have a humeral component made of al- loys such as cobalt-chromium-molyb- denum (Co-Cr-Mo) and titanium-alu- minum-vanadium (Ti-6Al-4V) alloys, and a glenoid resurfacing component made of ultra-high molecular weight polyethylene, or a combination of an articulating  ultra-high  molecular weight bearing surface fixed in a metal shell made of alloys such as Co-Cr-Mo and Ti-6Al-4V. The humeral component and glenoid backing have a porous coating made of, in the case  of  Co-Cr- Mo components, beads of  the  same alloy or commercially pure titanium powder, and in the case of Ti-6Al-4V components, beads or fibers of commer- cially pure titanium or Ti-6Al-4V alloy, or commercially pure titanium powder. The porous coating has a volume poros- ity between 30 and 70 percent, an aver- age pore size between 100 and 1,000 mi- crons, interconnecting porosity, and a porous coating thickness between 500 and 1,500 microns. This generic type of device is designed to achieve biological fixation to bone without  the  use  of bone cement.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Class II Special Con- trols Guidance: Shoulder Joint Metal/ Polymer/Metal Nonconstrained or Semi-Constrained	Porous-Coated Uncemented Prosthesis.’’

[66 FR 12737, Feb. 28, 2001]

#### § 888.3680

##### § 888.3680 Shoulder joint glenoid (hemi-shoulder) metallic cemented prosthesis.

  - Identification. A shoulder joint glenoid (hemi-shoulder) metallic ce- mented prosthesis is a device that has a glenoid (socket) component made of alloys, such as cobalt-chromium-mo- lybdenum, or alloys with ultra-high molecular weight polyethylene and in- tended to be implanted to replace part of a shoulder joint. This generic type of device is limited to those prostheses intended  for  use   with   bone   cement (§ 888.3027).

  - Classification. Class III.

  - Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any shoulder joint glenoid (hemi- shoulder) metallic cemented prosthesis that was in commercial distribution before May  28, 1976,  or that  has, on  or before December 26, 1996 been found to be substantially equivalent to a shoul- der joint glenoid (hemi-shoulder) me- tallic cemented prosthesis that was in commercial distribution before May 28, 1976. Any other shoulder joint glenoid (hemi-shoulder) metallic cemented prosthesis shall have an  approved  PMA or a declared completed PDP in effect before being placed in commercial dis- tribution.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50711, Sept. 27, 1996]

##### § 888.3690 Shoulder joint humeral (hemi-shoulder)	metallic uncemented prosthesis.

- Identification. A shoulder joint hu- meral	(hemi-shoulder)	metallic uncemented prosthesis is  a  device made of alloys, such as cobalt-chro- mium-molybdenum. It has an intramedullary stem and is intended to be implanted to replace the articular surface of the proximal end of the hu- merus and to be fixed without bone ce- ment (§ 888.3027). This device is not in- tended for biological fixation.

- Classification. Class II.

##### § 888.3720 Toe joint polymer con- strained prosthesis.

- Identification. A toe joint polymer constrained prosthesis is a device made

#### 21 CFR Ch. I (4–1–17 Edition)

of silicone elastomer or polyester rein- forced silicone elastomer intended to be implanted to replace the first metatarsophalangeal (big toe) joint. This generic type of device consists of a single flexible across-the-joint com- ponent that prevents dislocation in more than one anatomic plane.

- Classification. Class II.

##### § 888.3730   Toe joint phalangeal (hemi- toe) polymer prosthesis.

- Identification. A toe joint phalan- geal (hemi-toe) polymer prosthesis is a device made of silicone elastomer in- tended to be implanted to replace the base of the proximal phalanx of the toe.

- Classification. Class II.

##### § 888.3750   Wrist   joint   carpal   lunate polymer prosthesis.

- Identification. A wrist joint carpal lunate prosthesis is a one-piece device made of silicone elastomer intended to be implanted to replace the carpal lunate bone of the wrist.

- Classification. Class II.

##### § 888.3760   Wrist  joint  carpal  scaphoid polymer prosthesis.

- Identification. A wrist joint carpal scaphoid polymer prosthesis is a one- piece device made of silicone elastomer intended to be implanted to replace the carpal scaphoid bone of the wrist.

- Classification. Class II.

##### § 888.3770   Wrist joint carpal trapezium polymer prosthesis.

- Identification. A wrist joint carpal trapezium polymer prosthesis is a one- piece device made of silicone elastomer or silicone elastomer/polyester mate- rial intended to be implanted to re- place the carpal trapezium bone of the wrist.

- Classification. Class II.

##### § 888.3780   Wrist   joint   polymer   con- strained prosthesis.

- Identification. A wrist joint poly- mer constrained prosthesis is a device made of polyester-reinforced silicone elastomer intended  to  be implanted to replace a wrist joint. This generic type of  device  consists  of  a  single  flexible

#### Food and Drug Administration, HHS	§ 888.4200

across-the-joint component that pre- vents dislocation in more than one anatomic plane.

- Classification. Class II.

##### § 888.3790 Wrist joint metal con- strained cemented prosthesis.

- Identification. A wrist joint metal constrained cemented prosthesis is a device intended to be implanted to re- place a wrist joint. The device prevents dislocation in more than one anatomic plane and consists of either a single flexible across-the-joint component or two components linked together. This generic type of device is limited to a device which is made of alloys, such as cobalt-chromium-molybdenum, and is limited to  those  prostheses  intended for use with bone cement (§ 888.3027).

- Classification. Class III.

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any wrist joint metal constrained cemented prosthesis that was in com- mercial distribution before May 28, 1976, or that has, on or before Decem- ber 26, 1996 been found to be substan- tially equivalent to a wrist joint metal constrained cemented prosthesis that was in commercial distribution before May 28, 1976. Any other wrist joint metal constrained  cemented  prosthesis shall have an approved PMA or a de- clared completed PDP in effect before being placed in commercial distribu- tion.

[52 FR 33702, Sept. 4, 1987, as amended at 61

FR 50711, Sept. 27, 1996]

##### § 888.3800 Wrist joint metal/polymer semi-constrained cemented pros- thesis.

- Identification. A wrist joint metal/ polymer semi-constrained cemented prosthesis is a device intended to be implanted to replace a wrist joint. The device limits translation and  rotation in one or more planes via the geometry of its articulating surfaces. It has no linkage across-the-joint. This generic type of device includes prostheses that have either a one-part radial compo- nent made of alloys, such as cobalt- chromium-molybdenum, with an ultra- high    molecular    weight    polyethylene

bearing surface, or a two-part radial component made of alloys and an ultra-high molecular weight poly- ethylene ball that is mounted on the radial component with a trunnion bear- ing. The metallic portion of the two- part radial component is inserted into the radius. These devices have a meta- carpal component(s) made of alloys, such as cobalt-chromium-molybdenum. This generic type of  device  is  limited to those prostheses intended for use with bone cement (§ 888.3027).

- Classification. Class II.

##### § 888.3810   Wrist    joint    ulnar    (hemi- wrist) polymer prosthesis.

- Identification. A wrist joint ulnar (hemi-wrist) polymer prosthesis is a mushroom-shaped device made of a medical grade silicone elastomer or ultra-high molecular weight poly- ethylene intended to be implanted into the intramedullary canal of the bone and held in place by a suture. Its pur- pose is to cover the resected end of the distal ulna to control bone overgrowth and to provide an articular surface for the radius and carpus.

- Classification. Class II.

## Subpart E—Surgical Devices

##### § 888.4150   Calipers for clinical use.

- Identification. A caliper for clin- ical use is a compass-like device in- tended for use in measuring the thick- ness or diameter of a part of the body or the distance between two body sur- faces, such as for measuring an excised skeletal specimen to determine the proper replacement size of a prosthesis.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

##### § 888.4200   Cement dispenser.

- Identification.  A  cement  dispenser is a nonpowered syringe-like device in- tended for use in  placing  bone  cement (§ 888.3027) into surgical sites.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

§ 888.4210

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR  52953,  Dec.  29,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38815, July 25, 2001]

##### § 888.4210   Cement   mixer   for   clinical use.

  - Identification. A cement mixer for clinical use is a device consisting of a container intended for use  in  mixing bone cement (§ 888.3027).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR  52953,  Dec.  29,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38815, July 25, 2001]

##### § 888.4220   Cement	monomer	vapor evacuator.

- Identification. A cement monomer vapor  evacuator  is  a  device  intended for use  during surgery  to contain or  re- move undesirable fumes, such as mon- omer     vapor     from      bone      cement (§ 888.3027).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR 52954, Dec. 29, 1988; 66 FR 38815, July 25,

2001]

##### § 888.4230   Cement ventilation tube.

- Identification. A  cement  ventila- tion tube is a tube-like device usually made of plastic intended to be inserted into a surgical cavity to allow the re- lease of air  or  fluid  from  the  cavity  as it  is  being   filled   with   bone   cement (§ 888.3027).

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR  52954,  Dec.  29,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38815, July 25, 2001]

##### § 888.4300   Depth gauge for clinical use.

- Identification. A depth gauge for clinical use is a measuring device in-

#### 21 CFR Ch. I (4–1–17 Edition)

tended for various medical purposes, such as to determine the proper length of screws for fastening the ends of a fractured bone.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

##### § 888.4540   Orthopedic manual surgical instrument.

- Identification. An orthopedic man- ual surgical instrument is a nonpow- ered hand-held device intended for medical purposes to manipulate tissue, or for use with other devices in ortho- pedic surgery. This generic type of de- vice includes the cerclage applier, awl, bender, drill brace, broach, burr, cork- screw, countersink, pin crimper, wire cutter, prosthesis driver, extractor, file, fork, needle holder, impactor, bending or contouring instrument, compression instrument, passer, socket positioner, probe, femoral neck punch, socket pusher, reamer, rongeur, scis- sors, screwdriver,  bone  skid,  staple driver, bone screw starter, surgical stripper, tamp, bone tap, trephine, wire twister, and wrench.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38815,  July 25,

2001]

##### § 888.4580   Sonic   surgical   instrument and accessories/attachments.

- Identification. A sonic surgical in- strument is a hand-held device with various accessories or attachments, such as a cutting tip that vibrates at high frequencies, and is intended for medical purposes to cut bone or other materials, such as acrylic.

- Classification. Class II.

##### § 888.4600   Protractor for clinical use.

- Identification. A protractor for clinical use is a device intended for use in measuring the angles of bones, such as on x-rays or in surgery.

#### Food and Drug Administration, HHS	§ 888.5960

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

##### § 888.4800   Template for clinical use.

- Identification. A template for clin- ical use is a device that consists of a pattern or guide intended for medical purposes, such as selecting or posi- tioning orthopedic implants or guiding the marking of tissue before cutting.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

##### § 888.5850   Nonpowered	orthopedic traction apparatus and accessories.

- Identification. A nonpowered or- thopedic traction apparatus is a device that consists of a rigid frame with non- powered traction accessories, such as cords, pulleys, or weights, and that is intended to apply a therapeutic pulling force to the skeletal system.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[52 FR 33702, Sept. 4, 1987, as amended at 66

FR 38815, July 25, 2001]

##### § 888.5890   Noninvasive   traction   com- ponent.

- Identification. A noninvasive trac- tion component is a device, such as a head halter, pelvic belt, or a traction splint, that does not penetrate the skin and is intended to assist in connecting a patient to a traction apparatus so that a therapeutic pulling force may be applied to the patient’s body.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR 52954, Dec. 29, 1988; 66 FR 38815, July 25,

2001]

##### § 888.5940   Cast component.

- Identification. A cast component is a device intended for medical purposes to protect or support a cast. This ge- neric type of device includes the cast heel, toe cap, cast support, and walking iron.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR  52954,  Dec.  29,  1988;  59  FR  63014,  Dec.  7,

1994; 66 FR 38815, July 25, 2001]

##### § 888.5960   Cast removal instrument.

- Identification. A cast removal in- strument is an AC-powered, hand-held device intended to remove a cast from a patient. This generic type of device includes the electric cast cutter and cast vacuum.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9.

[55 FR 48443, Nov. 20, 1990, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38816,  July 25,

2001]

#### § 888.5980

21 CFR Ch. I (4–1–17 Edition)

| § 888.5980 Manual cast application and removal instrument. (a) Identification. A manual cast ap- plication and removal instrument is a nonpowered hand-held device intended | 890.3100 | Mechanical chair. |
| --- | --- | --- |
|  | 890.3110 | Electric positioning chair. |
|  | 890.3150 | Crutch. |
|  | 890.3175 | Flotation cushion. |
|  | 890.3410 | External limb orthotic component. |
|  | 890.3420 | External limb prosthetic compo- |
| to be used in applying or removing a cast. This generic type of device in- | nent 890.3450 | . Upper extremity prosthesis includ- |

cludes the cast knife, cast spreader, plaster saw, plaster dispenser, and casting  stand.

(b) Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 888.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records,  and

§ 820.198, regarding complaint files.

[52 FR 33702, Sept. 4, 1987, as amended at 53

FR 52954, Dec. 29, 1988; 66 FR 38816, July 25,

2001]

## PART 890—PHYSICAL MEDICINE DEVICES

##### Subpart A—General Provisions

Sec.

890.1    Scope.

890.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Physical Medicine Diagnostic Devices

890.1175   Electrode cable.

890.1225     Chronaximeter.

890.1375    Diagnostic electromyograph. 890.1385    Diagnostic  electromyograph  needle

electrode.

890.1450   Powered reflex hammer. 890.1575    Force-measuring platform.

890.1600    Intermittent pressure measurement

ing a simultaneously powered elbow and/ or shoulder with greater than two simul- taneous powered degrees of freedom and controlled by non-implanted electrical components.

890.3475   Limb orthosis.

890.3480  Powered	lower	extremity exoskeleton.

890.3490    Truncal orthosis.

890.3500   External     assembled    lower     limb prosthesis.

890.3520   Plinth.

890.3610    Rigid pneumatic structure orthosis. 890.3640  Arm sling.

890.3665    Congenital   hip   dislocation   abduc- tion   splint.

890.3675   Denis Brown splint.  890.3690   Powered wheeled stretcher.

890.3700   Nonpowered    communication    sys- tem.

890.3710    Powered communication system. 890.3725    Powered environmental control sys-

tem.

890.3750   Mechanical table.

890.3760    Powered table.

890.3790   Cane,  crutch,  and  walker  tips  and pads.

890.3800   Motorized three-wheeled vehicle. 890.3825   Mechanical walker.

890.3850   Mechanical wheelchair.

890.3860  Powered wheelchair. 890.3880   Special grade wheelchair. 890.3890    Stair-climbing wheelchair.

890.3900   Standup wheelchair.

890.3910   Wheelchair accessory.

890.3920   Wheelchair component.

890.3930   Wheelchair elevator. 890.3940   Wheelchair platform scale.

##### Subpart E [Reserved]

Subpart F—Physical Medicine Therapeutic Devices

890.5050   Daily activity assist device. 890.5100 Immersion hydrobath.

| system. | 890.5110 | Paraffin bath. |
| --- | --- | --- |
| 890.1615 Miniature pressure transducer. | 890.5125 | Nonpowered sitz bath. |
| 890.1850 Diagnostic muscle stimulator. | 890.5150 | Powered patient transport. |
| 890.1925 Isokinetic testing and evaluation | 890.5160 | Air-fluidized bed. |

system.

##### Subpart C [Reserved]

Subpart D—Physical Medicine Prosthetic Devices

890.3025    Prosthetic and orthotic accessory. 890.3075 Cane.

890.5170    Powered flotation therapy bed. 890.5180    Manual patient rotation bed. 890.5225    Powered patient rotation bed. 890.5250    Moist steam cabinet.

890.5275   Microwave diathermy.

890.5290   Shortwave diathermy.

890.5300    Ultrasonic diathermy.

890.5350   Exercise component.

890.5360   Measuring exercise equipment.

Food and Drug Administration, HHS	§ 890.3

890.5370    Nonmeasuring exercise equipment. 890.5380   Powered exercise equipment.

890.5410   Powered finger exerciser. 890.5500  Infrared lamp.

890.5525  Iontophoresis device.

890.5575   Powered   external   limb   overload warning device.

890.5650    Powered inflatable tube massager. 890.5660   Therapeutic massager.

890.5700   Cold pack.

890.5710   Hot or cold disposable pack.

890.5720   Water circulating hot or cold pack. 890.5730  Moist heat pack.

890.5740    Powered heating pad.

890.5760   Nonpowered  lower  extremity  pres- sure wrap.

890.5765   Pressure-applying device. 890.5850   Powered muscle stimulator.

890.5860    Ultrasound and muscle stimulator. 890.5880    Multi-function      physical      therapy

table.

890.5900   Powered traction equipment. 890.5925   Traction accessory.

890.5940   Chilling unit.

890.5950   Powered heating unit. 890.5975   Therapeutic vibrator.

AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 360l, 371.

SOURCE:  48  FR  53047,  Nov.  23,  1983,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 890 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 890.1   Scope.

    - This part sets forth the classifica- tion of physical medicine devices in- tended for human use that are in com- mercial distribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 may not show merely that the device is accurately described by the section title and iden- tification provisions of a regulation in this part, but shall state why the de- vice is substantially  equivalent  to other devices, as required by § 807.87.

    - To avoid duplicative listings, a physical medicine device that has two or more types of uses (e.g., used both as a diagnostic device and as a thera- peutic device) is listed only in one sub- part.

    - References in this part to regu- latory sections of the Code of Federal

Regulations are to chapter I of title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[52 FR 17741, May 11, 1987, as amended at 73

FR 34860, June 19, 2008; 78 FR 18233, Mar. 26,

2013]

##### § 890.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap- proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application of premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act FDA must promulgate a regulation under section 515(b) of the act requir- ing such approval, except  as  provided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

#### § 890.9

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, includiing a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows  that a  device  being com- mercially distributed may be a ‘‘new’’ device as defined in this section be- cause of any new intended use or other reasons, FDA may codify the statutory classification of the device into  class III for such new use. Accordingly, the regulation for such a class III device states that as of the enactment date of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

[52 FR 17741, May 11, 1987]

##### § 890.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

- The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or  the  device  is  intended  for  lay  use

#### 21 CFR Ch. I (4–1–17 Edition)

where the former intended use was by health care professionals only;

- The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

- The device is an in vitro device that is intended:

  - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

  - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

  - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

  - For assessing the risk of cardio- vascular diseases;

  - For use in diabetes management;

  - For identifying or inferring the identity of a microorganism directly from clinical material;

  - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

  - For noninvasive testing as defined in § 812.3(k) of this chapter; and

  - For near patient testing (point of care).

[65 FR 2321, Jan. 14, 2000]

#### Food and Drug Administration, HHS	§ 890.1615

Subpart B—Physical Medicine Diagnostic Devices

##### § 890.1175   Electrode cable.

    - Identification. An electrode cable is a device composed of strands of insu- lated electrical conductors laid to- gether around a central core and in- tended for medical purposes to connect an electrode from a patient to a diag- nostic machine.

    - Classification. Class II (special controls). The special controls consist of:

- The performance standard under part 898 of this chapter, and

- The guidance document entitled ‘‘Guidance on the Performance Stand- ard for Electrode Lead Wires and Pa- tient Cables.’’ This device is  exempt from the premarket notification proce- dures of subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR  63014,  Dec.  7,  1994;  65  FR  19319,  Apr.  11,

2000]

##### § 890.1225   Chronaximeter.

  - Identification. A chronaximeter is a device intended for medical purposes to measure neuromuscular excitability by means of a strength-duration curve that provides a basis for diagnosis and prognosis of neurological dysfunction.

  - Classification. Class II (perform- ance standards).

##### § 890.1375   Diagnostic electromyograph.

- Identification. A diagnostic electromyograph is a device intended for medical purposes, such as to mon- itor and display the bioelectric signals produced by muscles, to stimulate pe- ripheral nerves, and to monitor and display the electrical activity produced by nerves, for the diagnosis and prog- nosis of neuromuscular disease.

- Classification. Class II (perform- ance standards).

##### § 890.1385   Diagnostic  electromyograph needle electrode.

- Identification. A diagnostic electromyograph needle electrode is a monopolar or bipolar needle  intended to be inserted into muscle or nerve tis- sue  to  sense  bioelectrical  signals.  The

device is intended for medical purposes for use in connection with electromyography (recording the in- trinsic electrical properties of skeletal muscle).

- Classification. Class II (perform- ance standards).

##### § 890.1450   Powered reflex hammer.

- Identification. A powered reflex hammer is a motorized device intended for medical purposes to elicit and de- termine controlled deep tendon re- flexes.

- Classification. Class II (perform- ance standards).

##### § 890.1575   Force-measuring platform.

- Identification. A force-measuring platform is a device intended for med- ical purposes that converts pressure applied upon a planar surface into ana- log mechanical or electrical signals. This device is used  to  determine ground reaction force, centers of per- cussion, centers of torque, and their variations in both magnitude and di- rection with  time.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38816,  July 25,

2001]

##### § 890.1600   Intermittent pressure meas- urement system.

- Identification. An intermittent pressure measurement system is an evaluative device intended for medical purposes, such as to measure the ac- tual pressure between the body surface and the supporting media.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38816,  July 25,

2001]

##### § 890.1615   Miniature   pressure   trans- ducer.

- Identification. A miniature pres- sure transducer is a device intended for

#### § 890.1850

medical purposes to measure the pres- sure between a device and soft tissue by converting mechanical inputs to analog electrical signals.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38816,  July 25,

2001]

##### § 890.1850   Diagnostic	muscle	stimu- lator.

- Identification. A diagnostic muscle stimulator is a device used mainly with an electromyograph machine to ini- tiate muscle activity. It is intended for medical purposes, such as to diagnose motor nerve or sensory neuromuscular disorders and neuromuscular function.

- Classification. Class II (perform- ance standards).

##### § 890.1925   Isokinetic  testing  and  eval- uation system.

- Identification. An isokinetic test- ing and evaluation system is a rehabili- tative exercise device intended for medical purposes, such as to measure, evaluate, and increase the strength of muscles and the range of motion of joints.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59230, Nov. 3, 1998]

## Subpart C [Reserved]

Subpart D—Physical Medicine Prosthetic Devices

##### § 890.3025   Prosthetic  and  orthotic  ac- cessory.

- Identification. A prosthetic and orthotic accessory is a device intended for medical purposes to support, pro- tect, or aid in the use of a cast, ortho- sis (brace), or prosthesis. Examples of prosthetic and orthotic accessories in- clude the following: A pelvic support band and belt, a cast shoe, a cast ban- dage, a limb cover, a prosthesis align-

21 CFR Ch. I (4–1–17 Edition)

ment device, a postsurgical pylon, a transverse rotator, and a temporary training splint.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3075   Cane.

- Identification. A cane is a device intended for medical purposes that is used to provide minimal weight sup- port while walking. Examples of canes include the following: A standard cane, a forearm cane, and a cane with a tri- pod, quad, or retractable stud on the ground end.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3100   Mechanical chair.

- Identification. A mechanical chair is a manually operated device intended for medical purposes that is used to as- sist a disabled person in performing an activity that the person would other- wise find difficult to do or be unable to do. Examples of mechanical chairs in- clude the following: A chair with an elevating seat used to raise a person from a sitting position to a standing position and a chair with casters used by a person to move from one place to another while sitting.

- Classification. Class I (general con- trols).  The  device  is  exempt  from  the

#### Food and Drug Administration, HHS	§ 890.3150

premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38816,  July 25,

2001]

##### § 890.3110   Electric positioning chair.

  - Identification. An electric posi- tioning chair is a device with a motor- ized positioning control that is in- tended for medical purposes and that can be adjusted to various positions. The device is used to provide stability for patients with athetosis (involun- tary spasms) and to alter postural posi- tions.

  - Classification. Class II. The elec- tric positioning chair is exempt from premarket notification procedures in subpart E of part 807 of this chapter, subject to § 890.9 and the following con- ditions for exemption:

- Appropriate analysis and non-clin- ical testing must demonstrate that the safety controls are adequate to ensure safe use of the device and prevent user falls from the device in the event of a device failure;

- Appropriate analysis and non-clin- ical testing must demonstrate the abil- ity of the device to withstand the rated user weight load with an appropriate factor of safety;

- Appropriate analysis and non-clin- ical testing must demonstrate the lon- gevity of the device to withstand exter- nal forces applied to the device and provide the user with an expected serv- ice life of the device;

- Appropriate analysis and non-clin- ical testing must demonstrate proper environments of use and storage of the device  to  maximize  the  longevity  of the device;

- Appropriate analysis and non-clin- ical  testing  (such  as  that  outlined   in the  currently  FDA-recognized  editions of ANSI/AAMI/ES60601–1, ‘‘Medical Electrical Equipment—Part 1: General Requirements for Basic Safety and Es- sential Performance,’’ and ANSI/AAMI/ IEC 60601–1–2, ‘‘Medical Electrical Equipment—Part 1–2: General Require- ments for Basic Safety and Essential Performance—Collateral	Standard: Electromagnetic	Disturbances—Re- quirements   and   Tests’’)   must   validate

electromagnetic	compatibility	and electrical safety;

- Appropriate analysis and non-clin- ical testing (such as that outlined  in the currently FDA-recognized editions of ANSI/AAMI/ISO 10993–1, ‘‘Biological Evaluation of Medical Devices—Part 1: Evaluation and Testing Within a Risk Management Process,’’ ANSI/AAMI/ISO 10993–5, ‘‘Biological Evaluation of Med- ical Devices—Part 5: Tests for In Vitro Cytotoxicity,’’ and ANSI/AAMI/ISO 10993–10, ‘‘Biological Evaluation of Medical Devices—Part 10: Tests for Ir- ritation and Skin Sensitization’’) must validate that the skin-contacting com- ponents of the device are biocompat- ible;

- Appropriate analysis and non-clin- ical  testing  (such  as  that  outlined   in the  currently  FDA-recognized  editions of IEC 62304, ‘‘Medical Device Soft- ware—Software Life Cycle Processes’’) must validate  the  software  life  cycle and that all processes, activities,  and tasks are implemented and  docu- mented;

- Appropriate analysis and non-clin- ical testing must validate that the de- vice components  are  found  to  be  non- flammable;

- Appropriate analysis and non-clin- ical testing must validate that the bat- tery in the device (if applicable) per- forms as intended over the anticipated service life of the device; and

- Adequate patient labeling is pro- vided to the user to document proper use and maintenance of the device to ensure safe use of the device by the pa- tient in the intended use environment.

[48 FR 53047, Nov. 23, 1983, as amended at 80

FR 72950, Nov. 20, 2015]

##### § 890.3150   Crutch.

- Identification. A crutch is a device intended for medical purposes for use by disabled persons to provide minimal to moderate weight support while walking.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation

#### § 890.3175

in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements    concerning    records    and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3175   Flotation cushion.

  - Identification. A flotation cushion is a device intended for medical pur- poses that is made of plastic, rubber, or other type of covering, that is filled with water, air, gel, mud, or any other substance allowing a flotation media, used on a seat to lessen the likelihood of skin ulcers.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38816,  July 25,

2001]

##### § 890.3410   External limb orthotic com- ponent.

- Identification. An external limb orthotic component is a device in- tended for medical purposes for use in conjunction with an orthosis (brace) to increase the function of the orthosis for a patient’s particular needs. Exam- ples of external limb orthotic compo- nents include the following: A brace- setting twister and  an external  brace stirrup.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3420   External	limb	prosthetic component.

- Identification. An external limb prosthetic component is a device in- tended for medical purposes that, when put  together  with  other  appropriate

#### 21 CFR Ch. I (4–1–17 Edition)

components, constitutes a total pros- thesis. Examples of external limb pros- thetic components include the fol- lowing: Ankle, foot, hip, knee, and socket components; mechanical or powered hand, hook, wrist unit, elbow joint, and shoulder joint components; and  cable  and  prosthesis  suction valves.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3450 Upper extremity prosthesis including a simultaneously pow- ered elbow and/or shoulder with greater than two simultaneous pow- ered degrees of freedom and con- trolled by  non-implanted electrical components.

- Identification. A upper extremity prosthesis including a simultaneously powered elbow and/or shoulder with greater than two simultaneous powered degrees of freedom and controlled by non-implanted electrical components, is a prescription device intended for medical purposes, and is intended to re- place a partially or fully amputated or congenitally absent upper extremity. It uses electronic inputs (other than sim- ple, manually controlled electrical components such as switches) to pro- vide greater than two independent and simultaneously powered degrees of freedom and includes a simultaneously powered elbow and/or shoulder. Pros- thetic arm components that are in- tended to be used as a system with other arm components must include all degrees of freedom of the total upper extremity prosthesis system.

- Classification. Class II (special controls). The special controls for this device are:

#### Food and Drug Administration, HHS	§ 890.3480

  - Appropriate analysis/testing must validate electronic compatibility, elec- trical safety, thermal safety, mechan- ical safety, battery performance and safety, and wireless performance, if ap- plicable.

  - Appropriate software verification, validation,  and  hazard  analysis  must be performed.

  - Non-clinical performance data must demonstrate that the device per- forms as intended under anticipated conditions of use. Performance testing must include:

- Mechanical bench data, including durability testing, to demonstrate that the device will withstand forces, condi- tions, and environments encountered during use.

- Simulated use testing to dem- onstrate performance of arm com- mands and available safeguard(s) under worst case conditions and after dura- bility   testing.

- Verification and validation of force sensors and hand release button, if applicable, are necessary.

- Device functionality in terms of flame retardant materials, liquid/par- ticle ingress prevention, sensor and ac- tuator performance, and motor and brake performance.

- The accuracy of the device fea- tures and safeguards.

  - Non-clinical and clinical perform- ance testing must demonstrate the ac- curacy of device features and safe- guards.

  - Elements of the device that may contact the patient must be dem- onstrated to be biocompatible.

  - Documented clinical experience and human factors testing must dem- onstrate safe and effective use, capture any adverse events observed during clinical use and demonstrate the accu- racy of device features and safeguards.

  - Labeling for the Prosthetist and User Guide must include:

    - Appropriate instructions, warning, cautions, limitations, and information related to the necessary safeguards of the device, including warning against activities that may put the user at greater risk (e.g., driving).

    - Specific instructions and the clin- ical training needed for the safe use of the device, which includes:

- Instructions on assembling the device in all available configurations,

- Instructions on fitting the pa- tient,

- Instructions and explanations of all available programs and how to pro- gram the device,

- Instructions and explanation of all controls, input, and outputs,

- Instructions on all available modes or states of the device,

- Instructions on all safety features of the device, and

- Instructions for maintaining the device.

    - Information on the patient popu- lation for which the device has been demonstrated to be effective.

    - A detailed summary of the non- clinical and clinical testing  pertinent to use of the device.

[81 FR 71612, Oct. 18, 2016]

##### § 890.3475   Limb orthosis.

- Identification. A limb orthosis (brace) is a device intended for medical purposes that is worn on the upper or lower extremities to support, to cor- rect, or to prevent deformities or to align body structures for functional improvement. Examples of limb orthoses include the following: A whole limb and joint brace, a hand splint, an elastic stocking, a knee cage, and a corrective shoe.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3480   Powered    lower    extremity exoskeleton.

- Identification. A powered lower ex- tremity exoskeleton is a prescription device that is composed of an external, powered, motorized orthosis that is placed over a person’s paralyzed or weakened limbs for medical purposes.

#### § 890.3480

- Classification. Class II (special controls). The special controls for this device are:

- Elements of the device materials that may contact the patient must be demonstrated to be biocompatible.

- Appropriate analysis/testing must validate electromagnetic compat- ibility/interference (EMC/EMI), elec- trical safety, thermal safety, mechan- ical safety, battery performance and safety, and wireless performance, if ap- plicable.

- Appropriate software verification, validation,  and  hazard  analysis  must be performed.

- Design characteristics must en- sure geometry and materials composi- tion are consistent with intended use.

- Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. Performance testing must include:

  - Mechanical bench testing (includ- ing durability testing) to demonstrate that the device will withstand forces, conditions, and environments encoun- tered during use;

  - Simulated use testing (i.e., cyclic loading testing) to demonstrate per- formance of device commands and safe- guard under worst case conditions and after  durability  testing;

  - Verification and validation of manual override controls are nec- essary, if present;

  - The accuracy of device features and safeguards; and

  - Device functionality in terms of flame retardant materials, liquid/par- ticle ingress prevention, sensor and ac- tuator performance, and motor per- formance.

- Clinical testing  must  demonstrate a reasonable assurance of safe and ef- fective use and capture any adverse events observed during clinical use when used under the proposed condi- tions of use, which must include con- siderations for:

  - Level of supervision necessary, and

  - Environment of use (e.g., indoors and/or outdoors) including obstacles and terrain representative of the in- tended use environment.

- A training program must be in- cluded with sufficient educational ele-

#### 21 CFR Ch. I (4–1–17 Edition)

ments so that upon completion of training program, the clinician, user, and companion can:

- Identify the safe environments for device use,

- Use all safety features of device, and

- Operate the device in  simulated or actual use environments representa- tive of indicated environments and use.

- Labeling for the Physician and User must include the following:

- Appropriate instructions, warning, cautions, limitations, and information related to the necessary safeguards of the device, including warning against activities and environments that may put the user at greater risk.

- Specific instructions and the clin- ical training needed for the safe use of the device, which includes:

- Instructions on assembling the device in all available configurations;

- Instructions on fitting the pa- tient;

- Instructions and explanations of all available programs and how to pro- gram the device;

- Instructions and explanation of all controls, input, and outputs;

- Instructions on all available modes or states of the device;

- Instructions on all safety features of the device; and

- Instructions for properly main- taining the device.

- Information on the patient popu- lation for which the device has been demonstrated to have a reasonable as- surance of safety and effectiveness.

- Pertinent non-clinical testing in- formation (e.g., EMC, battery lon- gevity).

- A detailed summary of the clin- ical testing including:

- Adverse events encountered under use conditions,

- Summary of study outcomes and endpoints, and

- Information pertinent to use of the device including the conditions under which the device was studied (e.g., level of supervision or assistance, and environment of use (e.g., indoors and/or outdoors) including obstacles and terrain).

[80 FR 25529, May 4, 2015]

#### Food and Drug Administration, HHS	§ 890.3640

##### § 890.3490   Truncal orthosis.

  - Identification. A truncal orthosis is a device intended for medical pur- poses to support or to immobilize frac- tures, strains, or sprains of the neck or trunk of the body. Examples of truncal orthoses are the following: Abdominal, cervical, cervical-thoracic, lumbar, lumbo-sacral, rib fracture, sacroiliac, and thoracic orthoses and clavicle splints.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38816, July 25, 2001]

##### § 890.3500   External   assembled   lower limb prosthesis.

- Identification. An external assem- bled lower limb prosthesis is a device that is intended for medical purposes and is a preassembled external artifi- cial limb for the lower extremity. Ex- amples of external assembled lower limb prostheses are the following: Knee/shank/ankle/foot assembly and thigh/knee/shank/ankle/foot  assembly.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.3520   Plinth.

- Identification. A plinth is a flat, padded board with legs that is intended for medical purposes. A patient is placed on the device for treatment or examination.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good   manufacturing   practice   require-

ments of the quality  system  regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements    concerning    records    and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.3610   Rigid   pneumatic   structure orthosis.

- Identification. A rigid pneumatic structure orthosis is a device intended for medical purposes to provide whole body support by means of a pressurized suit to help thoracic paraplegics walk.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of a PDP is required. A PMA or a notice of completion of a PDP is required to be filed with the Food and Drug Adminis- tration on or before December 26, 1996 for any rigid pneumatic structure or- thosis that was in commercial distribu- tion before May 28, 1976, or that has, on or before December 26, 1996 been found to  be  substantially  equivalent   to   a rigid pneumatic structure orthosis that was in commercial distribution  before May 28, 1976. Any other  rigid  pneu- matic structure orthosis shall have an approved PMA or a declared completed PDP in effect before being placed in commercial     distribution.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR 17742, May 11, 1987; 61 FR 50711, Sept. 27,

1996]

##### § 890.3640   Arm sling.

- Identification. An arm sling is a device intended for medical purposes to immobilize the arm, by means of a fab- ric band suspended from around the neck.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

#### § 890.3665

##### § 890.3665   Congenital   hip   dislocation abduction splint.

  - Identification. A congenital hip dislocation abduction splint is a device intended for medical purposes to sta- bilize the hips of a young child with dislocated hips in an abducted position (away from the midline).

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.3675   Denis Brown splint.

- Identification. A Denis Brown splint is a device intended for medical purposes to immobilize the foot. It is used on young children with tibial tor- sion (excessive rotation of the lower leg) or club foot.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.3690   Powered wheeled stretcher.

- Identification. A  powered wheeled stretcher is a battery-powered table with wheels that is intended for med- ical purposes for use by patients  who are unable to propel themselves inde- pendently and who must maintain a prone or supine position for prolonged periods because of skin ulcers or con- tractures  (muscle  contractions).

- Classification. Class II (perform- ance standards).

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 890.3700   Nonpowered	communica- tion system.

- Identification. A nonpowered com- munication system is a mechanical de- vice intended for medical purposes that is used to assist a patient in commu- nicating when physical impairment prevents writing, telephone use, read- ing, or talking. Examples of nonpow- ered communications systems include an alphabet board and a page turner.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 54

FR 25052, June 12, 1989; 66 FR 38817, July 25,

2001]

##### § 890.3710   Powered	communication system.

- Identification. A powered commu- nication system is an AC- or battery- powered device intended for medical purposes that is used to transmit or re- ceive information. It is used by persons unable to use  normal communication methods because of physical impair- ment. Examples of powered commu- nication systems include the following: a specialized typewriter, a reading ma- chine, and a video picture and word screen.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.3725   Powered	environmental control system.

- Identification. A powered environ- mental control system is an AC- or battery-powered device intended for medical purposes that is used by a pa- tient to operate an environmental con- trol function. Examples of environ- mental  control  functions  include  the

#### Food and Drug Administration, HHS	§ 890.3850

following: to control room tempera- ture, to answer a doorbell or telephone, or to sound an alarm for assistance.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.3750   Mechanical table.

- Identification. A mechanical table is a device intended for medical pur- poses that has a flat surface that can be inclined or adjusted to various posi- tions. It is used by patients with cir- culatory, neurological, or musculo- skeletal conditions to increase toler- ance to an upright or standing posi- tion.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38817,  July 25,

2001]

##### § 890.3760   Powered table.

- Identification. A powered table is a device intended for medical purposes that is an electrically operated flat surface table that can be adjusted to various positions. It is used by patients with circulatory, neurological, or mus- culoskeletal conditions to increase tol- erance to an upright or standing posi- tion.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38817,  July 25,

2001]

##### § 890.3790   Cane,   crutch,   and   walker tips and pads.

- Identification. Cane, crutch, and walker tips and pads are rubber (or rubber substitute) device accessories intended for medical purposes that are applied to the ground end of mobility aids to prevent skidding or that are ap- plied to the body contact area of the

device for comfort or as an aid in using an ambulatory assist device.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.3800   Motorized three-wheeled ve- hicle.

- Identification. A motorized three- wheeled vehicle is a gasoline-fueled or battery-powered device intended for medical purposes that is used for out- side transportation by disabled per- sons.

- Classification. Class II (perform- ance standards).

##### § 890.3825   Mechanical walker.

- Identification. A mechanical walk- er is a four-legged device with a metal frame intended for medical purposes to provide moderate weight support while walking. It is used by disabled persons who lack strength, good balance, or en- durance.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.3850   Mechanical wheelchair.

- Identification. A mechanical wheelchair is a manually operated de- vice with wheels that is intended for medical purposes to provide mobility to persons restricted to a sitting posi- tion.

#### § 890.3860

- Classification. Class I (general con- trols).

##### § 890.3860   Powered wheelchair.

- Identification. A powered wheel- chair is a battery-operated device with wheels that is intended for medical purposes to provide mobility to persons restricted to a sitting position.

- Classification. Class II (perform- ance standards).

##### § 890.3880   Special grade wheelchair.

- Identification. A special grade wheelchair is a device with wheels that is intended for medical purposes to pro- vide mobility to persons restricted to a sitting position. It is intended to be used in all environments for long-term use, e.g., for paraplegics, quadraplegics, and amputees.

- Classification. Class II (perform- ance standards).

##### § 890.3890   Stair-climbing wheelchair.

- Identification. A stair-climbing wheelchair is a device with wheels that is intended for medical purposes to pro- vide mobility to persons restricted to a sitting position. The device is intended to climb stairs.

- Classification. Class II (special controls). The special controls for this device are:

- The design characteristics of the device must ensure that the geometry and material composition are con- sistent with the intended use.

- Performance testing must dem- onstrate adequate mechanical perform- ance under simulated  use  conditions and environments. Performance testing must include the following:

- Fatigue testing;

- Resistance to dynamic loads (im- pact testing);

- Effective use of the braking mechanism and how the device stops in case of an electrical brake failure;

- Demonstration of adequate sta- bility of the device on inclined planes (forward, backward, and lateral);

- Demonstration of the ability of the device to safely ascend and descend obstacles (i.e., stairs, curb); and

- Demonstration of ability to ef- fectively use the device during adverse temperatures and following storage in

#### 21 CFR Ch. I (4–1–17 Edition)

adverse    temperatures    and    humidity conditions.

- The skin-contacting  components of the device must be demonstrated to be biocompatible.

- Software design, verification, and validation must demonstrate that the device controls, alarms, and user inter- faces function as intended.

- Appropriate analysis and perform- ance testing must be conducted to verify electrical safety and electro- magnetic compatibility of the device.

- Performance testing must dem- onstrate battery safety and evaluate longevity.

- Performance testing must evalu- ate the flammability of device compo- nents.

- Patient labeling must bear all in- formation required for the safe and ef- fective use of the device, specifically including the following:

  - A clear description of the techno- logical features of the device and the principles of how the device works;

  - A clear description of the appro- priate use environments/conditions, in- cluding prohibited environments;

  - Preventive maintenance rec- ommendations;

  - Operating specifications for prop- er use of the device such as patient weight limitations, device width, and clearance for maneuverability; and

  - A detailed summary of the device- related adverse events and how to re- port any complications.

- Clinician labeling must include all the information in the Patient labeling noted in paragraph (b)(8) of this section but must also include the following:

  - Identification of patients who can effectively operate the device; and

  - Instructions on how to fit, mod- ify, or calibrate the device.

- Usability studies of the device must demonstrate that the device  can be used by the patient in the intended use environment with the instructions for use and user training.

[79 FR 20782, Apr. 14, 2014]

##### § 890.3900   Standup wheelchair.

- Identification. A standup wheel- chair is a device with wheels that is in- tended for medical purposes to provide mobility to persons restricted to a sit- ting position. The device incorporates

#### Food and Drug Administration, HHS	§ 890.3930

an external manually controlled me- chanical system that is intended to raise a paraplegic to an upright posi- tion by means of an elevating seat.

- Classification. Class II (perform- ance standards).

##### § 890.3910   Wheelchair accessory.

- Identification. A wheelchair acces- sory is a device intended for medical purposes that is sold separately from a wheelchair and is intended to meet the specific needs of a patient who uses a wheelchair. Examples of wheelchair ac- cessories include but are not limited to the following: armboard, lapboard, pusher cuff, crutch and cane holder, overhead suspension sling, head and trunk support, and blanket and leg rest strap.

- Classification. Class I (general con- trols). If the device is not intended for use as a protective restraint as defined in § 880.6760 of this chapter, it is exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter,  subject  to  the  limitations  in

§ 890.9. The device is also exempt  from the current good  manufacturing  prac- tice requirements  of  the  quality  sys- tem regulation in part 820 of this chap- ter, with the exception of § 820.180, re- garding general requirements  con- cerning records, and § 820.198, regarding complaint    files.

[61 FR 8439, Mar. 4, 1996, as amended at 66 FR

38817, July 25, 2001]

##### § 890.3920   Wheelchair component.

- Identification. A wheelchair com- ponent is a device intended for medical purposes that  is generally  sold  as an integral part of a wheelchair, but may also be sold separately as a replace- ment part. Examples of wheelchair components are the following: Arm- rest, narrowing attachment, belt, ex- tension brake, curb climber, cushion, antitip device, footrest, handrim, hill holder, leg rest, heel loops, and toe loops.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR 63014,  Dec. 7,  1994;  66  FR 38817,  July 25,

2001]

##### § 890.3930   Wheelchair elevator.

- Permanently mounted wheelchair platform lift—(1) Identification. A perma- nently mounted wheelchair platform lift is a motorized vertical or inclined platform lift device permanently in- stalled in one location that is intended for use in mitigating mobility impair- ment caused by injury or other disease by providing a guided platform to move a person from one level to another, with or without a wheelchair.

- Classification. Class II. The perma- nently mounted wheelchair  platform lift is exempt from premarket  notifica- tion procedures in subpart E of part 807 of this chapter, subject to § 890.9 and the following conditions for exemption:

  - Appropriate analysis and nonclin- ical testing (such as that outlined  in the currently FDA-recognized edition of ASME A18.1 ‘‘Safety Standard for Platform Lifts and Stairway Chair Lifts’’) must demonstrate that the safety controls are adequate to prevent a free fall of the platform in the event of a device failure;

  - Appropriate analysis and nonclin- ical testing (such as that outlined  in the currently FDA-recognized edition of ASME A18.1 ‘‘Safety Standard for Platform Lifts and Stairway Chair Lifts’’) must demonstrate the ability of the device to withstand the rated load with an appropriate factor of safety;

  - Appropriate analysis and non- clinical testing (such as that outlined in the currently FDA-recognized edi- tion of ASME A18.1 ‘‘Safety Standard for Platform  Lifts and  Stairway Chair Lifts’’) must demonstrate the ability of the enclosures to prevent the user from falling from the device; and

  - Appropriate analysis and non- clinical testing (such as  that  outlined in the currently FDA-recognized edi- tions of AAMI/ANSI/IEC 60601–1–2, ‘‘Medical Electrical Equipment—Part 1–2: General Requirements for Safety— Collateral Standard: Electromagnetic Compatibility—Requirements	and Tests,’’     and     ASME     A18.1     ‘‘Safety

#### § 890.3940

Standard for Platform Lifts and Stair- way Chair Lifts’’) must validate elec- tromagnetic compatibility and elec- trical  safety.

- Portable wheelchair elevators—(1) Identification. A portable wheelchair el- evator is a motorized lift device that is not permanently  mounted in  one loca- tion and that is intended for use in mitigating   mobility    impairment caused by injury or other disease by providing a means to move a person, with or without a wheelchair, from one level to another (e.g., portable  plat- form lifts, attendant-operated stair climbing devices for wheelchairs).

(2) Classification. Class II.

[78 FR 14015, Mar. 4, 2013]

##### § 890.3940   Wheelchair platform scale.

- Identification. A wheelchair plat- form scale is a device with a base de- signed to accommodate a wheelchair. It is intended for medical purposes to weigh a person  who is confined  to a wheelchair.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR 63015,  Dec. 7,  1994;  66  FR 38817,  July 25,

2001]

## Subpart E [Reserved]

Subpart F—Physical Medicine Therapeutic Devices

##### § 890.5050   Daily activity assist device.

- Identification. A daily activity as- sist device is a modified adaptor or utensil (e.g., a dressing, grooming, rec- reational activity, transfer, eating, or homemaking aid) that is intended for medical purposes to assist a patient to perform a specific function.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

21 CFR Ch. I (4–1–17 Edition)

subpart E of part 807 of this chapter, subject to the limitations in § 890.9.  If the device is not labeled or otherwise represented as sterile,  the  device  is also exempt from the current good manufacturing   practice    requirements of  the  quality   system   regulation   in part 820 of this chapter, with the excep- tion of § 820.180, regarding general re- quirements     concerning     records     and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38817, July 25, 2001]

##### § 890.5100   Immersion hydrobath.

- Identification. An immersion hydrobath is a device intended for med- ical purposes that consists of water agitators and that may  include  a  tub to be filled with water. The water tem- perature may be measured by a gauge. It is used in hydrotherapy to relieve pain and itching and as an aid in the healing process of inflamed and trau- matized tissue, and it serves as a set- ting for removal of contaminated tis- sue.

- Classification. Class II (perform- ance standards).

##### § 890.5110   Paraffin bath.

- Identification. A paraffin bath is a device intended for medical purposes that consists of a tub to be filled with liquid paraffin (wax) and maintained at an elevated temperature in which the patient’s appendages (e.g., hands or fin- gers) are placed to relieve pain and stiffness.

- Classification. Class II (perform- ance standards).

##### § 890.5125   Nonpowered sitz bath.

- Identification. A nonpowered sitz bath is a device intended for medical purposes that consists of a tub to be filled with water for use in external hydrotherapy to relieve pain or pruritis and to accelerate the healing of inflamed or traumatized tissues of the perianal and perineal areas.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation

#### Food and Drug Administration, HHS	§ 890.5170

in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements    concerning    records    and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 54

FR 25052, June 12, 1989; 66 FR 38818, July 25,

2001]

##### § 890.5150   Powered patient transport.

  - Powered patient stairway chair lifts—(1) Identification. A powered pa- tient stairway chair lift is a motorized lift equipped with a seat and perma- nently mounted in one location that is intended for use in mitigating mobility impairment caused by injury or other disease by moving a person up and down a stairway.

- Classification. Class II. The stair- way chair lift is exempt from pre- market notification procedures in sub- part E of part 807 of this chapter, sub- ject to § 890.9 and the following condi- tions for exemption:

  - Appropriate analysis and nonclin- ical testing (such as that outlined  in the currently FDA-recognized edition of American Society of Mechanical En- gineers (ASME) A18.1 ‘‘Safety Standard for Platform  Lifts and  Stairway Chair Lifts’’) must demonstrate that the safety controls are adequate to prevent a free fall of the chair in the event of a device failure;

  - Appropriate analysis and nonclin- ical testing must demonstrate the abil- ity of the device,  including  armrests, to withstand the rated load with an ap- propriate factor of safety;

  - Appropriate restraints must be provided to prevent the user from fall- ing from the device (such as that out- lined in the currently FDA-recognized edition of ASME A18.1 ‘‘Safety Stand- ard for Platform Lifts and Stairway Chair  Lifts’’);

  - Appropriate analysis and non- clinical testing (such as  that  outlined in the currently FDA-recognized edi- tions of AAMI/ANSI/IEC 60601–1–2, ‘‘Medical Electrical Equipment—Part 1–2: General Requirements for Safety— Collateral Standard: Electromagnetic Compatibility—Requirements	and Tests,’’ and ASME A18.1 ‘‘Safety Standard for Platform Lifts and Stair- way Chair Lifts’’) must validate elec- tromagnetic compatibility and elec- trical  safety;  and

  - Appropriate analysis and nonclin- ical testing must demonstrate the re- sistance of the device upholstery to ig- nition.

  - All other powered patient trans- port—(1) Identification. A powered pa- tient transport is a motorized device intended for use in mitigating mobility impairment caused by injury or other disease by moving a person from one location or level to another, such as up and down flights of stairs (e.g., attend- ant-operated portable stair-climbing chairs). This generic type of device does not include motorized three- wheeled vehicles or wheelchairs.

- Classification. Class II.

[78 FR 14017, Mar. 4, 2013]

##### § 890.5160   Air-fluidized bed.

  - Identification. An air-fluidized bed is a device employing the circulation of filtered air through ceramic spherules (small, round ceramic objects) that is intended for medical purposes to treat or prevent bedsores, to treat severe or extensive burns, or to aid circulation.

  - Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.5170   Powered   flotation   therapy bed.

- Identification. A powered flotation therapy bed is a device that is equipped with a mattress that contains a large volume of constantly moving water, air, mud, or sand. It is intended for medical purposes to treat or prevent a patient’s bedsores, to treat severe or extensive burns, or to aid circulation. The mattress may be electrically heat- ed.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

#### § 890.5180

##### § 890.5180   Manual	patient	rotation bed.

  - Identification. A manual patient rotation bed is a device that turns a patient who is restricted to a reclining position. It is intended for medical pur- poses to treat or prevent bedsores, to treat severe and extensive burns, or to aid circulation.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1963, as amended at 65

FR 2322, Jan. 14, 2000]

##### § 890.5225   Powered    patient    rotation bed.

- Identification. A powered patient rotation bed is a device that turns a patient who is restricted to a reclining position. It is intended for medical pur- poses to treat or prevent bedsores, to treat severe and extensive burns, uri- nary tract blockage, and to aid circula- tion.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.5250   Moist steam cabinet.

- Identification. A moist steam cabi- net is a device intended for medical purposes that delivers a flow of heated, moisturized air to a patient in an en- closed unit. It is used to treat arthritis and fibrosis (a formation of fibrosis tis- sue) and to increase local blood flow.

- Classification. Class II (perform- ance standards).

##### § 890.5275   Microwave diathermy.

- Microwave diathermy for use in ap- plying therapeutic deep heat for selected medical conditions—(1) Identification. A microwave diathermy for use in apply- ing therapeutic deep heat for selected medical conditions is a device that ap- plies to specific areas of the body elec- tromagnetic energy in the microwave frequency bands of 915 megahertz to 2,450 megahertz and that is intended to generate  deep  heat  within  body  tissues

#### 21 CFR Ch. I (4–1–17 Edition)

for the treatment of selected medical conditions such as relief of pain, mus- cle spasms, and joint contractures, but not for the treatment of malignancies.

(2) Classification. Class II (perform- ance standards).

- Microwave diathermy for all other uses—(1) Identification. A microwave di- athermy  for  all  other  uses  except   for the treatment of malignancies is a de- vice that applies to the body electro- magnetic energy in the microwave fre- quency bands of 915 megahertz to 2,450 megahertz and that is intended for the treatment of medical conditions by means other than the  generation  of deep heat within body tissues as de- scribed  in  paragraph  (a)  of  this  section.

- Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or a notice of completion of a PDP for a device de- scribed in paragraph (b) of this section is required to be filed with the  Food and Drug Administration on or before July 13, 1999, for any microwave dia- thermy described in paragraph (b) of this section that was in commercial distribution before May 28, 1976, or that has, on or before July 13, 1999, been found to be substantially equivalent to a microwave diathermy described in paragraph (b) of this  section  that  was in commercial distribution before May 28, 1976. Any other microwave dia- thermy described in paragraph (b) of this section shall have  an  approved PMA or declared completed PDP in ef- fect before being placed in commercial distribution.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR  17742,  May  11,  1987;  64  FR  18331,  Apr.  14,

1999]

##### § 890.5290   Shortwave diathermy.

- Shortwave diathermy for use in ap- plying therapeutic deep heat for selected medical conditions—(1) Identification. A shortwave diathermy for use in apply- ing therapeutic deep heat for selected medical conditions is a device that ap- plies to specific areas of the body elec- tromagnetic energy in the radio- frequency (RF) bands of 13.56 mega- hertz (MHz) or 27.12 MHz and that is in- tended to generate deep heat within body tissues for the treatment of se- lected medical conditions such as relief

#### Food and Drug Administration, HHS	§ 890.5300

of pain, muscle spasms, and joint con- tractures, but not for the treatment of malignancies.

(2) Classification. Class II (perform- ance standards).

- Nonthermal shortwave therapy—(1) Identification. A nonthermal shortwave therapy is a prescription device that applies to the body pulsed electro- magnetic  energy  in  the  RF  bands  of

13.56 MHz or 27.12 MHz and that is in- tended for adjunctive use in the pallia- tive treatment of  postoperative  pain and edema of soft tissue  by  means other than the generation of deep heat within body tissues as described in paragraph (a) of this section.

(2) Classification: Class II (special con- trols). The  device  is  classified  as  class

- The special controls for this device are:

  - Components of the device that come into human contact must be dem- onstrated to be biocompatible.

  - Appropriate analysis/testing must demonstrate that the device is elec- trically safe and electromagnetically compatible in its intended use environ- ment.

  - Non-clinical performance testing must demonstrate that the device per- forms as intended under anticipated conditions of use. Non-clinical per- formance testing must characterize the output waveform of the device and demonstrate that the device meets ap- propriate output performance speci- fications. The output characteristics and the methods used to determine these characteristics, including the fol- lowing, must be determined:

- Peak output power;

- Pulse width;

- Pulse frequency;

- Duty cycle;

- Characteristics of other types of modulation that may be used;

- Average measured output pow- ered into the RF antenna/applicator;

- Specific absorption rates in sa- line gel test load or other appropriate model;

- Characterization of the electrical and magnetic fields in saline gel test load or other appropriate model for each RF antenna and prescribed RF an- tenna orientation/position; and

  - Characterization of the deposited energy density in saline gel test load or other appropriate model.

  - A detailed summary of the clin- ical testing pertinent to use of the de- vice to demonstrate the  effectiveness of the device in its intended use.

  - Labeling must include the fol- lowing:

    - Output characteristics of the de- vice;

    - Recommended treatment re- gimes, including duration of use; and

    - A detailed summary of the clin- ical testing pertinent to the use of the device and a summary of the adverse events and complications.

  - Nonthermal shortwave therapy devices marketed prior to the effective date of this reclassification must sub- mit an amendment to their previously cleared premarket notification (510(k)) demonstrating compliance with these special controls.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR  17742,  May  11,  1987;  80  FR  61302,  Oct.  13,

2015]

##### § 890.5300   Ultrasonic diathermy.

- Ultrasonic diathermy for use in ap- plying therapeutic deep heat for selected medical conditions—(1) Identification. An ultrasonic diathermy for use in apply- ing therapeutic deep heat for selected medical conditions is a device that ap- plies to specific areas of the body ultra- sonic energy at a frequency beyond 20 kilohertz and that is intended to gen- erate deep heat within body tissues for the treatment of selected medical con- ditions such as relief of pain, muscle spasms, and joint contractures, but not for the treatment of malignancies.

(2) Classification. Class II (perform- ance standards).

- Ultrasonic diathermy for all other uses—(1) Identification. An ultrasonic diathermy for all other uses except for the treatment of malignancies is a de- vice that applies to the body ultrasonic energy at a frequency beyond 20 kilo- hertz and that is intended for the treatment of medical conditions by means other than the  generation  of deep heat within body tissues as de- scribed in paragraph (a) of this section.

(2) Classification. Class III (premarket approval).

#### § 890.5350

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP for a device de- scribed in paragraph (b) of this section is required to be filed with the  Food and Drug Administration on or before July 13, 1999, for any ultrasonic dia- thermy described in paragraph (b) of this section that was in commercial distribution before May 28, 1976, or that has, on or before July 13, 1999, been found to be substantially equivalent to an ultrasonic diathermy described in paragraph (b) of this  section  that  was in commercial distribution before May 28, 1976. Any other ultrasonic dia- thermy described in paragraph (b) of this section shall have  an  approved PMA or declared completed PDP in ef- fect before being placed in commercial distribution.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR  17742,  May  11,  1987;  64  FR  18331,  Apr.  14,

1999]

##### § 890.5350   Exercise component.

- Identification. An exercise compo- nent is a device that is used in conjunc- tion with other forms of exercise and that is intended for medical purposes, such as to redevelope muscles or re- store motion to joints or for use as an adjunct treatment for obesity. Exam- ples include weights, dumbbells, straps, and adaptive hand mitts.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38818, July 25, 2001]

##### § 890.5360   Measuring   exercise   equip- ment.

- Identification. Measuring exercise equipment consist of manual devices intended for medical purposes, such as to redevelop muscles or restore motion to joints or for use as an adjunct treat- ment for obesity. These devices also in- clude   instrumentation,   such   as   the

#### 21 CFR Ch. I (4–1–17 Edition)

pulse rate monitor, that provide infor- mation used for physical evaluation and physical planning purposes., Exam- ples include a therapeutic exercise bi- cycle with measuring instrumentation, a manually propelled treadmill with measuring instrumentation, and a row- ing machine with measuring instru- mentation.

- Classification. Class II (perform- ance standards).

##### § 890.5370   Nonmeasuring	exercise equipment.

- Identification. Nonmeasuring exer- cise equipment consist of devices in- tended for medical purposes, such as to redevelop muscles or restore motion to joints or for use as an adjunct treat- ment for obesity. Examples include a prone scooter board, parallel bars, a mechanical treadmill,  an  exercise table, and a manually propelled exer- cise  bicycle.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38818, July 25, 2001]

##### § 890.5380   Powered	exercise	equip- ment.

- Identification. Powered exercise equipment consist of powered devices intended for medical purposes, such as to redevelop muscles or restore motion to joints or for use as an adjunct treat- ment for obesity. Examples include a powered treadmill, a powered bicycle, and powered parallel bars.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

#### Food and Drug Administration, HHS	§ 890.5525

##### § 890.5410   Powered finger exerciser.

  - Identification. A powered finger ex- erciser is a device intended for medical purposes to increase flexion and the ex- tension range of motion of the joints of the second to the fifth fingers of the hand.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

##### § 890.5500   Infrared lamp.

- Identification. An  infrared  lamp  is a device intended for medical purposes that emits energy at infrared fre- quencies (approximately 700 nano- meters to 50,000 nanometers) to provide topical    heating.

- Classification. Class II (perform- ance standards).

##### § 890.5525   Iontophoresis device.

- Iontophoresis device intended for certain specified uses—(1) Identification. An iontophoresis device is a device that is intended to use a direct current to introduce ions of soluble salts or other drugs into the body and induce sweating for use in the diagnosis of cystic fibrosis or for other uses if the labeling of the drug intended for use with the device bears adequate direc- tions for the device’s use with that drug. When used in the diagnosis of cystic fibrosis, the sweat is collected and its composition and weight are de- termined.

(2) Classification. Class II (perform- ance standards).

- Iontophoresis device intended for any other purposes—(1) Identification. An iontophoresis device intended for any other purposes is a prescription device that is intended to use a current to in- troduce ions of drugs or non-drug solu- tions into the body for medical pur- poses other than those specified in paragraph (a) of this section, meaning that the device is not intended for use in diagnosis of cystic fibrosis, or a spe- cific drug is not specified in the label- ing of the iontophoresis device.

(2) Classification. Class II (special con- trols).  The  device  is  classified  as  class

- The special controls for this device are:

  - The following performance testing must be conducted:

- Testing using a drug approved for iontophoretic delivery, or a solution, if identified in the labeling, to dem- onstrate safe use of the device as in- tended;

- Testing of the ability of the de- vice to maintain a safe pH level; and

- If used in the ear, testing of the device to demonstrate mechanical safe- ty.

  - Labeling must include adequate instructions for use, including suffi- cient information for the health care provider to determine the device char- acteristics that affect delivery of the drug or solution and to select appro- priate drug or solution dosing informa- tion for administration by ionto- phoresis. This includes the following:

    - A description and/or graphical representation  of  the  electrical  output;

    - A description of the electrode ma- terials and pH buffer;

    - When intended for general drug delivery, language referring the user to drug	labeling	approved	for iontophoretic delivery to determine if the drug they intend to deliver is spe- cifically approved for use with that type of device and to obtain relevant dosing information; and

    - A detailed summary of the de- vice-related and procedure-related complications pertinent to  use of  the device, and appropriate warnings and contraindications, including the fol- lowing warning:

Warning: Potential systemic adverse effects may result from use of this de- vice. Drugs or solutions delivered with this device have the potential to reach the blood stream and cause systemic effects. Carefully read all labeling of the drug or solution used with this de- vice to understand all potential ad- verse effects and to ensure appropriate dosing information. If systemic mani- festations occur, refer to the drug or solution labeling for appropriate ac- tion.

  - Appropriate analysis/testing must     demonstrate     electromagnetic

#### § 890.5575

compatibility,  electrical  safety,  ther- mal safety, and mechanical safety.

  - Appropriate	software verification, validation, and hazard analysis must be performed.

  - The elements of the device that may contact the patient must be dem- onstrated to be biocompatible.

  - The elements of the device that may contact the patient must be as- sessed for sterility, for  devices  labeled as  sterile.

  - Performance data must support the shelf life of the elements of the de- vice that may be affected by aging by demonstrating continued package in- tegrity and device functionality over the stated shelf life.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR 17742,  May  11,  1987;  81  FR 48706,  July 26,

2016]

##### § 890.5575   Powered external limb over- load warning device.

- Identification. A powered external limb overload warning device is a de- vice intended for medical purposes to warn a patient of an overload or an underload in the amount of pressure placed on a leg.

- Classification. Class II (perform- ance standards).

##### § 890.5650   Powered	inflatable	tube massager.

- Identification. A powered inflat- able tube massager is a powered device intended for medical purposes, such as to relieve minor muscle aches and pains and to increase circulation. It simulates kneading and stroking of tis- sues with the hands by use of an inflat- able pressure cuff.

- Classification. Class II (perform- ance standards).

##### § 890.5660   Therapeutic massager.

- Identification. A therapeutic mas- sager is an electrically powered device intended for medical purposes, such as to relieve minor muscle aches and pains.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

##### § 890.5700   Cold pack.

- Identification. A cold pack is a de- vice intended for medical purposes that consists of a compact fabric envelope containing a specially hydrated pliable silicate gel capable of forming to the contour of the body and that provides cold therapy for body surfaces.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807. The device also is exempt from the current good manu- facturing practice requirements of the quality system regulation in part 820, with the exception of § 820.180, with re- spect to general requirements con- cerning records, and § 820.198, with re- spect  to  complaint  files.

##### § 890.5710   Hot or cold disposable pack.

- Identification. A hot or cold dis- posable pack is a device intended for medical purposes that consists of a sealed plastic bag incorporating chemi- cals that, upon activation, provides hot or cold therapy for body surfaces.

- Classification. Class I (general con- trols). Except when intended for use on infants, the device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1963, as amended at 65

FR 2322, Jan. 14, 2000]

##### § 890.5720   Water   circulating   hot   or cold pack.

- Identification. A water circulating hot or cold pack is a device intended for medical purposes that operates by pumping heated or chilled water through a plastic bag and that provides hot or cold therapy for body surfaces.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

#### Food and Drug Administration, HHS	§ 890.5860

##### § 890.5730   Moist heat pack.

  - Identification. A moist heat  pack is a device intended for medical pur- poses that consists of silica gel in a fabric container used to retain an ele- vated temperature and that provides moist heat therapy for body surfaces.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 66

FR 38818, July 25, 2001]

##### § 890.5740   Powered heating pad.

- Identification. A powered heating pad is an electrical device intended for medical purposes that provides dry heat therapy for body surfaces. It is ca- pable of maintaining an elevated tem- perature during use.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E part 807 of this chapter subject to § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 63

FR 59231, Nov. 3, 1998]

##### § 890.5760   Nonpowered  lower  extrem- ity pressure wrap.

- Identification. A nonpowered lower extremity pressure wrap is a prescrip- tion device that applies mechanical pressure by wrapping around the lower extremity, such as the leg or foot, and is intended for primary Restless Leg Syndrome.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[79 FR 37950, July 3, 2014]

##### § 890.5765   Pressure-applying device.

- Identification. A pressure-applying device is a device intended for medical purposes to apply continuous pressure

to the paravertebral tissues for mus- cular relaxation and  neuro-inhibition. It consists of a table with an adjustable overhead weight that, in place of the therapist’s hands, presses on the back of a prone patient.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 59

FR 63015,  Dec. 7,  1994;  66  FR 38818,  July 25,

2001]

##### § 890.5850   Powered muscle stimulator.

- Identification. A powered muscle stimulator is an electrically powered device intended for medical purposes that repeatedly contracts muscles by passing electrical currents through electrodes contacting the affected body area.

- Classification. Class II (perform- ance standards).

##### § 890.5860   Ultrasound	and	muscle stimulator.

- Ultrasound and muscle stimulator for use in applying therapeutic deep heat for selected medical conditions—(1) Identi- fication. An ultrasound and muscle stimulator for use in applying thera- peutic deep heat for selected medical conditions is a device that applies to specific areas of the body ultrasonic energy at a frequency beyond 20 kilo- hertz and that is intended to generate deep heat within body tissues for the treatment of selected medical condi- tions such as relief of pain, muscle spasms, and joint contractures, but not for the treatment of malignancies. The device also passes electrical currents through the body area to stimulate or relax  muscles.

(2) Classification. Class II (perform- ance standards).

- Ultrasound and muscle stimulator for all other uses—(1) Identification. An ultrasound and muscle stimulator for all other uses except for the treatment of malignancies is a device that applies to the body ultrasonic energy at a fre- quency beyond 20 kilohertz and applies to the body  electrical  currents  and that is intended for the treatment of medical   conditions   by   means   other

#### § 890.5880

than the generation of deep heat with- in body tissues and the stimulation or relaxation of muscles as described in paragraph (a) of this section.

(2) Classification. Class III (premarket approval).

- Date PMA or notice of completion of PDP is required. A PMA or notice of completion of a PDP for a device de- scribed in paragraph (b) of this section is required to be filed with the Food and Drug Administration on or before July 13, 1999 for  any  ultrasound  and muscle stimulator described in para- graph (b) of this section that was in commercial distribution before May 28, 1976, or that has, on or before July 13, 1999, been found to be substantially equivalent to an ultrasound and muscle stimulator described in paragraph (b) of this section that was in commercial distribution before May 28, 1976. Any other ultrasound and muscle stimu- lator described in paragraph (b) of this section shall have an approved PMA or declared completed PDP in effect  be- fore being placed in commercial dis- tribution.

[48 FR 53047, Nov. 23, 1983, as amended at 52

FR  17742,  May  11,  1987;  64  FR  18331,  Apr.  14,

1999]

##### § 890.5880   Multi-function	physical therapy table.

- Identification. A multi-function physical therapy  table  is  a  device in- tended for medical purposes that con- sists of a motorized table equipped to provide patients with heat, traction, and muscle relaxation therapy.

- Classification. Class II (perform- ance standards).

##### § 890.5900   Power traction equipment.

- Identification. Powered traction equipment consists of powered devices intended for medical purposes for use in conjunction with traction acces- sories, such as belts and harnesses, to exert therapeutic pulling forces on the patient’s body.

- Classification. Class II (perform- ance standards).

##### § 890.5925   Traction accessory.

- Identification. A traction acces- sory is a nonpowered accessory device intended for medical purposes to be used with powered traction equipment

#### 21 CFR Ch. I (4–1–17 Edition)

to aid in exerting therapeutic pulling forces on the patient’s body. This ge- neric type of device includes the pul- ley, strap, head halter, and pelvic belt.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, regarding general requirements  concerning  records  and

§ 820.198, regarding complaint files.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

##### § 890.5940   Chilling unit.

- Identification. A chilling unit is a refrigerative device intended for med- ical purposes to chill and maintain cold packs at a reduced temperature.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

##### § 890.5950   Powered heating unit.

- Identification. A powered heating unit is a device intended for medical purposes that consists of an encased cabinet containing hot water and that is intended to heat and maintain hot packs at an elevated temperature.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

##### § 890.5975   Therapeutic vibrator.

- Identification. A therapeutic vibra- tor is an electrically powered device in- tended for medical purposes that incor- porates various kinds of pads and that is held in the hand or attached to the hand or  to  a table.  It  is intended  for

#### Food and Drug Administration, HHS	Pt. 892

various  uses,  such  as  relaxing  muscles and relieving minor aches and pains.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 890.9.

[48 FR 53047, Nov. 23, 1983, as amended at 61

FR 1125,  Jan. 16,  1996;  66  FR 38818,  July 25,

2001]

## PART 892—RADIOLOGY DEVICES

##### Subpart A—General Provisions

Sec.

892.1    Scope.

892.3   Effective dates of requirement for pre- market  approval.

  - Limitations of exemptions from sec- tion 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

##### Subpart B—Diagnostic Devices

892.1000    Magnetic  resonance  diagnostic  de- vice.

892.1100   Scintillation (gamma) camera. 892.1110  Positron camera.

892.1130   Nuclear whole body counter. 892.1170   Bone densitometer.

892.1180  Bone sonometer.

892.1200   Emission     computed     tomography system.

892.1220   Fluorescent scanner. 892.1300    Nuclear rectilinear scanner.

892.1310    Nuclear tomography system. 892.1320   Nuclear uptake probe.

892.1330     Nuclear whole body scanner. 892.1350     Nuclear scanning bed.

892.1360    Radionuclide dose calibrator. 892.1370    Nuclear anthropomorphic phantom. 892.1380   Nuclear flood source phantom.

892.1390    Radionuclide rebreathing system. 892.1400    Nuclear sealed calibration source. 892.1410    Nuclear     electrocardiograph     syn-

chronizer.

892.1420    Radionuclide test pattern phantom. 892.1540    Nonfetal ultrasonic monitor.

892.1550    Ultrasonic  pulsed  doppler  imaging system.

892.1560    Ultrasonic pulsed echo imaging sys- tem.

892.1570    Diagnostic ultrasonic transducer. 892.1600    Angiographic x-ray system.

892.1610    Diagnostic  x-ray  beam-limiting  de- vice.

892.1620   Cine   or   spot   fluorographic   x-ray camera.

892.1630 Electrostatic x-ray imaging system. 892.1640 Radiographic film marking system. 892.1650    Image-intensified    fluoroscopic    x-

ray system.

892.1660    Non-image-intensified    fluoroscopic x-ray   system.

892.1670   Spot-film device.

892.1680  Stationary x-ray system.

892.1700    Diagnostic  x-ray  high  voltage  gen- erator.

892.1710    Mammographic x-ray system. 892.1715    Full-field     digital     mammography

system.

892.1720    Mobile x-ray system.

892.1730    Photofluorographic x-ray system. 892.1740    Tomographic x-ray system.

892.1750   Computed   tomography   x-ray   sys- tem.

892.1760    Diagnostic  x-ray  tube  housing  as- sembly.

892.1770    Diagnostic x-ray tube mount. 892.1820    Pneumoencephalographic chair. 892.1830    Radiologic patient cradle.

892.1840   Radiographic film.

892.1850    Radiographic film cassette.

892.1860    Radiographic film/cassette changer. 892.1870    Radiographic  film/cassette  changer

programmer.

892.1880    Wall-mounted radiographic cassette holder.

892.1890     Radiographic film illuminator. 892.1900    Automatic  radiographic  film  proc-

essor.

892.1910   Radiographic grid.  892.1920   Radiographic head holder.

892.1940    Radiologic    quality    assurance    in- strument.

892.1950  Radiographic	anthropomorphic phantom.

892.1960    Radiographic intensifying screen. 892.1970   Radiographic   ECG/respirator   syn-

chronizer.

892.1980   Radiologic table.

892.1990    Transilluminator for breast evalua- tion.

892.2010    Medical image storage device. 892.2020    Medical  image  communications  de-

vice.

892.2030    Medical image digitizer.  892.2040    Medical image hardcopy device.

892.2050    Picture  archiving  and  communica- tions   system.

##### Subparts C–E [Reserved]

Subpart F—Therapeutic Devices

892.5050    Medical   charged-particle   radiation therapy   system.

892.5300   Medical  neutron  radiation  therapy system.

892.5650    Manual radionuclide applicator sys- tem.

892.5700    Remote controlled radionuclide ap- plicator    system.

892.5710    Radiation     therapy     beam-shaping block.

892.5730    Radionuclide brachytherapy source. 892.5740    Radionuclide teletherapy source.

892.5750    Radionuclide radiation therapy sys- tem.

892.5770    Powered  radiation  therapy  patient support   assembly.

§ 892.1

892.5780    Light  beam  patient  position  indi- cator.

892.5840    Radiation  therapy  simulation  sys- tem.

892.5900    X-ray radiation therapy system. 892.5930    Therapeutic  x-ray  tube  housing  as-

sembly.

##### Subpart G—Miscellaneous Devices

892.6500   Personnel protective shield. AUTHORITY:  21  U.S.C.  351,  360,  360c,  360e,

360j, 371.

SOURCE: 53  FR 1567,  Jan. 20,  1988,  unless

otherwise noted.

EDITORIAL NOTE: Nomenclature changes  to part 892 appear at 73 FR 35341, June 23, 2008.

## Subpart A—General Provisions

##### § 892.1   Scope.

    - This part sets forth the classifica- tion of radiology devices intended for human use that are in commercial dis- tribution.

    - The identification of a device in a regulation in this part is not a precise description of every device that is, or will be, subject to the regulation. A manufacturer who submits a pre- market notification submission for a device under part 807 cannot show merely that the device is accurately described by the section title and iden- tification provision of a regulation in this part but shall state why the device is substantially equivalent to other de- vices, as required by § 807.87.

    - To avoid duplicative listings, a ra- diology device that has two or more types of uses (e.g., use both as a diag- nostic device and a therapeutic device) is listed in one subpart only.

    - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of this title 21, unless otherwise noted.

    - Guidance documents referenced in this  part  are  available  on  the  Internet..

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 73

FR 40969, July 17, 2008; 78 FR 18233, Mar. 26,

2013]

##### § 892.3   Effective  dates  of  requirement for premarket approval.

A device included in this part that is classified into class III (premarket ap-

21 CFR Ch. I (4–1–17 Edition)

proval) shall not be commercially dis- tributed after the date shown in the regulation classifying the device unless the manufacturer has an  approval under section 515 of the act (unless an exemption has been granted under sec- tion 520(g)(2) of the act). An approval under section 515 of the act consists of FDA’s issuance of an order  approving an application for premarket approval (PMA) for the device or declaring com- pleted a product development protocol (PDP) for the device.

- Before FDA requires that a device commercially distributed before the enactment date of the amendments, or a device that has been found substan- tially equivalent to such a  device,  has an approval under section 515 of the act, FDA must promulgate a  regula- tion under section 515(b) of the act re- quiring such approval, except as pro- vided in paragraph (b) of this section. Such a regulation under section 515(b) of the act shall not be effective during the grace period ending on the 90th day after its promulgation or on the last day of the 30th full calendar month after the regulation that classifies the device into class III is effective, which- ever is later. See section 501(f)(2)(B) of the act. Accordingly, unless an effec- tive date of the requirement for pre- market approval is shown in the regu- lation for a device  classified  into  class III in this part, the device may be com- mercially distributed without FDA’s issuance of an order approving a PMA or declaring completed a PDP for the device. If FDA promulgates a regula- tion under section 515(b) of the act re- quiring premarket approval for a de- vice, section 501(f)(1)(A) of the act ap- plies to the device.

- Any new, not substantially equiv- alent, device introduced into commer- cial distribution on or after May 28, 1976, including a device formerly mar- keted that has been substantially al- tered, is classified by statute (section 513(f) of the act) into class III without any grace period and FDA must have issued an order approving a PMA or de- claring completed a PDP for the device before the device is commercially dis- tributed unless it  is  reclassified.  If FDA knows that a device being com- mercially  distributed  may  be  a  ‘‘new’’

#### Food and Drug Administration, HHS	§ 892.1000

device as defined in this  section  be- cause of any new intended use or other reasons, FDA may codify the statutory classification  of  the  device   into   class III for such new use. Accordingly, the regulation for such a class  III  device states that as of the enactment  date  of the amendments, May 28, 1976, the de- vice must have an approval under sec- tion 515 of the act before commercial distribution.

##### § 892.9 Limitations of exemptions from section 510(k) of the Federal Food, Drug, and Cosmetic Act (the act).

The exemption from the requirement of premarket notification (section 510(k) of the act) for a generic type of class I or II device is only to the extent that the device has existing or reason- ably foreseeable characteristics of commercially distributed devices with- in that generic type or, in the case of in vitro diagnostic devices, only to the extent that misdiagnosis as a result of using the device would not be associ- ated with high morbidity or mortality. Accordingly, manufacturers of any commercially distributed class I or II device for which FDA has granted an exemption from the requirement of premarket notification must still sub- mit a premarket notification to FDA before introducing or delivering for in- troduction into interstate commerce for commercial distribution the device when:

  - The device is intended for a use different from the intended use of a le- gally marketed device in that generic type of device; e.g., the device is in- tended for a different medical purpose, or the device is intended for lay use where the former intended use was by health care professionals only;

  - The modified device operates using a different fundamental  sci- entific technology than a legally mar- keted device in that generic type of de- vice; e.g., a surgical instrument cuts tissue with a laser beam rather than with a sharpened metal blade, or an in vitro diagnostic device detects or iden- tifies infectious agents by using deoxyribonucleic acid (DNA) probe or nucleic acid hybridization technology rather than culture or immunoassay technology; or

  - The device is an in vitro device that is intended:

    - For use in the diagnosis, moni- toring, or screening of neoplastic dis- eases with the exception of immunohistochemical  devices;

    - For use in  screening or diagnosis of familial or acquired genetic dis- orders, including inborn errors of me- tabolism;

    - For measuring an analyte that serves as a surrogate marker for screening, diagnosis, or monitoring life-threatening diseases such as ac- quired immune deficiency syndrome (AIDS), chronic or active hepatitis, tu- berculosis, or myocardial infarction or to monitor therapy;

    - For assessing the risk of cardio- vascular diseases;

    - For use in diabetes management;

    - For identifying or inferring the identity of a microorganism directly from clinical material;

    - For detection of antibodies to microorganisms	other	than immunoglobulin G (IgG) or IgG assays when the results are not qualitative, or are used to determine immunity, or the assay is  intended for use  in matrices other than serum or plasma;

    - For noninvasive testing as defined in § 812.3(k) of this chapter; and

    - For near patient testing (point of care).

[65 FR 2322, Jan. 14, 2000]

## Subpart B—Diagnostic Devices

##### § 892.1000   Magnetic    resonance    diag- nostic device.

- Identification. A magnetic reso- nance diagnostic device is intended for general diagnostic use to present im- ages which reflect the spatial distribu- tion and/or magnetic resonance spectra which reflect frequency and distribu- tion of nuclei exhibiting nuclear mag- netic resonance. Other physical param- eters derived from the images and/or spectra may also be produced. The de- vice includes hydrogen-1 (proton) imag- ing, sodium-23 imaging, hydrogen-1 spectroscopy, phosphorus-31 spectros- copy, and chemical shift imaging (pre- serving simultaneous frequency and spatial  information).

§ 892.1100

- Classification. Class II.

[53 FR 5078, Feb. 1, 1989]

##### § 892.1100   Scintillation  (gamma)  cam- era.

- Identification. A scintillation (gamma) camera is a device intended to image the distribution of radio- nuclides in the body by means of a pho- ton radiation detector. This generic type of device may include signal anal- ysis and display equipment, patient and equipment supports, radionuclide anatomical markers, component parts, and accessories.

- Classification. Class I (general con- trols).

[55 FR 48443, Nov. 20, 1990, as amended at 66

FR 46953, Sept. 10, 2001]

##### § 892.1110   Positron camera.

- Identification. A positron camera is a device intended to image the dis- tribution of positron-emitting radio- nuclides in the body. This generic type of device may include signal analysis and display equipment, patient and equipment supports, radionuclide ana- tomical markers, component parts, and accessories.

- Classification. Class I (general con- trols).

[55 FR 48444, Nov. 20, 1990, as amended at 66

FR 46953, Sept. 10, 2001]

##### § 892.1130   Nuclear	whole	body counter.

- Identification. A nuclear whole body counter is a device intended to measure the amount of radionuclides in the entire body. This generic type of device may include signal analysis and display equipment, patient and equip- ment supports, component parts, and accessories.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 59

FR 63015,  Dec. 7,  1994;  66  FR 38818,  July 25,

2001]

[55 FR 48444, Nov. 20, 1990]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 892.1170   Bone densitometer.

- Identification. A bone densitom- eter is a device intended for medical purposes to measure bone density and mineral content by x-ray or gamma ray	transmission	measurements through the bone and adjacent tissues. This generic type of device may in- clude signal analysis and display equip- ment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

##### § 892.1180   Bone sonometer.

- Identification. A bone sonometer is a device that transmits ultrasound en- ergy into the human body to measure acoustic properties of bone that indi- cate overall bone health and fracture risk. The primary components of the device are a voltage generator, a trans- mitting transducer, a receiving trans- ducer, and hardware and software for reception and processing of the re- ceived ultrasonic signal.

- Classification. Class II (special controls). The special control for this device is FDA’s ‘‘Guidance for Industry and FDA  Staff; Class  II Special  Con- trols Guidance Document: Bone Sonometers.’’ See § 892.1(e) for the availability  of  this  guidance  document.

[73 FR 40969, July 17, 2008]

##### § 892.1200   Emission  computed  tomog- raphy system.

- Identification. An emission com- puted tomography system is a device intended to detect the location and dis- tribution of gamma ray- and positron- emitting radionuclides in the body and produce cross-sectional images through computer reconstruction of the data. This generic type of device may in- clude signal analysis and display equip- ment, patient and equipment supports, radionuclide anatomical markers, com- ponent parts, and accessories.

- Classification. Class II.

##### § 892.1220   Fluorescent scanner.

- Identification. A fluorescent scan- ner is a device intended to measure the induced fluorescent radiation in the body by exposing the body to certain x- rays or low-energy gamma rays. This

#### Food and Drug Administration, HHS	§ 892.1370

generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts and accessories.

- Classification. Class II.

##### § 892.1300   Nuclear rectilinear scanner.

- Identification. A nuclear recti- linear scanner is a device intended to image the distribution of radionuclides in the body by means of a detector (or detectors) whose position moves in two directions with respect to the patient. This generic type of device may in- clude signal analysis and display equip- ment, patient and equipment supports, radionuclide anatomical markers, com- ponent parts, and accessories.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2322,  Jan. 14,  2000;  66  FR 38818,  July 25,

2001]

##### § 892.1310   Nuclear tomography system.

- Identification. A nuclear tomog- raphy system is a device intended to detect nuclear radiation in the body and produce images of a specific cross- sectional plane of the body by blurring or eliminating detail from other planes. This generic type of devices may include signal analysis and dis- play equipment, patient and equipment supports, radionuclide anatomical markers, component parts, and acces- sories.

- Classification. Class II.

##### § 892.1320   Nuclear uptake probe.

- Identification. A nuclear uptake probe is a device intended to measure the amount of radionuclide taken  up by a particular organ or body region. This generic type of device may in- clude a single or multiple detector probe, signal analysis and display equipment, patient and equipment sup- ports, component parts, and acces- sories.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

subpart  E  of  part  807  of  this  chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2322, Jan. 14, 2000]

##### § 892.1330   Nuclear  whole  body  scan- ner.

- Identification. A nuclear whole body scanner is a device intended to measure and image the distribution of radionuclides in the body by means of a wide-aperture detector whose posi- tion moves in one direction with re- spect to the patient. This generic type of device may include signal analysis and display equipment, patient and equipment supports, radionuclide ana- tomical markers, component parts, and accessories.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2322, Jan. 14, 2000]

##### § 892.1350   Nuclear scanning bed.

- Identification. A nuclear scanning bed is an adjustable bed intended to support a patient during a nuclear medicine procedure.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 59

FR  63015,  Dec.  7,  1994;  65  FR  2322,  Jan.  14,

2000]

##### § 892.1360   Radionuclide	dose	cali- brator.

- Identification. A radionuclide dose calibrator is a radiation detection de- vice intended to assay radionuclides before their administration to patients.

- Classification. Class II.

##### § 892.1370   Nuclear	anthropomorphic phantom.

- Identification. A nuclear anthropomorphic phantom is a human tissue facsimile that contains a radio- active source or a cavity in which a ra- dioactive sample can be inserted. It is intended to calibrate nuclear uptake probes or other medical instruments.

#### § 892.1380

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 54

FR 13832,  Apr. 5,  1989;  66  FR 38818,  July 25,

2001]

##### § 892.1380   Nuclear  flood  source  phan- tom.

- Identification. A nuclear flood source phantom is a device that con- sists of a radiolucent container filled with a uniformly distributed  solution of a desired radionuclide. It is intended to calibrate a medical gamma camera- collimator system for uniformity of re- sponse.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 54

FR 13832,  Apr. 5,  1989;  66  FR 38819,  July 25,

2001]

##### § 892.1390   Radionuclide	rebreathing system.

- Identification. A radionuclide re- breathing system is a device intended to be used to contain a gaseous or vola- tile radionuclide or a radionuclide-la- beled aerosol and permit it to be re- spired by the patient during nuclear medicine ventilatory tests (testing process of exchange between the lungs and the atmosphere). This generic type of device may include signal analysis and display equipment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

##### § 892.1400   Nuclear   sealed   calibration source.

- Identification. A nuclear sealed calibration source is a device that con- sists of an encapsulated reference radionuclide intended for calibration of medical nuclear radiation detectors.

- Classification. Class I (general con- trols). The device is exempt from the premarket   notification   procedures   in

#### 21 CFR Ch. I (4–1–17 Edition)

subpart  E  of  part  807  of  this  chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 54

FR 13832,  Apr. 5,  1989;  66  FR 38819,  July 25,

2001]

##### § 892.1410   Nuclear   electrocardiograph synchronizer.

- Identification. A nuclear electro- cardiograph synchronizer is a device intended for use in nuclear radiology to relate the time of image formation to the cardiac cycle during the production of dynamic cardiac images.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2322, Jan. 14, 2000]

##### § 892.1420   Radionuclide   test   pattern phantom.

- Identification. A radionuclide test pattern phantom is a device that con- sists of an arrangement of radiopaque or radioactive material sealed  in  a solid pattern intended to serve as a test for a performance characteristic of a nuclear medicine imaging device.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 54

FR 13832,  Apr. 5,  1989;  66  FR 38819,  July 25,

2001]

##### § 892.1540   Nonfetal ultrasonic monitor.

- Identification. A nonfetal ultra- sonic monitor is a device that projects a continuous high-frequency sound wave into body tissue  other  than  a fetus to determine frequency changes (doppler shift) in the reflected wave and is intended for use in the investiga- tion of nonfetal blood flow and other nonfetal body tissues in motion. This generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts, and accessories.

- Classification. Class II.

#### Food and Drug Administration, HHS	§ 892.1640

##### § 892.1550   Ultrasonic   pulsed   doppler imaging system.

  - Identification. An ultrasonic pulsed doppler imaging system is a de- vice that combines the features of con- tinuous wave doppler-effect technology with pulsed-echo effect technology and is intended to determine stationary body tissue characteristics, such as depth or location of tissue interfaces or dynamic tissue characteristics such as velocity of blood or tissue motion. This generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts, and accessories.

  - Classification. Class II.

##### § 892.1560   Ultrasonic  pulsed  echo  im- aging system.

- Identification. An ultrasonic pulsed echo imaging system is a device intended to project a pulsed sound beam into body tissue to determine the depth or location of the tissue inter- faces and to measure the duration of an acoustic pulse from the transmitter to the tissue interface and back to the re- ceiver. This generic type of device may include signal analysis and display equipment, patient and equipment sup- ports, component parts, and acces- sories.

- Classification. Class II.

##### § 892.1570   Diagnostic ultrasonic trans- ducer.

- Identification. A diagnostic ultra- sonic transducer is a device made of a piezoelectric material that converts electrical signals into acoustic signals and acoustic signals into electrical sig- nals and intended for use in diagnostic ultrasonic medical devices. Accessories of this generic type of device may in- clude transmission media for acous- tically coupling the transducer to the body surface, such as acoustic gel, paste, or a flexible fluid container.

- Classification. Class II.

##### § 892.1600   Angiographic x-ray system.

- Identification. An angiographic x- ray system is a device intended for radiologic visualization of the heart, blood vessels, or lymphatic system dur- ing or after injection of a contrast me- dium. This generic type of device may include   signal   analysis   and   display

equipment, patient and equipment sup- ports, component parts, and acces- sories.

- Classification. Class II.

##### § 892.1610   Diagnostic  x-ray  beam-lim- iting device.

- Identification. A diagnostic x-ray beam-limiting device is a device such as a collimator, a cone, or an aperture intended to restrict the dimensions of a diagnostic x-ray field by limiting the size of the primary x-ray beam.

- Classification. Class II.

##### § 892.1620   Cine  or  spot  fluorographic x-ray camera.

- Identification. A cine or spot fluo- rographic x-ray camera is a device in- tended to photograph diagnostic im- ages produced by x-rays with an image intensifier.

- Classification. Class II.

##### § 892.1630   Electrostatic  x-ray  imaging system.

- Identification. An electrostatic x- ray imaging system is a device in- tended for medical purposes that uses an electrostatic field across a semiconductive plate,  a gas-filled chamber, or other similar device to convert a pattern of x-radiation into an electrostatic image and, subsequently, into a visible image. This generic type of device may include signal analysis and display equipment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

##### § 892.1640   Radiographic  film  marking system.

- Identification. A radiographic film marking system is a device intended for medical purposes to add identifica- tion and other information onto radio- graphic film by means of exposure to visible light.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 59

FR 63015,  Dec. 7,  1994;  66  FR 38819,  July 25,

2001]

#### § 892.1650

##### § 892.1650   Image-intensified fluoroscopic x-ray system.

  - Identification. An image-intensi- fied fluoroscopic x-ray system is a de- vice intended to visualize anatomical structures by converting a pattern  of x-radiation into a visible  image through electronic amplification. This generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts, and accessories.

  - Classification. Class II. When in- tended as an accessory to the device described in paragraph (a) of this sec- tion, the fluoroscopic compression de- vice is exempt from the premarket no- tification procedures in subpart E of part   807   of   this   chapter   subject   to

§ 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 66

FR 57369, Nov. 15, 2001]

##### § 892.1660   Non-image-intensified fluoroscopic x-ray system.

- Identification. A non-image-inten- sified fluoroscopic x-ray system is a de- vice intended to be used to visualize anatomical structures by using a fluo- rescent screen to convert a pattern of x-radiation into a visible image. This generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts, and accessories.

- Classification. Class II.

##### § 892.1670   Spot-film device.

- Identification. A spot-film device is an electromechanical component of a fluoroscopic x-ray system that is in- tended to be used for medical purposes to position a radiographic film cassette to obtain radiographs during fluoros- copy.

- Classification. Class II.

##### § 892.1680   Stationary x-ray system.

- Identification. A stationary x-ray system is a permanently installed diag- nostic system intended to generate and control x-rays for examination of var- ious anatomical regions. This generic type of device may include signal anal- ysis and display equipment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 892.1700   Diagnostic  x-ray  high  volt- age generator.

- Identification. A diagnostic x-ray high voltage generator is a device that is intended to supply and control the electrical energy applied to a diag- nostic x-ray tube for medical purposes. This generic type of device may in- clude a converter that changes alter- nating current to direct current, fila- ment transformers for the x-ray tube, high voltage switches, electrical pro- tective devices, or other appropriate elements.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

##### § 892.1710   Mammographic   x-ray   sys- tem.

- Identification. A mammographic x- ray system is a device intended to be used to produce radiographs of the breast. This generic type of device may include signal analysis and display equipment, patient and equipment sup- ports, component parts, and acces- sories.

- Classification. Class II.

##### § 892.1715   Full-field   digital   mammog- raphy system.

- Identification. A full-field digital mammography system is a device in- tended to produce planar digital x-ray images of the entire breast. This ge- neric type of device may include dig- ital mammography acquisition soft- ware, full-field digital image receptor, acquisition workstation, automatic ex- posure control, image processing and reconstruction programs, patient and equipment supports, component parts, and accessories.

- Classification. Class II (special controls). The special control for the device is FDA’s guidance document en- titled ‘‘Class II Special Controls Guid- ance Document: Full-Field Digital Mammography System.’’ See § 892.1(e) for the availability of this guidance document.

[75 FR 68203, Nov. 5, 2010]

#### Food and Drug Administration, HHS	§ 892.1840

##### § 892.1720   Mobile x-ray system.

  - Identification. A mobile x-ray sys- tem is a transportable device system intended to be used to generate and control x-ray for diagnostic procedures. This generic type  of device may  in- clude signal analysis and display equip- ment, patient and equipment supports, component parts, and accessories.

  - Classification. Class II.

##### § 892.1730   Photofluorographic	x-ray system.

- Identification.	A photofluorographic x-ray system is a device that includes a fluoroscopic x- ray unit and a camera intended to be used to produce, then photograph, a fluoroscopic image of the body. This generic type of device may include sig- nal analysis and display equipment, pa- tient and equipment supports, compo- nent parts, and accessories.

- Classification. Class II.

##### § 892.1740   Tomographic x-ray system.

- Identification. A tomographic x- ray system is an x-ray device intended to be used to produce radiologic images of a specific cross-sectional plane of the body by blurring or eliminating de- tail from other planes. This generic type of device may include signal anal- ysis and display equipment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

##### § 892.1750   Computed tomography x-ray system.

- Identification. A computed tomog- raphy x-ray system is a diagnostic x- ray system intended to produce cross- sectional images of the body by com- puter reconstruction of x-ray trans- mission data from the same axial plane taken at different angles. This generic type of device may include signal anal- ysis and display equipment, patient and equipment supports, component parts, and accessories.

- Classification. Class II.

##### § 892.1760   Diagnostic x-ray tube hous- ing assembly.

- Identification. A diagnostic x-ray tube housing assembly is an x-ray gen- erating tube encased in a radiation- shielded  housing  that  is  intended  for

diagnostic purposes. This generic  type of device may include high voltage and filament transformers or other appro- priate components.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

##### § 892.1770   Diagnostic	x-ray	tube mount.

- Identification. A diagnostic x-ray tube mount is a device intended to sup- port and to position the diagnostic x- ray tube housing assembly for a med- ical radiographic procedure.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

##### § 892.1820   Pneumoencephalographic chair.

- Identification.		A pneumoencephalographic chair is a chair intended to support and position a	patient	during pneumoencephalography (x-ray imag- ing of the brain).

- Classification. Class II.

##### § 892.1830   Radiologic patient cradle.

- Identification. A radiologic patient cradle is a support device intended to be used for rotational positioning about the longitudinal axis of a patient during radiologic procedures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

##### § 892.1840   Radiographic film.

- Identification. Radiographic film is a device that consists of a thin sheet of radiotransparent  material  coated  on

#### § 892.1850

one or both sides with a photographic emulsion intended to record images during diagnostic radiologic proce- dures.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 66

FR 38819, July 25, 2001]

##### § 892.1850   Radiographic film cassette.

- Identification. A radiographic film cassette is a device intended for use during diagnostic x-ray procedures to hold a radiographic film in close con- tact with an x-ray intensifying screen and to provide a light-proof enclosure for direct exposure of radiographic film.

- Classification. Class II.

##### § 892.1860   Radiographic	film/cassette changer.

- Identification. A radiographic film/ cassette changer is a device intended to be used during a radiologic procedure to move a radiographic film or cassette between x-ray exposures and to posi- tion it during the exposure.

- Classification. Class II.

##### § 892.1870   Radiographic	film/cassette changer programmer.

- Identification. A radiographic film/ cassette changer programmer is a de- vice intended to be used to control the operations of a film or cassette changer during serial medical radiography.

- Classification. Class II.

##### § 892.1880   Wall-mounted   radiographic cassette holder.

- Identification. A wall-mounted ra- diographic cassette holder is a device that is a support intended to hold and position radiographic cassettes for a radiographic exposure for medical use.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 892.1890   Radiographic	film	illu- minator.

- Identification. A radiographic film illuminator is a device containing a visible light  source  covered  with a translucent front that is intended to be used to view medical radiographs.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2323, Jan. 14, 2000]

##### § 892.1900   Automatic radiographic film processor.

- Identification. An automatic radio- graphic film processor is a device in- tended to be used to develop, fix, wash, and dry  automatically  and  continu- ously film exposed for medical pur- poses.

- Classification. Class II.

[55 FR 48444, Nov. 20, 1990]

##### § 892.1910   Radiographic grid.

- Identification. A radiographic grid is a device that consists of alternating radiolucent and radiopaque strips in- tended to  be  placed  between the  pa- tient and the image receptor to reduce the amount of scattered radiation reaching the image receptor.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 65

FR 2323, Jan. 14, 2000]

##### § 892.1920   Radiographic head holder.

- Identification. A radiographic head holder is a device intended to position the patient’s head during a radio- graphic procedure.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral  requirements  concerning  records,

#### Food and Drug Administration, HHS	§ 892.1990

and § 820.198, with respect to complaint files.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 66

FR 38819, July 25, 2001]

##### § 892.1940   Radiologic	quality	assur- ance instrument.

  - Identification. A radiologic quality assurance instrument is a device in- tended for medical purposes to measure a physical characteristic associated with another radiologic device.

  - Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 66

FR 38819, July 25, 2001]

##### § 892.1950   Radiographic anthropomorphic phantom.

- Identification. A radiographic anthropomorphic phantom is a device intended for medical purposes to simu- late a human body for positioning radi- ographic equipment.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9. The device is also exempt from the current good manufacturing practice require- ments of the quality system regulation in part 820 of this chapter, with the ex- ception of § 820.180, with respect to gen- eral requirements concerning records, and § 820.198, with respect to complaint files.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 66

FR 38819, July 25, 2001]

##### § 892.1960   Radiographic	intensifying screen.

- Identification. A radiographic in- tensifying screen is a device that is a thin radiolucent sheet coated with a luminescent material that transforms incident   x-ray   photons   into   visible

light and intended for medical purposes to expose radiographic film.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 65

FR 2323, Jan. 14, 2000]

##### § 892.1970   Radiographic	ECG/res- pirator synchronizer.

- Identification. A radiographic ECG/ respirator synchronizer is a device in- tended to be used to coordinate an  x- ray film exposure with the signal from an electrocardiograph (ECG) or res- pirator at a predetermined phase of the cardiac or respiratory cycle.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[55 FR 48444, Nov. 20, 1990, as amended at 65

FR 2323, Jan. 14, 2000]

##### § 892.1980   Radiologic table.

- Identification. A radiologic table is a device intended for medical purposes to support a patient during radiologic procedures. The table may be fixed or tilting and may be electrically pow- ered.

- Classification. Class II (special controls).  The  device  is  exempt  from the  premarket  notification  procedures in subpart E of part 807 of this chapter subject to § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 63

FR 59231, Nov. 3, 1998]

##### § 892.1990   Transilluminator  for  breast evaluation.

- Identification. A transilluminator, also known as a diaphanoscope or lightscanner, is an electrically powered device that uses low intensity  emis- sions of visible light and near-infrared radiation (approximately 700–1050 nano- meters (nm)), transmitted through the breast, to  visualize  translucent  tissue for the diagnosis of cancer, other con- ditions,  diseases,  or  abnormalities.

- Classification. Class III (premarket approval).

#### § 892.2010

- Date premarket approval (PMA) or notice of completion of product develop- ment protocol (PDP) is required. A PMA or notice of completion of a PDP is re- quired to be filed with FDA by April 17, 2014, for any transilluminator for breast evaluation that was in commer- cial distribution before May 28, 1976, or that has, by April 17, 2014, been found to be substantially equivalent to any transilluminator for breast evaluation that was in commercial distribution before May 28, 1976. Any other transilluminator for breast evaluation shall have an approved PMA or de- clared completed PDP in effect before being placed in commercial distribu- tion.

[60 FR 36639, July 18, 1995, as amended at 79

FR 3094, Jan. 17, 2014]

##### § 892.2010   Medical  image  storage  de- vice.

- Identification. A medical image storage device is a device that provides electronic storage and retrieval func- tions for medical images. Examples in- clude devices employing magnetic and optical discs, magnetic tape, and dig- ital memory.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[63 FR 23387, Apr. 29, 1998; 63 FR 44998, Aug.

24, 1998, as amended at 65 FR 2323, Jan. 14,

2000]

##### § 892.2020   Medical  image  communica- tions device.

- Identification. A medical image communications device provides elec- tronic transfer of medical image data between medical devices. It may in- clude a physical communications me- dium, modems, interfaces, and a com- munications protocol.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[63 FR 23387, Apr. 29, 1998; 63 FR 44998, Aug.

24, 1998, as amended at 65 FR 2323, Jan. 14,

2000]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 892.2030   Medical image digitizer.

- Identification. A medical image digitizer is a device intended to con- vert an analog medical image into a digital format. Examples include Iystems employing video  frame grab- bers, and scanners which use lasers or charge-coupled devices.

- Classification. Class II (special controls; voluntary standards—Digital Imaging and Communications in Medi- cine (DICOM) Std., Joint Photographic Experts Group (JPEG) Std.).

[63 FR 23387, Apr. 29, 1998]

##### § 892.2040   Medical image hardcopy de- vice.

- Identification. A medical image hardcopy device is a device that pro- duces a visible printed record of a med- ical image and associated identifica- tion information. Examples include multiformat cameras and laser print- ers.

- Classification. Class II (special controls; voluntary standards—Digital Imaging and Communications in Medi- cine (DICOM) Std., Joint Photographic Experts Group (JPEG) Std., Society of Motion Picture and Television Engi- neers (SMPTE) Test Pattern).

[63 FR 23387, Apr. 29, 1998]

##### § 892.2050   Picture archiving and com- munications system.

- Identification. A picture archiving and communications system is a device that provides one or more capabilities relating to the acceptance, transfer, display, storage, and digital processing of medical images. Its hardware com- ponents may include workstations, digitizers, communications devices, computers, video monitors, magnetic, optical disk, or other digital data stor- age devices, and hardcopy devices. The software components may provide functions for performing operations re- lated to image manipulation, enhance- ment, compression or quantification.

- Classification. Class II (special controls; voluntary standards—Digital Imaging and Communications in Medi- cine (DICOM) Std., Joint Photographic Experts Group (JPEG) Std., Society of

#### Food and Drug Administration, HHS	§ 892.5740

Motion  Picture  and  Television  Engi- neers (SMPTE) Test Pattern).

[63 FR 23387, Apr. 29, 1998]

## Subparts C–E [Reserved]

Subpart F—Therapeutic Devices

##### § 892.5050   Medical charged-particle ra- diation therapy system.

  - Identification. A medical charged- particle radiation therapy system is a device that produces by acceleration high energy charged particles (e.g., electrons and protons) intended for use in radiation therapy. This generic type of device may include signal analysis and display equipment, patient and equipment supports, treatment plan- ning computer programs, component parts, and accessories.

  - Classification. Class II. When in- tended for use as a quality control sys- tem, the film dosimetry system (film scanning system) included as an acces- sory to the device described in para- graph (a) of this  section,  is  exempt from the premarket notification proce- dures in subpart E of part 807 of this chapter   subject   to   the   limitations   in

§ 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 64

FR 1125, Jan. 8, 1999]

##### § 892.5300   Medical   neutron   radiation therapy system.

- Identification. A medical neutron radiation therapy system is a device intended to generate high-energy neu- trons for radiation therapy. This ge- neric type of device may include signal analysis and display equipment, pa- tient and equipment support, treat- ment planning computer programs, component parts, and accessories.

- Classification. Class II.

##### § 892.5650   Manual  radionuclide  appli- cator system.

- Identification. A manual radio- nuclide applicator system is a manu- ally operated device intended to apply a radionuclide source into the body or to the surface of the body for radiation therapy. This generic type of device may include patient and equipment supports, component parts, treatment

planning  computer  programs,  and  ac- cessories.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 65

FR 2323, Jan. 14, 2000]

##### § 892.5700   Remote	controlled	radio- nuclide applicator system.

- Identification. A remote controlled radionuclide applicator system is an electromechanical or pneumatic device intended to enable an operator to apply, by remote control, a radio- nuclide source into the body or to the surface of the body for radiation ther- apy. This generic type of device may include patient and equipment sup- ports, component parts, treatment planning computer programs, and ac- cessories.

- Classification. Class II.

##### § 892.5710   Radiation	therapy	beam- shaping block.

- Identification. A radiation therapy beam-shaping block is a device made of a highly attenuating material (such as lead) intended for medical purposes to modify the shape of a beam from a ra- diation therapy source.

- Classification. Class II.

##### § 892.5730   Radionuclide brachytherapy source.

- Identification. A  radionuclide brachytherapy source is a device that consists of a radionuclide which  may be enclosed in a sealed container made of gold, titanium, stainless steel, or platinum and intended for medical pur- poses to be placed onto a body surface or into a body cavity or tissue as a source of nuclear radiation for therapy.

- Classification. Class II.

##### § 892.5740   Radionuclide	teletherapy source.

- Identification. A radionuclide tele- therapy source is a device consisting of a radionuclide enclosed in a sealed con- tainer. The device is intended for radi- ation therapy, with the radiation source located at a distance from the patient’s body.

§ 892.5750

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 59

FR 63015,  Dec. 7,  1994;  66  FR 38819,  July 25,

2001]

##### § 892.5750   Radionuclide radiation ther- apy system.

- Identification. A radionuclide radi- ation therapy system is a device in- tended to permit an operator to admin- ister gamma radiation therapy, with the radiation source located at a dis- tance from the patient’s body. This ge- neric type of device may include signal analysis and display equipment, pa- tient and equipment supports, treat- ment planning computer programs, component parts (including beam-lim- iting devices), and accessories.

- Classification. Class II.

##### § 892.5770   Powered  radiation  therapy patient support assembly.

- Identification. A powered radiation therapy patient support assembly is an electrically powered  adjustable  couch intended to support a patient during radiation therapy.

- Classification. Class II.

##### § 892.5780   Light beam patient position indicator.

- Identification. A light beam pa- tient position indicator is a device that projects a beam of light (incoherent light or laser) to determine the align- ment of the patient with a radiation beam. The beam of light is intended to be used during radiologic procedures to ensure proper positioning of the pa- tient and to monitor alignment of the radiation beam with the patient’s anat- omy.

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter, subject to the limitations in § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR 1125,  Jan. 16,  1996;  66  FR 38819,  July 25,

2001]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 892.5840   Radiation   therapy   simula- tion system.

- Identification. A radiation therapy simulation system is a fluoroscopic or radiographic x-ray system intended for use in localizing the volume to be ex- posed during radiation therapy and confirming the position and size of the therapeutic irradiation field produced. This generic type of device may in- clude signal analysis and display equip- ment, patient and equipment supports, treatment planning computer pro- grams, component parts, and acces- sories.

- Classification. Class II.

##### § 892.5900   X-ray radiation therapy sys- tem.

- Identification. An x-ray radiation therapy system is a device intended to produce and control x-rays used for ra- diation therapy. This generic type of device may include signal analysis and display equipment, patient and equip- ment supports, treatment planning computer programs, component parts, and accessories.

- Classification. Class II.

##### § 892.5930   Therapeutic	x-ray	tube housing assembly.

- Identification. A therapeutic x-ray tube housing assembly is an x-ray gen- erating tube encased in a radiation- shielded housing intended for use in ra- diation therapy. This generic type of device may include high-voltage and filament transformers or other appro- priate components when contained in radiation-shielded  housing.

- Classification. Class II.

## Subpart G—Miscellaneous Devices

##### § 892.6500   Personnel protective shield.

- Identification. A personnel protec- tive shield is a device intended for medical purposes to protect the pa- tient, the operator, or other persons from unnecessary exposure to radiation during radiologic procedures by pro- viding an attenuating barrier to radi- ation. This generic type of device may include articles of clothing, furniture, and movable or stationary structures.

Food and Drug Administration, HHS	§ 895.21

- Classification. Class I (general con- trols). The device is exempt from the premarket notification procedures in subpart E of part 807 of this chapter subject to § 892.9.

[53  FR 1567,  Jan. 20,  1988,  as  amended  at 61

FR  1125,  Jan.  16,  1996;  65  FR  2323,  Jan.  14,

2000]

## PART 895—BANNED DEVICES

##### Subpart A—General Provisions

Sec.

895.1    Scope.

  - General.

  - Procedures for banning a device.

  - Submission of data and information by the manufacturer, distributor, or im- porter.

895.25   Labeling.

895.30    Special effective date.

##### Subpart B—Listing of Banned Devices

  - Prosthetic hair fibers.

  - Powdered surgeon’s glove.

  - Powdered	patient	examination glove.

  - Absorbable  powder  for  lubricating  a surgeon’s glove.

AUTHORITY: 21 U.S.C. 352, 360f, 360h, 360i, 371.

SOURCE:  44  FR  29221,  May  18,  1979,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 895.1   Scope.

    - This part describes the procedures by which the Commissioner may insti- tute proceedings to make a device in- tended for human use that presents substantial deception or an unreason- able and substantial risk of illness or injury a banned device.

    - This part applies to any ‘‘device’’, as defined in section 201(h) of the Fed- eral Food, Drug, and  Cosmetic  Act (act) that is intended for human use.

    - A device that is made a banned device in accordance with this part is adulterated under section 501(g) of the act. A restricted device that is banned may also be misbranded under section 502(q) of the act.

    - Although this part does not cover devices intended for animal use, the manufacturer, distributor, importer, or any other person(s) responsible for the labeling  of  the  device  that  is  banned

cannot avoid the ban by relabeling the device for veterinary use. A device that has been banned from human use but that also has a valid veterinary use may be marketed for use as a veteri- nary device  only  under the  following conditions: The device shall comply with all requirements applicable to veterinary devices under the Federal Food, Drug, and Cosmetic Act and this chapter, and the label for the device shall bear the following statement: ‘‘For Veterinary Use Only. Caution: Federal law prohibits the distribution of this device for human use.’’ A device so labeled, however, that is determined by the Food and Drug Administration to be intended for human use, will be considered to be a banned device. In de- termining whether such a device is in- tended for human use, the Food and Drug Administration will consider, among other things, the ultimate des- tination of the device.

##### § 895.20   General.

The Commissioner may initiate a proceeding to make a device a banned device whenever the Commissioner finds, on the basis of all available data and information, that the device pre- sents substantial deception or an un- reasonable and substantial risk of ill- ness or injury that the Commissioner determines cannot be, or has not been, corrected or eliminated by labeling  or by a change in labeling, or by a change in advertising if the device is a re- stricted  device.

[44 FR 29221, May 18, 1979, as amended at 57

FR 58405, Dec. 10, 1992]

##### § 895.21   Procedures for banning a de- vice.

- Before initiating a proceeding to make a device a banned device, the Commissioner shall find that the con- tinued marketing of the device  pre- sents a substantial deception or an un- reasonable and substantial risk of ill- ness or injury.

  - In determining whether the decep- tion or risk of illness or injury is sub- stantial, the Commissioner will con- sider whether the deception or risk posed by continued marketing of the device,  or  continued  marketing  of  the

§ 895.21

device as presently labeled, is impor- tant, material, or significant in rela- tion to the benefit to the public health from its continued marketing.

  - In determining whether a device is deceptive, the Commissioner will con- sider whether users of the device may be deceived or otherwise harmed by the device. The  Commissioner is  not re- quired to determine that there was an intent on the part of the manufacturer, distributor, importer, or any other re- sponsible person(s) to mislead or other- wise harm users of the device or that there exists any actual proof of decep- tion of, or injury to, an individual.

  - In determining whether a device presents deception or risk of illness or injury, the Commissioner will consider all available data and information, in- cluding data and information that the Commissioner may obtain under other provisions of the act, data and informa- tion that may be supplied by the manu- facturer, distributor, or importer of the device under § 895.22, and data and in- formation  voluntarily  submitted  by any other interested persons.

- Before initiating a proceeding to make a device a banned device, the Commissioner of Food and Drugs (the Commissioner) may consult with the panel established under section 513 of the act that has expertise with respect to the type of device under consider- ation. The consultation with the panel may occur at a regular or specially scheduled panel meeting or may be ac- complished by correspondence or tele- phone conversation with panel mem- bers. The Commissioner may request that the panel submit in writing any advice on the device under consider- ation. The Commissioner will record in written memoranda any oral commu- nications with a panel or its members.

- If the Commissioner determines that any substantial deception or un- reasonable and substantial risk of ill- ness or injury or any unreasonable, di- rect, and substantial danger to the health of individuals presented by a de- vice can be corrected or eliminated by labeling or change in  labeling,  or change in advertising if the device is a restricted  device,  the  Commissioner will notify the responsible person of the required labeling or change in la- beling  or  change  in  advertising  in  ac-

#### 21 CFR Ch. I (4–1–17 Edition)

cordance with § 895.25.  If such  required relabeling  or  change  in  advertising  is not   accomplished   in   accordance   with

§ 895.25, the  Commissioner  may  initiate a proceeding to ban the device in ac- cordance with § 895.21(d) and, when ap- propriate, may establish a special ef- fective date in accordance with § 895.30.

- If the Commissioner decides to initiate a proceeding to make a device a banned device, a notice of proposed rulemaking will be published in the FEDERAL REGISTER to this effect. The notice will briefly summarize—

- The Commissioner’s finding under paragraph (a) of this section that the device   presents   substantial   deception or  an  unreasonable  and   substantial risk of illness or injury, and, when ap- propriate, the Commissioner’s deter- mination under § 895.30 that the decep- tion or risk of illness  or  injury  presents an unreasonable,  direct,  and  substan- tial  danger  to  the  health  of  individuals;

- The reasons why the Commis- sioner initiated the proceeding;

- The evaluation of data and infor- mation obtained under other provisions of the act, submitted by the manufac- turer, distributer, or importer of the device, or voluntarily  submitted  by any other interested persons under paragraph (a)(3) of this section, if any;

- The consultation with the panel, if any, under paragraph (b) of this sec- tion;

- The determination as to whether the deception or risk of illness or in- jury or the danger to the health of in- dividuals could be corrected by label- ing or change in labeling, or change in advertising if the device is a restricted device;

- The determination of whether the required labeling or change of labeling, or change in advertising if the device is a restricted device, if any, has been made in accordance with paragraph (c) of this section;

- The determination as to whether, and the reasons why, the banning should apply to devices already in com- mercial distribution or those already sold to the ultimate user, or both; and

- Any other data and information that the Commissioner believes are pertinent to the proceeding. The notice will afford all interested persons an op- portunity  to  submit  written  comments

#### Food and Drug Administration, HHS	§ 895.22

within 30 days after the date of publi- cation of the proposed regulation. All nonconfidential information  upon which the proposed finding is based, in- cluding the recommendations of the panel, will be available for public re- view in the Division of Dockets Man- agement, Food and Drug Administra- tion.

(e)(1) If, after reviewing the adminis- trative record of the regulatory hear- ing before the Food and Drug Adminis- tration, if any, the written comments received on the  proposed  regulation, and any additional available data and information, the Commissioner deter- mines to ban a device, a final regula- tion to this effect will be  published  in the FEDERAL REGISTER. The final regu- lation will amend subpart B by adding the name or description  of  the  device, or both, to the list of banned devices.

(2) If the Commissioner determines not to ban the device, a notice of with- drawal and termination of rulemaking proceedings and reasons therefor will be  published  in  the  FEDERAL  REGISTER.

- The effective date of a final regu- lation to make a device a banned de- vice, promulgated under paragraph (e) of this section, will be the date of pub- lication of the final regulation in the FEDERAL REGISTER unless the Commis- sioner, for reasons stated, determines that the effective date should be later than the date of the publication and specifies that date in the notice. Each such regulation will specify whether devices already in commercial distribu- tion or sold to the ultimate user or both are banned.

- A regulation promulgated under paragraph (e) of this section is final agency action, subject to  judicial  re- view under section 517 of the act.

- Upon petition of any interested person  submitted  in  accordance  with

§ 10.30 of this chapter, or as a matter of discretion, the Commissioner may in- stitute proceedings  to  amend  or  revoke a regulation that  made  a  device  a banned device  if  the  Commissioner finds that the conditions that  con- stituted the basis for the regulation banning the device are no longer appli- cable.    When    appropriate,    the    proce-

dures  in  this  section  will  be  employed in such proceedings.

[44 FR 29221, May 18, 1979, as amended at 53

FR  11254,  Apr.  6,  1988;  57  FR  58405,  Dec.  10,

1992; 65 FR 43690, July 14, 2000]

##### § 895.22 Submission of data and infor- mation by the manufacturer, dis- tributor, or importer.

  - A manufacturer, distributor, or importer of a device may be required to submit to the Food and Drug Adminis- tration all relevant and available data and information to enable the Commis- sioner to determine whether the device presents substantial deception, unrea- sonable and substantial risk  of  illness or injury, or unreasonable, direct, and substantial danger to the health of in- dividuals. The data and information re- quired by the Commissioner may in- clude scientific  or  test data, reports, records, or other information, includ- ing data and information on whether the device is safe and effective for its intended use or when used as directed, whether the device performs according to the claims made for the device, and information on adulteration or mis- branding. Any  relevant  information that is voluntarily submitted will  also be reviewed.

  - A manufacturer, distributor, or importer of a device required to submit data and information as provided in paragraph (a) of this section will be no- tified in writing by the Food and Drug Administration that such data and in- formation shall be submitted. The written notification will advise the manufacturer, distributor, or importer of the device that the purpose for the request is to enable the Commissioner to determine whether any of the condi- tions listed in paragraph (a) of this sec- tion or § 895.30(a)(1) exists with respect to the device such that a proceeding should be initiated to make the device a banned device. When the required data and information can be identified by the Food and  Drug  Administration at the time of the notification, the agency will provide such  identification to the manufacturer, distributor, or importer of the device.

  - The required data and informa- tion shall be submitted to the Food and Drug  Administration  no  more  than  30

#### § 895.25

days after the date of receipt of the re- quest, unless the Commissioner deter- mines that the data and information shall be submitted by some other date and so informs the manufacturer, dis- tributor, or importer, in which case the data and information shall be sub- mitted on the date specified by the Commissioner.

  - If the data or information sub- mitted to the Food and Drug Adminis- tration is sufficient to persuade the Commissioner that the deception  or risk of illness or  injury  or  the  danger to the health of individuals  presented by a device could be corrected or elimi- nated by labeling or change in labeling, or change in advertising if the device is a restricted device, the Commissioner will proceed in accordance with § 895.25.

  - If the data or information sub- mitted to the Food and Drug Adminis- tration is insufficient to show that the device does not present a substantial deception or an unreasonable and sub- stantial risk of illness or injury, or an unreasonable, direct, and substantial danger to the  health  of  individuals,  or if the manufacturer, distributor, or im- porter fails to submit the required in- formation, the Commissioner may rely upon this insufficiency or failure to submit the required information in considering whether to initiate a pro- ceeding to make the device a banned device under § 895.21(d) and, when ap- propriate, to establish a special effec- tive date in accordance with § 895.30. The Commissioner may also initiate other regulatory action as provided in the act or this chapter.

##### § 895.25   Labeling.

- If the Commissioner determines that the substantial deception or un- reasonable and substantial risk of ill- ness or injury or the unreasonable, di- rect, and substantial danger to the health of individuals presented by a de- vice can be corrected or eliminated by labeling or a change in labeling, or change in advertising if the device is a restricted  device,  the  Commissioner will provide written notice to the man- ufacturer, distributor, importer, or any other person(s) responsible  for  the  la- beling or advertising of the device specifying:

#### 21 CFR Ch. I (4–1–17 Edition)

- The deception or risk of illness or injury or the danger to the health of individuals,

- The labeling or change in label- ing, or change in advertising if the de- vice is a restricted device, necessary to correct the deception or eliminate or reduce such risk or danger, and

- The period of time within which the labeling, change in labeling, or change in advertising must be accom- plished.

- In specifying the labeling or change in labeling or change in adver- tising to correct the deception or to eliminate or reduce the risk of illness or injury or the danger to the health of individuals, the Commissioner may re- quire the manufacturer, distributor, importer, or any other person(s) re- sponsible for the labeling or adver- tising of the device to include in label- ing for the device, and in advertising if the device is a restricted device, a statement, notice,  or warning.  Such statement, notice, or warning shall be in the manner and form prescribed by the Commissioner and shall identify the deception or risk of illness or in- jury or the unreasonable, direct, and substantial danger to the health of in- dividuals associated with the device as previously labeled. Such statement, no- tice, or warning shall be used in the la- beling and advertising of the device for a time period specified by the Commis- sioner on the basis of the degree of de- ception, risk of illness or injury, or danger to health; the frequency of sale of the device; the length of time the de- vice has  been  on  the  market;  the  in- tended uses of the device; the method of its use; and any other factors that the Commissioner considers pertinent.

- The Commissioner will allow a manufacturer, distributor, importer, or any other person(s) responsible for the labeling or advertising of the device a reasonable time, considering the decep- tion or risk of illness or injury or the danger to the health of individuals pre- sented by the device, within which to accomplish the required labeling, change in labeling, and, if the device is a restricted device, any change in ad- vertising. The Commissioner may, however, request that no additional de- vices   be   introduced   into   commerce

#### Food and Drug Administration, HHS	§ 895.30

until the labeling or change in label- ing, or change in advertising is accom- plished by the manufacturer, dis- tributor, importer, or other person(s) responsible for the labeling or adver- tising of the device.

- If such voluntary action is not taken, the Commissioner may take ac- tion under other sections of the act to prevent the introduction of the devices into commerce.  The  Commissioner may consider the failure of a manufac- turer, distributor, importer,  or  any other person(s) responsible  for  the  la- beling or advertising of the device to accomplish the required labeling or change in labeling, or change in adver- tising in accordance with  this  section as a basis for initiating a proceeding to make a device a banned device in ac- cordance with § 895.21(d) and when ap- propriate to establish a special effec- tive date in accordance with § 895.30.

##### § 895.30   Special effective date.

- The Commissioner may declare a proposed regulation under § 895.21(d) to be effective upon its publication in the FEDERAL REGISTER and until the effec- tive date of any final action taken re- specting the regulation if:

  - The Commissioner determines, on the basis of all available data and in- formation, that the deception or risk of illness or injury associated with use of the device that is subject to the regula- tion presents an unreasonable, direct, and substantial danger to the health of individuals,  and

  - Before the date of the publication of such regulation, the Commissioner notifies the domestic manufacturer and importer, if any, of the device that the regulation is to be made so effective. If necessary, the Commissioner may also notify the distributor or any other re- sponsible person(s). In addition, the Commissioner will attempt to notify any foreign manufacturer when the name and address of the foreign manu- facturer are readily available.

- This procedure may be used when the Commissioner determines that the potential or actual injury involved is a serious one that the Commissioner be- lieves will endanger the health of indi- viduals who have been, or will be, ex- posed to the device. In assessing the de- gree of danger, the Commissioner need

not find that the danger is immediate, and it shall be sufficient for the Com- missioner to determine that the danger may involve a serious long-term risk.

- If the Commissioner makes a pro- posed regulation effective in accord- ance with this section, the Commis- sioner will, as expeditiously as  pos- sible, give interested persons prompt notice of this action in the FEDERAL REGISTER and will provide an oppor- tunity for an informal hearing in ac- cordance with part 16 of this chapter.

- After the hearing, if  any,  and after considering any written com- ments submitted on the proposal and any additional available information and data, the Commissioner will as ex- peditiously as possible either affirm, modify, or revoke the proposed regula- tion making the device a banned de- vice. If the Commissioner decides to af- firm or modify the proposed regulation to make a device a banned device, the Commissioner will amend subpart B by adding the name or description of the device, or both, to the list of banned devices. If the Commissioner decides to revoke a proposed regulation making a device a banned device, a notice of ter- mination of rulemaking proceedings and reasons therefor will be published in  the  FEDERAL   REGISTER.

- The  Commissioner  may   declare the special  effective  date  provided  by this section to be in effect after the publication of a proposed  regulation under § 895.21(d), if, based on new infor- mation, or upon reconsideration of pre- viously available information, the Commissioner makes  the  determina- tion and provides the appropriate no- tices and an opportunity  for  a  hearing in  accordance  with  paragraphs  (a)  and

- of this section.

  - Those devices that have been named banned devices  under  § 895.30 and that have already been sold to the public may be subject to relabeling  by the  manufacturer,  distributor,  im- porter, or any other person(s) respon- sible for the labeling of the device or may be subject to the provisions of sec- tion 518(a) or (b) of the act.

[44 FR 29221, May 18, 1979, as amended at 57

FR  58405,  Dec.  10,  1992;  80  FR  31300,  June  2,

2015]

#### § 895.101

Subpart B—Listing of Banned Devices

##### § 895.101   Prosthetic hair fibers.

Prosthetic hair fibers are devices in- tended for implantation into the human scalp to  simulate natural  hair or conceal baldness. Prosthetic hair fi- bers may consist of various materials; for example, synthetic fibers, such as modacrylic, polyacrylic, and polyester; and natural fibers, such as processed human hair. Excluded from the banned device are natural hair transplants, in which a person’s  hair and its  sur- rounding tissue are surgically removed from one location on the person’s scalp and then grafted onto another area of the person’s scalp.

[48 FR 25136, June 3, 1983]

##### § 895.102   Powdered surgeon’s glove.

    - Identification. A powdered sur- geon’s glove is a device intended to be worn on the hands of operating room personnel to protect a surgical wound from contamination. A powdered sur- geon’s glove incorporates powder for purposes other than manufacturing.

    - [Reserved]

[81 FR 91731, Dec. 19, 2016]

##### § 895.103   Powdered   patient   examina- tion glove.

- Identification. A powdered patient examination glove is a disposable de- vice intended for medical purposes that is worn on the examiner’s hand or fin- ger to prevent contamination between patient and examiner. A powdered pa- tient examination glove incorporates powder for purposes other than manu- facturing.

- [Reserved]

[81 FR 91731, Dec. 19, 2016]

##### § 895.104   Absorbable powder for lubri- cating a surgeon’s glove.

Absorbable powder for lubricating a surgeon’s glove is a powder made from cornstarch that meets the specifica- tions for absorbable powder in the United States Pharmacopeia (U.S.P.) and that is intended to be used to lu- bricate the surgeon’s hand before put-

#### 21 CFR Ch. I (4–1–17 Edition)

ting on a surgeon’s glove. The device is absorbable through biological degrada- tion.

[81 FR 91731, Dec. 19, 2016]

## PART 898—PERFORMANCE STAND- ARD FOR ELECTRODE LEAD WIRES AND PATIENT CABLES

Sec.

  - Applicability.

  - Performance  standard.

  - Compliance dates.

  - Exemptions and variances.

AUTHORITY:  21  U.S.C.  351,  352,  360c,  360d,

360gg–360ss, 371, 374; 42 U.S.C. 262, 264.

SOURCE:  62  FR  25497,  May  9,  1997,  unless

otherwise noted.

##### § 898.11   Applicability.

Electrode lead wires and patient ca- bles intended for use with a medical de- vice shall be subject to the perform- ance standard set forth in § 898.12.

##### § 898.12   Performance standard.

    - Any connector in a cable or elec- trode lead wire having a conductive connection to a patient shall be con- structed in such a manner as to comply with subclause 56.3(c) of the following standard:

International Electrotechnical Commis- sion (IEC)

601–1: Medical Electrical Equipment

601–1 (1988) Part 1: General requirements for  safety

Amendment No. 1 (1991)

Amendment No. 2 (1995).

    - Compliance with the standard shall be determined by inspection  and by applying the test requirements and test methods of subclause 56.3(c) of the standard set forth in paragraph (a) of this  section.

##### § 898.13   Compliance dates.

The dates for compliance with the standard set forth in § 898.12(a) shall be as  follows:

- For electrode lead wires and pa- tient cables used with, or intended for use with, the following   devices,  the date for which  compliance  is  required is May 11, 1998:

Food and Drug Administration, HHS	§ 898.14

LISTING OF DEVICES FOR WHICH COMPLIANCE IS REQUIRED EFFECTIVE

May 11, 1998

| Phase | Product code | 21 CFR sec- tion | Class | Device name |
| --- | --- | --- | --- | --- |
| 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... 1 ................... | 73 BZQ FLS DPS 74 DRG 74 DRT 74 DRX 74 DSA 74 DSH 74 DSI 74 DXH | 868.2375 868.2375 870.2340 870.2910 870.2300 870.2360 870.2900 870.2800 870.1025 870.2920 | II II II II II II II II III II | Monitor, Breathing Frequency. Monitor (Apnea Detector), Ventilatory Effort. Electrocardiograph. Transmitters and Receivers, Physiological Signal, Radio Frequency. Monitor, Cardiac (including Cardiotachometer and Rate Alarm). Electrode, Electrocardiograph. Cable, Transducer and Electrode, Patient (including Connector). Recorder, Magnetic Tape, Medical. Detector and Alarm, Arrhythmia. Transmitters and Receivers, Electrocardiograph, Tele- phone. |

- For electrode lead wires and pa- tient cables used with, or intended for use with, any other device, the date for which compliance is required is May 9, 2000.

##### § 898.14   Exemptions and variances.

- A request for an exemption or variance shall be submitted  in  the  form of a  petition under § 10.30 of this  chap- ter and shall comply with the require- ments  set  out  therein.   The   petition shall  also  contain  the  following:

- The name of the device, the class in which the device has been classified, and representative labeling  showing the intended uses(s) of the device;

- The reasons why compliance with the performance standard is unneces- sary or unfeasible;

- A complete description of alter- native steps that are available, or  that the petitioner has already taken, to en- sure that a patient will not be inad- vertently connected to hazardous voltages via an unprotected patient cable or electrode lead wire for in- tended use with the device; and

- Other information justifying the exemption or variance.

- An exemption or variance is not effective until the agency approves the request under § 10.30(e)(2)(i) of this chapter.

EFFECTIVE DATE NOTE: At 62 FR 25477, May

9, 1997, § 898.14 was stayed pending Office of Management and Budget approval of infor- mation collection and recordkeeping require- ments.

# SUBCHAPTER I—MAMMOGRAPHY QUALITY STANDARDS ACT

Sec.

## PART 900—MAMMOGRAPHY

##### Subpart A—Accreditation

to perform screening or diagnostic mammography services. Subpart A fur- ther establishes requirements and standards for accreditation bodies to ensure  that  all  mammography  facili-

  - Scope.

  - Definitions.

  - Application for approval as an accredi- tation body.

  - Standards for accreditation bodies.

  - Evaluation.

  - Withdrawal of approval.

  - Hearings. 900.8–900.9  [Reserved]

##### Subpart B—Quality Standards and Certification

  - Applicability.

  - Requirements for certification.

  - Quality standards.

  - Revocation of accreditation and rev- ocation of accreditation body approval.

  - Suspension  or  revocation  of  certifi- cates.

  - Appeals of adverse accreditation or reaccreditation decisions that preclude certification or recertification.

  - Appeals of denials of certification.

  - [Reserved]

  - Alternative  requirements   for   § 900.12 quality    standards.

##### Subpart C—States as Certifiers

  - Scope.

  - Application for approval as a certifi- cation agency.

  - Standards for certification agencies.

  - Evaluation.

  - Withdrawal of approval.

  - Hearings and appeals.

AUTHORITY: 21 U.S.C. 360i, 360nn, 374(e); 42

U.S.C. 263b.

SOURCE:  62  FR  55976,  Oct.  28,  1997,  unless

otherwise   noted.   Republished   and   corrected at 62 FR 60614, Nov. 10, 1997.

## Subpart A—Accreditation

##### § 900.1   Scope.

The regulations set forth in this part implement the Mammography Quality Standards Act (MQSA) (42 U.S.C. 263b). Subpart A of this part establishes pro- cedures whereby an entity can apply to become a Food and Drug Administra- tion  (FDA)-approved   accreditation body to accredit facilities to be eligible

ties under the  jurisdiction  of  the United States are adequately and con- sistently  evaluated  for   compliance with national quality standards for mammography. Subpart B of this part establishes minimum national quality standards for mammography facilities to ensure safe, reliable, and accurate mammography. The regulations set forth in this part do not apply to facili- ties of the Department of Veterans Af- fairs.

##### § 900.2   Definitions.

The following definitions apply to subparts A, B, and C of this part:

- Accreditation body or body means an entity that has been approved by FDA under § 900.3(d) to accredit mam- mography   facilities.

- Action limits or action levels means the minimum  and  maximum  values  of a quality assurance measurement that can be interpreted as representing ac- ceptable performance with respect to the parameter being tested. Values less than the minimum or greater than the maximum action limit or level indicate that corrective action must be  taken by the facility. Action limits or  levels are also sometimes called control  lim- its or levels.

- Adverse event means an undesir- able experience associated with mam- mography activities within the  scope of 42 U.S.C. 263b. Adverse events in- clude but are not limited to:

- Poor image quality;

- Failure to send mammography re- ports within 30 days to the referring physician or in a timely manner to the self-referred patient; and

- Use of personnel that do not meet the       applicable       requirements       of

§ 900.12(a).

- Air kerma means kerma in a given mass of air. The unit used to measure the quantity of air kerma is the Gray (Gy).  For  X-rays  with  energies  less

Food and Drug Administration, HHS	§ 900.2

than 300 kiloelectron volts (keV), 1 Gy

= 100 rad. In air, 1 Gy of absorbed dose is delivered by 114 roentgens (R) of ex- posure.

- Breast implant means a prosthetic device implanted in the breast.

- Calendar quarter means any one of the following time periods during a given year: January 1 through March 31, April 1 through June 30, July 1 through September 30, or October 1 through December 31.

- Category I means medical edu- cational activities that have been des- ignated as Category I by the Accredita- tion Council for Continuing Medical Education (ACCME), the American Os- teopathic Association (AOA), a state medical society, or an equivalent orga- nization.

- Certificate means the certificate described in § 900.11(a).

- Certification means  the  process  of approval of a facility by FDA or a cer- tification agency to provide mammog- raphy services.

- Clinical image means a mammo- gram.

- Consumer means an individual who chooses to comment  or  complain in reference to a mammography exam- ination, including the patient or rep- resentative of the patient (e.g., family member or referring physician).

- Continuing education unit or con- tinuing education credit means one con- tact hour of training.

- Contact hour means an hour of training received through direct in- struction.

- Direct instruction means:

  - Face-to-face interaction between instructor(s) and student(s), as when the instructor provides a lecture, con- ducts demonstrations, or reviews stu- dent performance; or

  - The administration and correction of student examinations by an instruc- tor(s) with subsequent feedback to the student(s).

- Direct supervision means that:

- During joint interpretation of mammograms, the supervising inter- preting physician reviews, discusses, and confirms the diagnosis of the phy- sician being supervised and signs the resulting report before it is entered into the patient’s records; or

- During the performance of a mam- mography examination or survey of the facility’s equipment and quality assur- ance program, the supervisor is present to observe and correct, as needed, the performance of the individual being su- pervised who is performing the exam- ination or conducting the survey.

- Established operating level means the value of a particular quality assur- ance parameter that has been estab- lished as an acceptable normal level by the facility’s quality assurance pro- gram.

- Facility means a hospital, out- patient department, clinic, radiology practice, mobile unit, office of a physi- cian, or other facility that conducts mammography activities, including the following: Operation of equipment to produce a mammogram, processing of the mammogram, initial interpreta- tion of the mammogram, and main- taining viewing conditions for that in- terpretation. This term does not in- clude a facility of the Department of Veterans Affairs.

- First allowable time means the ear- liest time a resident physician is eligi- ble to take the diagnostic radiology boards from an FDA-designated certi- fying body. The ‘‘first allowable time’’ may vary with the certifying body.

- FDA means the Food and Drug Administration.

- Interim regulations means the regu- lations  entitled  ‘‘Requirements  for  Ac- crediting  Bodies  of  Mammography  Fa- cilities’’ (58 FR 67558–67565) and ‘‘Qual- ity    Standards    and    Certification    Re- quirements   for   Mammography   Facili- ties’’ (58 FR 67565–67572), published by FDA on December 21, 1993, and amend- ed on September 30, 1994 (59 FR 49808– 49813).   These   regulations   established the  standards  that  had  to  be  met  by mammography   facilities   in   order   to lawfully   operate   between   October   1, 1994, and April 28, 1999.

- Interpreting physician means a li- censed physician who interprets mam- mograms and who meets the require- ments set forth in § 900.12(a)(1).

- Kerma means  the sum  of  the ini- tial energies  of  all  the charged par- ticles liberated by uncharged ionizing particles in a material of given mass.

- Laterality means the designation of either the right or left breast.

#### § 900.2

- Lead interpreting physician means the interpreting physician assigned the general responsibility for ensuring that a facility’s quality assurance program meets   all   of   the   requirements   of

§ 900.12(d) through (f). The administra- tive title and other supervisory respon- sibilities of the individual, if any, are left to the discretion of the facility.

- Mammogram means a radiographic image produced through mammog- raphy.

- Mammographic modality means a technology,   within   the   scope   of   42

    - 263b, for radiography of the breast. Examples are screen-film mam- mography and xeromammography.

(aa) Mammography means radiog- raphy of the breast, but, for the pur- poses of this part, does not include:

      - Radiography of the breast per- formed during invasive interventions for localization or biopsy procedures; or

      - Radiography of the breast per- formed with an investigational mam- mography device as part of a scientific study conducted in accordance with FDA’s investigational  device  exemp- tion regulations in part 812 of this chapter.

(bb) Mammography equipment evalua- tion means an onsite assessment of mammography unit or image processor performance by a medical physicist for the purpose of making a preliminary determination as to whether the equip- ment meets all of the applicable stand- ards in § 900.12(b) and (e).

(cc) Mammography medical outcomes audit means a systematic collection of mammography results and the com- parison of those results with outcomes data.

(dd) Mammography unit or  units means an assemblage of components for the production of X-rays for use during mammography, including, at a minimum: An X-ray generator, an X- ray control, a tube housing assembly, a beam limiting device, and the sup- porting structures for these compo- nents.

(ee) Mean optical density means the average of the optical densities meas- ured using phantom thicknesses of 2, 4, and 6 centimeters with values of kilo- volt peak (kVp) clinically appropriate for those thicknesses.

#### 21 CFR Ch. I (4–1–17 Edition)

(ff) Medical physicist means a person trained in evaluating the performance of mammography equipment and facil- ity quality assurance programs and who meets the qualifications for a medical      physicist      set      forth      in

§ 900.12(a)(3).

(gg) MQSA means the Mammography Quality Standards Act.

(hh) Multi-reading means two or more physicians, at least one of whom is an interpreting   physician,    interpreting the same mammogram.

(ii) Patient means any individual who undergoes a mammography evaluation in a facility, regardless of whether the person is referred by a physician or is self-referred.

(jj) Phantom means a test object used to simulate radiographic characteris- tics of compressed breast tissue and containing components that radio- graphically model aspects of breast dis- ease and cancer.

(kk) Phantom image means a radio- graphic image of a phantom.

(ll) Physical science means physics, chemistry, radiation science (including medical physics and health physics), and engineering.

(mm) Positive mammogram means a mammogram that has an overall as- sessment of findings that are either ‘‘suspicious’’ or ‘‘highly suggestive of malignancy.’’

(nn) Provisional certificate means the provisional    certificate    described    in

§ 900.11(b)(2).

(oo) Qualified instructor means an in- dividual whose training and experience adequately prepares him or her  to carry out specified training assign- ments.	Interpreting	physicians, radiologic technologists, or medical physicists who meet the  requirements of § 900.12(a) would be considered quali- fied instructors in  their  respective areas of mammography. Other exam- ples of individuals who may be quali- fied instructors for the purpose of pro- viding training to meet the regulations of this part include, but are not limited to, instructors in a post-high school training institution and manufactur- er’s representatives.

(pp) Quality control technologist  means an  individual   meeting   the    require- ments of § 900.12(a)(2) who is responsible

#### Food and Drug Administration, HHS	§ 900.3

for those quality assurance responsibil- ities not assigned to the lead inter- preting physician or to the medical physicist.

(qq) Radiographic equipment means X- ray equipment used for the production of  static  X-ray  images.

(rr) Radiologic technologist means an individual specifically trained  in  the use of radiographic equipment and the positioning of patients for radiographic examinations and who meets the re- quirements set forth in § 900.12(a)(2).

(ss) Serious adverse event means an ad- verse advent that may significantly compromise clinical outcomes, or an adverse event for which a facility fails to take appropriate corrective  action in a timely manner.

(tt) Serious complaint means a report of a serious adverse event.

(uu) Standard breast means a 4.2 cen- timeter (cm) thick compressed breast consisting of 50 percent glandular  and 50 percent adipose tissue.

(vv) Survey means an onsite physics consultation and evaluation of a facil- ity quality assurance program per- formed by a medical physicist.

(ww) Time cycle means the film devel- opment  time.

(xx) Traceable to a national standard means an instrument is calibrated at either the National Institute of Stand- ards and Technology (NIST) or at a calibration  laboratory  that  partici- pates in a proficiency program  with NIST at least once every 2  years  and the results of the proficiency test con- ducted within 24 months of calibration show agreement within 3 percent of the national standard in the mammog- raphy energy range.

(yy) Review physician means a physi- cian who, by meeting the requirements set out in § 900.4(c)(5), is qualified to re- view clinical images on behalf of the accreditation  body.

(zz) Certification agency means a State that has  been  approved  by FDA  under

§ 900.21 to certify mammography facili- ties.

(aaa) Performance indicators mean the measures used to evaluate the certifi- cation agency’s ability to conduct cer- tification, inspection, and compliance activities.

(bbb) Authorization means obtaining approval  from  FDA  to  utilize  new  or

changed State regulations or proce- dures during the issuance, mainte- nance, and withdrawal of certificates by the certification agency.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10,  1997,  as  amended  at  63  FR  56558,  Oct. 22,

1998;  64  FR 32407,  June 17,  1999;  67  FR 5467,

Feb. 6, 2002]

##### § 900.3   Application for approval as an accreditation body.

- Eligibility. Private nonprofit orga- nizations or State agencies capable of meeting the requirements of this sub- part A may apply for approval as ac- creditation  bodies.

- Application for initial approval. (1) An applicant seeking initial FDA ap- proval as an accreditation body shall inform the Division of Mammography Quality and Radiation Programs (DMQRP), Center for Devices and Radi- ology Health  (HFZ–240), Food and Drug Administration, 1350 Piccard Dr., Rockville, MD 20850, marked Attn: Mammography  Standards   Branch,   of its desire to be approved as an accredi- tation body and of  its  requested  scope of   authority.

- Following receipt of the request, FDA will provide the applicant with additional information to aid in sub- mission of an application for approval as an accreditation body.

- The applicant shall furnish to FDA, at the address in § 900.3(b)(1), three copies of an application con- taining the following information, ma- terials,  and  supporting  documentation:

- Name, address, and phone number of the applicant and, if the applicant is not a State agency, evidence of non- profit status (i.e., of fulfilling Internal Revenue Service requirements as a nonprofit  organization);

- Detailed description of the ac- creditation standards the applicant will require facilities to meet and a discussion substantiating their equiva- lence to FDA standards required under

§ 900.12;

- Detailed description of the appli- cant’s accreditation review and deci- sionmaking process, including:

  - Procedures for performing ac- creditation and reaccreditation clinical image    review    in    accordance    with

§ 900.4(c), random clinical image re- views in accordance with § 900.4(f), and

#### § 900.3

additional  mammography  review  in  ac- cordance with § 900.12(j);

  - Procedures for performing phan- tom image review;

  - Procedures for assessing mam- mography equipment evaluations and surveys;

  - Procedures for initiating and per- forming onsite visits to facilities;

  - Procedures for assessing facility personnel  qualifications;

  - Copies of the accreditation appli- cation forms, guidelines, instructions, and other materials the applicant will send to facilities during the accredita- tion process, including an accredita- tion history form that requires each fa- cility to provide a complete history of prior accreditation activities and a statement that all information and data submitted in the application is true and accurate, and that no mate- rial fact has been omitted;

  - Policies and procedures for noti- fying facilities of deficiencies;

  - Procedures for monitoring correc- tions of deficiencies by facilities;

  - Policies and procedures for sus- pending or revoking a facility’s accred- itation;

  - Policies and procedures that will ensure processing of accreditation ap- plications and renewals within a time- frame approved by FDA and assurances that the body will adhere to such poli- cies and procedures; and

  - A description of the applicant’s appeals process for facilities contesting adverse accreditation status decisions.

- Education, experience, and train- ing requirements for the applicant’s professional staff, including reviewers of clinical or phantom images;

- Description of the applicant’s electronic data management and anal- ysis system with respect to accredita- tion review and decision processes and the applicant’s ability to provide elec- tronic data in a  format  compatible with FDA data systems;

- Resource analysis that dem- onstrates that the applicant’s staffing, funding, and other resources are ade- quate to perform the required accredi- tation  activities;

- Fee schedules with supporting cost data;

- Statement of policies and pro- cedures  established  to  avoid  conflicts

#### 21 CFR Ch. I (4–1–17 Edition)

of interest or the appearance of con- flicts of interest by the applicant’s board members, commissioners, profes- sional personnel (including reviewers of clinical and phantom images), con- sultants, administrative personnel, and other  representatives  of  the  applicant;

- Statement of policies and proce- dures established to protect confiden- tial information the applicant will col- lect or receive in its role as an accredi- tation body;

- Disclosure of any specific brand of imaging system or component, meas- uring device, software  package,  or other commercial product used in mammography that the applicant de- velops, sells, or distributes;

- Description of  the  applicant’s consumer complaint mechanism;

- Satisfactory assurances that the applicant shall comply with the re- quirements of § 900.4; and

- Any other information as may be required by FDA.

- Application for renewal of approval. An approved accreditation body that intends to continue to serve as an ac- creditation body beyond its current term shall apply to FDA for renewal or notify FDA of its plans not to apply for renewal in accordance with the fol- lowing procedures and schedule:

  - At least 9 months  before  the  date of expiration of a body’s approval, the body shall inform FDA, at the address given in § 900.3(b)(1), of its intent to seek  renewal.

  - FDA will notify the applicant of the  relevant  information,  materials, and supporting documentation required under § 900.3(b)(3) that the applicant shall submit as part of the renewal pro- cedure.

  - At least 6 months before the date of expiration of a body’s approval, the applicant shall furnish to FDA, at the address in § 900.3(b)(1), three copies of a renewal application containing the in- formation, materials, and supporting documentation requested by FDA in accordance with § 900.3(c)(2).

  - No later than July 28, 1998, any ac- creditation body approved under the interim regulations published in the FEDERAL REGISTER of December 21, 1993 (58 FR 67558), that  desires  to  continue to  serve  as  an  accreditation  body  under

#### Food and Drug Administration, HHS	§ 900.4

the final regulations shall apply for re- newal of approval in accordance with the procedures set forth in paragraphs (c)(1) through (c)(3) of this section.

  - Any accreditation body that does not plan to renew its approval shall so notify FDA at the address given in paragraph (b)(1) of this section at least

9 months before the expiration of the body’s term of approval.

- Rulings on applications for initial and renewed approval. (1) FDA will con- duct a review and evaluation to deter- mine whether the applicant substan- tially meets the applicable require- ments of this subpart and whether the accreditation standards the applicant will require facilities to meet are sub- stantially the same as the quality standards published under subpart B of this part.

  - FDA will notify the applicant of any deficiencies in the application and request that those deficiencies be rec- tified within a specified time period. If the deficiencies are not rectified to FDA’s satisfaction within the specified time period, the application for ap- proval as an accreditation body may be rejected.

  - FDA shall  notify the  applicant whether the application has been ap- proved or denied. That notification shall list any conditions associated with approval or state the bases for any denial.

  - The review of any application may include a meeting between FDA and representatives of the applicant at a time and location mutually  acceptable to FDA and the applicant.

  - FDA will advise the applicant of the circumstances under which a de- nied application may be resubmitted.

  - If FDA does not reach a final deci- sion on a renewal application in ac- cordance with this paragraph before the expiration of an accreditation body’s current term of approval, the approval will be deemed extended until the agency reaches a final decision on the application, unless an accredita- tion body does not rectify deficiencies in the application within the specified time period, as required in paragraph (d)(2) of this section.

- Relinquishment of authority. An ac- creditation body that decides to relin- quish  its  accreditation  authority  be-

fore expiration of the body’s term of approval shall submit a letter of such intent   to   FDA,   at   the   address   in

§ 900.3(b)(1), at least 9 months before re- linquishing   such   authority.

- Transfer of records. An accredita- tion body that does not apply for re- newal of accreditation body approval, is denied such approval by FDA, or re- linquishes its accreditation authority and duties before expiration of its term of approval, shall:

  - Transfer facility records and other related information as required  by FDA to a location and according to a schedule approved by FDA.

  - Notify, in a manner and time pe- riod approved by FDA, all facilities ac- credited or seeking accreditation by the body that the body will no longer have accreditation authority.

- Scope of authority. An accredita- tion body’s term of approval is for a pe- riod not to exceed 7 years. FDA may limit the scope of accreditation au- thority.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10, 1997]

##### § 900.4   Standards	for	accreditation bodies.

- Code of conduct and general respon- sibilities. The accreditation body shall accept the following responsibilities in order to ensure safe and accurate mam- mography at the facilities it accredits and shall perform these responsibilities in a manner that ensures the integrity and impartiality of accreditation body actions.

(1)(i) When an accreditation body re- ceives or discovers information that suggests inadequate image quality, or upon request by FDA, the accredita- tion body shall review a facility’s clin- ical images or other aspects of a facili- ty’s practice to assist FDA in deter- mining whether or not the facility’s practice poses a serious risk to human health. Such reviews are in addition to the evaluation an accreditation body performs as part of the initial accredi- tation or renewal process for facilities.

(ii) If review by the accreditation body demonstrates that a problem does exist with respect to image quality or other aspects of a facility’s compliance with quality standards, or upon request by  FDA,  the  accreditation  body  shall

#### § 900.4

require or monitor corrective actions, or suspend or revoke accreditation of the  facility.

- The accreditation body shall in- form FDA as soon as possible but in no case longer than 2 business days after becoming aware of equipment or prac- tices that pose a serious risk to human health.

- The accreditation body shall es- tablish and administer a quality assur- ance (QA) program that has been ap- proved  by  FDA  in  accordance  with

§ 900.3(d) or paragraph (a)(8) of this sec- tion. Such quality assurance program shall:

  - Include requirements for clinical image review and phantom image re- view;

  - Ensure that clinical and phantom images are evaluated consistently and accurately; and

  - Specify the methods and fre- quency of training and evaluation for clinical and phantom image reviewers, and the bases and procedures for re- moval of such reviewers.

- The accreditation body shall es- tablish measures that FDA has ap- proved in accordance with § 900.3(d) or paragraph (a)(8) of this section to re- duce the possibility of conflict of inter- est or facility bias on the part of indi- viduals acting on the body’s behalf. Such individuals who review clinical or phantom images under the provisions of paragraphs (c) and (d) of this section or who visit facilities under the provi- sions of paragraph (f) of this section shall not review clinical or phantom images from or visit a facility with which such individuals maintain a rela- tionship, or when it would otherwise be a conflict of interest for them to do so, or when they have a bias in favor of or against the facility.

- The accreditation body may re- quire specific equipment performance or design characteristics that FDA has approved. However, no accreditation body shall require, either explicitly or implicitly, the use of any specific brand of imaging system or component, measuring device, software package, or other commercial product as a condi- tion for accreditation by the body, un- less FDA determines that it is in the best interest of public health to do so.

#### 21 CFR Ch. I (4–1–17 Edition)

- Any representation, actual or im- plied, either orally, in  sales  literature, or in any other form of representation, that the purchase or use of a particular product brand is required in order for any facility to  be  accredited  or cer- tified under § 900.11(b), is prohibited, unless FDA approves such representa- tion.

- Unless FDA has approved the ex- clusive use and promotion of a par- ticular commercial product in accord- ance with this section, all products produced, distributed, or sold by an ac- creditation body or an organization that has a financial or other relation- ship with the accreditation body that may be  a conflict  of interest  or have the appearance of a conflict of interest with the body’s accreditation func- tions, shall bear a disclaimer stating that the purchase or use of such prod- ucts is not required for  accreditation or  certification  of  any  facility  under

§ 900.11(b). Any representations about such products shall include a similar disclaimer.

- When an accreditation body de- nies accreditation to a facility, the ac- creditation body shall notify the facil- ity in writing and explain the bases for its decision. The notification shall also describe the appeals process available from the accreditation body for the fa- cility to contest the decision.

- No accreditation body may estab- lish requirements that preclude facili- ties    from    being    accredited    under

§ 900.11(b) by any other accreditation body, or require accreditation by itself under MQSA if another accreditation body is available to a facility.

- The accreditation body shall ob- tain  FDA  authorization  for   any changes it proposes to make in any standards that FDA has previously ac- cepted under § 900.3(d).

- An accreditation body shall estab- lish procedures to protect confidential information it collects or receives  in its role as an accreditation body.

  - Nonpublic information collected from facilities for the purpose of car- rying out accreditation body respon- sibilities shall not be used for any other purpose or disclosed, other than to FDA or its duly designated rep- resentatives, including State  agencies, without the consent of the facility;

#### Food and Drug Administration, HHS	§ 900.4

  - Nonpublic information that FDA or its duly designated representatives, including State agencies, share with the accreditation body concerning a fa- cility that is accredited or undergoing accreditation by that body shall not be further disclosed except with the writ- ten permission of FDA.

- Monitoring facility compliance with quality standards. (1) The accreditation body shall require that each facility it accredits meet standards for the per- formance of quality mammography that are substantially the same as those in this subpart and in subpart B of this part.

- The accreditation body shall no- tify a facility regarding equipment, personnel, and other aspects of the fa- cility’s practice that do not meet such standards and advise the facility that such equipment, personnel, or other as- pects of the practice should not be used by the facility for activities within the scope of part 900.

- The accreditation body shall specify the actions that facilities shall take to correct deficiencies in equip- ment, personnel, and other aspects of the practice to ensure facility compli- ance with applicable standards.

- If deficiencies cannot be corrected to ensure compliance with standards or if a facility is unwilling to take correc- tive actions, the accreditation body shall immediately so notify FDA, and shall suspend or revoke the facility’s accreditation in accordance with the policies and procedures described under

§ 900.3(b)(3)(iii)(I).

- Clinical image review for accredita- tion and reaccreditation—(1)  Frequency of review. The accreditation body shall review clinical images from each  facil- ity accredited by the body at least once every 3 years.

- Requirements for clinical image at- tributes. The accreditation body shall use the following attributes for all clinical image reviews, unless FDA has approved other attributes:

  - Positioning. Sufficient breast tis- sue shall be imaged to ensure that can- cers are not likely to be missed because of inadequate positioning.

  - Compression. Compression shall be applied in a manner that minimizes the potential obscuring effect of overlying breast tissue and motion artifact.

  - Exposure level. Exposure level shall be adequate to visualize breast structures. Images shall be neither un- derexposed nor overexposed.

  - Contrast. Image contrast shall permit differentiation of subtle tissue density differences.

  - Sharpness. Margins of normal breast structures shall be distinct and not blurred.

  - Noise. Noise in the image shall not obscure breast structures or sug- gest the appearance of structures not actually present.

  - Artifacts. Artifacts due to lint, processing, scratches, and other factors external to the breast shall not obscure breast structures or suggest the ap- pearance of structures not actually present.

  - Examination identification. Each image shall have the following infor- mation indicated on it in a permanent, legible, and unambiguous manner and placed so as not to obscure anatomic structures:

- Name of the patient and an addi- tional patient identifier.

- Date of examination.

- View and laterality. This informa- tion shall  be placed on  the image in  a position near the axilla. Standardized codes specified by the accreditation body and approved by FDA in accord- ance with § 900.3(d) or paragraph (a)(8) of this section shall be used to identify view  and  laterality.

- Facility name and location. At a minimum, the location shall include the city, State, and zip code of the fa- cility.

- Technologist identification.

- Cassette/screen  identification.

- Mammography unit identifica- tion, if there is more than one unit in the  facility.

- Scoring of clinical images. Accredi- tation bodies shall establish and ad- minister a system for scoring clinical images using  all  attributes  specified in paragraphs (c)(2)(i) through (c)(2)(viii) of this section or an alternative system that FDA has approved in accordance with § 900.3(d) or paragraph (a)(8) of this section. The scoring system shall in- clude  an  evaluation  for  each  attribute.

  - The accreditation body shall es- tablish and employ criteria for accept- able and nonacceptable results for each

#### § 900.4

of the 8 attributes as well as an overall pass-fail system for clinical image re- view that has been approved by FDA in accordance with § 900.3(d) or paragraph (a)(8) of this section.

  - All clinical  images  submitted  by a facility to the  accreditation  body shall be reviewed independently by two or more review physicians.

- Selection of clinical images for re- view. Unless otherwise specified by FDA, the accreditation body shall re- quire that for each mammography unit in  the  facility:

- The facility shall submit craniocaudal (CC) and mediolateral ob- lique (MLO) views from two mammo- graphic examinations that  the facility produced during a time period specified by the accreditation body;

- Clinical images submitted from one such mammographic examination for each unit shall be of dense breasts (predominance of glandular tissue) and the other shall be  of  fat-replaced breasts (predominance of adipose tis- sue);

- All clinical images submitted shall be images that the facility’s in- terpreting physician(s) interpreted as negative or benign.

- If the facility has no clinical im- ages meeting the requirements in para- graphs (c)(4)(i) through (c)(4)(iii) of this section, it shall so notify the accredita- tion body, which shall specify alter- native clinical image  selection  meth- ods that do not compromise care of the patient.

- Review physicians. Accreditation bodies shall ensure that all of their re- view physicians:

  - Meet the interpreting physician requirements specified in § 900.12(a)(1) and meet such additional requirements as have been established by the accred- itation body and approved by FDA;

  - Are trained and evaluated in the clinical image review process, for the types of clinical images to be evaluated by a review physician, by the accredi- tation body before designation as  re- view physicians and periodically there- after; and

  - Clearly document their findings and reasons for assigning a particular score to any clinical image and provide information  to  the  facility  for  use  in

#### 21 CFR Ch. I (4–1–17 Edition)

improving the attributes for which sig- nificant deficiencies were identified.

- Image management. The accredita- tion body’s QA program shall include a tracking system to ensure the security and return to the facility of all clinical images received and to ensure comple- tion of all clinical image reviews  by the body in a timely manner. The ac- creditation body shall return all clin- ical images to the facility within 60 days of their receipt by the body, with the following exceptions:

  - If the clinical images are needed earlier by the facility for clinical pur- poses, the accreditation body shall co- operate with the facility to accommo- date such needs.

  - If a review physician identifies a suspicious abnormality on an image submitted for clinical image review, the accreditation body shall  ensure that this information is provided to the facility and that the clinical images are returned to the facility. Both shall occur no later than 10-business days after identification of the suspected ab- normality.

- Notification of unsatisfactory image quality. If the accreditation body deter- mines that the clinical images received from a facility are of unsatisfactory quality, the body shall notify the facil- ity of the nature of the problem and its possible causes.

- Phantom image review for accredita- tion and reaccreditation—(1)  Frequency of review. The accreditation body shall review phantom images from each fa- cility accredited by the body at least once every 3 years.

- Requirements for the phantom used. The accreditation body shall  require that each facility submit for review phantom images that the facility pro- duced using a phantom and methods of use specified by the body and approved by FDA in accordance with § 900.3(d) or paragraph (a)(8) of this section.

- Scoring phantom images. The ac- creditation body shall use a system for scoring phantom images that has been approved  by  FDA  in  accordance  with

§ 900.3(b) and (d) or paragraph (a)(8) of this   section.

- Phantom images selected for review. For each mammography unit in the fa- cility, the accreditation body shall re- quire  the  facility  to  submit  phantom

#### Food and Drug Administration, HHS	§ 900.4

images that the facility produced dur- ing a time period specified by the body.

- Phantom image reviewers. Accredi- tation bodies shall ensure that all of their phantom image reviewers:

  - Meet the requirements specified  in

§ 900.12(a)(3) or alternative require- ments established by the accreditation body and approved by FDA in accord- ance with § 900.3 or paragraph (a)(8) of this   section;

  - Are trained and evaluated in the phantom image review process, for the types of phantom images to be evalu- ated by a phantom image reviewer, by the accreditation body before designa- tion as phantom image reviewers and periodically thereafter; and

  - Clearly document their findings and reasons for assigning a particular score to any phantom image and pro- vide information to the facility for use in improving its phantom image qual- ity with regard to the significant defi- ciencies identified.

- Image management. The accredita- tion body’s QA program shall include a tracking system to ensure the security of all phantom images received and to ensure completion of all phantom image reviews by the body in a timely manner. All phantom images that re- sult in a failure of accreditation shall be returned to the facility.

- Notification measures for unsatisfac- tory image quality. If the accreditation body determines that the phantom im- ages received from a facility are of un- satisfactory quality, the body shall no- tify the facility of the nature of the problem and its possible causes.

- Reports of mammography equipment evaluation, surveys, and quality control. The  following  requirements  apply  to all facility equipment covered by the provisions of subparts A and B:

  - The accreditation body shall re- quire every facility applying for ac- creditation to submit:

    - With its initial accreditation ap- plication, a mammography equipment evaluation that was performed by a medical physicist no earlier than 6 months before the date of application for accreditation by the facility. Such evaluation shall demonstrate compli- ance of the facility’s equipment with the requirements in § 900.12(e).

    - Prior to accreditation, a survey that was performed no earlier than 6 months before the date of application for accreditation by the facility. Such survey shall assess the facility’s com- pliance with the facility standards ref- erenced in paragraph (b) of this sec- tion.

  - The accreditation body shall re- quire that all facilities undergo an an- nual survey to ensure continued com- pliance with the standards referenced in paragraph (b) of this section and to provide continued oversight of facili- ties’ quality control programs as they relate to such standards. The accredi- tation body shall require for all facili- ties that:

    - Such surveys be conducted annu- ally;

    - Facilities  take  reasonable  steps to ensure that they receive reports of such surveys within 30 days of survey completion; and

    - Facilities submit the results of such surveys and any other informa- tion that the body may require to the body at least annually.

  - The accreditation body shall re- view and analyze the information re- quired in this section and use it to identify necessary corrective measures for facilities and to determine whether facilities should remain accredited by the body.

- Accreditation body onsite visits and random clinical image reviews. The ac- creditation body shall conduct onsite visits and random clinical image re- views of a sample of facilities to mon- itor and assess their compliance with standards established by the body for accreditation. The accreditation body shall submit annually to FDA, at the address given in § 900.3(b)(1), 3 copies of a summary report describing all  facil- ity assessments the body conducted under the provisions of this section for the year being reported.

  - Onsite visits—(i) Sample size. Annu- ally,  each  accreditation  body   shall visit at  least  5 percent of the  facilities it accredits. However, a minimum of 5 facilities shall be visited, and  visits  to no more than 50 facilities are required, unless problems identified in paragraph (f)(1)(i)(B) of this section indicate a need to visit more than 50 facilities.

#### § 900.4

- At least 50 percent of the facili- ties visited shall be selected randomly.

- Other facilities visited shall be selected based on problems identified through State or FDA inspections, se- rious complaints received from con- sumers or others, a previous history of noncompliance, or any other informa- tion in the possession of the accredita- tion body, inspectors, or FDA.

- Before, during, or after any facil- ity visit, the accreditation body may require that the facility submit to the body for review clinical images, phan- tom images, or any other information relevant to applicable standards in this subpart and in subpart B of this part.

(ii) Visit plan. The accreditation body shall conduct facility onsite visits ac- cording to a visit plan that has been approved by FDA in accordance with

§ 900.3(d) or paragraph (a)(8) of this sec- tion, unless otherwise directed by FDA in particular circumstances. At a min- imum, such a plan shall provide for:

- Assessment of overall clinical image QA activities of the facility;

- Review of facility documentation to determine if appropriate mammog- raphy reports are sent to patients and physicians as required;

- Selection of a sample of clinical images for clinical image review by the accreditation body. Clinical images shall be selected in a manner specified by the accreditation body and approved by FDA that does not compromise care of the patient as a result of the absence of the selected images from the facil- ity;

- Verification that  the  facility  has a medical audit system in place and is correlating films and pathology reports for positive cases;

- Verification that personnel speci- fied by the facility are the ones actu- ally performing designated personnel functions;

- Verification that equipment spec- ified by the facility is the equipment that is actually being used to perform designated equipment functions;

- Verification that a consumer complaint mechanism is in place and that the facility is following its proce- dures; and

- Review of all factors related to previously identified concerns or con- cerns identified during that visit.

#### 21 CFR Ch. I (4–1–17 Edition)

  - Clinical image review for random sample of facilities—(i) Sample size. In addition to conducting clinical image reviews for accreditation and re- accreditation for all facilities, the ac- creditation body shall conduct clinical image reviews annually for a randomly selected sample as specified by FDA, but to include at least 3 percent of the facilities the body accredits. Accredita- tion bodies may count toward this ran- dom sample requirement all facilities selected randomly for the onsite visits described in paragraph (f)(1)(i)(A) of this section. Accreditation bodies shall not count toward the random sample requirement any facilities described in paragraph (f)(1)(i)(B) of this  section that were selected for a visit because of previously identified concerns.

- Random clinical image review. In performing clinical image reviews of the random sample of facilities, accred- itation bodies shall evaluate the same attributes as those in paragraph (c) of this section for review of clinical im- ages for accreditation and reaccredita- tion.

- Accreditation bodies should not schedule random clinical image re- views at facilities that have received notification of the need to begin the accreditation renewal process or that have completed the accreditation re- newal process within the previous 6 months.

- Selection of the random sample of clinical images for clinical image review by the accreditation body. Clinical im- ages shall be selected in a manner, specified by the accreditation body and approved by FDA under § 900.3(d) or paragraph (a)(8) of this section, that does not compromise care of the pa- tient as a result of the absence of the selected images from the facility.

- Consumer complaint mechanism. The accreditation body shall  develop and administer a written and docu- mented system, including timeframes, for collecting and resolving serious consumer complaints that could not be resolved at a facility.  Such  system shall have been approved by FDA in ac- cordance with§ 900.3(d) or paragraph (a)(8) of this section. Accordingly, all accreditation bodies shall:

#### Food and Drug Administration, HHS	§ 900.6

  - Provide a mechanism for all fa- cilities it accredits to file serious unre- solved complaints with the accredita- tion body;

  - Maintain a record of every serious unresolved complaint received by the body on all facilities it accredits for a period of at least 3 years from the date of receipt of each such complaint;

- Reporting and recordkeeping. All reports to FDA specified in paragraphs (h)(1) through (h)(4) of this section shall be prepared and submitted in a format and medium prescribed by FDA and shall be submitted to a location and according to a schedule specified by FDA. The accreditation body shall:

- Collect and submit to FDA the in- formation required by 42 U.S.C. 263b(d) for each facility when the facility is initially accredited and at least annu- ally when updated, in a manner and at a time specified by FDA.

- Accept applications  containing the information required in 42 U.S.C. 263b(c)(2) for provisional certificates and in § 900.11(b)(3) for extension of pro- visional certificates, on behalf of FDA, and notify FDA of the receipt of such information;

- Submit to FDA the name, identi- fying information, and other informa- tion relevant to 42 U.S.C. 263b and spec- ified by FDA for any facility for which the accreditation body denies, sus- pends, or revokes accreditation, and the reason(s) for such action;

- Submit to FDA an annual report summarizing all serious complaints re- ceived during the previous calendar year, their resolution status, and any actions taken in response to them;

- Provide to FDA other information relevant to 42 U.S.C. 263b and required by FDA about any  facility  accredited or undergoing accreditation by the body.

  - Fees. Fees charged to facilities for accreditation shall be reasonable. Costs of accreditation  body  activities  that are not related to accreditation func- tions under 42 U.S.C. 263b are not re- coverable through fees established for accreditation.

- The accreditation body shall make public its fee structure, includ- ing those factors, if  any,  contributing to variations in fees for different facili- ties.

- At FDA’s request, accreditation bodies shall provide financial records or other material to assist FDA in as- sessing the reasonableness of accredi- tation body fees. Such material  shall be provided to FDA in a manner and time period specified by the agency.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10, 1997, as amended at 64 FR 32407, June 17,

1999]

##### § 900.5   Evaluation.

FDA shall evaluate annually the per- formance of each accreditation body. Such evaluation shall include an as- sessment of the reports of FDA or State inspections of facilities accred- ited by the body as well as any addi- tional information deemed relevant by FDA that has been provided by the ac- creditation body or other sources or has been required by FDA as part of its oversight initiatives. The evaluation shall include a determination of wheth- er there are major deficiencies in the accreditation body’s performance that, if not corrected, would warrant with- drawal of the approval of the accredita- tion   body   under   the   provisions   of

§ 900.6.

##### § 900.6   Withdrawal of approval.

If FDA  determines,  through the  eval- uation activities of § 900.5, or through other means,  that  an  accreditation body is not in substantial compliance with  this  subpart,  FDA  may  initiate the following actions:

  - Major deficiencies. If  FDA  deter- mines that an accreditation body has failed to perform a major accreditation function satisfactorily, has dem- onstrated willful disregard for public health, has violated the code of con- duct, has committed fraud, or has sub- mitted material  false  statements  to the agency, FDA may withdraw its ap- proval of that accreditation body.

    - FDA shall notify the accredita- tion body of the agency’s action  and the grounds on which the approval was withdrawn.

    - An accreditation body that has lost its  approval shall  notify facilities accredited or seeking accreditation by it that its approval has been with- drawn. Such notification shall be made within a time period and in a manner approved by FDA.

#### § 900.7

  - Minor deficiencies. If  FDA  deter- mines that an accreditation body has demonstrated deficiencies in per- forming accreditation functions and re- sponsibilities that are less serious or more limited than the deficiencies in paragraph (a) of this section, FDA shall notify the body that it has a specified period of time to take particular cor- rective measures directed by FDA or to submit to FDA for approval the body’s own plan of corrective action address- ing the minor deficiencies. FDA may place the body on probationary status for a period of time determined by FDA, or may withdraw approval of the body as an accreditation body if correc- tive action is not taken.

    - If FDA places an accreditation body on probationary status, the body shall notify all facilities accredited or seeking accreditation by it of its pro- bationary status within a time period and in a manner approved by FDA.

    - Probationary status shall remain in effect until such time as the  body can demonstrate to the satisfaction of FDA that it has successfully imple- mented or is implementing the correc- tive action plan within the established schedule, and that the corrective ac- tions have substantially eliminated all identified problems.

    - If FDA determines that an accred- itation body  that  has  been  placed  on probationary status is not imple- menting corrective actions satisfac- torily or within the established sched- ule, FDA may withdraw approval of the accreditation body. The accreditation body shall notify all facilities accred- ited or seeking accreditation by  it  of its loss of FDA approval, within a time period and in a manner approved by FDA.

  - Reapplication by accreditation bod- ies that have had their approval with- drawn. (1) A former accreditation body that has had its approval withdrawn may submit a new application for ap- proval if the body can provide informa- tion to FDA to establish that the prob- lems that were grounds for withdrawal of approval have been resolved.

- If FDA determines that the new application demonstrates that the body satisfactorily has addressed the  causes of  its  previous  unacceptable  perform-

#### 21 CFR Ch. I (4–1–17 Edition)

ance,  FDA  may  reinstate  approval  of the accreditation body.

- FDA may request additional infor- mation or establish additional condi- tions that must be met by a former ac- creditation body before FDA approves the reapplication.

- FDA may refuse to accept an ap- plication from a former accreditation body whose approval was withdrawn because of fraud or willful disregard of public health.

##### § 900.7   Hearings.

  - Opportunities to challenge final adverse actions taken by FDA regard- ing approval or reapproval of accredi- tation bodies, withdrawal of  approval of accreditation  bodies,  or  rejection  of a proposed fee for accreditation  shall be communicated through notices of opportunity for informal hearings in accordance with part 16 of this chapter.

  - A facility that has been denied ac- creditation is entitled to an appeals process from the accreditation  body. The appeals process shall  be  specified in writing by the accreditation body and shall have been  approved  by  FDA in     accordance     with     § 900.3(d)     or

§ 900.4(a)(8).

  - A facility that cannot achieve sat- isfactory resolution of an adverse ac- creditation decision through the ac- creditation body’s appeals process may appeal to FDA for reconsideration in accordance with § 900.15.

##### §§ 900.8–900.9   [Reserved]

Subpart B—Quality Standards and Certification

§ 900.10   Applicability.

The provisions of subpart B are appli- cable to all facilities under the regu- latory jurisdiction of the United States that provide mammography services, with the exception of  the  Department of  Veterans  Affairs.

##### § 900.11   Requirements	for	certifi- cation.

- General. After October 1, 1994, a certificate  issued  by  FDA  is  required

#### Food and Drug Administration, HHS	§ 900.11

for lawful operation of all mammog- raphy facilities subject to the  provi- sions of this subpart. To obtain a cer- tificate from FDA, facilities are re- quired to meet the quality standards in

§ 900.12 and to be accredited by an ap- proved accreditation body or other en- tity as designated by FDA.

- Application—(1) Certificates. (i) In order to qualify for a certificate, a fa- cility must apply to an FDA-approved accreditation body, or to another enti- ty designated by FDA.  The  facility shall submit to such body or entity the information required in 42 U.S.C. 263b(d)(1).

  - Following the agency’s receipt of the accreditation body’s decision to ac- credit a facility, or an equivalent deci- sion by another entity designated by FDA, the agency may issue a certifi- cate to the facility, or renew an exist- ing certificate, if the agency deter- mines that the facility has satisfied the requirements for certification or recertification.

- Provisional certificates. (i) A new facility beginning operation  after Octo- ber 1, 1994, is eligible to apply for a pro- visional certificate. The provisional certificate will enable the facility to perform mammography and to obtain the clinical images needed to complete the accreditation process. To apply for and receive a provisional certificate, a facility must meet the requirements of 42 U.S.C. 263b(c)(2) and submit the nec- essary information to an approved ac- creditation body or other entity des- ignated  by  FDA.

(ii) Following the agency’s receipt of the  accreditation   body’s   decision   that a facility has submitted the required information, FDA may issue a provi- sional certificate to a facility upon de- termination that  the  facility  has  satis- fied the requirements of § 900.11(b)(2)(i). A provisional certificate shall be effec- tive for up to 6 months from the date of issuance. A provisional certificate can- not be renewed, but a  facility  may apply for a 90-day extension of the pro- visional   certificate.

- Extension of  provisional  certificate.

- To apply for a 90-day extension to a provisional certificate, a facility shall submit to its accreditation body, or other entity designated by FDA, a statement of what the facility is doing

to obtain certification and evidence that there would be a significant ad- verse impact on access to mammog- raphy in the geographic area served if such facility did not obtain an exten- sion.

- The accreditation body shall for- ward the request, with its rec- ommendation, to FDA within 2 busi- ness days after receipt.

- FDA may issue a 90-day exten- sion for a provisional certificate upon determination  that  the   extension meets the criteria set forth in 42 U.S.C. 263b(c)(2).

- There can be no renewal of a pro- visional certificate beyond the 90-day extension.

- Reinstatement policy. A previously certified facility that has allowed its certificate to expire, that has been re- fused a renewal of its certificate by FDA, or that has had its certificate suspended or revoked by FDA, may apply to have the certificate reinstated so that the facility may be considered to be a new facility and thereby be eli- gible for a provisional certificate.

  - Unless prohibited from reinstate- ment under § 900.11(c)(4), a facility ap- plying for reinstatement shall:

    - Contact an FDA-approved accredi- tation body or other entity designated by FDA to determine the requirements for reapplication for accreditation;

    - Fully document its history as a previously provisionally certified or certified mammography facility, in- cluding the following information:

      - Name and address of the facility under which it was previously provi- sionally certified or certified;

      - Name of previous owner/lessor;

      - FDA facility identification num- ber assigned to the facility under its previous certification; and

      - Expiration date of the most re- cent FDA provisional certificate or certificate; and

    - Justify application for reinstate- ment of accreditation by submitting to the accreditation body or other entity designated by FDA, a corrective action plan that details how the facility has corrected deficiencies that contributed to the lapse of, denial of renewal, or revocation of its certificate.

  - FDA may issue a provisional cer- tificate to the facility if:

#### § 900.12

    - The accreditation body or other entity designated by FDA notifies the agency that the facility has adequately corrected, or is in the process of cor- recting, pertinent deficiencies; and

    - FDA determines that the facility has taken sufficient corrective action since the lapse of, denial of renewal, or revocation of its previous certificate.

  - After receiving the provisional certificate, the facility may lawfully resume performing mammography services while completing the require- ments for certification.

  - If a facility’s certificate was re- voked on the basis of an act described in 41 U.S.C. 263b(i)(1), no person who owned or operated that facility at the time the act occurred may own or oper- ate a mammography facility within 2 years of the date of revocation.

##### § 900.12   Quality standards.

- Personnel. The following require- ments apply to all  personnel  involved in any aspect of mammography, includ- ing the production, processing, and in- terpretation of mammograms and re- lated  quality  assurance  activities:

  - Interpreting physicians. All physi- cians interpreting mammograms shall meet the following qualifications:

    - Initial qualifications. Unless the ex- emption in paragraph (a)(1)(iii)(A) of this section applies, before  beginning to interpret mammograms independ- ently, the interpreting physician shall:

(A) Be licensed to practice medicine in a State;

(B)(1) Be certified in an appropriate specialty area by a body determined by FDA to have procedures and require- ments adequate to ensure that physi- cians certified by the body are com- petent to interpret radiological proce- dures, including mammography; or

  - Have had at least 3 months of doc- umented formal training in the inter- pretation of mammograms and in  top- ics related to mammography. The training shall include instruction in ra- diation physics, including radiation physics specific to mammography, ra- diation effects, and radiation protec- tion. The mammographic interpreta- tion component shall be under the di- rect supervision of a physician who meets the requirements of paragraph (a)(1) of this section;

#### 21 CFR Ch. I (4–1–17 Edition)

- Have a minimum of 60 hours of documented medical education in mammography, which shall include: In- struction in the interpretation of mam- mograms and education in basic breast anatomy, pathology, physiology, tech- nical aspects of mammography, and quality assurance and quality  control in mammography. All 60 of these hours shall be category I and at  least  15  of the category I hours shall have been acquired within the 3 years imme- diately prior to the date that the phy- sician qualifies as an interpreting phy- sician. Hours spent in residency specifi- cally devoted to mammography will be considered as equivalent to Category I continuing medical education credits and will be accepted if documented in writing by the appropriate representa- tive of the training institution; and

- Unless the exemption in para- graph (a)(1)(iii)(B) of this section ap- plies, have interpreted or multi-read at least 240 mammographic examinations within the 6-month period immediately prior to the date that the physician qualifies as an interpreting physician. This interpretation or  multi-reading shall be under the direct supervision of an  interpreting  physician.

- Continuing experience and edu- cation. All interpreting physicians shall maintain their qualifications by meet- ing the following requirements:

- Following the second anniversary date of the end of the calendar quarter in which the requirements of paragraph (a)(1)(i) of this section were completed, the interpreting physician shall have interpreted or multi-read at least 960 mammographic examinations  during the 24 months immediately preceding the date of the facility’s annual MQSA inspection or the last day of the cal- endar quarter preceding the inspection or any date in-between the two. The fa- cility will choose one of these dates to determine the 24-month period.

- Following the third anniversary date of the end of the calendar quarter in which the requirements of paragraph (a)(1)(i) of this section were completed, the interpreting physician shall have taught or completed at least 15 cat- egory I continuing medical education units in mammography during the 36 months    immediately    preceding    the

#### Food and Drug Administration, HHS	§ 900.12

date of the facility’s annual MQSA in- spection or the last day of the calendar quarter preceding the inspection or any date in between the two. The facility will choose one of these dates to deter- mine the 36-month period. This train- ing shall include at least six category I continuing  medical  education  credits in each mammographic modality used by the interpreting physician in his or her practice; and

- Before an interpreting physician may begin independently interpreting mammograms produced by a new mam- mographic modality, that is, a mam- mographic modality in which the phy- sician has not previously been trained, the  interpreting  physician  shall   have at least 8 hours of training in the new mammographic  modality.

- Units earned through teaching a specific course can be  counted  only once towards the 15 required by para- graph (a)(1)(ii)(B) of this section, even if the course is taught multiple times during the previous 36 months.

- Exemptions. (A) Those physicians who qualified as interpreting physi- cians under paragraph (a)(1) of this sec- tion of FDA’s interim regulations prior to April 28, 1999, are considered to have met the initial requirements of para- graph (a)(1)(i) of this  section.  They may continue to interpret mammo- grams provided they continue to meet the licensure requirement of paragraph (a)(1)(i)(A) of this section and the con- tinuing experience and education re- quirements of paragraph (a)(1)(ii) of this  section.

(B) Physicians who have  interpreted or multi-read at least 240 mammo- graphic examinations under the direct supervision of an interpreting  physi- cian in any 6-month period during the last 2 years of a diagnostic radiology residency and who become appro- priately board certified at the first al- lowable time, as defined by an eligible certifying body, are otherwise exempt from paragraph (a)(1)(i)(D) of this sec- tion.

- Reestablishing qualifications. In- terpreting physicians who fail to main- tain the required continuing experi- ence or continuing education require- ments shall reestablish their qualifica- tions before resuming the independent

interpretation of mammograms, as fol- lows:

- Interpreting physicians who fail to meet the continuing experience re- quirements of paragraph (a)(1)(ii)(A) of this section shall:

  - Interpret or multi-read at  least 240 mammographic examinations under the direct supervision of an inter- preting physician, or

  - Interpret or multi-read a suffi- cient number of mammographic exami- nations, under the direct supervision of an interpreting physician, to bring the physician’s total up to 960 examina- tions for the prior 24 months, which- ever is less.

  - The interpretations required under paragraph (a)(1)(iv)(A)(1) or (a)(1)(iv)(A)(2) of this section shall be done within the 6 months immediately prior to resuming independent inter- pretation.

- Interpreting physicians who fail to meet the continuing education re- quirements of paragraph (a)(1)(ii)(B) of this section shall obtain a sufficient number of additional category I con- tinuing medical education credits in mammography to bring their  total  up to the required 15 credits in the pre- vious 36 months before resuming inde- pendent  interpretation.

- Radiologic technologists. All mam- mographic examinations shall be per- formed by radiologic technologists who meet the following general require- ments, mammography requirements, and continuing education and experi- ence requirements:

  - General requirements. (A) Be li- censed to perform general radiographic procedures in a State; or

- Have general certification from one of the bodies determined by FDA to have procedures and requirements adequate to ensure that radiologic technologists certified by the body are competent to perform radiologic ex- aminations; and

  - Mammography requirements. Have, prior to April 28, 1999, qualified as a radiologic technologist under para- graph (a)(2) of this section of FDA’s in- terim regulations of December 21, 1993, or completed at least 40 contact hours of documented training specific to mammography under the supervision of  a  qualified  instructor.  The  hours  of

#### § 900.12

documented training shall include, but not necessarily be limited to:

    - Training in breast anatomy and physiology, positioning and compres- sion, quality assurance/quality control techniques, imaging of patients with breast  implants;

    - The performance of a minimum of 25 examinations under the direct super- vision of an individual qualified under paragraph (a)(2) of this section; and

    - At least 8 hours of training in each mammography modality to be used by the technologist in performing mammography exams; and

  - Continuing education require- ments. (A) Following the third anniver- sary date of the end of the calendar quarter in which the requirements of paragraphs (a)(2)(i) and (a)(2)(ii) of this section were completed, the radiologic technologist shall have taught or com- pleted at least 15 continuing education units in mammography during the 36 months immediately preceding  the date of the facility’s annual MQSA in- spection or the last day of the calendar quarter preceding the inspection or any date in between the two. The facility will choose one of these dates to deter- mine the 36-month period.

  - Units earned through teaching a specific course can be  counted  only once towards the 15 required in para- graph (a)(2)(iii)(A) of this section, even if the course is taught multiple times during the previous 36 months.

  - At least six of the continuing edu- cation units required in paragraph (a)(2)(iii)(A) of this section shall be re- lated to each mammographic modality used by the technologist.

  - Requalification. Radiologic tech- nologists who fail to meet the con- tinuing education requirements of paragraph (a)(2)(iii)(A) of this section shall obtain a sufficient number of con- tinuing education units in mammog- raphy to bring their total up to at least 15 in the previous 3 years, at least 6 of which shall be related to each modality used by the technologist in mammog- raphy. The technologist may not  re- sume performing unsupervised mam- mography examinations until the con- tinuing education requirements are completed.

  - Before a radiologic technologist may   begin   independently   performing

#### 21 CFR Ch. I (4–1–17 Edition)

mammographic examinations using a mammographic  modality  other  than one of those for which the technologist received training under paragraph (a)(2)(ii)(C) of this section, the tech- nologist shall have at least 8 hours of continuing education units in the new modality.

  - Continuing experience require- ments. (A) Following the second anni- versary date of the end of the calendar quarter in which the requirements of paragraphs (a)(2)(i) and (a)(2)(ii) of this section were completed or of April 28, 1999, whichever is later, the radiologic technologist shall have performed a minimum of 200 mammography exami- nations during the 24 months imme- diately preceding the date of the facili- ty’s annual inspection or  the  last  day of the calendar quarter preceding the inspection or any date in between the two. The facility will  choose  one  of these dates to determine the 24-month period.

(B) Requalification. Radiologic tech- nologists who fail to meet the con- tinuing experience requirements of paragraph (a)(2)(iv)(A) of this section shall perform a minimum of 25 mam- mography examinations under the di- rect supervision of a  qualified radiologic technologist, before resum- ing the performance of unsupervised mammography  examinations.

- Medical physicists. All medical physicists conducting surveys of mam- mography facilities and providing over- sight of the facility quality assurance program under paragraph (e) of this section shall meet the following:

- Initial qualifications. (A) Be State licensed or approved or have certifi- cation in an appropriate specialty area by one of the bodies determined by FDA to have procedures and require- ments to ensure that medical physi- cists certified by the body are com- petent to perform physics survey; and

(B)(1) Have a masters degree or high- er in a physical science from an accred- ited institution, with no less than 20 semester hours or equivalent (e.g., 30 quarter hours) of college under- graduate or graduate level physics;

- Have 20 contact hours of docu- mented specialized training in con- ducting surveys of mammography fa- cilities;  and

#### Food and Drug Administration, HHS	§ 900.12

- Have the experience of conducting surveys of at least 1 mammography fa- cility and a total of at least 10 mam- mography units. No more than one sur- vey of a specific unit within a period of

60 days can  be  counted  towards  the total mammography unit survey re- quirement. After April 28, 1999, experi- ence conducting surveys must be ac- quired under the direct supervision of a medical physicist who meets all the re- quirements of paragraphs (a)(3)(i) and (a)(3)(iii) of this section; or

- Alternative  initial  qualifications.

- Have qualified as a medical physi- cist under paragraph (a)(3) of this sec- tion of FDA’s interim regulations and retained that qualification by mainte- nance of the active status of any licen- sure, approval, or certification re- quired under the interim regulations; and

- Prior to the April 28, 1999, have:

  - A bachelor’s degree or higher in a physical science from an accredited in- stitution with no less than 10 semester hours or equivalent of college under- graduate or graduate level physics,

  - Forty contact hours of docu- mented specialized training in con- ducting surveys of mammography fa- cilities  and,

  - Have the experience of conducting surveys of at least 1 mammography fa- cility and a total of at least 20 mam- mography units. No more than one sur- vey of a specific unit within a period of

60 days can be counted towards the total mammography unit survey re- quirement. The training and experience requirements must be met after ful- filling the degree requirement.

- Continuing qualifications. (A) Continuing education. Following the third anniversary date of the end of the calendar quarter in which the require- ments of paragraph (a)(3)(i) or (a)(3)(ii) of this section were completed, the medical physicist shall have taught or completed at least 15 continuing edu- cation units in mammography during the 36 months immediately preceding the date of the facility’s annual inspec- tion or the last day of the calendar quarter preceding the inspection or any date in between the two. The facility shall choose one of these dates to de- termine the 36-month period. This con- tinuing  education  shall  include  hours

of training appropriate to each mam- mographic modality evaluated by the medical physicist during his or her sur- veys or oversight of quality assurance programs. Units earned through teach- ing a specific course can be counted only once towards the required 15 units in a 36-month period, even if the course is taught multiple times during the 36 months.

- Continuing experience.  Following the second anniversary date  of  the  end of the calendar quarter in which the re- quirements of paragraphs (a)(3)(i) and (a)(3)(ii) of this section were completed or of  April 28,  1999,  whichever is  later, the medical physicist shall have sur- veyed at least two mammography fa- cilities and a total of at least six mam- mography units during the 24 months immediately preceding the date of the facility’s  annual   MQSA   inspection   or the last day of the calendar quarter preceding the inspection or any date in between the two.  The  facility  shall choose one of these dates to  determine the 24-month period. No more than one survey of a specific facility within a 10- month period  or  a  specific  unit  within a period of 60 days can be counted to- wards  this  requirement.

- Before a medical physicist may begin independently performing mam- mographic surveys of a new mammo- graphic modality, that is, a mammo- graphic modality other than one for which the physicist received  training to qualify under paragraph (a)(3)(i) or (a)(3)(ii) of this section, the physicist must receive at least  8 hours  of train- ing in surveying units of the new mam- mographic  modality.

- Reestablishing qualifications. Med- ical physicists who fail to maintain the required continuing qualifications of paragraph (a)(3)(iii) of this section may not perform the MQSA surveys without the supervision of a qualified medical physicist. Before independently  sur- veying another facility, medical physi- cists must reestablish their qualifica- tions, as follows:

  - Medical physicists who fail to meet the continuing educational re- quirements of paragraph (a)(3)(iii)(A) of this section shall obtain a sufficient number of continuing education  units to bring their total units up to the re- quired 15 in the previous 3 years.

#### § 900.12

  - Medical physicists who fail to meet the continuing experience re- quirement of paragraph (a)(3)(iii)(B) of this section shall complete a sufficient number of surveys under the direct su- pervision of a medical physicist who meets the qualifications of paragraphs (a)(3)(i) and (a)(3)(iii) of this section to bring their total surveys up to the re- quired two facilities and six units  in the previous 24 months. No more than one survey of a specific unit within a period of 60 days can be counted to- wards the total mammography unit survey requirement.

  - Retention of personnel records. Fa- cilities shall maintain records to docu- ment the qualifications of all personnel who worked at the facility as inter- preting physicians, radiologic tech- nologists, or medical physicists. These records must be available for review by the MQSA inspectors. Records of per- sonnel no longer employed by the facil- ity should not be discarded until the next annual inspection has been com- pleted and FDA has determined that the facility is in compliance with the MQSA personnel requirements.

- Equipment. Regulations published under §§ 1020.30, 1020.31, and 900.12(e) of this chapter that are relevant to equip- ment performance should also be con- sulted for a more complete under- standing of the equipment performance requirements.

  - Prohibited equipment. Radiographic equipment designed for general purpose or special  nonmammography  proce- dures shall not be used for mammog- raphy. This prohibition includes sys- tems that  have been modified or equipped with special attachments for mammography. This requirement su- persedes  the  implied  acceptance  of such systems in § 1020.31(f)(3) of this chapter.

  - General. All radiographic equip- ment used for mammography shall be specifically  designed   for   mammog- raphy  and  shall  be   certified   pursuant to § 1010.2 of this chapter as  meeting the applicable requirements of §§ 1020.30 and 1020.31 of this chapter in  effect  at the  date  of  manufacture.

  - Motion of tube-image receptor assem- bly. (i) The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any

#### 21 CFR Ch. I (4–1–17 Edition)

such position, it shall not undergo un- intended motion.

(ii) The mechanism ensuring compli- ance with paragraph (b)(3)(i) of this section shall not fail in the event of power interruption.

  - Image receptor sizes. (i) Systems using screen-film image receptors shall provide, at a minimum,  for  operation with image receptors of 18 24 centi- meters (cm) and 24 30 cm.

- Systems using screen-film image receptors shall be equipped with mov- ing grids matched to all image receptor sizes provided.

- Systems used for magnification procedures shall be capable of oper- ation with the grid removed from be- tween the source and image receptor.

  - Light fields. For any mammog- raphy system with a light beam that passes through the x-ray beam-limiting device, the light shall provide an aver- age illumination  of  not less than  160 lux (15 foot candles) at 100 cm or the maximum source-image receptor dis- tance (SID), whichever is less.

  - Magnification. (i) Systems used to perform noninterventional problem solving procedures shall have radio- graphic magnification capability avail- able for use by the operator.

(ii) Systems used for magnification procedures shall provide, at a min- imum, at least one magnification value within the range of 1.4 to 2.0.

  - Focal spot selection. (i) When more than one focal spot is provided, the sys- tem shall indicate, prior to exposure, which focal spot is selected.

- When more than one target mate- rial is provided, the system shall indi- cate, prior to exposure, the preselected target   material.

- When the target material and/or focal spot is selected by a system algo- rithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure.

  - Compression. All mammography systems shall incorporate a compres- sion device.

    - Application of compression. Effec- tive October 28, 2002, each system shall provide:

      - An initial power-driven compres- sion  activated  by  hands-free  controls

#### Food and Drug Administration, HHS	§ 900.12

operable from both sides of the patient; and

      - Fine adjustment compression controls operable from both sides of the patient.

    - Compression paddle. (A) Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system. Compression paddles for special purposes, including those smaller than the full size of the image receptor (for ‘‘spot compres- sion’’) may be provided. Such compres- sion paddles for special purposes are not subject to the requirements of paragraphs (b)(8)(ii)(D) and (b)(8)(ii)(E) of this section.

- Except as provided in paragraph (b)(8)(ii)(C) of this section, the com- pression paddle shall be flat and par- allel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when com- pression is applied.

- Equipment intended by the manu- facturer’s design to not be flat and par- allel to the breast support table during compression shall meet the manufac- turer’s design specifications and main- tenance  requirements.

- The chest wall edge of the com- pression paddle shall be straight and parallel to the edge of the image recep- tor.

- The chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image.

  - Technique factor selection and dis- play. (i) Manual selection of milli- ampere seconds (mAs) or at least one of its component parts (milliapere (mA) and/or time) shall be available.

  - The technique factors (peak tube potential in kilovolt (kV)  and  either tube current in mA and exposure time in seconds or the product of tube cur- rent and exposure time in mAs) to be used during an exposure shall be indi- cated before the exposure begins, ex- cept when automatic exposure controls (AEC) are used, in which case the tech- nique factors that are set prior to the exposure shall be indicated.

  - Following AEC mode use, the system shall indicate the actual kilovoltage peak (kVp) and mAs used

during  the  exposure.  The  mAs  may  be displayed as mA and time.

  - Automatic exposure control. (i) Each screen-film system shall provide an AEC mode that is operable in all combinations of equipment configura- tion provided, e.g., grid, nongrid; mag- nification, nonmagnification; and var- ious target-filter combinations.

- The positioning or selection of the detector shall permit flexibility in the placement of the detector under the target tissue.

  - The size and available  positions of the detector shall be clearly indi- cated at the X-ray input surface of the breast compression paddle.

  - The selected position of the de- tector shall be clearly indicated.

- The system shall provide means for the operator to vary the selected optical density from the normal (zero) setting.

  - X-ray film. The facility shall use X-ray film for mammography that has been designated by the film manufac- turer as appropriate for mammog- raphy.

  - Intensifying screens. The facility shall use intensifying screens for mam- mography that have been designated by the screen manufacturer as appro- priate for mammography and shall use film that is matched to the screen’s spectral output as  specified by the manufacturer.

  - Film processing solutions. For processing mammography films, the fa- cility shall use chemical solutions that are capable of developing the films used by the facility in a manner equiv- alent to the minimum requirements specified by the film manufacturer.

  - Lighting. The facility shall make special lights  for  film  illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the inter- preting physicians.

  - Film masking devices. Facilities shall ensure that film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are avail- able to all interpreting physicians in- terpreting for the facility.

- Medical records and mammography reports—(1) Contents and terminology. Each   facility   shall   prepare   a   written

#### § 900.12

report of the results of each mammog- raphy examination performed under its certificate. The mammography report shall include the following informa- tion:

- The name of the patient and an additional patient identifier;

- Date of examination;

- The name of the interpreting physician who interpreted the mammo- gram;

- Overall final assessment of find- ings, classified in one of the following categories:

  - ‘‘Negative:’’ Nothing to comment upon (if the interpreting physician is aware of clinical findings or symptoms, despite the negative assessment, these shall be explained);

  - ‘‘Benign:’’ Also a negative assess- ment;

  - ‘‘Probably Benign:’’ Finding(s) has a high probability of being benign;

  - ‘‘Suspicious:’’ Finding(s) without all the characteristic morphology of breast cancer but indicating a definite probability of being malignant;

  - ‘‘Highly suggestive of malig- nancy:’’ Finding(s) has a high prob- ability of being malignant;

- In cases where no final assess- ment category can be assigned due to incomplete work-up, ‘‘Incomplete: Need additional imaging evaluation’’ shall be assigned as an assessment and reasons why no assessment can be made shall be stated by the inter- preting physician; and

- Recommendations made to the health care provider about what addi- tional actions, if any, should  be  taken. All clinical questions raised by the re- ferring health care provider shall be addressed in the report to the extent possible, even if the assessment is neg- ative or benign.

- Communication of mammography re- sults to the patients. Each facility shall send each patient a summary of the mammography report written in lay terms within 30 days of the mammo- graphic  examination.  If  assessments are ‘‘Suspicious’’ or ‘‘Highly  suggestive of  malignancy,’’  the   facility   shall make reasonable attempts to  ensure that the results are communicated to the patient as soon as possible.

- Patients who do not name a health   care   provider   to   receive   the

#### 21 CFR Ch. I (4–1–17 Edition)

mammography report shall be sent the report described in paragraph (c)(1) of this section within 30  days,  in  addition to the written notification of results  in lay   terms.

- Each facility that accepts pa- tients who do not have a health care provider shall maintain a system for referring such patients to a health care provider when clinically indicated.

- Communication of mammography re- sults to health care providers. When the patient has a referring health care pro- vider or the patient has named a health care provider, the facility shall:

  - Provide a written report of the mammography examination, including the items listed in paragraph (c)(1) of this section, to that health care pro- vider as  soon as  possible, but no  later than 30 days from the date of the mam- mography examination; and

  - If the assessment is ‘‘Suspicious’’ or ‘‘Highly suggestive of malignancy,’’ make reasonable attempts to commu- nicate with the health care provider as soon as possible, or if the health care provider is unavailable, to a respon- sible designee of the health care pro- vider.

- Recordkeeping. Each facility that performs mammograms:

  - Shall (except as provided in para- graph (c)(4)(ii) of this section) main- tain mammography films and  reports in a permanent medical record of the patient for a period of not  less than 5 years, or not less than 10 years if no ad- ditional mammograms  of  the  patient are performed at the facility, or a longer period if mandated by State or local law; and

  - Shall upon request by, or on be- half of, the patient, permanently or temporarily transfer the original mam- mograms and copies of the patient’s re- ports to a medical institution, or to a physician or health care provider of the patient, or to the patient directly;

  - Any fee charged to the patients for providing the services in paragraph (c)(4)(ii) of this section shall not exceed the documented costs associated with this  service.

- Mammographic image identification. Each mammographic image shall have the  following  information  indicated  on

#### Food and Drug Administration, HHS	§ 900.12

it in a permanent, legible, and unam- biguous manner and placed so as not to obscure anatomic structures:

  - Name of patient and an additional patient identifier.

  - Date of examination.

  - View and laterality. This informa- tion shall  be placed on  the image in  a position near the axilla. Standardized codes specified by the accreditation body and approved by FDA in accord- ance with § 900.3(b) or § 900.4(a)(8) shall be used to identify view and laterality.

  - Facility name and location. At a minimum, the location shall include the city, State, and zip code of the fa- cility.

  - Technologist identification.

  - Cassette/screen  identification.

  - Mammography unit identifica- tion, if there is more than one unit in the  facility.

- Quality assurance—general. Each facility shall establish and maintain a quality assurance program to ensure the safety, reliability, clarity, and ac- curacy of mammography services per- formed at the facility.

  - Responsible individuals. Responsi- bility for the quality assurance  pro- gram and for each of its elements shall be assigned to individuals who are qualified for their assignments and who shall be allowed adequate time to per- form these duties.

- Lead interpreting physician. The fa- cility shall identify a lead interpreting physician who shall have the general responsibility of ensuring that the quality assurance program meets all requirements of paragraphs (d) through

(f) of this section. No other individual shall be assigned or shall retain respon- sibility for quality assurance tasks un- less the lead interpreting physician has determined that the individual’s quali- fications for, and performance of, the assignment are adequate.

- Interpreting physicians. All inter- preting physicians interpreting mam- mograms for the facility shall:

  - Follow the facility procedures for corrective action when the images they are asked to interpret are of poor qual- ity, and

  - Participate in the facility’s med- ical outcomes audit program.

- Medical physicist. Each facility shall  have  the  services  of  a  medical

physicist available to survey mammog- raphy equipment and oversee the equipment-related quality assurance practices of the facility. At a  min- imum, the medical physicist(s) shall be responsible for performing the surveys and mammography equipment evalua- tions and  providing  the  facility with the reports described in paragraphs (e)(9) and (e)(10) of this section.

- Quality control technologist. Re- sponsibility for all individual tasks within the quality assurance program not assigned to the lead interpreting physician or the medical physicist shall be assigned to a quality control technologist(s). The tasks are to be performed by the quality control tech- nologist or by other personnel qualified to perform the tasks. When other per- sonnel are utilized for these tasks, the quality control technologist shall en- sure that the tasks are completed in such a way as to meet the require- ments of paragraph (e) of this section.

  - Quality assurance records. The lead interpreting physician, quality control technologist, and medical physicist shall ensure that records concerning mammography technique and proce- dures, quality control (including moni- toring data, problems detected by anal- ysis of that data, corrective actions, and the effectiveness of the correction actions), safety, protection, and em- ployee qualifications to meet assigned quality assurance tasks are properly maintained and updated. These quality control records shall be kept for each test specified in paragraphs (e) and (f) of this section until the next annual in- spection has been completed and FDA has determined that the facility is in compliance with the quality assurance requirements or until the test has been performed two additional times at the required frequency, whichever is longer.

- Quality assurance—equipment—(1) Daily quality control tests. Film proc- essors used to develop mammograms shall be adjusted and maintained  to meet the technical development speci- fications for the mammography film in use. A processor performance test shall be performed on each day that clinical films are processed before any clinical films are processed that day. The test shall   include   an   assessment   of   base

#### § 900.12

plus fog density, mid-density, and den- sity difference,  using  the  mammog- raphy film used clinically at the facil- ity.

- The base plus fog density shall be within + 0.03 of the established oper- ating  level.

- The  mid-density  shall  be  within

0.15 of the established operating level.

- The density difference shall be within 0.15 of the established oper- ating  level.

- Weekly quality control tests. Facili- ties with screen-film systems shall per- form an image quality evaluation test, using an FDA-approved phantom, at least  weekly.

  - The optical density of the film at the center of an image of a standard FDA-accepted phantom shall be at least 1.20 when exposed under a typical clinical condition.

  - The optical density of the film at the center of the phantom image shall not change by more than 0.20 from the established operating level.

  - The phantom image shall achieve at least the minimum score es- tablished by the accreditation  body and accepted by  FDA  in  accordance with § 900.3(d) or § 900.4(a)(8).

  - The density difference between the background of the phantom and an added test object, used to assess image contrast, shall be measured and shall not vary by more than 0.05 from the established operating level.

- Quarterly quality control tests. Fa- cilities with screen-film systems shall perform the following quality control tests at least quarterly:

  - Fixer retention in film. The residual fixer shall be no more than 5 micrograms per square cm.

  - Repeat analysis. If the total repeat or reject rate changes from the pre- viously  determined  rate  by  more  than

2.0 percent of the  total  films  included in the analysis, the reason(s) for the change shall be determined. Any cor- rective actions shall be recorded and the results of these corrective actions shall be assessed.

- Semiannual quality control tests. Facilities with screen-film  systems shall perform the following quality control tests at least semiannually:

- Darkroom fog. The optical density attributable to darkroom fog shall not

#### 21 CFR Ch. I (4–1–17 Edition)

exceed 0.05 when a mammography film of the type used in the facility,  which has a mid-density of no less than  1.2 OD, is exposed to typical darkroom conditions for 2  minutes  while  such film is placed on the counter top emul- sion side up. If the darkroom has a safelight used for  mammography  film, it shall be on during this test.

- Screen-film contact. Testing for screen-film contact shall be conducted using 40 mesh copper screen. All cas- settes used in the facility for mammog- raphy shall be tested.

- Compression   device   performance.

- A compression force of at least 111 newtons (25 pounds) shall be provided.

- Effective October 28, 2002, the maximum compression force  for  the initial power drive shall be between 111 newtons (25 pounds) and 200 newtons (45 pounds).

- Annual quality control tests. Facili- ties with screen-film systems shall per- form the following quality control tests at least annually:

  - Automatic exposure control perform- ance. (A) The AEC shall be capable of maintaining film optical density  with- in 0.30 of the mean optical density when thickness of a homogeneous ma- terial is varied over a range  of  2  to  6 cm and the kVp is varied appropriately for such thicknesses over the kVp range used clinically in the facility. If this requirement cannot be met, a technique chart shall be developed showing appropriate techniques (kVp and density control settings) for dif- ferent breast thicknesses and composi- tions that must be used so that optical densities within 0.30 of the average under phototimed conditions can be produced.

- After October 28, 2002, the AEC shall be capable of maintaining film optical density (OD) within 0.15 of the mean optical density when thickness of a homogeneous material is varied over a range of 2 to 6 cm and the kVp is var- ied appropriately for such thicknesses over the kVp range used clinically  in the   facility.

- The optical density of the film in the center of the phantom image shall not be less than 1.20.

  - Kilovoltage peak (kVp) accuracy and  reproducibility.  (A)  The  kVp  shall

#### Food and Drug Administration, HHS	§ 900.12

be accurate within 5 percent of the in- dicated or selected kVp at:

    - The lowest clinical kVp that can be measured by a kVp test device;

    - The most commonly used clinical kVp;

    - The   highest   available   clinical kVp, and

(B) Focal spot dimensions. Measured values of the focal spot length (dimen- sion parallel to the anode cathode axis) and width (dimension perpendicular to the anode cathode axis) shall be within the tolerance limits specified  in  table 1.

TABLE 1

(B) At the most commonly used clin-

ical settings of kVp, the coefficient of variation of reproducibility of the kVp shall be equal to or less than 0.02.

  - Focal spot condition. Until Octo- ber 28, 2002, focal spot condition shall be evaluated either by determining sys- tem resolution or by measuring focal spot dimensions. After October 28, 2002, facilities shall evaluate focal spot con- dition only by determining the system resolution.

- System resolution. (1) Each X-ray system used for mammography, in combination with the mammography screen-film combination used in the fa- cility, shall provide a minimum resolu- tion of 11 Cycles/millimeter (mm) (line- pairs/mm) when a high contrast resolu- tion bar test pattern is oriented  with the  bars  perpendicular  to  the  anode-

Focal Spot Tolerance Limit

| Nominal Focal Spot Size (mm) | Maximum Measured Dimensions |
| --- | --- |
|  | Width(mm) | Length(mm) |
| 0.10 | 0.15 | 0.15 |
| 0.15 | 0.23 | 0.23 |
| 0.20 | 0.30 | 0.30 |
| 0.30 | 0.45 | 0.65 |
| 0.40 | 0.60 | 0.85 |
| 0.60 | 0.90 | 1.30 |

  - Beam quality and half-value layer (HVL). The HVL shall meet the speci- fications of § 1020.30(m)(1) of this chap- ter for the minimum HVL. These val- ues, extrapolated to the mammo- graphic range, are shown in table 2. Values not shown in table 2 may be de- termined by linear interpolation or ex- trapolation.

TABLE 2

cathode axis, and a minimum resolu-

tion of 13 line-pairs/mm when the bars are parallel to that axis.

- The bar pattern shall be placed 4.5 cm above the breast support surface, centered with respect to the chest wall edge of the image receptor, and  with the edge of the pattern within 1 cm of the chest wall edge of the image recep- tor.

- When more than one target mate- rial is provided, the measurement in paragraph (e)(5)(iii)(A) of this section shall be made using the appropriate focal spot for each target material.

- When more than one SID is pro- vided, the test shall be performed at SID most commonly used clinically.

- Test kVp shall be set at the value used clinically by the facility for a standard breast and shall be performed in the AEC mode, if available. If nec- essary, a suitable absorber may be placed in the beam to increase expo- sure times. The screen-film cassette combination used by the facility shall be used to test for this requirement and shall be placed in the normal loca- tion used for clinical procedures.

X-ray Tube Voltage (kilovolt peak) and Minimum HVL

| Designed Op- erating Range (kV) | Measured Operating Voltage (kV) | Minimum HVL (milli- meters of aluminum) |
| --- | --- | --- |
| Below 50 | 20 25 30 | 0.20 0.25 0.30 |

  - Breast entrance air kerma and AEC reproducibility. The coefficient of vari- ation for both air kerma and mAs shall not exceed 0.05.

  - Dosimetry. The average glandular dose delivered during a single cranio- caudal view of an FDA-accepted phan- tom simulating a standard breast shall not exceed 3.0 milligray (mGy) (0.3 rad) per exposure. The dose shall be deter- mined with technique factors and con- ditions used clinically for a standard breast.

  - X-ray field/light field/image recep- tor/compression paddle alignment. (A) All systems shall have beam-limiting de- vices that allow the entire chest wall edge of the x-ray field to extend to the chest wall edge of the image receptor and provide means to assure that the x- ray field does not extend beyond any

#### § 900.12

edge  of  the  image  receptor  by  more than 2 percent of the SID.

- If a light field that passes through the X-ray beam limitation de- vice is provided,  it  shall  be  aligned with the X-ray field so that the total of any misalignment of the edges of the light field and the X-ray field along ei- ther the length or the width of the vis- ually defined field at the plane of the breast support surface shall not exceed 2 percent of the SID.

- The chest wall edge of the com- pression paddle shall not extend be- yond the chest wall edge of the image receptor by more than one percent of the SID when tested with the compres- sion paddle placed above the breast support surface  at a  distance equiva- lent to standard breast thickness. The shadow of the vertical edge of the com- pression paddle shall not be visible on the image.

  - Uniformity of screen speed. Uni- formity of screen speed of all the cas- settes in the facility shall be tested and the difference between the max- imum and minimum optical densities shall not exceed 0.30. Screen artifacts shall also be evaluated during this test.

  - System artifacts. System artifacts shall be evaluated with a high-grade, defect-free sheet of homogeneous mate- rial large enough to cover the mam- mography cassette and shall be per- formed for all cassette sizes used in the facility using a grid appropriate for the cassette size being tested. System arti- facts shall also be evaluated for all available focal spot sizes and target fil- ter combinations used clinically.

  - Radiation output. (A) The system shall be capable of producing  a  min- imum output of 4.5 mGy air kerma per second (513 milli Roentgen (mR) per second) when operating at 28 kVp  in the standard  mammography  (moly/ moly) mode at any SID where the  sys- tem is designed to operate and when measured by a detector with its center located 4.5 cm above the breast support surface with the compression paddle in place between the source and the detec- tor. After October 28, 2002, the system, under the same measuring conditions shall be capable of producing  a  min- imum output of 7.0 mGy air kerma per second (800 mR per second) when oper- ating  at  28  kVp  in  the  standard  (moly/

#### 21 CFR Ch. I (4–1–17 Edition)

moly) mammography mode at any SID where the system is designed to oper- ate.

(B) The system shall be capable of maintaining the required minimum ra- diation output averaged over a 3.0 sec- ond period.

  - Decompression. If the system is equipped with a provision for auto- matic decompression after completion of an exposure or interruption of power to the system, the system shall be test- ed to confirm that it provides:

- An override capability to allow maintenance of compression;

- A continuous display of the over- ride status; and

- A manual emergency compression release that can be activated  in  the event of power or automatic release failure.

- Quality control tests—other modali- ties. For systems with image receptor modalities other than screen-film, the quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor manufacturer, ex- cept that the maximum allowable dose shall not exceed the maximum allow- able dose for screen-film systems in paragraph (e)(5)(vi) of this section.

- Mobile units. The facility shall verify that mammography units  used to produce mammograms at more than one location meet the requirements in paragraphs (e)(1) through (e)(6) of this section. In addition, at each examina- tion location, before any examinations are conducted, the facility shall verify satisfactory performance of such units using a test method that establishes the adequacy of the image quality pro- duced by the unit.

- Use of test results. (i) After comple- tion of the tests specified in paragraphs (e)(1) through (e)(7) of this section, the facility shall compare the test results to the corresponding specified action limits; or, for nonscreen-film modali- ties, to the manufacturer’s rec- ommended action limits; or, for post- move, preexamination testing of mo- bile units, to the limits established in the test method used by the facility.

- If the test results fall outside of the action limits, the source of the problem shall be identified and correc- tive actions shall be taken:

#### Food and Drug Administration, HHS	§ 900.12

  - Before any further examinations are performed or any films are proc- essed using a component of the mam- mography system that failed any of the tests described in paragraphs (e)(1), (e)(2), (e)(4)(i), (e)(4)(ii), (e)(4)(iii), (e)(5)(vi), (e)(6), or (e)(7) of this section;

  - Within 30 days of the test date for all  other  tests  described  in  paragraph

- of this section.

- Surveys. (i) At least once a year, each facility shall undergo a survey by a medical physicist or by an individual under the direct supervision of a med- ical physicist. At a minimum, this sur- vey shall include the performance of tests to ensure that the facility  meets the quality assurance requirements of the annual tests described in para- graphs (e)(5) and (e)(6) of this section and the weekly phantom image quality test described in paragraph (e)(2) of this  section.

- The results of all tests conducted by the facility in accordance with para- graphs (e)(1) through (e)(7) of this sec- tion, as well as written documentation of any corrective actions taken and their results, shall be evaluated for adequacy by the medical physicist per- forming the survey.

- The medical physicist shall pre- pare a survey report that includes a summary of this review and rec- ommendations for necessary improve- ments.

- The survey report shall be sent to the facility within 30 days of the date of the survey.

- The survey report shall be dated and signed by the medical physicist performing or supervising  the  survey. If the survey was performed entirely or in part by another individual under the direct supervision of the medical physi- cist, that individual and the part of the survey that individual performed shall also be identified in the survey report.

- Mammography equipment evalua- tions. Additional evaluations of mam- mography units or image processors shall be conducted whenever a new unit or processor is installed, a unit or proc- essor is disassembled and reassembled at the same or a new location, or major components of a mammography unit or processor equipment are changed or re- paired. These evaluations shall be used to   determine   whether   the    new    or

changed equipment meets the require- ments of applicable standards in para- graphs (b) and (e) of this section. All problems shall be corrected before the new or changed equipment is put into service for examinations or film proc- essing. The mammography equipment evaluation shall be performed by a medical physicist or by an individual under the direct supervision of a med- ical   physicist.

- Facility cleanliness. (i) The facil- ity shall establish and implement ade- quate protocols for maintaining dark- room, screen, and view box cleanliness.

(ii) The facility shall document that all cleaning procedures are performed at the frequencies specified in the pro- tocols.

- Calibration of air kerma measuring instruments. Instruments used by med- ical physicists in their  annual  survey to measure the air kerma or air kerma rate from a mammography unit  shall be calibrated at least once  every  2 years and each time the instrument is repaired. The instrument calibration must be traceable to a national stand- ard and calibrated with an accuracy of

6 percent (95 percent confidence level) in  the  mammography  energy  range.

- Infection control. Facilities shall establish and comply with a system specifying procedures to be followed by the facility for cleaning and dis- infecting mammography equipment after contact with blood or other po- tentially infectious materials. This system shall specify the methods for documenting facility compliance  with the infection control procedures estab- lished and shall:

- Comply with all applicable Fed- eral, State, and local regulations per- taining to infection control; and

- Comply with the manufacturer’s recommended procedures for the clean- ing and disinfection of the mammog- raphy equipment used in the facility; or

- If adequate manufacturer’s rec- ommendations are not available, com- ply with generally accepted guidance on infection control, until such rec- ommendations become available.

- Quality assurance-mammography medical   outcomes   audit.   Each   facility

#### § 900.12

shall establish and maintain a mam- mography medical outcomes audit pro- gram to followup positive mammo- graphic assessments and to correlate pathology results with the interpreting physician’s  findings.  This   program shall be designed to ensure the reli- ability, clarity, and accuracy of the in- terpretation of mammograms.

  - General requirements. Each facility shall establish a system to collect and review outcome data for all mammo- grams performed, including followup on the disposition of all positive mam- mograms and correlation of pathology results with the interpreting physi- cian’s mammography report. Analysis of these outcome data shall be made in- dividually and collectively for all in- terpreting physicians at the facility. In addition, any cases of breast cancer among women imaged at the facility that subsequently become known to the facility shall prompt the facility to initiate followup on surgical and/or pa- thology results and review of the mam- mograms taken prior to the diagnosis of a malignancy.

  - Frequency of audit analysis. The fa- cility’s first audit analysis shall be ini- tiated  no  later  than  12  months   after the date the facility  becomes  certified, or 12 months after April 28, 1999, which- ever date is the latest. This audit  anal- ysis shall be completed within an addi- tional 12 months to  permit  completion of diagnostic procedures and data col- lection. Subsequent audit analyses  will be conducted at least once every 12 months.

  - Audit interpreting physician. Each facility shall designate at least one in- terpreting physician to review the medical outcomes audit data at least once every 12 months. This individual shall record the dates of the audit pe- riod(s) and shall be responsible for ana- lyzing results based on this audit. This individual shall also be responsible for documenting the results and for noti- fying other interpreting physicians of their results and the facility aggregate results. If followup actions are taken, the audit interpreting physician shall also be responsible for documenting the nature of the followup.

- Mammographic procedure and tech- niques for mammography of patients with breast  implants.  (1)  Each  facility  shall

#### 21 CFR Ch. I (4–1–17 Edition)

have a procedure to inquire whether or not the patient has breast  implants prior to the actual  mammographic exam.

(2) Except where contraindicated, or unless modified by a physician’s direc- tions, patients with breast implants undergoing mammography shall have mammographic views to maximize the visualization of breast tissue.

- Consumer   complaint   mechanism.

Each facility shall:

  - Establish a  written  and docu- mented system for collecting and re- solving consumer complaints;

  - Maintain a record of each serious complaint received by the  facility  for at least 3 years from the date the com- plaint was received;

  - Provide the consumer with ade- quate directions for filing serious com- plaints with the facility’s accreditation body if the facility is unable to resolve a serious complaint to the consumer’s satisfaction;

  - Report unresolved serious com- plaints to the accreditation body in a manner and timeframe specified by the accreditation body.

(i) Clinical image quality. Clinical im- ages produced by any certified facility must continue to comply with the standards for clinical image quality es- tablished by that facility’s accredita- tion body.

(j) Additional mammography review and patient notification. (1) If FDA be- lieves that mammography quality at a facility has been compromised and may present a serious risk to human health, the facility shall provide clinical im- ages and other relevant information, as specified by FDA, for review by the ac- creditation body or other entity des- ignated by FDA. This additional mam- mography review will help the agency to determine whether the facility is in compliance with this section and, if not, whether there is a need to notify affected patients, their physicians, or the public that the reliability, clarity, and accuracy of interpretation of mam- mograms has been compromised.

- If FDA determines that the qual- ity of mammography performed by a facility, whether or not certified under

§ 900.11, was so inconsistent with the quality standards established in this section  as  to  present  a  significant  risk

#### Food and Drug Administration, HHS	§ 900.14

to individual or public health, FDA may require such facility to notify pa- tients who received mammograms at such facility, and their referring physi- cians, of the deficiencies presenting such risk, the potential harm result- ing, appropriate remedial measures, and such other relevant information as FDA may require. Such notification shall occur within a timeframe and in a manner specified by FDA.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10,  1997,  as  amended  at  63  FR  56558,  Oct. 22,

1998; 64 FR 18333, Apr. 14, 1999; 64 FR 32408,

June 17, 1999; 65 FR 43690, July 14, 2000]

##### § 900.13 Revocation  of  accreditation and revocation of  accreditation body approval.

  - FDA action following revocation of accreditation. If a facility’s accredita- tion is revoked by an accreditation body, the agency may conduct an in- vestigation into the reasons for the revocation. Following such investiga- tion, the agency may determine that the facility’s certificate shall no longer be in effect or the agency may take whatever other action or combination of actions will best protect the public health, including the establishment and implementation of a corrective plan of action that will permit the cer- tificate to continue in effect while the facility seeks reaccreditation. A facil- ity whose certificate is no longer in ef- fect because it has lost its accredita- tion may not practice mammography.

  - Withdrawal of FDA approval of an accreditation body. (1) If FDA withdraws approval of an accreditation  body under § 900.6, the certificates of facili- ties previously accredited by such body shall remain in effect for up  to 1 year from the date of the withdrawal of ap- proval, unless FDA  determines,  in order to protect human health or be- cause the accreditation body fraudu- lently accredited facilities, that the certificates of some or all of the facili- ties should be revoked or suspended or that a shorter time period should be es- tablished for  the  certificates to remain in  effect.

(2) After 1 year from the date of with- drawal of approval of an accreditation body, or within any shorter period of time established by the agency, the af- fected  facilities  must  obtain  accredita-

tion from another accreditation body, or from another entity designated by FDA.

##### § 900.14   Suspension  or  revocation  of certificates.

- Except  as  provided  in  paragraph

- of this section, FDA may suspend or revoke a certificate if FDA finds, after providing the owner or operator of the facility with notice and opportunity for an informal hearing in accordance with part 16  of this chapter, that the owner, operator, or any employee of the  facility:

  - Has been guilty of misrepresenta- tion  in  obtaining  the  certificate;

  - Has failed to comply with the standards of § 900.12;

  - Has failed to comply with reason- able requests of the agency or the ac- creditation body for records, informa- tion, reports, or materials that FDA believes are necessary to determine the continued eligibility  of  the  facility  for a certificate or continued compliance with the standards of § 900.12;

  - Has refused a reasonable  request of a duly designated FDA  inspector, State inspector, or accreditation body representative for permission to in- spect the facility or the operations and pertinent records of the facility;

  - Has violated or aided and abetted in the violation of any provision of or regulation promulgated pursuant to 42

U.S.C. 263b; or

  - Has failed to comply with prior sanctions imposed by the agency under 42 U.S.C. 263b(h).

- FDA may suspend the certificate of a facility before holding a hearing if FDA makes a finding described in para- graph (a) of this section and also deter- mines that;

  - The failure to comply with re- quired standards presents a  serious risk to human health;

  - The refusal to permit inspection makes immediate suspension nec- essary; or

  - There is reason to believe that the violation or aiding and abetting of the violation was intentional or associated with fraud.

- If FDA suspends a certificate in accordance with paragraph (b) of this section:

#### § 900.15

  - The agency shall provide the facil- ity with an opportunity for an informal hearing under part 16 of this chapter not later than 60 days from the  effec- tive date of this suspension;

  - The suspension shall remain in ef- fect  until  the  agency  determines  that:

    - Allegations of violations or mis- conduct were not substantiated;

    - Violations of required standards have been corrected to the agency’s satisfaction; or

    - The facility’s certificate is re- voked in accordance with paragraph (d) of this section;

- After providing a hearing in ac- cordance with paragraph (c)(1) of this section, the agency may revoke the fa- cility’s certificate if the agency deter- mines that the facility:

  - Is unwilling or unable to correct violations that were the basis for sus- pension; or

  - Has engaged in fraudulent activ- ity to obtain or continue certification.

##### § 900.15 Appeals of adverse accredita- tion or reaccreditation decisions that preclude certification or recer- tification.

- The appeals procedures described in this section are available only for adverse accreditation  or  reaccredita- tion decisions that preclude  certifi- cation or recertification by  FDA.  Agen- cy decisions to suspend or revoke cer- tificates that are already  in  effect  will be handled in accordance with § 900.14.

- Upon learning that a facility has failed to become accredited or re- accredited, FDA will notify the facility that the agency is unable to certify that facility without proof of accredi- tation.

- A facility that has been denied ac- creditation or reaccreditation is enti- tled to an appeals process from the ac- creditation  body,  in  accordance  with

§ 900.7.  A  facility  must  avail  itself   of the accreditation body’s appeal process before requesting reconsideration from FDA.

- A facility that cannot achieve satisfactory resolution of  an adverse accreditation decision through the ac- creditation body’s appeal process is en- titled to further appeal  in  accordance with procedures set forth in this sec- tion and in regulations published in 42 CFR part 498.

#### 21 CFR Ch. I (4–1–17 Edition)

- References to the Health Care Fi- nancing Administration (HCFA) in 42 CFR part 498 should be read as the Di- vision of Mammography Quality and Radiation Programs (DMQRP), Center for Devices and Radiological Health, Food and Drug Administration.

- References to the Appeals Council of  the  Social  Security   Administration in 42 CFR  part 498  should be  read as references to the Departmental Ap- peals Board.

- In accordance with the procedures set forth in subpart B of 42 CFR  part 498, a facility that has been denied ac- creditation following appeal to the ac- creditation body may request reconsid- eration of that adverse decision from DMQRP.

  - A facility must request reconsid- eration  by  DMQRP  within  60  days   of the accreditation  body’s  adverse  ap- peals decision, at the following address: Food  and  Drug  Administration,  Center for Devices and Radiological Health, Division of  Mammography  Quality and Radiation Programs, Attn: Facility Ac- creditation Review Committee,  10903 New Hampshire Ave., Bldg. 66, rm. 4521, Silver Spring, MD 20993–0002.

  - The request for reconsideration shall include three copies of the fol- lowing records:

    - The accreditation body’s original denial of accreditation.

    - All information the facility sub- mitted to the accreditation body as part of the appeals process;

    - A copy of the accreditation body’s adverse appeals decision; and

    - A statement of the basis for the facility’s disagreement with the ac- creditation body’s decision.

  - DMQRP will conduct its recon- sideration in accordance with the pro- cedures set forth in subpart B of 42 CFR part 498.

- A facility that is dissatisfied with DMQRP’s decision following reconsid- eration is entitled to  a  formal  hearing in accordance with procedures  set  forth in subpart D of 42 CFR part 498.

- Either the facility or FDA may re- quest review of the hearing officer’s de- cision.  Such  review  will  be   conducted by the Departmental Appeals Board in accordance with subpart E of 42 CFR part 498.

#### Food and Drug Administration, HHS	§ 900.18

- A facility cannot perform mam- mography services while an adverse ac- creditation decision is being appealed.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10,  1997,  as  amended  at  75  FR  20916,  Apr.  22,

2010]

##### § 900.16   Appeals  of  denials  of  certifi- cation.

- The appeals procedures described in this section are available only to fa- cilities that are denied certification by FDA after they have been  accredited by an approved accreditation body. Ap- peals for facilities that have failed to become accredited are governed by the procedures set forth in § 900.15.

- FDA may deny the application if the agency has reason to believe that:

  - The facility will not  be  operated in accordance with standards estab- lished under § 900.12;

  - The facility will not permit in- spections or provide access to records or information in a timely fashion; or

  - The facility has been guilty of misrepresentation in obtaining the ac- creditation.

(c)(1) If FDA denies an application for certification by a faciity that has re- ceived accreditation from an approved accreditation body, FDA shall provide the facility with a statement of the grounds on which the denial is based.

- A facility that has been denied ac- creditation may request reconsider- ation and appeal of FDA’s determina- tion in accordance with the applicable provisions of § 900.15(d).

##### § 900.17   [Reserved]

§ 900.18   Alternative   requirements  for

§ 900.12 quality standards.

  - Criteria for approval of alternative standards. Upon application by a quali- fied party as defined in paragraph (b) of this section, FDA may approve an al- ternative to a quality standard under

§ 900.12,   when   the   agency   determines that:

    - The proposed alternative standard will be at least as effective in assuring quality mammography as the standard it proposes to replace, and

    - The proposed alternative:

      - Is too limited in  its  applicability to justify an amendment to the stand- ard; or

      - Offers an expected benefit to human health that is so great that the time required for amending the stand- ard would present an unjustifiable risk to the human health; and

    - The granting of the  alternative  is in   keeping   with   the   purposes   of   42

U.S.C. 263b.

  - Applicants for alternatives. (1) Mammography facilities and accredita- tion bodies may apply  for  alternatives to the quality standards of § 900.12.

- Federal agencies and State gov- ernments that are not accreditation bodies may apply for alternatives to the standards of § 900.12(a).

- Manufacturers and assemblers of equipment used for mammography may apply for alternatives to  the  standards of § 900.12(b) and (e).

  - Applications for approval of an al- ternative standard. An application for approval of an alternative standard or for an amendment or extension of the alternative standard shall be  submitted in an original  and  two  copies  to  the Food  and  Drug  Administration,  Center for Devices and Radiological Health, Director, Division of Mammography Quality and Radiation Programs, 10903 New Hampshire Ave., Bldg. 66, rm. 4521, Silver Spring, MD 20993–0002. The appli- cation for approval of an alternative standard shall include the following in- formation:

- Identification of the original standard for which the alternative standard is being proposed and an ex- planation of why the applicant is pro- posing the alternative;

- A description of the manner in which the alternative is proposed to de- viate from the original standard;

- A description, supported by data, of the advantages to be derived from such deviation;

- An explanation, supported by data, of how such a deviation would en- sure equal or greater quality of produc- tion, processing, or interpretation of mammograms than the original stand- ard;

- The suggested period of time that the proposed alternative  standard would be in effect; and

- Such other information required by the Director to evaluate and act on the application.

#### § 900.18

  - Ruling on applications. (1) FDA may approve or deny, in whole or in part, a request for approval of an alter- native standard or any amendment or extension thereof, and shall inform the applicant in writing of this action. The written notice shall state  the  manner in which the requested alternative standard differs from the agency stand- ard and a summary of the reasons for approval or denial of the request. If the request is approved, the written notice shall also  include  the  effective  date and the termination date of the ap- proval and a summary of the limita- tions and conditions attached to the approval and any other  information that may be relevant to the approved request. Each approved alternative standard shall be assigned an identi- fying  number.

  - Notice of an approved request for an alternative standard or any amend- ment or extension thereof shall be placed in the public docket file in the Division of Dockets Management and may also be in the form of a notice published in the  FEDERAL  REGISTER. The notice shall state the name of the applicant, a description of the pub- lished agency standard, and a descrip- tion of the approved alternative stand- ard, including limitations and condi- tions attached to the approval of the alternative standard.

  - Summaries of the approval of al- ternative standards, including informa- tion on their nature and number, shall be provided to the National Mammog- raphy Quality Assurance Advisory Committee.

  - All applications for approval of al- ternative standards and for amend- ments and extensions thereof and all correspondence (including written no- tices of approval) on these applications shall be available for public disclosure in the Division of Dockets Manage- ment, excluding patient identifiers and confidential commercial information.

  - Amendment or extension of an alter- native standard. An application for amending or extending approval of an alternative standard shall include the following information:

    - The approval number and the ex- piration date of the alternative stand- ard;

#### 21 CFR Ch. I (4–1–17 Edition)

    - The amendment or extension re- quested and the basis for the amend- ment or extension; and

    - An explanation, supported by data, of how such an amendment or ex- tension would ensure equal or greater quality of production, processing, or interpretation of mammograms  than the original standard.

  - Applicability of the alternative standards. (1) Except as provided in paragraphs (f)(2) and (f)(3) of this sec- tion, any approval of an alternative standard, amendment,  or  extension may be implemented only by the entity to which it was granted and under the terms under which it was granted. Other entities interested in similar or identical approvals must file their own application following the procedures of paragraph (c) of this section.

- When an alternative standard is approved for a manufacturer of equip- ment, any facility using that equip- ment will also be covered by the alter- native standard.

- The agency may extend the alter- native standard to other entities when FDA determines that expansion of the approval of the alternative standard would be an effective means of pro- moting the acceptance of measures to improve the quality of mammography. All such determinations will be pub- licized by appropriate means.

  - Withdrawal of approval of alter- native requirements. FDA shall amend or withdraw approval of an alternative standard whenever the agency deter- mines that this action is necessary to protect the human health or otherwise is justified by § 900.12. Such action will become effective on the  date  specified in the written notice of the action sent to the applicant, except that it will be- come effective immediately upon noti- fication of the applicant when FDA de- termines that such action  is  necessary to prevent an imminent health hazard.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10,  1997,  as  amended  at  75  FR  20916,  Apr.  22,

2010]

## Subpart C—States as Certifiers

SOURCE: 67 FR 5467, Feb. 6, 2002, unless oth-

erwise noted.

#### Food and Drug Administration, HHS	§ 900.21

##### § 900.20   Scope.

The regulations set forth in this part implement the Mammography Quality Standards Act (MQSA) (42 U.S.C. 263b). Subpart C of this part establishes pro- cedures whereby a State can apply to become a FDA-approved certification agency to certify facilities within the State to perform mammography serv- ices. Subpart  C of  this part further  es- tablishes requirements and standards for State certification agencies to en- sure that all mammography facilities under their jurisdiction are adequately and consistently evaluated for compli- ance with quality standards at least as stringent as the national quality standards established by FDA.

##### § 900.21   Application for approval as a certification agency.

- Eligibility. State agencies may apply for approval as a certification agency if they have  standards  at  least as stringent as those of § 900.12, quali- fied personnel, adequate resources to carry out the States as Certifiers’ re- sponsibilities, and the  authority  to enter into a legal agreement with FDA to accept these responsibilities.

- Application for approval. (1) An ap- plicant seeking FDA approval as a cer- tification agency shall inform the Food and Drug Administration, Center for Devices and Radiological Health, Di- rector, Division  of  Mammography Quality and Radiation Programs, Attn: States as Certifiers Coordinator, 10903 New Hampshire Ave., Bldg. 66, rm. 4521, Silver Spring, MD 20993–0002, in writ- ing, of its desire to be approved as a certification     agency.

  - Following receipt of the written request, FDA will provide the applicant with additional information to aid in the submission of an application for ap- proval as a certification agency.

  - The applicant shall furnish to FDA, at the address in paragraph (b)(1) of this section, three copies of an appli- cation containing the  following  infor- mation, materials, and supporting doc- umentation:

- Name, address, and phone number of the applicant;

- Detailed description of the mam- mography quality standards the appli- cant will require facilities to meet and, for    those    standards    different    from

FDA’s quality standards, information substantiating that they are at least as stringent   as   FDA   standards   under

§ 900.12;

- Detailed description of the appli- cant’s review and decisionmaking proc- ess for facility certification, including:

  - Policies and procedures for noti- fying facilities of certificate  denials and expirations;

  - Procedures for monitoring and enforcement of the correction of defi- ciencies  by  facilities;

  - Policies and procedures for sus- pending or revoking a facility’s certifi- cation;

  - Policies and procedures that will ensure processing certificates within a timeframe approved by FDA;

  - A description of the appeals proc- ess for facilities contesting adverse certification status decisions;

  - Education, experience, and train- ing requirements of the applicant’s professional and supervisory staff;

  - Description of the applicant’s electronic data management and anal- ysis system;

  - Fee schedules;

  - Statement of policies and proce- dures established to avoid conflict of interest;

  - Description of the applicant’s mechanism for handling facility in- quiries and complaints;

  - Description of a plan to ensure that certified mammography facilities will be inspected according to MQSA (42 U.S.C. 263b) and procedures and policies for notifying facilities of in- spection  deficiencies;

  - Policies and procedures for moni- toring and enforcing the correction of facility deficiencies discovered during inspections or by other means;

  - Policies and procedures for addi- tional mammography review and for requesting such reviews from accredi- tation bodies;

  - Policies and  procedures  for  pa- tient  notification;

  - If a State has regulations that are   more   stringent   than   those   of

§ 900.12, an explanation of how adverse actions  taken  against  a  facility  under the more stringent regulations will be distinguished from  those  taken  under the requirements of § 900.12; and

#### § 900.22

  - Any other information that FDA identifies as necessary to make a de- termination on the approval of the State as a certification agency.

- Rulings on applications for ap- proval. (1) FDA will conduct a review and evaluation  to  determine  whether the applicant substantially meets the applicable requirements of this subpart and whether the certification stand- ards the applicant will require facili- ties to meet are the quality standards published under subpart B of this part or at least as stringent as those of sub- part B.

- FDA will notify the applicant of any deficiencies in the application and request that those deficiencies be cor- rected within a specified time period. If the deficiencies are not corrected to FDA’s satisfaction within the specified time period, FDA may deny the appli- cation for approval as a certification agency.

- FDA shall  notify the  applicant whether the application has been ap- proved or denied. The notification shall list any conditions associated with ap- proval or state the bases for any de- nial.

- The review of any application may include a meeting between FDA and representatives of the applicant at a time and location mutually  acceptable to FDA and the applicant.

- FDA will advise the applicant of the circumstances under which a de- nied application may be resubmitted.

- Scope of authority. FDA may limit the scope of certification authority delegated to the State in accordance with MQSA.

[62  FR 55976,  Oct. 28,  1997;  62  FR 60614,  Nov.

10,  1997,  as  amended  at  75  FR  20916,  Apr.  22,

2010]

##### § 900.22   Standards	for	certification agencies.

The certification agency shall accept the following responsibilities in order to ensure quality mammography at the facilities it certifies and shall perform these responsibilities in a manner that ensures the integrity and impartiality of the certification agency’s actions:

- Conflict of interest. The certifi- cation agency shall establish and im- plement measures that FDA has ap- proved in accordance with § 900.21(b) to

#### 21 CFR Ch. I (4–1–17 Edition)

reduce the possibility of conflict of in- terest or facility bias on the part of in- dividuals acting on the certification agency’s behalf.

- Certification and inspection respon- sibilities. Mammography facilities shall be certified and inspected in accord- ance with statutory and regulatory re- quirements that are at least as strin- gent as those of MQSA and this part.

- Compliance with quality standards. The scope, timeliness, disposition, and technical accuracy of completed in- spections and related enforcement ac- tivities shall ensure compliance with facility quality  standards  required under § 900.12.

- Enforcement actions. (1) There shall be appropriate criteria and proc- esses for the suspension and revocation of  certificates.

(2) There shall be prompt investiga- tion of and appropriate enforcement action for facilities performing mam- mography  without  certificates.

- Appeals. There  shall  be processes for facilities to appeal inspection find- ings, enforcement  actions,  and  adverse certification decision or adverse ac- creditation decisions  after exhausting appeals to the accreditation body.

- Additional mammography review. There shall be a process for the certifi- cation agency to request additional mammography review from accredita- tion bodies for issues related to mam- mography image quality and clinical practice.  The  certification  agency should request additional mammog- raphy review only when it believes that mammography quality at a facility has been compromised and may present a serious risk to human health.

- Patient notification. There shall be processes for the certification agency to conduct, or cause to be conducted, patient notifications should the certifi- cation agency determine that mam- mography quality has been com- promised to such an extent that it may present a serious risk to human health.

- Electronic data transmission. There shall be processes to ensure the timeli- ness and accuracy of electronic trans- mission of inspection data and facility certification status information in a format and timeframe determined by FDA.

#### Food and Drug Administration, HHS	§ 900.24

- Changes to standards. A certifi- cation agency shall obtain FDA au- thorization for any changes it proposes to make in any standard that FDA has previously accepted under § 900.21 be- fore requiring facilities to comply with the changes as a condition of obtaining or  maintaining  certification.

##### § 900.23   Evaluation.

FDA shall evaluate annually the per- formance of each certification agency. The evaluation shall include the use of performance indicators that address the adequacy of program performance in certification, inspection, and en- forcement activities. FDA will also consider any additional information deemed relevant by FDA that has been provided by the certification body or other sources or has been required by FDA as part of its oversight mandate. The evaluation also shall include a re- view of any changes in the standards or procedures   in   the   areas   listed   in

§§ 900.21(b) and 900.22 that have taken place since the original application or the last evaluation, whichever is most recent. The evaluation shall include a determination of whether there are major deficiencies in the certification agency’s regulations or performance that, if not corrected, would warrant withdrawal of the  approval of the cer- tification agency under the  provisions of § 900.24, or minor deficiencies that would require corrective action.

##### § 900.24   Withdrawal of approval.

If FDA determines, through the eval- uation activities of § 900.23, or through other means, that a certification agen- cy is not in substantial compliance with this subpart, FDA may initiate the following actions:

- Major deficiencies. If, after pro- viding notice and opportunity for cor- rective action, FDA determines that a certification agency has demonstrated willful disregard for public health, has committed fraud, has failed to provide adequate resources for  the  program, has submitted material false  state- ments to the agency, has failed to achieve the MQSA goals of quality mammography and access, or has per- formed or failed to perform a delegated function in a manner that may cause serious   risk   to   human   health,   FDA

may withdraw its approval of that cer- tification agency. The certification agency shall notify, within a time pe- riod and in a manner approved by FDA, all facilities certified or seeking cer- tification by it that it has been re- quired to correct major deficiencies.

- FDA shall notify the certification agency of FDA’s action and the grounds on which the approval was withdrawn.

- A certification agency that has lost its  approval shall  notify facilities certified or seeking certification by it, as well as the appropriate accredita- tion bodies with jurisdiction in the State, that its approval has been with- drawn. Such notification shall be made within a timeframe and in a manner approved by FDA.

- Minor deficiencies. If  FDA  deter- mines that a certification agency has demonstrated deficiencies in per- forming certification functions and re- sponsibilities that are less serious or more limited than the deficiencies in paragraph (a) of this section, including failure to follow the certification agen- cy’s own procedures and policies as ap- proved by FDA, FDA shall notify the certification agency that it has a speci- fied period of time to take particular corrective measures as directed by FDA or to submit to FDA for approval the certification agency’s own plan of corrective action addressing the minor deficiencies. If the approved corrective actions are not being implemented sat- isfactorily or within the established schedule, FDA may place the  agency on probationary status for a period of time determined by FDA, or may with- draw approval of the certification agency.

- If FDA places a certification agen- cy on probationary status, the certifi- cation agency shall notify all facilities certified or seeking certification by it of its probationary status within a time period and in a manner approved by FDA.

- Probationary status shall remain in effect until such time as the certifi- cation agency can demonstrate to the satisfaction of FDA that it has success- fully implemented or is implementing the  corrective  action  plan  within  the

#### § 900.25

established schedule, and that the cor- rective actions have substantially eliminated all identified problems, or

- If FDA determines that a certifi- cation agency that has been placed on probationary status is not imple- menting corrective actions satisfac- torily or within the established sched- ule, FDA may withdraw approval of the certification agency. The certification agency shall notify all facilities cer- tified or seeking certification by it, as well as the appropriate accreditation bodies with jurisdiction in the State, of its loss of FDA approval, within a timeframe and in a manner approved by FDA.

- Transfer of records. A certification agency that has its approval with- drawn shall transfer facility records and  other  related  information  as  re-

#### 21 CFR Ch. I (4–1–17 Edition)

quired by FDA to a location and ac- cording to a schedule approved by FDA.

##### § 900.25   Hearings and appeals.

- Opportunities to challenge final adverse actions taken by FDA regard- ing approval of certification  agencies or withdrawal of approval of certifi- cation agencies shall be communicated through notices of opportunity for in- formal hearings in accordance  with part 16 of this chapter.

- A facility that has been denied certification is entitled to an appeals process from the  certification  agency. The appeals process shall  be  specified in writing  by  the  certification  agency and shall have been  approved  by  FDA in accordance with §§ 900.21 and 900.22.

# SUBCHAPTER J—RADIOLOGICAL HEALTH

Sec.

## PART 1000—GENERAL

##### Subpart A—General Provisions

- Chassis family means a group of one or more models with all of the fol- lowing common characteristics:

  - The same circuitry in the high voltage,    horizontal    oscillator,    and

1000.1   General.

1000.3   Definitions.

##### Subpart B—Statements of Policy and Interpretation

1000.15 Examples of electronic products sub- ject to the Radiation Control for Health and Safety Act of 1968.

##### Subpart C—Radiation Protection Recommendations

1000.50 Recommendation for the use of spe- cific area gonad shielding on patients during medical diagnostic x-ray proce- dures.

1000.55 Recommendation for quality assur- ance programs in diagnostic radiology fa- cilities.

- Recommendation on administra- tively required dental x-ray examina- tions.

AUTHORITY: 21 U.S.C. 360hh–360ss.

SOURCE:  38  FR  28624,  Oct.  15,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1000.1   General.

References in this subchapter J to regulatory sections of the Code of Fed- eral Regulations are to chapter I of title 21 unless otherwise noted.

[50 FR 33688, Aug. 20, 1985]

##### § 1000.3   Definitions.

As used in this subchapter J:

  - Accidental radiation occurrence means a single event or series of events that has/have resulted in injurious or potentially injurious exposure of any person to electronic product radiation as a result of the manufacturing, test- ing, or use of an electronic product.

  - Act  means  the   Federal   Food, Drug, and Cosmetic Act (21 U.S.C. 360hh–360ss).

power supply sections;

  - The same worst component fail- ures;

  - The same type of high voltage hold-down or safety circuits; and

  - The same design and installation.

- Commerce means:

  - Commerce between any place in any State and any place outside there- of, and

  - Commerce wholly within the Dis- trict of Columbia.

- Component, for the purposes of this part, means an essential func- tional part of a subassembly or of an assembled electronic product,  and which may affect the quantity, quality, direction, or radiation emission of the finished product.

- Dealer means a person engaged in the business of offering electronic products for sale to purchasers, with- out regard to whether such person is or has been primarily engaged in such business, and includes persons who offer such products for lease or  as prizes or awards.

- Director means the Director of the Center for Devices and Radiological Health.

- Distributor means a person en- gaged in the business of offering elec- tronic products for sale to dealers, without regard to whether such person is or has been primarily or customarily engaged in such business.

- Electromagnetic radiation includes the entire electromagnetic spectrum of radiation of any wavelength. The elec- tromagnetic spectrum illustrated in figure 1 includes, but is not limited to, gamma rays, x-rays, ultra-violet, visi- ble, infrared, microwave, radiowave, and low frequency radiation.

§ 1000.3

- Electronic product means:

- Any manufactured or assembled product which, when in operation:

- Contains or acts as part of an elec- tronic circuit and

- Emits (or in the absence of effec- tive shielding or other controls would emit) electronic product radiation, or

- Any manufactured or assembled article that is intended for use as a component, part, or accessory of a product described in paragraph (j)(1) of this section and which, when in oper- ation, emits (or in the absence of effec- tive shielding or other controls would emit) such radiation.

- Electronic product radiation means:

  - Any ionizing or nonionizing elec- tromagnetic  or  particulate  radiation, or

  - Any sonic, infrasonic, or ultra- sonic wave that is emitted from an electronic product as the result of the operation of an electronic circuit in such product.

- Federal standard means a perform- ance  standard  issued  pursuant  to  sec-

#### 21 CFR Ch. I (4–1–17 Edition)

tion 534 of the Federal Food, Drug, and Cosmetic Act.

- Infrasonic, sonic (or audible) and ultrasonic waves refer to energy trans- mitted as an alteration (pressure, par- ticle displacement or density) in a property of an elastic medium (gas, liq- uid or solid) that can be detected by an instrument  or  listener.

- Manufacturer means any person engaged in the business of manufac- turing, assembling, or importing elec- tronic products.

- Model means any identifiable, unique electronic product design, and refers to products having the same structural and electrical design charac- teristics and to which the manufac- turer has assigned a specific designa- tion to differentiate between it and other products produced by that manu- facturer.

- Model family means products hav- ing similar design and radiation char- acteristics but different manufacturer model numbers.

#### Food and Drug Administration, HHS	§ 1000.15

- Modified model means a product that is redesigned so that actual or po- tential radiation emission, the  manner of compliance with a standard, or the manner of radiation safety testing is affected.

- Particulate radiation is  defined  as:

  - Charged particles, such as pro- tons, electrons, alpha particles, or heavy particles, which have sufficient kinetic energy to produce ionization or atomic or electron excitation by colli- sion, electrical attractions or elec- trical repulsion; or

  - Uncharged particles, such as neu- trons, which can initiate a nuclear transformation or liberate charged par- ticles having sufficient kinetic  energy to produce ionization or atomic or electron excitation.

- Phototherapy product means any ultraviolet lamp, or product containing such lamp, that is intended for irradia- tion of any part of the living human body by light in  the  wavelength  range of 200 to 400 nanometers, in order to perform a therapeutic function.

- Purchaser means the first person who, for value, or as an award or prize, acquires an electronic product for pur- poses other than resale, and includes a person who leases an electronic prod- uct for purposes other than subleasing.

- State means a State,  the  District of Columbia, the Commonwealth of Puerto Rico, the Virgin Islands, Guam, and American Samoa.

[60 FR 48380, Sept. 19, 1995; 61 FR 13422, Mar.

27, 1996]

## Subpart B—Statements of Policy and Interpretation

##### § 1000.15 Examples of electronic prod- ucts subject to the Radiation Con- trol for Health and Safety Act of 1968.

The following listed electronic prod- ucts are intended to serve as illus- trative examples of sources of elec- tronic product radiation to which the regulations of this part apply.

- Examples of electronic products which may emit x-rays and other ion- izing electromagnetic radiation, elec- trons, neutrons, and other particulate radiation include:

Ionizing electromagnetic radiation:

Television receivers. Accelerators.

X-ray  machines  (industrial,  medical,  re- search, educational).

Particulate  radiation  and  ionizing  electro- magnetic radiation:

Electron microscopes. Neutron generators.

- Examples of electronic products which may emit ultraviolet, visible, in- frared, microwaves, radio and low fre- quency electromagnetic radiation in- clude:

Ultraviolet:

Biochemical and medical analyzers. Tanning and therapeutic lamps.

Sanitizing and sterilizing devices. Black light sources.

Welding equipment. Visible:

White light devices. Infrared:

Alarm systems. Diathermy units.

Dryers, ovens, and heaters. Microwave:

Alarm systems. Diathermy units.

Dryers, ovens, and heaters. Medico-biological  heaters.

Microwave power generating devices. Radar devices.

Remote control devices. Signal generators.

Radio and low frequency: Cauterizers.

Diathermy units.

Power generation  and  transmission  equip- ment.

Signal generators. Electromedical equipment.

- Examples of  electronic  products which may emit coherent electro- magnetic radiation produced by stimu- lated emission include:

Laser:

Art-form,   experimental   and   educational devices.

Biomedical analyzers.

Cauterizing, burning and welding devices. Cutting and drilling devices.

Communications transmitters. Rangefinding devices.

Maser:

Communications transmitters.

- Examples of electronic products which may emit infrasonic, sonic, and ultrasonic vibrations resulting from operation of an electronic circuit in- clude:

Infrasonic: Vibrators.

§ 1000.50

Sonic:

Electronic  oscillators.

Sound amplification equipment. Ultrasonic:

Cauterizers.

Cell and tissue disintegrators. Cleaners.

Diagnostic    and    nondestructive    testing equipment.

Ranging and detection equipment.

## Subpart C—Radiation Protection Recommendations

##### § 1000.50 Recommendation for the use of specific area gonad shielding on patients during medical diagnostic x-ray procedures.

Specific area gonad shielding covers an area slightly larger than the region of the gonads. It may therefore be used without interfering with the objectives of the examination to protect the ger- minal tissue of patients from radiation exposure that may cause genetic mutations during many medical x-ray procedures in which the gonads lie within or are in close proximity to the x-ray field. Such shielding should be provided when the following conditions exist:

  - The gonads will lie within the pri- mary x-ray field, or within close prox- imity (about 5 centimeters), despite proper beam limitation. Except as pro- vided in paragraph (b) or (c) of this sec- tion:

    - Specific area testicular shielding should always be used during those ex- aminations in which the testes usually are in the primary x-ray field, such as examinations of the pelvis, hip, and upper femur;

    - Specific area testicular shielding may also be warranted during other ex- aminations of the abdominal region in which the testes may lie within or in close proximity to the primary x-ray field, depending upon the size of the pa- tient and the examination techniques and equipment employed. Some exam- ples of these are: Abdominal, lumbar spine and lumbosacral spine examina- tions, intravenous pyelograms, and ab- dominal scout film for barium enemas and upper GI series. Each x-ray facility should evaluate its procedures, tech- niques, and equipment and compile a list of such examinations for which specific     area     testicular     shielding

21 CFR Ch. I (4–1–17 Edition)

should be routinely considered for use. As a basis for judgment, specific area testicular shielding  should  be consid- ered for all examinations of male pa- tients in which the pubic symphysis will be visualized on the film;

    - Specific area gonad shielding should never be used as a substitute for careful patient positioning, the use of correct technique factors and film processing, or proper beam limitation (confinement of the x-ray field to the area of diagnostic  interest),  because this could result in  unnecessary  doses to other sensitive tissues and could ad- versely affect the quality of the radiograph; and

    - Specific area gonad shielding should provide attenuation of x-rays at least equivalent to that afforded by 0.25 millimeter of lead.

  - The clinical objectives of the ex- amination will not be compromised.

    - Specific area testicular shielding usually does not obscure needed infor- mation except in a few cases such as oblique views of the hip, retrograde urethrograms	and	voiding cystourethrograms,   visualization   of the rectum and, occasionally, the pubic symphysis. Consequently, specific area testicular shielding should be consid- ered for use in the majority of x-ray ex- aminations of male patients in which the testes will lie within the primary beam or within 5 centimeters of its edge. It is not always possible to posi- tion shields on male patients  so  that no bone is obscured. Therefore, if all bone structure of the pelvic area must be visualized for a particular patient, the use of shielding should be carefully evaluated. The decision concerning the applicability of shielding for an indi- vidual patient is dependent upon con- sideration of the patient’s unique an- thropometric characteristics and the diagnostic information needs of the ex- amination.

    - The use of specific area ovarian shielding is frequently impractical at present because the exact location of the ovaries is difficult to estimate, and the shield may obscure visualization of portions of adjacent structures such as the spine, ureters, and small and large bowels. However, it may be possible for

#### Food and Drug Administration, HHS	§ 1000.55

practitioners to use specific area ovar- ian shielding during selected views in some examinations.

  - The patient has a reasonable re- productive  potential.

    - Specific area shielding need not be used on patients who cannot or are not likely to have children in the future.

    - The following table of statistical data regarding the average number of children expected by potential parents in various age categories during their remaining lifetimes is provided for x- ray facilities that wish to use it as a basis for judging reproductive poten- tial:

EXPECTED NUMBER OF FUTURE CHILDREN VERSUS AGE OF POTENTIAL PARENT 1

| Age | Male parent | Female parent |
| --- | --- | --- |
| Fetus ...................................................... | 2.6 | 2.6 |
| 0 to 4 ..................................................... | 2.6 | 2.5 |
| 5 to 9 ..................................................... | 2.7 | 2.5 |
| 10 to 14 ................................................. | 2.7 | 2.6 |
| 15 to 19 ................................................. | 2.7 | 2.6 |
| 20 to 24 ................................................. | 2.6 | 2.2 |
| 25 to 29 ................................................. | 2.0 | 1.4 |
| 30 to 34 ................................................. | 1.1 | .6 |
| 35 to 39 ................................................. | .5 | .2 |
| 40 to 44 ................................................. | .2 | .04 |
| 45 to 49 ................................................. | .07 | 0 |
| 50 to 54 ................................................. | .03 | 0 |
| 55 to 64 ................................................. | .01 | 0 |
| Over 65 .................................................. | 0 | 0 |

1 Derived  from data  published by  the  National Center  for Health Statistics, ‘‘Final Natality Statistics 1970,’’ HRA 74– 1120, vol. 22, No. 12, Mar. 20, 1974.

[41 FR 30328, July 23, 1976; 41 FR 31812, July

30, 1976]

##### § 1000.55 Recommendation for quality assurance programs in diagnostic radiology facilities.

- Applicability. Quality assurance programs as described in paragraph (c) of this section are recommended for all diagnostic radiology facilities.

- Definitions. As used in this sec- tion, the following definitions apply:

  - Diagnostic radiology facility means any facility in which an x-ray sys- tem(s) is used in any procedure that in- volves irradiation of any part of the human body for the purpose of diag- nosis or visualization. Offices of indi- vidual physicians, dentists, podiatrists, and chiropractors, as well as mobile laboratories, clinics, and hospitals are all examples of diagnostic radiology fa- cilities.

  - Quality assurance means the planned and systematic actions that provide adequate confidence that a di- agnostic x-ray facility will produce consistently high quality images with minimum exposure of the patients and healing arts personnel. The determina- tion of what constitutes high quality will be made by the facility producing the images. Quality assurance actions include both ‘‘quality control’’ tech- niques and ‘‘quality administration’’ procedures.

  - Quality assurance program means an organized entity designed to provide ‘‘quality assurance’’ for a diagnostic radiology facility. The nature and ex- tent of this program will vary with the size and type of the facility, the type of examinations conducted, and other fac- tors.

  - Quality control techniques are those techniques used in the monitoring (or testing) and maintenance of the com- ponents of an x-ray system. The qual- ity control techniques thus are con- cerned directly with the equipment.

  - Quality administration procedures are those management actions in- tended to guarantee that monitoring techniques are properly performed and evaluated and that necessary correc- tive measures are taken in response to monitoring results. These procedures provide the organizational framework for the quality assurance program.

  - X-ray system means an assemblage of components for the controlled pro- duction of diagnostic images with x- rays. It includes minimally  an  x-ray high voltage generator, an x-ray con- trol, a tube-housing assembly, a beam- limiting device, and the necessary sup- porting structures. Other components that function with the system, such as image receptors, image  processors, view boxes, and darkrooms, are also parts of the system.

- Elements. A quality assurance pro- gram should contain the elements list- ed in paragraphs (c)(1) through (10) of this section. The extent to which each element of the quality assurance pro- gram is implemented should be deter- mined by an analysis of the facility’s objectives and resources conducted by its qualified staff or by qualified out- side consultants. The extent of imple- mentation should be determined on the

#### § 1000.55

basis of whether the expected benefits in radiation exposure reduction, im- proved image quality, and/or financial savings will compensate for the re- sources required for the program.

  - Responsibility. (i) Responsibility and authority for the overall quality assurance program as well as for moni- toring, evaluation, and corrective measures should be specified and re- corded in a quality assurance manual.

- The owner or  practitioner in charge of the facility has primary re- sponsibility for implementing and maintaining the quality assurance pro- gram.

- Staff technologists will gen- erally be delegated a basic quality as- surance role by the practitioner in charge. Responsibility for specific quality control monitoring and main- tenance techniques or quality adminis- tration procedures may be assigned, provided that the staff technologists are qualified by training or experience for these duties. The staff tech- nologists should also be responsible for identifying problems or potential prob- lems requiring actions beyond the level of their training. They should bring these problems to the attention of the practitioner in charge, or his or her representative, so that assistance in solving the problems may be obtained from inside or outside the facility.

- In facilities where they are avail- able, physicists, supervisory tech- nologists, or quality control tech- nologists should have a major role in the quality assurance program. Such specialized personnel may be assigned responsibility for day-to-day adminis- tration of the program, may carry out monitoring duties beyond the level of training of the staff technologist or, if desired by the facility, may relieve the staff technologists of some or all of their basic monitoring duties. Staff service engineers may also be assigned responsibility for certain preventive or corrective maintenance actions.

- Responsibility for certain quality control techniques and corrective measures may be assigned to personnel qualified by training or experience, such as consultants or industrial rep- resentatives, from outside of the facil- ity, provided there is a written agree- ment clearly specifying these services.

#### 21 CFR Ch. I (4–1–17 Edition)

- In large facilities, responsibility for long-range planning of quality as- surance goals and activities should be assigned to a quality assurance com- mittee as described in paragraph (c)(9) of this section.

  - Purchase specifications. Before pur- chasing new equipment, the staff of the diagnostic radiology facility should de- termine the desired performance speci- fications for the equipment. Initially, these specifications may be stated in terms of the desired performance of the equipment, or prospective vendors may be informed solely of the functions the equipment should be able to perform and asked to provide the performance specifications of items from their equipment line that can perform these functions. In either case, the responses of the prospective vendors should serve as the basis for negotiations to estab- lish the final purchase specifications, taking into account the state of the art and balancing the need for the speci- fied performance levels with the cost of the equipment to meet them. The final purchase specifications should be in writing and should include perform- ance specifications. The availability of experienced service personnel should also be taken into consideration in making the final purchase  decisions. Any understandings with respect to service personnel should be incor- porated into the purchase specifica- tions. After the equipment is installed, the facility should conduct a testing program, as defined in its purchase specifications, to ensure that the equipment meets the agreed upon spec- ifications, including applicable Federal and State  regulatory  requirements. The equipment should not be formally accepted until any necessary correc- tions have been made by the vendor. The purchase specifications and the records of the acceptance testing should be retained throughout  the  life of the equipment for comparison with monitoring results in order to assess continued acceptability of perform- ance.

  - Monitoring and maintenance. A routine quality control monitoring and maintenance system incorporating state-of-the-art procedures should be established and conducted on a regular schedule. The purpose of monitoring is

#### Food and Drug Administration, HHS	§ 1000.55

to permit evaluation of the perform- ance of the facility’s x-ray system(s) in terms of the standards for image qual- ity established by the facility (as de- scribed in paragraph (c)(4) of this sec- tion) and compliance with applicable Federal and State regulatory require- ments. The maintenance program should include corrective maintenance to eliminate problems revealed by monitoring or other means before they have a serious deleterious impact on patient care. To the extent permitted by the training of the facility staff, the maintenance program should also in- clude preventive maintenance, which could prevent unexpected breakdowns of equipment and disruption of depart- mental routine.

    - The parameters to be monitored in a facility should be determined by that facility on the basis of an analysis of expected benefits and cost. Such fac- tors as the size and resources of the fa- cility, the type of examinations con- ducted, and the quality assurance prob- lems that have occurred in that or similar facilities should be taken into account in establishing the monitoring system. The monitoring frequency should also be based upon need and can be different for different parameters.

    - Although the parameters to be monitored will vary somewhat from fa- cility to facility, every diagnostic radi- ology facility should consider moni- toring the following five key compo- nents of the x-ray system:

- Film processing.

- Basic performance characteristics of the x-ray unit.

- Cassettes and grids.

- View boxes.

- Darkroom.

    - Examples of parameters of the above-named components and of more specialized equipment that may be monitored are as follows:

- For film processing:

An index of speed. An index of contrast. Base plus fog.

Solution temperatures. Film artifact identification.

- For basic performance character- istics of the x-ray unit:

  - For fluoroscopic x-ray  units:

Table-top exposure rates. Centering alignment.

Collimation.

kVp accuracy and reproducibility. mA accuracy and reproducibility.

Exposure time accuracy and reproducibility. Reproducibility of x-ray output.

Focal spot size consistency. Half-value layer.

Representative entrance skin exposures.

  - For image-intensified systems:

Resolution. Focusing. Distortion. Glare.

Low contrast performance.

Physical   alignment   of   camera   and   colli- mating lens.

  - For radiographic x-ray units:

Reproducibility of x-ray output.

Linearity   and   reproducibility   of   mA   sta- tions.

Reproducibility  and  accuracy  of  timer  sta- tions.

Reproducibility  and  accuracy  of  kVp  sta- tions.

Accuracy  of  source-to-film  distance  indica- tors.

Light/x-ray field congruence. Half-value layer.

Focal spot size consistency. Representative entrance skin exposures.

  - For  automatic  exposure  control devices:

Reproducibility. kVp compensation.

Field sensitivity matching. Minimum response time.

Backup timer verification.

- For cassettes and grids:

  - For cassettes:

Film/screen contact. Screen condition.

Light leaks.

Artifact identification.

  - For grids:

Alignment and focal distance. Artifact  identification.

- For view boxes:

Consistency of light output with time.

Consistency of light output from one box to another.

View box surface conditions.

- For darkrooms:

Darkroom integrity. Safe light conditions.

- For specialized equipment:

  - For tomographic systems:

Accuracy of depth and cut indicator.

#### § 1000.55

Thickness of cut plane. Exposure angle.

Completeness of tomographic motion. Flatness of tomographic field.

Resolution.

Continuity of exposure. Flatness of cassette.

Representative entrance skin exposures.

  - For computerized tomography:

Precision (noise). Contrast scale.

High and low contrast resolution. Alignment.

Representative entrance skin exposures.

    - The maintenance program should include both preventive and corrective aspects.

- Preventive maintenance. Preventive maintenance should be performed on a regularly scheduled basis with the goal of preventing breakdowns due to equip- ment failing without warning signs de- tectable by monitoring. Such actions have been found cost effective if re- sponsibility is assigned to facility staff members. Possible preventive mainte- nance procedures are visual inspection of the mechanical and electrical char- acteristics of the x-ray system (cov- ering such things as checking condi- tions of cables, watching the tomo- graphic unit for smoothness of motion, assuring cleanliness with respect to spilling of contaminants in the exam- ination room or the darkroom, and lis- tening for unusual noises in the mov- ing parts of the system), following the manufacturer’s recommended proce- dures for cleaning and maintenance of the equipment, and regular inspection and replacement of switches and parts that routinely wear out or fail. The procedures included would depend upon the background of the staff members available. Obviously, a large facility with its own service engineers can do more than an individual practitioner’s office.

- Corrective maintenance. For max- imum effectiveness, the quality assur- ance program should make provision, as described in paragraph (c)(5) of this section, for ascertaining whether po- tential problems are developing. If po- tential or actual problems are de- tected, corrective maintenance should be carried out to eliminate them before they cause a major impact on patient care.

#### 21 CFR Ch. I (4–1–17 Edition)

- Standards for image quality. Stand- ards of acceptable image quality should be established. Ideally, these should be objective, e.g., acceptability limits for the variations of parameter values, but they may be subjective, e.g., the opinions of professional per- sonnel, in cases where adequate objec- tive standards cannot be defined. These standards should be routinely reviewed and redefined as needed, as described in paragraph (c)(10) of this section.

- Evaluation. The facility’s quality assurance program should include means for two levels of evaluation.

  - On the first level, the results of the monitoring procedures should be used to evaluate the performance of the x-ray system(s) to determine whether corrective actions are needed to adjust the equipment so that the image quality consistently meets the standards for image quality. This eval- uation should include analysis of trends in the monitoring data as well as the use of the data to determine the need for corrective actions on a day-by- day basis. Comparison of monitoring data with the purchase specifications and acceptance testing results for the equipment in question is also useful.

  - On the second level, the facility quality assurance program should also include means for evaluating the effec- tiveness of the program itself. Possible means include ongoing studies of the retake rate and the causes of the re- peated radiographs, examination of equipment repair and replacement costs, subjective evaluation of the radiographs being produced, occurrence and reasons for complaints by radiolo- gists, and analysis of trends in the re- sults of monitoring procedures such as sensitometric studies. Of these, ongo- ing studies of the retake rate (reject rate) and its causes are often the most useful and may also provide informa- tion of value in the first level of eval- uation. Such studies can be used to evaluate potential for improvement, to make corrections, and to determine whether the corrective actions were ef- fective. The number of rejects should be recorded daily or weekly, depending on the facility’s analysis of its needs. Ideally,  the  reasons  for  the  rejection

#### Food and Drug Administration, HHS	§ 1000.55

should also be determined and re- corded. Should determining these rea- sons be impossible on a regular basis with the available staff, the analysis should be done for a 2-week period after major changes have occurred in diag- nostic procedures or the x-ray system and at least semi-annually.

- Records. The program should in- clude provisions for the keeping of records on the results of the  moni- toring techniques, any difficulties de- tected, the corrective measures applied to these difficulties, and the effective- ness of these measures. The extent and form of these records should be deter- mined by the facility on the basis of its needs. The facility should view these records as a tool for maintaining an ef- fective quality assurance program and not view the data in them as an end in itself but rather as a beginning. For ex- ample, the records should be made available to vendors to help them pro- vide better service. More importantly, the data should be the basis for the evaluation and the reviews  suggested in paragraphs (c)(5) and (10) of this sec- tion.

- Manual. A quality assurance man- ual should be written in a format per- mitting convenient revision as needed and should be made  readily  available to all  personnel.  The content  of  the manual should be determined by the fa- cility staff, but the following items are suggested as providing essential infor- mation:

  - A list of the individuals respon- sible for monitoring and maintenance techniques.

  - A list of the parameters to be monitored and the frequency of moni- toring.

  - A description of the standards, criteria of quality, or limits of accept- ability that have been established for each of the parameters monitored.

  - A brief description of the proce- dures to be used for monitoring each parameter.

  - A description of procedures to be followed when difficulties are detected to call these difficulties to the atten- tion of those responsible for correcting them.

  - A list of the publications in which detailed instructions for moni- toring and maintenance procedures can

be found. Copies of these publications should also be readily available to the entire staff, but they should be sepa- rate from the manual. (Publications providing these instructions can usu- ally be obtained from FDA or private sources, although the facility may wish to make some modifications to  meet its needs more effectively.)

  - A list of the records, with  sam- ple forms, that the facility staff has de- cided should  be kept. The  facility  staff should also determine and note in the manual the length of time each type of record should be kept before dis- carding.

  - A copy of each set of purchase specifications developed for new equip- ment and the results of the acceptance testing for that equipment.

- Training. The program should in- clude provisions for appropriate train- ing for all personnel with quality as- surance responsibilities. This should include both training provided before the quality assurance responsibilities are assumed and continuing education to keep the personnel up-to-date. Prac- tical experience with the techniques conducted under the supervision of ex- perienced instructors, either in the fa- cility or in a special program, is the most desirable type of training. The use of self-teaching materials can be an adequate substitute for supervised in- struction, especially in continuing edu- cation programs, if supervised instruc- tion is not available.

- Committee. A facility whose size would make it impractical for all staff members to meet for planning purposes should consider the establishment of a quality assurance committee whose primary function would be to maintain lines of communication among all groups with quality assurance and/or image production or interpretation re- sponsibilities. For maximum commu- nication, all departments of the  facil- ity with x-ray equipment should be represented. The committee may also be assigned policy-making duties such as some or all of the following:

Assign quality assurance responsibil- ities; maintain acceptable standards of quality; periodically review program effectiveness,   etc.   Alternatively,   the

#### § 1000.60

duties of this committee could be as- signed to an already-existing com- mittee such as the Radiation Safety Committee. In smaller facilities, all staff members should participate in the committee’s tasks. The Quality Assur- ance Committee should report directly to the head of the radiology depart- ment, or, in facilities where more than one department operates x-ray equip- ment, to the chief medical officer  of the facility. The committee should meet on a regular basis.

- Review. The facility’s quality as- surance program should be reviewed by the Quality Assurance  Committee and/ or the practitioner in charge to deter- mine whether its effectiveness could be improved. Items suggested for  inclu- sion in the review include:

- The reports of the monitoring and maintenance techniques to ensure that they are being performed on schedule and effectively. These reports should be reviewed at least quarterly.

- The monitoring and maintenance techniques and their schedules to en- sure that they continue to be appro- priate and in step with the latest devel- opments in quality assurance. They should be made current at least annu- ally.

- The standards for image quality to ensure that they are consistent with the state-of-the-art and the needs and resources of the facility. These stand- ards should be evaluated at least annu- ally.

- The results  of  the evaluations of the effectiveness of the quality assur- ance actions to determine whether changes need to be made. This deter- mination should be made at least annu- ally.

- The quality  assurance manual should also be reviewed at least annu- ally to determine whether revision is needed.

[44 FR 71737, Dec. 11, 1979]

##### § 1000.60 Recommendation on adminis- tratively required dental x-ray ex- aminations.

- The Food and Drug Administra- tion recommends that dental x-ray ex- aminations be performed only after careful consideration of the dental or other health needs of  the  patient,  that is, when the patient’s dentist or physi-

#### 21 CFR Ch. I (4–1–17 Edition)

cian judges them to be necessary for di- agnosis, treatment, or prevention of disease. Administratively required den- tal x-ray examinations are those re- quired by a remote third party for rea- sons not related to the patient’s imme- diate dental needs. These x-ray exami- nations are usually a source of unnec- essary radiation exposure to the pa- tient. Because any unnecessary radi- ation exposure should be avoided, third parties should not require dental x-ray examinations unless they can dem- onstrate that such examinations pro- vide a direct clinical benefit to the pa- tient, and the patient’s dentist or phy- sician agrees with that assessment.

- Some examples of administrative x-ray examinations that should not be required by third parties are those in- tended solely:

  - To monitor insurance claims or detect fraud;

  - To satisfy a prerequisite for reim- bursement;

  - To provide training or experience;

  - To certify qualifications or com- petence.

- This recommendation is not in- tended to preclude dental x-ray exami- nations ordered by the attending prac- titioner, based on the  patient’s  history or physical examination, or those per- formed on selected populations  shown to have significant yields  of previously undiagnosed disease. This rec- ommendation is also not intended to preclude the administrative use by third parties of dental radiographs that are taken on the order of the patient’s dentist or physician  as  a  necessary part of the patient’s clinical care.

[45 FR 40978, June 17, 1980]

## PART 1002—RECORDS AND REPORTS

##### Subpart A—General Provisions

Sec.

- Applicability.

- [Reserved]

- Notification  to  user  of  performance and technical data.

- Confidentiality of information.

1002.7    Submission of data and reports.

Food and Drug Administration, HHS	§ 1002.1

##### Subpart B—Required Manufacturers’ Reports for Listed Electronic Products

- Product reports.

- Supplemental reports.

- Abbreviated reports.

- Annual reports.

##### Subpart C—Manufacturers’ Reports on Accidental Radiation Occurrences

1002.20    Reporting of accidental radiation oc- currences.

##### Subpart D—Manufacturers’ Records

- Records to be maintained by manu- facturers.

- Preservation	and	inspection	of records.

##### Subpart E—Dealer and Distributor Records

- Records  to  be  obtained  by  dealers and distributors.

- Disposition  of  records  obtained  by dealers and distributors.

- Confidentiality  of  records  furnished by dealers and distributors.

##### Subpart F—Exemptions From Records and Reports Requirements

- Special exemptions.

- Exemptions for manufacturers of products intended for the U.S. Govern- ment.

AUTHORITY:  21  U.S.C.  352,  360,  360i,  360j,

360hh–360ss, 371, 374.

SOURCE:  38  FR  28625,  Oct.  15,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1002.1   Applicability.

The provisions of this part are appli- cable as follows:

  - All    manufacturers    of    electronic products are subject to § 1002.20.

  - Manufacturers, dealers, and dis- tributors of electronic products are subject to the provisions of part 1002 as set forth in table 1 of this  section,  un- less excluded by paragraph (c) of this section, or unless an  exemption  has been granted under § 1002.50 or § 1002.51.

  - The requirements of part 1002 as specified in table 1 of this section  are not applicable to:

    - Manufacturers of electronic prod- ucts intended solely for export if such product is labeled or tagged to show that the product meets all the applica- ble requirements of the country to which such product is intended for ex- port.

    - Manufacturers of electronic prod- ucts listed in table 1 of this section if such product is sold exclusively to other manufacturers for use as compo- nents of electronic products to be sold to purchasers, with the exception that the provisions are applicable to those manufacturers certifying  components of diagnostic x-ray systems  pursuant to provisions of § 1020.30(c) of this chap- ter.

    - Manufacturers of electronic prod- ucts that are intended for use by the

U.S. Government and whose  function or design cannot be divulged by the manufacturer for reasons of national security, as evidenced by government security   classification.

    - Assemblers of diagnostic x-ray equipment subject to the provisions of

§ 1020.30(d) of this chapter, provided the assembler has submitted the report re- quired by § 1020.30(d)(1) or (d)(2) of this chapter and retains a copy of such re- port for a period of 5 years from its date.

TABLE 1—RECORD AND REPORTING REQUIREMENTS BY PRODUCT

| Manufacturer | Dealer & Distributor |
| --- | --- |
| Products | Product reports § 1002.10 | Supple- mental reports § 1002.11 | Abbre- viated re- ports § 1002.12 | Annual reports § 1002.13 | Test records § 1002.30(a) 1 | Distribution records § 1002.30(b) 2 |  |
|  |  |  |  |  |  |  | Distribution records §§ 1002.40 and 1002.41 |
| DIAGNOSTIC X-RAY 3 (1020.30, | X | X |  | X | X | X | X |
| 1020.31, 1020.32, 1020.33) |  |  |  |  |  |  |  |
| Computed tomography |  |  |  |  |  |  |  |
| X-ray system 4 | X | X |  | X | X | X | X |
| Tube housing assembly | X | X |  | X | X | X |  |
| X-ray control | X | X |  | X | X | X | X |
| X-ray high voltage generator | X | X |  | X | X | X | X |

§ 1002.1

21 CFR Ch. I (4–1–17 Edition)

TABLE 1—RECORD AND REPORTING REQUIREMENTS BY PRODUCT—Continued

| Manufacturer | Dealer & Distributor |
| --- | --- |
| Products | Product reports § 1002.10 | Supple- mental reports § 1002.11 | Abbre- viated re- ports § 1002.12 | Annual reports § 1002.13 | Test records § 1002.30(a) 1 | Distribution records § 1002.30(b) 2 |  |
|  |  |  |  |  |  |  | Distribution records §§ 1002.40 and 1002.41 |
| X-ray table or cradle | X | X | X | X | X | X | X |
| X-ray film changer |  |  | X |  | X | X |  |
| Vertical cassette holders mount- |  |  | X |  | X | X | X |
| ed in a fixed location and cas- |  |  |  |  |  |  |  |
| sette holders with front panels |  |  |  |  |  |  |  |
| Beam-limiting devices |  |  |  |  | X | X | X |
| Spot-film devices and image in- | X | X |  | X | X | X | X |
| tensifiers manufactured after |  |  |  |  |  |  |  |
| April 26, 1977 |  |  |  |  |  |  |  |
| Cephalometric devices manufac- |  |  | X |  | X | X |  |
| tured after February 25, 1978 |  |  |  |  |  |  |  |
| Image receptor support devices |  |  | X |  | X | X | X |
| for mammographic X-ray sys- |  |  |  |  |  |  |  |
| tems manufactured after Sep- |  |  |  |  |  |  |  |
| tember 5, 1978 |  |  |  |  |  |  |  |
| CABINET X RAY (§ 1020.40) |  |  |  |  |  |  |  |
| Baggage inspection | X | X |  | X | X | X | X |
| Other | X | X |  | X | X | X |  |
| PRODUCTS INTENDED TO |  |  |  |  |  |  |  |
| PRODUCE PARTICULATE |  |  |  |  |  |  |  |
| RADIATION OR X-RAYS |  |  |  |  |  |  |  |
| OTHER THAN DIAGNOSTIC |  |  |  |  |  |  |  |
| OR CABINET DIAGNOSTIC |  |  |  |  |  |  |  |
| X-RAY |  |  |  |  |  |  |  |
| Medical |  |  | X | X | X | X |  |
| Analytical |  |  | X | X | X | X |  |
| Industrial |  |  | X | X | X | X |  |
| TELEVISION PRODUCTS |  |  |  |  |  |  |  |
| (§ 1020.10) |  |  |  |  |  |  |  |
| <25 kilovolt (kV) and <0.1 |  |  | X | X 6 |  |  |  |
| milliroentgen per hour (mR/hr |  |  |  |  |  |  |  |
| IRLC 5 6 |  |  |  |  |  |  |  |
| 25kV and <0.1mR/hr IRLC 5 | X | X |  | X |  |  |  |
| 0.1mR/hr IRLC 5 | X | X |  | X | X | X |  |
| MICROWAVE/RF |  |  |  |  |  |  |  |
| MW ovens (§ 1030.10) | X | X |  | X | X | X |  |
| MW diathermy |  |  | X |  |  |  |  |
| MW heating, drying, security sys- |  |  | X |  |  |  |  |
| tems |  |  |  |  |  |  |  |
| RF sealers, electromagnetic in- |  |  | X |  |  |  |  |
| duction and heating equip- |  |  |  |  |  |  |  |
| ment, dielectric heaters (2–500 |  |  |  |  |  |  |  |
| megahertz) |  |  |  |  |  |  |  |
| OPTICAL |  |  |  |  |  |  |  |
| Phototherapy products | X | X |  |  |  |  |  |
| Laser products (§§ 1040.10, |  |  |  |  |  |  |  |
| 1040.11) |  |  |  |  |  |  |  |
| Class I lasers and products con- | X |  |  | X | X |  |  |
| taining such lasers 7 |  |  |  |  |  |  |  |
| Class I laser products containing | X |  |  | X | X | X |  |
| class IIa, II, IIIa, lasers 7 |  |  |  |  |  |  |  |
| Class IIa, II, IIIa lasers and prod- | X | X |  | X | X | X | X |
| ucts other than class I prod- |  |  |  |  |  |  |  |
| ucts containing such lasers 7 |  |  |  |  |  |  |  |
| Class IIIb and IV lasers and | X | X |  | X | X | X | X |
| products containing such la- |  |  |  |  |  |  |  |
| sers 7 |  |  |  |  |  |  |  |
| Sunlamp products (§ 1040.20) |  |  |  |  |  |  |  |
| Lamps only | X |  |  |  |  |  |  |
| Sunlamp products | X | X |  | X | X | X | X |
| Mercury vapor lamps (§ 1040.30) |  |  |  |  |  |  |  |
| T lamps | X | X |  | X |  |  |  |
| R lamps |  |  | X |  |  |  |  |
| ACOUSTIC |  |  |  |  |  |  |  |
| Ultrasonic therapy (1050.10) | X | X |  | X | X | X | X |
| Diagnostic ultrasound |  |  | X |  |  |  |  |

#### Food and Drug Administration, HHS	§ 1002.7

TABLE 1—RECORD AND REPORTING REQUIREMENTS BY PRODUCT—Continued

| Manufacturer | Dealer & Distributor |
| --- | --- |
| Products | Product reports § 1002.10 | Supple- mental reports § 1002.11 | Abbre- viated re- ports § 1002.12 | Annual reports § 1002.13 | Test records § 1002.30(a) 1 | Distribution records § 1002.30(b) 2 |  |
|  |  |  |  |  |  |  | Distribution records §§ 1002.40 and 1002.41 |
| Medical ultrasound other than therapy or diagnostic Nonmedical ultrasound | X | X | X |  |  |  |  |

1However, authority to inspect all appropriate documents supporting the adequacy of a manufacturer’s compliance testing pro- gram is retained.

2The requirement includes §§ 1002.31 and 1002.42, if applicable.

3Report of Assembly (Form FDA 2579) is required for diagnostic x-ray components; see 21 CFR 1020.30(d)(1) through (d)(3).

4Systems records and reports are required if a manufacturer exercises the option and certifies the system as permitted in 21 CFR 1020.30(c).

5Determined using the isoexposure rate limit curve (IRLC) under phase III test conditions (1020.10(c)(3)(iii)).

6Annual report is for production status information only.

7Determination of the applicable reporting category for a laser product shall be based on the worst-case hazard present within the laser product.

[60 FR 48382, Sept. 19, 1995; 61 FR 13423, Mar. 27, 1996]

##### § 1002.2   [Reserved]

§ 1002.3   Notification   to   user   of   per- formance and technical data.

The Director and Deputy Director of the Center for Devices and Radio- logical Health, as authorized under del- egated authority, may require a manu- facturer of a radiation emitting elec- tronic product to provide to the ulti- mate purchaser, at the time of original purchase, such performance data and other technical data related  to  safety of the product as the Director or Dep- uty Director finds necessary.

[69 FR 17292, Apr. 2, 2004]

##### § 1002.4   Confidentiality of information.

The Secretary or his representative shall not disclose any information re- ported to or otherwise  obtained  by him, pursuant to this part, which con- cerns or relates to a trade secret or other matter referred to in section 1905 of title 18 of the United States Code, except that such information may be disclosed to other officers or employees of the  Department and  of the  other agencies concerned with carrying out the requirements of the  Act.  Nothing in this section shall authorize the withholding of information by the Sec- retary, or by any officers or employees under his control, from the duly au- thorized committees of the Congress.

##### § 1002.7   Submission   of   data   and   re- ports.

All submissions such as reports, test data, product descriptions, and other information required by this part, or voluntarily submitted to the Director, Center for Devices and Radiological Health, shall be filed  with  the  number of copies as prescribed by the Director, Center for Devices and Radiological Health, and shall be signed by the per- son making the submission. The sub- missions required by this part shall be addressed to the Food and Drug Admin- istration, Center for Devices and Radi- ological Health, ATTN:  Electronic Product Reports, Document Mail Cen- ter, 10903 New Hampshire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993– 0002.

- In addition to the requirements of this part, all material submitted to the Director, Center for Devices and Radio- logical Health, shall be submitted pur- suant to the provisions of part 20—Pub- lic Information, of this chapter.

- Where guides or instructions have been issued by the Director for the sub- mission of  material  required by this part, such as test data, product reports, abbreviated reports, supplemental re- ports, and annual reports, the material submitted shall conform to the appli- cable reporting guides or instructions. Where it is not feasible or where it would not be appropriate to conform to any  portion  of  a  prescribed  reporting

#### § 1002.10

guide or instruction, an alternate for- mat for providing the information re- quested by that portion of the guide or instruction may be used provided the submitter of such information submits adequate explanation and justification for use of an alternate format. If the Director, Center for Devices and Radio- logical Health, determines that such justification is inadequate  and  that  it is feasible or appropriate to conform to the prescribed reporting guide or in- struction, he may require resubmission of the information in conformance with the reporting guide or instruction.

- Where the submission of quality control and testing information is com- mon to more than one model, or model family of the same product category, a ‘‘common aspects report’’ consoli- dating similar information may be pro- vided, if applicable.

[42  FR  18062,  Apr.  5,  1977,  as  amended  at  53

FR  11254,  Apr.  6,  1988;  60  FR  48385,  Sept.  19,

1995;  72  FR  17400,  Apr.  9,  2007;  75  FR  20916,

Apr. 22, 2010]

## Subpart B—Required Manufactur- ers’ Reports for Listed Elec- tronic Products

SOURCE: 60 FR 48386, Sept. 19, 1995, unless

otherwise noted.

##### § 1002.10   Product reports.

Every manufacturer of a product or component  requiring  a  product   report as set forth in table 1 of § 1002.1 shall submit  a  product  report  to  the  Food and Drug Administration, Center for Devices   and   Radiological   Health, ATTN: Electronic Product Reports, Document Mail Center, 10903 New Hampshire Ave., Bldg. 66, rm. G609,  Sil- ver Spring, MD 20993–0002, prior to the introduction of such product into com- merce. The report shall be distinctly marked ‘‘Radiation Safety Product  Re- port of (name of manufacturer)’’  and shall:

- Identify which listed product is being reported.

- Identify each model of the listed product together with sufficient infor- mation concerning the manufacturer’s code or other system of labeling to en- able the Director to determine  the place of manufacture.

21 CFR Ch. I (4–1–17 Edition)

- Include information on all compo- nents and accessories provided in, on, or with the listed product that may af- fect the quantity, quality, or direction of the radiation emissions.

- Describe the function, operational characteristics affecting radiation emissions, and intended and known uses of each model of the listed prod- uct.

- State the standard or design spec- ifications, if any, for each model with respect to electronic product radiation safety. Reference may be made to a Federal standard, if applicable.

- For each model, describe the phys- ical or electrical characteristics, such as shielding or electronic circuitry, in- corporated into the product in order to meet the standards or specifications re- ported pursuant to paragraph (e) of this section.

- Describe the methods and proce- dures employed, if any, in testing and measuring each model with respect to electronic product radiation safety, in- cluding the control of unnecessary, sec- ondary, or leakage electronic product radiation, the applicable  quality  con- trol procedures used for each model, and the basis for selecting such testing and quality control procedures.

- For those products which may produce increased radiation with aging, describe the methods and procedures used, and frequency of testing of each model for durability and stability with respect to electronic product radiation safety. Include the basis for selecting such methods and procedures, or for de- termining that such testing and  qual- ity control procedures are not nec- essary.

- Provide sufficient results of the testing, measuring, and quality control procedures described in  accordance with paragraphs (g) and (h) of this sec- tion to enable the Director to deter- mine the effectiveness of those test methods and procedures.

- Report for each model all warning signs, labels, and instructions for in- stallation, operation, and use that re- late to electronic product radiation safety.

- Provide, upon request, such other information as the Director may rea- sonably require to enable him/her to determine  whether  the  manufacturer

#### Food and Drug Administration, HHS	§ 1002.20

has acted or is acting in compliance with the Act and any standards pre- scribed thereunder, and to enable the Director to carry out the purposes  of the  Act.

[60 FR 48386, Sept. 19, 1995, as amended at 72

FR  17400,  Apr.  9,  2007;  75  FR  20916,  Apr.  22,

2010]

##### § 1002.11   Supplemental reports.

Prior to the introduction into com- merce of a new or modified model with- in a model or chassis family  of  a  prod- uct listed in table 1 of § 1002.1 for which a report under § 1002.10 is required, each manufacturer  shall  submit  a  report with respect to such new or modified model describing any changes in the in- formation previously submitted in the product report. Reports  will  be  re- quired  for  changes  that:

  - Affect actual or potential radi- ation emission.

  - Affect the manner of compliance with a standard or manner of testing for radiation safety.

##### § 1002.12   Abbreviated reports.

Manufacturers of products requiring abbreviated reports  as  specified  in table 1 of § 1002.1 shall submit, prior to the introduction of such product, a re- port  distinctly  marked   ‘‘Radiation Safety   Abbreviated   Report’’    which shall   include:

- Firm and model identification.

- A brief description of operational characteristics that affect radiation emissions, transmission, or leakage or that control exposure.

- A list of applications or uses.

- Radiation emission, transmission, or leakage levels.

- If necessary, additional informa- tion as may be requested to determine compliance with the Act and this part.

##### § 1002.13   Annual reports.

- Every manufacturer of products requiring an annual report  as  specified in table 1 of § 1002.1 shall submit an an- nual  report  summarizing  the   contents of the records required to be  main- tained by § 1002.30(a) and providing the volume of products produced, sold, or installed.

- Reports are due annually by Sep- tember 1. Such reports shall cover the

12-month period ending on June 30 pre- ceding the due date of the report.

- New models of a model family that do not involve changes in radi- ation emission or requirements of a performance standard do not require supplemental reports prior to introduc- tion into commerce. These model num- bers should be reported in quarterly updates to the annual report.

## Subpart C—Manufacturers’ Re- ports on Accidental Radiation Occurrences

##### § 1002.20 Reporting of accidental radi- ation occurrences.

- Manufacturers of electronic prod- ucts shall, where reasonable grounds for suspecting that such an incident has occurred, immediately report to the Director, Center for Devices and Radiological Health, all accidental ra- diation occurrences reported to or oth- erwise known to the manufacturer and arising from the manufacturing, test- ing, or use of any product introduced or intended to be introduced into com- merce by such manufacturer. Reason- able grounds include, but are not nec- essarily limited to, professional, sci- entific, or medical facts or opinions documented or otherwise, that con- clude or lead to the conclusion that such an incident has occurred.

- Such reports shall be  addressed to Food  and  Drug  Administration,  Center for Devices and Radiological Health, ATTN: Accidental  Radiation  Occur- rence Reports, Document Mail Center, 10903 New Hampshire Ave., Bldg. 66, rm. G609, Silver Spring, MD 20993–0002, and the reports and their envelopes shall be distinctly marked ‘‘Report on 1002.20’’ and shall contain all of the following information where known to the manu- facturer:

  - The nature of the accidental radi- ation occurrence;

  - The location at which the acci- dental radiation occurrence occurred;

  - The manufacturer, type, and model number of the electronic prod- uct or products involved;

  - The circumstances surrounding the accidental radiation occurrence, including causes;

  - The number of persons involved, adversely  affected,  or  exposed  during

§ 1002.30

the  accidental  radiation   occurrence, the nature and magnitude of their ex- posure and/or injuries and, if requested by the Director, Center for Devices and Radiological Health, the names of the persons involved;

  - The actions, if any, which may have been taken by the manufacturer, to control, correct, or eliminate the causes and to prevent reoccurrence; and

  - Any other pertinent information with respect to the accidental radi- ation occurrence.

- If a manufacturer is required to report to the Director under paragraph

- of this section and  also  is  required to report under part 803 of this chapter, the manufacturer shall report in ac- cordance with part 803. If a manufac- turer is required to report to the Direc- tor under paragraph (a) of this section and is not required to report under part 803, the manufacturer shall report in accordance with paragraph (a) of this section. A manufacturer need not file a separate report under this section if an incident involving an accidental radi- ation occurrence is associated with a defect or noncompliance  and  is  re- ported pursuant to § 1003.10 of this chapter.

[38 FR 28625, Oct. 15, 1973, as amended at 49

FR  36351,  Sept.  14,  1984;  53  FR  11254,  Apr.  6,

1988; 60 FR 48386, Sept. 19, 1995; 72 FR 17400,

Apr. 9, 2007; 75 FR 20916, Apr. 22, 2010]

## Subpart D—Manufacturers’ Records

##### § 1002.30   Records to be maintained by manufacturers.

  - Manufacturers of products listed under table 1 of § 1002.1 shall establish and  maintain  the  following  records with  respect  to  such  products:

    - Description of the quality control procedures with respect to electronic product radiation safety.

    - Records of the results of tests for electronic product radiation safety, in- cluding the control of unnecessary, sec- ondary or leakage electronic product radiation, the methods, devices, and procedures used in such tests, and the basis for selecting such methods, de- vices, and procedures.

    - For those products displaying aging effects which may increase elec-

21 CFR Ch. I (4–1–17 Edition)

tronic product radiation emission, records of the results of tests for dura- bility and stability of the product, and the basis for selecting these tests.

    - Copies of all written communica- tions between the manufacturer and dealers, distributors, and purchasers concerning radiation safety including complaints, investigations, instruc- tions, or explanations affecting  the use, repair, adjustment, maintenance, or testing of the listed product.

    - Data on production and sales vol- ume levels if available.

  - In addition to the records re- quired by paragraph (a) of this section, manufacturers of  products  listed  in table 1 of § 1002.1 shall establish and maintain the following records with re- spect  to  such  products:

    - A record of the manufacturer’s distribution of products in a form which will enable the tracing of spe- cific products or production lots to dis- tributors or to dealers in those in- stances in which the manufacturer dis- tributes directly to dealers.

    - Records received from dealers or distributors pursuant to § 1002.41.

[38 FR 28625, Oct. 15, 1973, as amended at 60

FR  48386,  Sept.  19,  1995;  75  FR  16352,  Apr.  1,

2010]

##### § 1002.31   Preservation  and  inspection of records.

      - Every manufacturer required to maintain records pursuant to this part, including records received pursuant to

§ 1002.41, shall preserve such records for a period of 5 years from the date of the record.

      - Upon reasonable notice by an offi- cer or employee duly designated by the Department, manufacturers shall per- mit such officer or employee to inspect appropriate books, records, papers, and documents as are relevant to deter- mining whether the manufacturer has acted or is acting in compliance with Federal standards.

      - Upon request of the Director, Cen- ter  for   Devices   and   Radiological Health, a manufacturer  of  products listed in table 1 of § 1002.1 shall submit to  the  Director,  copies  of  the  records

#### Food and Drug Administration, HHS	§ 1002.50

required   to   be   maintained   by   para- graph (b) of § 1002.30.

[38 FR 28625, Oct. 15, 1973, as amended at 53

FR  11254,  Apr.  6,  1988;  60  FR  48386,  Sept.  19,

1995]

## Subpart E—Dealer and Distributor Records

##### § 1002.40   Records  to  be  obtained  by dealers and distributors.

        - Dealers and distributors of elec- tronic products for which there are per- formance standards and for which the retail price is $50 or more shall obtain such information as is necessary to identify and locate first purchasers if the product is subject to  this  section by virtue of table 1 of § 1002.1.

        - Such information shall include:

          - The name and mailing address of the distributor, dealer, or purchaser to whom the product was transferred.

          - Identification and brand name of the product.

          - Model number and serial or other identification number of the product.

          - Date of sale, award, or lease.

        - The information obtained  pursu- ant to this section shall be forwarded immediately to the appropriate manu- facturer of the electronic product, or preserved as prescribed in § 1002.41.

[38 FR 28625, Oct. 15, 1973, as amended at 42

FR  18063,  Apr.  5,  1977;  60  FR  48386,  Sept.  19,

1995]

##### § 1002.41   Disposition   of   records   ob- tained by dealers and distributors.

- Information obtained by dealers and distributors pursuant to § 1002.40 shall immediately be forwarded to the appropriate   manufacturer   unless:

  - The dealer or distributor elects to hold and preserve such information and to immediately furnish it to the manu- facturer when advised by the manufac- turer or the Director, Center for De- vices and Radiological Health, that such information is required for pur- poses of section 535 of the Act; and

  - The dealer or distributor, upon making the election under paragraph (a)(1) of this section, promptly notifies the manufacturer of such election; such notification shall be in writing and shall identify the dealer or dis- tributor and the electronic product or

products  for  which  the  information  is being accumulated and preserved.

- Every dealer or distributor who elects to hold and preserve information required pursuant to § 1002.40 shall pre- serve the information for a period of 5 years from the date  of  the  sale,  award, or lease of the product,  or  until  the dealer or distributor discontinues deal- ing in, or distributing the product, whichever is sooner. If the dealer or distributor discontinues dealing in, or distributing the product, such informa- tion as obtained pursuant to § 1002.40 shall be furnished at that time, or be- fore, to the manufacturer of the prod- uct.

[38 FR 28625, Oct. 15, 1973, as amended at 42

FR  18063,  Apr.  5,  1977;  53  FR  11254,  Apr.  6,

1988; 75 FR 16352, Apr. 1, 2010]

##### § 1002.42 Confidentiality of records furnished by dealers and distribu- tors.

All information furnished to manu- facturers by dealers and distributors pursuant to this part shall  be  treated by such manufacturers as confidential information which may be used only as necessary to notify persons pursuant to section 535 of the Act.

[75 FR 16353, Apr. 1, 2010]

## Subpart F—Exemptions From Records and Reports Require- ments

##### § 1002.50   Special exemptions.

- Manufacturers of electronic prod- ucts may submit to the Director a re- quest, together with accompanying jus- tification, for exemption from any re- quirements listed in table 1 of § 1002.1. The request must specify each require- ment from which an exemption is re- quested. In addition to other informa- tion that is required, the justification must contain documented evidence showing that the product or product type for which the exemption is re- quested does not pose a public health risk and meets at least one of the fol- lowing  criteria:

  - The products cannot emit elec- tronic product radiation in sufficient intensity or of such quality, under any conditions  of  operation,  maintenance,

§ 1002.51

service,  or  product  failure,  to  be  haz- ardous;

  - The products are produced in small quantities;

  - The products are used by trained individuals and are to be used by the same manufacturing corporation or for research, investigation, or training.

  - The products are custom designed and used by trained individuals knowl- edgeable of the hazards; or

  - The products are produced in such a way that the requirements are inap- propriate or unnecessary.

- The Director may, subject to any conditions that the Director deems necessary to protect the public health, exempt manufacturers from all or part of the record and reporting require- ments of this part on the basis of infor- mation submitted  in  accordance  with paragraph (a) of this section or such other information which the Director may possess if the Director determines that such exemption is in keeping with the purposes of the Act.

- The Director will provide written notification of the reason for any de- nial. If the exemption is granted, the Director will provide written notifica- tion of:

  - The electronic product or prod- ucts for which the exemption has been granted;

  - The requirements from which the product is exempted; and

  - Such conditions as are deemed necessary to protect the public health and safety. Copies of exemptions  shall be  available  upon  request  from   the Food  and  Drug  Administration,  Center for Devices and Radiological Health, Division of  Mammography  Quality and Radiation Programs, 10903 New Hamp- shire Ave., Bldg. 66, rm. 4521, Silver Spring, MD 20993–0002.

- The Director may, on  the  Direc- tor’s own motion, exempt  certain  class- es of products from the reporting re- quirements listed in table 1 of § 1002.1, provided that the Director  finds  that such exemption is in keeping with the purposes of the act.

- Manufacturers of products for which there is no applicable perform- ance standard under parts 1020 through 1050 of this chapter and for which an investigational device exemption  has been   approved   under   § 812.30   of   this

#### 21 CFR Ch. I (4–1–17 Edition)

chapter or for which a premarket ap- proval application has been approved in accordance with § 814.44(d) of this chap- ter are exempt from submitting all re- ports listed in table 1 of § 1002.1.

[60 FR 48387, Sept. 19, 1995, as amended at 72

FR  17401,  Apr.  9,  2007;  75  FR  20916,  Apr.  22,

2010]

##### § 1002.51   Exemptions  for  manufactur- ers  of  products  intended  for  the

U.S. Government.

Upon application therefor by the manufacturer, the Director, Center for Devices and Radiological Health, may exempt from the provisions of this part a manufacturer of any electronic prod- uct intended for use by departments or agencies of the United States provided such department or agency has pre- scribed procurement specifications governing emissions of electronic prod- uct radiation and provided further that such product is of a type used solely or predominantly by departments or agencies of the United States.

[38 FR 28625, Oct. 15, 1973, as amended at 53

FR 11254, Apr. 6, 1988]

## PART 1003—NOTIFICATION OF DEFECTS OR FAILURE TO COMPLY

##### Subpart A—General Provisions

Sec.

- Applicability.

- Defect in an electronic product.

1003.5    Effect of regulations on other laws.

##### Subpart B—Discovery of Defect or Failure To Comply

- Discovery of defect or failure of com- pliance by manufacturer; notice require- ments.

- Determination   by   Secretary   that product fails to comply or has a defect.

##### Subpart C—Notification

- Notification by the manufacturer to the Secretary.

- Notification by the manufacturer to affected persons.

- Copies   of   communications   sent   to purchasers, dealers, or distributors.

##### Subpart D—Exemptions from Notification Requirements

- Application for exemption from noti- fication requirements.

- Granting the exemption.

Food and Drug Administration, HHS	§ 1003.11

AUTHORITY: 21 U.S.C. 360hh–360ss.

SOURCE:  38  FR  28628,  Oct.  15,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1003.1   Applicability.

The provisions of this part are appli- cable to electronic products which were manufactured after October 18, 1968.

##### § 1003.2   Defect  in  an  electronic  prod- uct.

For the purpose of this part, an elec- tronic product shall be considered to have a defect which relates to the safe- ty of use by reason of the emission of electronic product radiation if:

  - It is a product which does not uti- lize the emission of electronic product radiation in order to accomplish its purpose, and from which such emis- sions are unintended, and as a result of its design, production or assembly;

    - It emits electronic product radi- ation which creates a risk of injury, in- cluding genetic injury, to any person, or

    - It fails to conform to its design specifications relating to electronic ra- diation emissions; or

  - It is a product which utilizes elec- tronic product radiation to accomplish its primary purpose and from which such emissions are intended, and as a result of its design, production or as- sembly it;

    - Fails to conform to its design specifications relating to the emission of electronic product radiation; or

    - Without regard to the design spec- ifications of the product, emits elec- tronic product radiation unnecessary to the accomplishment of its primary purpose which creates a risk of injury, including genetic injury to any person; or

    - Fails to accomplish the intended purpose.

##### § 1003.5   Effect of regulations on other laws.

The remedies provided for in this subchapter shall be in addition to and not in substitution for any other rem- edies provided by law and shall not re- lieve any person from liability at com- mon law or under statutory law.

## Subpart B—Discovery of Defect or Failure To Comply

##### § 1003.10 Discovery of defect or failure of compliance by manufacturer; no- tice requirements.

Any manufacturer who discovers that any electronic product produced, assembled, or imported by him, which product has left its place of manufac- ture, has a defect or fails to comply with an applicable Federal standard shall:

- Immediately  notify  the  Secretary in accordance with § 1003.20, and

- Except as authorized by § 1003.30, furnish notification with reasonable promptness  to  the  following  persons:

  - The dealers or distributors to whom such product was delivered by the manufacturer; and

  - The purchaser of such product and any subsequent transferee of such prod- uct (where known to the manufacturer or where the manufacturer upon rea- sonable inquiry to dealers, distribu- tors, or purchasers can identify the present user).

- If a manufacturer is required to notify the  Secretary  under  paragraph

- of this section and  also  is  required to report to the Food and Drug Admin- istration under part 803 of this chapter, the manufacturer shall report in ac- cordance with part 803. If a manufac- turer is required to notify the  Sec- retary under paragraph (a) of this sec- tion and is not required to  report  to the  Food  and  Drug   Administration under part 803, the manufacturer shall notify   the   Secretary   in   accordance with  paragraph  (a)  of  this  section.

[38  FR  28628,  Oct.  15,  1973  and  49  FR  36351,

Sept. 14, 1984]

##### § 1003.11 Determination by Secretary that product fails to comply or has a defect.

  - If, the Secretary, through testing, inspection, research, or examination of reports or other data, determines that any electronic product does not comply with an applicable Federal standard issued pursuant to the Act or has a de- fect, he shall immediately notify the manufacturer of the product in writing specifying:

    - The defect in the product or the manner  in  which  the  product  fails  to

§ 1003.20

comply with the applicable Federal standard;

    - The Secretary’s findings, with ref- erences to the tests, inspections, stud- ies, or reports upon which such find- ings are based;

    - A reasonable period of time during which the manufacturer may present his views and evidence to establish that there is no failure of compliance or that the alleged defect does not exist or does not relate to safety of use of the product by reason of the emis- sion of electronic product radiation.

The manufacturer shall have an oppor- tunity for a regulatory hearing before the Food and Drug Administration pur- suant to part 16 of this chapter.

  - Every manufacturer who receives a notice under paragraph (a) of this section shall immediately advise the Secretary in writing of the total num- ber of such product units produced and the approximate number of such prod- uct units which have left the place of manufacture.

  - If, after the expiration of the pe- riod of time specified in the notice, the Secretary determines that the product has a defect or does not comply with an applicable Federal standard and the manufacturer has not applied for an ex- emption, he shall direct the manufac- turer to furnish the notification to the persons specified in § 1003.10(b) in the manner specified in § 1003.21. The manu- facturer shall within 14 days from the date of receipt of such directive furnish the  required  notification.

[38 FR 28628, Oct. 15, 1973, as amended at 41

FR  48269,  Nov.  2,  1976;  42  FR  15676,  Mar.  22,

1977]

## Subpart C—Notification

##### § 1003.20   Notification by the manufac- turer to the Secretary.

The notification to the Secretary re- quired by § 1003.10(a) shall be confirmed in writing and, in addition to other rel- evant information which the Secretary may require, shall include the fol- lowing:

- Identification of the product or products involved;

- The total number of such product units so produced, and the approximate number of such product units which have left the place of manufacture;

21 CFR Ch. I (4–1–17 Edition)

- The expected usage for the prod- uct if known to the manufacturer;

- A description of the defect in the product or the manner in which the product fails to comply with an appli- cable Federal standard;

- An evaluation of the hazards rea- sonably related to defect or the failure to comply with the Federal standard;

- A statement of the measures to be taken to repair such defect or to bring the product into compliance with the Federal standard;

- The date and circumstances under which the defect was discovered; and

- The identification of any trade se- cret information which the manufac- turer desires kept confidential.

##### § 1003.21   Notification by the manufac- turer to affected persons.

- The notification to the persons specified in § 1003.10(b) shall be in writ- ing and, in addition to other relevant information which the Secretary may require, shall include:

  - The   information   prescribed   by

§ 1003.20 (a), (d), and instructions with respect to the use of the  product  pend- ing  the  correction  of  the  defect;

  - A clear evaluation in nontechnical terms of the hazards reasonably related to any defect or failure to comply; and

  - The following statement:

The manufacturer will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be ap- proved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you.

Provided, That if at the time the notifi- cation is sent, the Secretary has ap- proved a plan for the repair, replace- ment or refund of the product, the no- tification may include the details of the approved plan in lieu of the above statement.

- The envelope containing the no- tice shall not contain advertising or other extraneous material, and such mailings will be made in accordance with this section.

  - No. 10 white envelopes shall be used, and the name and address of the manufacturer shall appear in the upper left corner of the envelope.

#### Food and Drug Administration, HHS	§ 1003.31

  - The following statement is to ap- pear in the far left third of the enve- lope in the type and size indicated and in reverse printing, centered in a red rectangle 33⁄4 inches wide and 21⁄4 inches high:

IMPORTANT—ELECTRONIC PRODUCT RADIATION WARNING

The statement shall be in three lines, all capitals, and centered. ‘‘Important’’ shall be in 36-point Gothic Bold type. ‘‘Electronic Product’’ and ‘‘Radiation Warning’’ shall be in 36-point Gothic Condensed type.

  - Envelopes with markings similar to those prescribed in this section shall not be used by manufacturers for mail- ings other than those required by this part.

- The notification shall be sent:

  - By certified mail to purchasers of the product and to subsequent trans- ferees.

  - By certified mail or other more expeditious means to dealers and dis- tributors.

- Where products were sold under a name other than that of the manufac- turer of the product, the name of the individual or company under whose name the product was sold may be used in the notification required by this sec- tion.

##### § 1003.22 Copies of  communications sent to  purchasers, dealers  or  dis- tributors.

- Every manufacturer of electronic products shall furnish to the Secretary a copy of all notices, bulletins, or other communications sent to the dealers or distributors of such manufacturers or to purchasers (or subsequent trans- ferees) of electronic products of such manufacturer regarding any defect in such product or any failure of such product to comply with an applicable Federal standard.

- In the event the Secretary deems the content of such notices to be insuf- ficient to protect the public health and safety, the Secretary may require addi- tional notice to such recipients, or may elect to make or cause to be made such notification by whatever means he deems appropriate.

## Subpart D—Exemptions From Notification Requirements

##### § 1003.30   Application	for	exemption from notification requirements.

- A manufacturer may at  the  time of giving the written confirmation re- quired by § 1003.20 or within 15 days of the receipt of any notice from the Sec- retary pursuant to § 1003.11(a), apply for an exemption from the requirement of notice    to    the    persons    specified    in

§ 1003.10(b).

- The application for  exemption shall  contain  the  information  required by § 1003.20 and in addition shall set forth in detail the grounds upon  which the  exemption  is  sought.

##### § 1003.31   Granting the exemption.

- If, in the judgment of the Sec- retary,  the  application  filed   pursuant to § 1003.30 states reasonable grounds for an exemption from  the  requirement of notice, the Secretary shall give the manufacturer  written   notice   specifying a reasonable  period  of  time  during which he may present his views and evidence  in  support  of  the  application.

- Such views and evidence shall be confined to matters relevant  to  wheth- er the defect in the product or  its  fail- ure to comply with an applicable Fed- eral standard is such as to create a sig- nificant  risk  of  injury,  including   ge- netic injury, to any person and shall be presented in writing unless the Sec- retary determines that an oral presen- tation is desirable. Where such evi- dence includes nonclinical laboratory studies, the data submitted shall in- clude, with respect to each such study, either a statement that the study was conducted in compliance with the re- quirements set forth in part 58 of this chapter, or, if the study was not con- ducted in compliance with such regula- tions, a brief statement of the  reason for the noncompliance. When such evi- dence includes clinical investigations involving human subjects, the data submitted  shall  include,  with   respect to each clinical investigation either a statement that each investigation was conducted in compliance with the re- quirements set forth in part 56 of this chapter, or a statement that the inves- tigation is not subject to such require- ments   in   accordance   with   § 56.104   or

Pt. 1004

§ 56.105, and a statement that each in- vestigation was conducted in compli- ance with the requirements set forth in part 50 of this chapter.

- If, during the  period of time  af- forded the manufacturer to present his views and evidence, the manufacturer proves to the Secretary’s satisfaction that the defect or failure to comply does not create a significant risk of in- jury, including genetic injury, to any person, the Secretary shall issue an ex- emption from the requirement of noti- fication to the manufacturer and shall notify the manufacturer in writing specifying:

  - The electronic product or prod- ucts for which the exemption has been issued; and

  - Such conditions as the Secretary deems necessary to protect the public health and safety.

- Any person who contests denial of an exemption shall have an oppor- tunity for a regulatory hearing before the Food and Drug Administration pur- suant to part 16 of this chapter.

[38 FR 28628, Oct. 15, 1973, as amended at 41

FR  48269,  Nov.  2,  1976;  42  FR  15676,  Mar.  22,

1977; 50 FR 7518, Feb. 22, 1985]

## PART 1004—REPURCHASE, REPAIRS, OR REPLACEMENT OF ELEC- TRONIC PRODUCTS

Sec.

- Manufacturer’s obligation to repair, replace, or refund cost of electronic prod- ucts.

- Plans for the repair of electronic products.

- Plans for the replacement of elec- tronic products.

- Plans for refunding the cost of elec- tronic products.

- Approval  of  plans. AUTHORITY: 21 U.S.C. 360hh–360ss.

SOURCE:  38  FR  28629,  Oct.  15,  1973,  unless

otherwise noted.

##### § 1004.1 Manufacturer’s obligation to repair, replace, or refund cost of electronic products.

  - If any electronic product fails to comply with an applicable Federal standard or has a defect and the notifi- cation specified in § 1003.10(b) of this chapter is required to be furnished, the manufacturer of such product shall;

21 CFR Ch. I (4–1–17 Edition)

- Without charge, bring such prod- uct into conformity with such standard or remedy such defect and provide re- imbursement for any expenses for transportation of such product in- curred in connection with having such product brought into conformity or having such defect remedied; or

- Replace such product with  a  like or equivalent product which complies with each applicable Federal standard and which has no defect relating to the safety of its use; or

- Make a refund of the cost of the product to the purchaser.

  - The manufacturer shall take the action required by this section in ac- cordance with a plan approved by the Secretary pursuant to § 1004.6.

##### § 1004.2   Plans  for  the  repair  of  elec- tronic products.

Every plan for bringing an electronic product into conformity with applica- ble Federal standards or for remedying any defect in such product shall be sub- mitted to the Secretary in writing, and in addition to other relevant informa- tion which the Secretary may require, shall include:

- Identification of the product in- volved.

- The approximate number of defec- tive product units which have left the place of manufacture.

- The specific modifications, alter- ations, changes, repairs, corrections, or adjustments to be made to bring the product into conformity or remedy any defect.

- The manner in which the oper- ations described in paragraph (c)  will be accomplished, including the proce- dure for obtaining access to, or posses- sion of, the products and the location where such operations will be per- formed.

- The technical data, test results or studies demonstrating the effective- ness of the proposed remedial action.

- A time limit, reasonable  in  light of the circumstances, for completion of the operations.

- The system by which the manu- facturer will provide reimbursement for any transportation expenses in- curred in connection with having such product brought into conformity or having any defect remedied.

#### Food and Drug Administration, HHS	§ 1004.4

- The text of the statement which the manufacturer will send to the per- sons specified in § 1003.10(b)  of this chapter informing such persons;

  - That the manufacturer, at his ex- pense, will repair the electronic prod- uct involved,

  - Of the method by which the manu- facturer will obtain access to or posses- sion of the product to make such re- pairs,

  - That the manufacturer will reim- burse such persons for any transpor- tation expenses incurred in connection with making such repairs, and

  - Of the manner in which such reim- bursement will be effected.

    - An assurance that the manufac- turer will provide the Secretary with progress reports on the effectiveness of the plan, including the number of elec- tronic products repaired.

##### § 1004.3   Plans  for  the  replacement  of electronic products.

Every plan for replacing an elec- tronic product with a like or equiva- lent product shall be submitted to the Secretary in writing, and in addition to other relevant information which the Secretary may require, shall include:

- Identification of the product to be replaced.

- A description of the replacement product in sufficient detail to support the manufacturer’s contention that the replacement product is like or equiva- lent to the product being replaced.

- The approximate number of defec- tive product units which have left the place of manufacture.

- The manner in which the replace- ment operation will be effected includ- ing the procedure for obtaining posses- sion of the product to be replaced.

- A time limit, reasonable, in light of the circumstances for completion of the replacement.

- The steps which the manufacturer will take to insure that the defective product will not be reintroduced into commerce, until it complies with each applicable Federal standard and has no defect relating to the safety of its use.

- The system by which the manu- facturer will provide reimbursement for any expenses for transportation of such product incurred in connection with effecting the replacement.

- The text of the statement which the manufacturer will send to the per- sons specified in § 1003.10(b)  of this chapter informing such persons;

  - That the manufacturer, at its ex- pense, will replace the electronic prod- uct involved,

  - Of the method by which the manu- facturer will obtain possession of the product and effect the replacement,

  - That the manufacturer will reim- burse such persons for any transpor- tation expenses incurred in connection with effecting such replacement, and

  - Of the manner in which such reim- bursement will be made.

    - An assurance that the manufac- turer will provide the Secretary with progress reports on the effectiveness of the plan, including the number of elec- tronic products replaced.

##### § 1004.4   Plans for refunding the cost of electronic products.

Every plan for refunding the cost of an electronic product shall be sub- mitted to the Secretary in writing, and in addition to other relevant informa- tion which the Secretary may require, shall include:

- Identification of the product in- volved.

- The approximate number of defec- tive product units which have left the place of manufacture.

- The manner in which the refund operation will be effected including the procedure for obtaining possession of the product for which the refund is to be made.

- The steps which the manufacturer will take to insure that the defective products will not be reintroduced into commerce, until it complies with each applicable Federal standard and has no defect relating to the safety of its use.

- A time limit, reasonable  in  light of the circumstances, for obtaining the product and making the refund.

- A statement that the manufac- turer will refund the cost of such prod- uct together with the information the manufacturer has used to determine the amount of the refund.

- The text of the statement which the manufacturer will send to the per- sons specified in § 1003.10(b)  of this chapter informing such persons;

#### § 1004.6

  - That the manufacturer, at his ex- pense, will refund the cost of the elec- tronic product plus any transportation costs,

  - Of the amount to be refunded ex- clusive of transportation costs,

  - Of the method by which the manu- facturer will obtain possession of the product and make the refund.

- An assurance that the manufac- turer will provide the Secretary with progress reports on the effectiveness of the plan, including the number of re- funds made.

##### § 1004.6   Approval of plans.

If, after review of any plan submitted pursuant to this subchapter, the Sec- retary determines that the action to be taken by the manufacturer will expedi- tiously and effectively fulfill the man- ufacturer’s obligation under § 1004.1 in a manner designed to encourage the public to respond to the proposal, the Secretary will send written notice of his approval of such plan to the manu- facturer. Such approval may be condi- tioned upon such additional terms as the Secretary deems necessary to pro- tect the public health and safety. Any person who contests denial of a plan shall have an opportunity for a regu- latory hearing before the Food and Drug Administration pursuant to part 16 of this chapter.

[38 FR 28629, Oct. 15, 1973, as amended at 41

FR  48269,  Nov.  2,  1976;  42  FR  15676,  Mar.  22,

1977]

## PART 1005—IMPORTATION OF ELECTRONIC PRODUCTS

##### Subpart A—General Provisions

Sec.

- Applicability.

- Definitions.

- Importation   of   noncomplying   goods prohibited.

##### Subpart B—Inspection and Testing

- Notice of sampling.

- Payment for samples.

##### Subpart C—Bonding and Compliance Procedures

- Hearing.

- Application for permission to bring product into compliance.

21 CFR Ch. I (4–1–17 Edition)

- Granting permission to bring prod- uct into compliance.

- Bonds.

- Costs of bringing product into com- pliance.

- Service of process on manufacturers. AUTHORITY: 21 U.S.C. 360ii, 360mm.

SOURCE:  38  FR  28630,  Oct.  15,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1005.1   Applicability.

  - The provisions of §§ 1005.1 through

1005.24 are applicable to electronic products which are subject to the standards prescribed under this sub- chapter and are offered for importation into  the  United  States.

  - Section 1005.25 is applicable to every manufacturer of electronic prod- ucts offering an electronic product for importation  into  the  United  States.

[38 FR 28630, Oct. 15, 1973, as amended at 45

FR 81739, Dec. 12, 1980]

##### § 1005.2   Definitions.

As used in this part:

The term owner or consignee means the person who makes entry under the provisions of section 484 of the Tariff Act of 1930, as amended (19 U.S.C. 1484), namely, the ‘‘importer of record.’’

[81 FR 85973, Nov. 29, 2016]

##### § 1005.3   Importation  of  noncomplying goods prohibited.

The importation of any electronic product for which standards have been prescribed under section 534 of the Fed- eral Food, Drug, and Cosmetic Act (the act) (21 U.S.C. 360kk) shall be refused admission into the United States un- less there is affixed to such product a certification in the form of a  label  or tag in conformity with section 534(h) of the act (21 U.S.C. 360kk(h)). Merchan- dise refused admission shall be de- stroyed or exported under regulations prescribed by the Secretary of the Treasury unless a timely and adequate petition for permission to bring the product into compliance is filed and granted under §§ 1005.21 and 1005.22.

[69 FR 11314, Mar. 10, 2004]

#### Food and Drug Administration, HHS	§ 1005.22

Subpart B—Inspection and Testing

##### § 1005.10   Notice of sampling.

When a sample of a product to be of- fered for importation has been re- quested by the Secretary, the District Director of Customs having  jurisdic- tion over the shipment shall, upon the arrival of the shipment, procure the sample and shall give to its owner or consignee prompt notice of the deliv- ery or  of  the  intention to deliver such sample to the Secretary.  If  the  notice so requires, the owner or consignee will hold the shipment of which the sample is typical and not release such ship- ment until he receives notice of the re- sults of the tests of the  sample  from the Secretary, stating  that  the  product is in compliance with the requirements of the Act. The District Director of Customs will be given the results of the tests. If the Secretary notifies the Dis- trict Director of Customs that  the product does not meet the require- ments of the Act,  the  District  Director of Customs shall require the expor- tation or destruction of  the  shipment in accordance with customs laws.

##### § 1005.11   Payment for samples.

The Department  of  Health  and Human Services will pay for all import samples of electronic products ren- dered unsalable as a  result  of  testing, or will pay the reasonable costs of re- packaging such samples for sale, if the samples are found to be in compliance with the requirements of Subchapter C—Electronic Product Radiation Con- trol of the Federal Food, Drug, and Cosmetic Act (formerly the Radiation Control for Health and Safety Act of 1968). Billing for reimbursement shall be made by the owner or consignee to the Center for Devices and Radio- logical Health, 5600 Fishers Lane, Rockville, MD 20857. Payment for sam- ples will not  be made  if the sample  is found to be in violation of the Act, even though subsequently brought into compliance pursuant to terms specified in a notice of permission issued under

§ 1005.22.

[73 FR 34860, June 19, 2008]

## Subpart C—Bonding and Compliance Procedures

##### § 1005.20   Hearing.

- If, from an examination of the sample or otherwise, it appears that the product may be subject to a refusal of admission, the Secretary shall give the owner or consignee a written no- tice to that effect, stating the reasons therefor. The notice shall specify a place and a period of time during which the owner or consignee shall have an opportunity to introduce testimony unless the owner or consignee indicates his intention to bring the product into compliance. Upon timely request, such time and place may be changed. Such testimony shall be confined to matters relevant to the admissibility of the ar- ticle and may be introduced orally or in writing.

- If the owner or consignee submits or indicates his intention to submit an application for permission to perform such action as is necessary to bring the product into compliance with the Act, such application shall include the in- formation required by § 1005.21.

- If the application is not submitted at or prior to the hearing, the Sec- retary may allow a reasonable time for filing such application.

##### § 1005.21   Application for permission to bring product into compliance.

Application for permission to per- form such action as is necessary to bring the product into compliance with the Act may be filed only by the owner, consignee, or manufacturer and, in ad- dition to any other information which the Secretary may reasonably require, shall:

- Contain a detailed proposal for bringing the product into compliance with the Act;

- Specify the time and place where such operations will be effected and the approximate time for their completion; and

- Identify the bond required to be filed pursuant to § 1005.23.

##### § 1005.22   Granting permission to bring product into compliance.

- When permission contemplated by

§ 1005.21  is  granted,  the  Secretary  shall

§ 1005.23

notify the applicant in writing, speci- fying:

  - The procedure to be followed;

  - The disposition of the rejected ar- ticles or portions thereof;

  - That the operations are to be car- ried out under the supervision of a rep- resentative of the Department of Health and Human Services;

  - A reasonable time limit for com- pleting the operations; and

  - Such other conditions as he finds necessary to maintain adequate super- vision and control over the product.

- Upon receipt of a written request for an extension of time to complete the operations necessary to bring the product into compliance, the Secretary may grant such additional time as he deems necessary.

- The notice of permission may be amended upon a showing of reasonable grounds thereof and the filing of an amended application for  permission with  the  Secretary.

- If ownership of a product included in a notice of permission changes be- fore the operations specified in the no- tice have been completed, the original owner will remain responsible under its bond, unless the new owner has exe- cuted a superseding bond on customs Form 7601 and obtained a new notice.

- The Secretary will notify the Dis- trict Director of Customs having juris- diction over the shipment involved, of the determination as to whether or not the product has in fact been brought into compliance with the Act.

##### § 1005.23   Bonds.

The bond required under section 360(b) of the Act shall be  executed  by the owner or consignee on the appro- priate form of a customs single-entry bond, customs Form 7551 or term bond, customs Form  7553  or  7595,  containing a condition for the redelivery of the shipment or any part thereof not com- plying with the laws and regulations governing its admission into the com- merce of the United States  upon  de- mand of the District  Director  of  Cus- toms and containing a provision for the performance of any action necessary to bring the product into  compliance  with all  applicable  laws   and   regulations. The bond shall be filed with the Dis- trict  Director  of  Customs.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 1005.24   Costs   of   bringing   product into compliance.

The costs of supervising the oper- ations necessary to  bring  a  product into compliance with the Act shall be paid by  the  owner  or  consignee  who files an application pursuant to § 1005.21 and executes a bond  under  section 360(b) of the Act. Such costs shall in- clude:

- Travel expenses of the supervising officer;

- Per diem in lieu of subsistence of the supervising officer when away from his home station, as provided by law;

- Service fees: (1) The charge for the services  of  the  supervising   officer, which shall include administrative sup- port, shall be computed at a rate per hour equal to 266 percent of the hourly rate of regular pay of a grade GS–11/4 employee, except that such services performed by a customs officer and subject to the provisions of the act of February 13, 1911, as amended (sec. 5, 36 Stat. 901, as amended (19 U.S.C. 267)), shall be calculated as provided in that act.

- The charge for the services of the analyst, which shall include adminis- trative and laboratory support, shall be computed at a rate per hour equal to

266  percent  of  the  hourly  rate  of  reg- ular pay of a grade GS–12/4 employee.

- The rate per hour equal to 266 per- cent of the equivalent hourly rate of regular pay of the supervising officer (GS–11/4) and the analyst (GS–12/4) is computed as follows:

Hours

Gross  number  of  working  hours  in  52  40-hour

weeks ...................................................................	2,080

Less:

Nine legal public  holidays—New Years Day, Washington’s Birthday, Memorial Day, Inde- pendence Day, Labor Day, Columbus Day, Veterans   Day,   Thanksgiving   Day,   and

Christmas Day ..............................................	72

Annual Leave—26 days ...................................	208

Sick Leave—13 days  .......................................	104

Total ...................................................	384

Net number of working hours ............	1,696 Gross  number  of  working  hours  in  52  40-hour

weeks ...................................................................	2,080

Working hour equivalent of Government contribu- tions for employee retirement, life insurance, and health benefits computed at 81⁄2% of annual

rate of pay of employee .......................................	176

Equivalent annual working hours ............................	2,256

#### Food and Drug Administration, HHS	§ 1005.25

|  | Hours |
| --- | --- |
| Support required to equal to 1 man-year ................ Equivalent gross annual working hours charged to Food and Drug appropriation ............................... | 2,256 |
|  | 4,512 |

NOTE: Ratio of equivalent gross annual number  of  working  hours  charged  to  Food and Drug appropriation to net number of an- nual working hours (4512/1696) = 266 pct.

- The minimum charge for services of supervising officers shall be not less than the charge for 1 hour and time after the first hour  shall  be  computed in multiples of 1 hour, disregarding fractional parts  less  than  one-half hour.

[38 FR 28630, Oct. 15, 1973, as amended at 42

FR  55207,  Oct.  14,  1977;  42  FR  62130,  Dec.  9,

1977]

##### § 1005.25   Service of process on manu- facturers.

- Every manufacturer of electronic products, prior to offering such product for importation into the United States, shall designate a permanent resident of the United States as the manufactur- er’s agent upon whom service of all processes, notices, orders, decisions, and requirements may be made for and on behalf of the manufacturer as pro- vided in section 536(d)  of Subchapter C—Electronic Product Radiation Con- trol of the Federal Food, Drug, and Cosmetic Act (formerly the Radiation Control for Health and Safety Act of 1968) (21 U.S.C. 360mm(d)) and this sec- tion. The agent may be an individual, a firm, or a domestic corporation. For purposes of this section, any number of manufacturers may designate the same agent.

- A manufacturer designating an agent must address the designation to the Center for Devices and Radio- logical Health, 10903  New Hampshire Ave., Document Mail Center—WO66– G609, Silver Spring, MD 20993–0002. It must be in writing and dated; all signa- tures must be in ink. The designation must be made in the legal form re- quired to make it valid and binding on the manufacturer under the laws, cor- porate bylaws, or other requirements governing the making of the designa- tion by the manufacturer at the place and time where it is made, and the per- sons or person signing the designation

shall certify that it is so made. The designation must disclose the manufac- turer’s full legal name and the name(s) under which the manufacturer con- ducts the business, if applicable, the principal place of business, and mailing address. If any of the products of the manufacturer do not bear his legal name, the designation must identify the marks, trade names, or other des- ignations of origin which these prod- ucts bear. The designation must pro- vide that it will remain in effect until withdrawn or replaced by the manufac- turer and shall bear a declaration of acceptance duly signed by the des- ignated agent. The full legal name and mailing address of the agent must be stated. Until rejected by the Secretary, designations are binding on the manu- facturer even when not in compliance with all the requirements of this sec- tion. The designated agent may not as- sign performance of his function under the designation to another.

- Service of any process, notice, order, requirement, or decision speci- fied in section 536(d) of Subchapter C— Electronic  Product  Radiation  Control of the Federal Food, Drug, and Cos- metic Act (formerly the Radiation Control for Health and Safety Act of 1968) (21 U.S.C. 360mm(d)) may be made by registered or certified mail ad- dressed to the agent with return re- ceipt requested, or in any other man- ner authorized by law.  In  the  absence of such a designation or if for any rea- son service on the designated agent cannot be effected, service may be made as provided in section 536(d) by posting such process, notice, order, re- quirement, or decision in the Office  of the Director, Center for  Devices and Radiological Health and publishing a notice that such service was made  in the   FEDERAL    REGISTER.

[38 FR 28630, Oct. 15, 1973, as amended at 53

FR  11254,  Apr.  6,  1988;  65  FR  17137,  Mar.  31,

2000;  72  FR  17401,  Apr.  9,  2007;  73  FR  34860,

June 19, 2008; 75 FR 16353, Apr. 1, 2010; 78 FR

18234, Mar. 26, 2013]

#### Pt. 1010

PART	1010—PERFORMANCE STANDARDS FOR ELECTRONIC PRODUCTS: GENERAL

##### Subpart A—General Provisions

Sec.

- Scope.

- Certification.

- Identification.

- Variances.

- Exemptions for products intended for United States Government use.

##### Subpart B—Alternate Test Procedures

1010.13    Special test procedures.

##### Subpart C—Exportation of Electronic Products

- Electronic  products  intended  for  ex- port.

AUTHORITY: 21 U.S.C. 351, 352, 360, 360e–360j,

360hh–360ss, 371, 381.

SOURCE:  38  FR  28631,  Oct.  15,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1010.1   Scope.

The standards listed in this sub- chapter are prescribed pursuant to sec- tion 534 of Subchapter C—Electronic Product Radiation Control of the Fed- eral Food, Drug, and Cosmetic Act (for- merly the Radiation Control for Health and Safety Act of 1968) (21 U.S.C. 360kk) and are applicable to electronic products as specified herein, to control electronic product radiation from such products. Standards so prescribed are subject to amendment or revocation and additional standards may be pre- scribed as are determined necessary for the protection of the public health and safety.

[73 FR 34861, June 19, 2008]

##### § 1010.2   Certification.

  - Every manufacturer of an elec- tronic product for which an applicable standard is in effect under this sub- chapter shall furnish to the dealer or distributor, at the time of delivery of such product, the certification that such product conforms to all applicable standards under this subchapter.

  - The certification shall be in the form of a label or tag permanently af-

21 CFR Ch. I (4–1–17 Edition)

fixed to or inscribed on such product so as to be legible and readily accessible to view when the product is fully as- sembled for use, unless the applicable standard prescribes some other manner of certification. All such labels or tags shall be in the English language.

  - Such certification shall be based upon a test, in accordance with the standard, of the individual article to which it is attached or upon a testing program which is in accordance with good manufacturing practices. The Di- rector, Center for Devices and Radio- logical Health may disapprove such a testing program on the grounds that it does not assure the adequacy of safe- guards against hazardous electronic product radiation or that it does not assure that electronic products comply with the standards prescribed under this subchapter.

  - In the case of products for which it is not feasible to certify in accord- ance with paragraph (b) of this section, upon application by the manufacturer, the Director, Center for Devices and Radiological Health may approve an alternate means by which such certifi- cation may be provided.

[38 FR 28631, Oct. 15, 1973, as amended at 40

FR 32257,  July 31,  1975;  42  FR 18063,  Apr. 5,

1977; 53 FR 11254, Apr. 6, 1988]

##### § 1010.3   Identification.

- Every manufacturer of an elec- tronic product to which a standard under this subchapter is applicable shall set forth the information speci- fied in paragraphs (a)(1) and (2) of this section. This information shall be pro- vided in the form of a tag or label per- manently affixed or inscribed on such product so as to be legible and readily accessible to view when the product is fully assembled for use or in such other manner as may be prescribed in the ap- plicable standard. Except for foreign equivalent abbreviations as authorized in paragraph (a)(1) of this section all such labels or tags shall be in  the English  language.

  - The full name and address of the manufacturer of the product; abbrevia- tions such as ‘‘Co.,’’ ‘‘Inc.,’’ or their foreign equivalents and the first and middle initials of individuals may be used. Where products are sold under a

#### Food and Drug Administration, HHS	§ 1010.4

name other than that of the manufac- turer of the product, the full name and address of the individual or company under whose name the product was sold may be set forth, provided such indi- vidual or company has previously suppled the Director, Center for De- vices and Radiological Health with suf- ficient information to identify the manufacturer of the product.

  - The place and month and year of manufacture:

    - The place of manufacture may be expressed in code provided the manu- facturer has previously supplied the Di- rector, Center for Devices and Radio- logical Health with the key to such code.

    - The month and year of manufac- ture shall be provided clearly and leg- ibly, without abbreviation, and with the year shown as a four-digit number as follows:

MANUFACTURED: (INSERT MONTH AND

YEAR OF MANUFACTURE.)

- In the case of products for which it is not feasible to affix identification labeling in accordance with paragraph

(a) of this section, upon application by the manufacturer, the Director, Center for Devices and Radiological Health may approve an alternate means by which such identification may be pro- vided.

(c) Every manufacturer of an elec- tronic product to which a standard under this subchapter is applicable shall provide to the Director, Center for Devices and Radiological Health a list identifying each brand name which is applied to the product together with the full name and address of the indi- vidual or company for whom each prod- uct so branded is manufactured.

[40 FR 32257, July 31, 1975, as amended at 42

FR  18063,  Apr.  5,  1977;  53  FR  11254,  Apr.  6,

1988]

##### § 1010.4   Variances.

- Criteria for variances. (1) Upon ap- plication by a manufacturer (including an assembler), the Director, Center for Devices and Radiological Health, Food and Drug Administration, may grant a variance from one or  more  provisions of any performance standard under subchapter J of this chapter for an electronic    product    subject    to    such

standard when the Director determines that granting such a variance is in keeping with the purposes of Sub- chapter C—Electronic Product Radi- ation Control of the Federal Food, Drug, and Cosmetic Act (formerly the Radiation Control for Health  and  Safe- ty Act of 1968), and:

- The scope of the requested vari- ance is so limited in its applicability as not to justify an amendment to the standard, or

- There is not sufficient time for the promulgation of an amendment to the standard.

- The issuance of the variance shall be based upon a determination that:

  - The product utilizes an alternate means for providing radiation safety or protection equal to  or  greater  than that provided by products meeting all requirements of the applicable stand- ard, or

  - The product performs a function or is intended for a  purpose  which could not be performed or accom- plished if required to meet the applica- ble standards, and suitable means for assuring radiation safety or protection are provided, or

  - One or more requirements of the applicable standard are not appro- priate, and suitable means for assuring radiation safety or protection are pro- vided.

- Applications for variances. If you are submitting an application for variances or for amendments or exten- sions thereof, you must submit  an original and two copies  to  the  Division of  Dockets  Management  (HFA–305), Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852.

  - The application for variance shall include the following information:

    - A description of the product  and its intended use.

    - An explanation of how compli- ance with the applicable  standard would restrict or be inappropriate for this intended use.

    - A description of the manner in which it is proposed to deviate  from the requirements of the applicable standard.

    - A description of the  advantages to be derived from such deviation.

#### § 1010.4

    - An explanation of how alternate or suitable means of radiation protec- tion will be provided.

    - The period of time it is desired that the variance be in effect, and, if appropriate, the number of units the applicant wishes to manufacture.

    - In the case of prototype or ex- perimental equipment, the proposed lo- cation of each unit.

    - Such other information re- quired by regulation or by the Direc- tor, Center for Devices and Radio- logical Health, to evaluate and act on the application.

    - With respect to each nonclinical laboratory study contained in the ap- plication, either a statement that the study was conducted in compliance with the good laboratory practice regu- lations set forth in part 58 of this chap- ter, or, if the study was not conducted in compliance with such regulations, a brief statement of the reason for the noncompliance.

    - [Reserved]

    - If the electronic product is used in a clinical investigation involving human subjects, is subject to the re- quirements for institutional review set forth in part 56 of this chapter, and is subject to the requirements for in- formed consent set forth in part 50 of this chapter, the investigation shall be conducted in compliance with such re- quirements.

  - The application for amendment or extension of a variance shall include the following information:

    - The variance number and expira- tion  date.

    - The amendment or extension re- quested and basis for  the  amendment or extension.

    - A description of the effect of the amendment or extension on protection from radiation produced by the prod- uct.

    - An explanation of how alternate or suitable means of protection will be provided.

- Ruling on applications. (1) The Di- rector, Center for Devices and Radio- logical Health, may approve or deny, in whole or in part, a requested variance or any amendment or extension there- of, and the director shall inform the applicant in writing of this action on a requested  variance  or  amendment  or

#### 21 CFR Ch. I (4–1–17 Edition)

extension. The  written  notice  will state the manner in which the variance differs from the standard, the effective date and the termination date of the variance, a  summary  of the require- ments and conditions attached to the variance, any other information  that may be relevant to the application or variance, and,  if  appropriate, the num- ber of units or other similar limita- tions for which the variance is ap- proved. Each variance will be assigned an identifying number.

- The Director, Center for Devices and Radiological Health, shall  amend or withdraw a variance whenever the Director determines that this action is necessary to protect the public health or otherwise is justified by this sub- chapter. Such action will become effec- tive on the date specified in the writ- ten notice of the action sent to the ap- plicant, except that it will become ef- fective immediately upon notification to the applicant when the Director de- termines that such action is necessary to prevent an imminent health hazard.

- All applications for variances and for amendments and extensions thereof and all correspondence (including writ- ten notices of approval) on these appli- cations will be available for public dis- closure in the office of the Division of Dockets Management, except for infor- mation regarded as confidential under section 537(e) of the act.

- Certification of equipment covered by variance. The manufacturer of any product for which a variance is granted shall modify the tag,  label, or other certification required by § 1010.2 to state:

  - That the product is in conformity with the applicable standard, except with respect to those characteristics covered by the variance;

  - That the product is in conformity with the provisions of the variance; and

  - The assigned number and effective date of the variance.

[39 FR 13879, Apr. 18, 1974, as amended at 44

FR  48191,  Aug.  17,  1979;  50  FR  7518,  Feb.  22,

1985;  50  FR  13565,  Apr.  5,  1985;  53  FR  11254,

Apr. 6, 1988; 53 FR 52683, Dec. 29, 1988; 59 FR

14365, Mar. 28, 1994; 65 FR 17137, Mar. 31, 2000;

73 FR 34861, June 19, 2008; 75 FR 16353, Apr. 1,

2010]

#### Food and Drug Administration, HHS	§ 1010.5

##### § 1010.5 Exemptions for products in- tended for United States Govern- ment use.

- Criteria for exemption. Upon appli- cation by a manufacturer (including assembler) or by a U.S. department or agency, the Director, Center for  De- vices and Radiological  Health,  Food and  Drug  Administration,  may  grant an exemption from any performance standard under subchapter J of this chapter for an electronic product, or class of products, otherwise subject to such standard when he determines that such electronic product or class is in- tended for use by departments or agen- cies of the United States and meets the criteria set forth in paragraph (a) (1) or

- of this section.

  - The procuring agency shall pre- scribe procurement specifications for the product or class of products gov- erning emissions of electronic product radiation, and the product or class shall be of a type used solely or pre- dominantly by a department or agency of the United States.

  - The product or class of products is intended for research, investigations, studies, demonstration, or training, or for reasons of national security.

- Consultation between the procuring agency and the Food and Drug Adminis- tration. The United States department or agency that intends to procure or manufacture a product or class of prod- ucts subject to electronic product radi- ation safety standards  contained  in this subchapter should consult with the Center for Devices and Radio- logical Health, Food and Drug Admin- istration, whenever it is anticipated that the specifications for the product or class must deviate from, or be in conflict with, such applicable stand- ards. Such consultation should  occur as early as possible during development of such specifications. The department or agency should include in the speci- fications all requirements of such standards that are not in conflict with, or are not inappropriate for, the special or unique uses for which the product is intended. The procuring agency should indicate to the Center for Devices and Radiological Health if it desires to be notified of the approval,  amendment, or withdrawal of the exemption.

- Application for  exemption.  If  you are submitting an application for ex- emption, or for  amendment  or  exten- sion thereof, you must submit an  origi- nal and two copies to the Division of Dockets Management (HFA–305), Food and Drug Administration, 5630 Fishers Lane, rm.  1061,  Rockville,  MD  20852. For an exemption under the criteria prescribed in paragraph (a)(1) of  this section,  the   application   shall   include the information prescribed in para- graphs (c)(1) through (c)(13) of this sec- tion. For an exemption under the cri- teria prescribed in paragraph (a)(2) of this section, the application shall in- clude the information prescribed in paragraphs (c)(3) through (c)(13) of this section.  An  application  for   exemption, or for amendment or extension thereof, and correspondence relating to such application shall be made available for public disclosure in the Division of Dockets Management, except for con- fidential or proprietary information submitted in accordance with part 20 of this chapter. Information classified for reasons  of  national  security  shall   not be  included  in  the   application.   Except as indicated in this paragraph, the ap- plication for  exemption  shall  include the   following:

  - The procurement specifications for the product or class of products that govern emissions of electronic product radiation.

  - Evidence that the product or class of products is of a type used solely or predominantly by departments or agencies of the United States.

  - Evidence that such product or class of products is intended for use by a department or agency of the United States.

  - A description of the product or class of products and its intended use.

  - An explanation of how compliance with the applicable standard would re- strict or be inappropriate for this in- tended use.

  - A description of the manner in which it is proposed that  the  product or class of products shall deviate from the requirements of the applicable standard.

  - An explanation of the  advantages to be derived from such deviation.

  - An explanation of how means of radiation  protection  will  be  provided

#### § 1010.5

where the product or class of products deviates from the requirements of the applicable  standard.

  - The period of time it is desired that the exemption be in effect, and, if appropriate, the number of units to be manufactured under the exemption.

  - The name, address, and telephone number of the manufacturer or his agent.

  - The name, address, and telephone number of the appropriate office of the United States department or agency purchasing the product or class of products.

  - Such other  information  required by regulation or by the  Director,  Cen- ter  for   Devices   and   Radiological Health, to evaluate and act on the ap- plication. Where such information in- cludes  nonclinical  laboratory   studies, the information shall include, with re- spect to   each  nonclinical  study,   either a statement that each study was con- ducted in compliance with the require- ments set forth in part 58 of  this  chap- ter, or, if the  study  was  not  conducted in compliance with such regulations, a statement that describes in detail all differences between the practices  used in the study and those required in the regulations. When such information in- cludes clinical investigations involving human subjects, the information shall include, with respect to each clinical investigation, either a statement that each investigation was conducted in compliance with the requirements set forth in part 56 of this chapter, or a statement that the investigation is not subject to such requirements in accord- ance with § 56.104 or § 56.105 and a state- ment that each investigation was con- ducted in compliance with the require- ments set forth in part 50 of  this  chap- ter.

  - With respect to each nonclinical laboratory study contained in the ap- plication, either a statement that the study was conducted in compliance with the requirements set forth in part 58 of this chapter, or, if the study was not conducted in compliance with such regulations, a brief statement of the reason for the noncompliance.

- Amendment or extension of an ex- emption. An exemption is granted on the basis of the information contained in the orginal applicaion. Therefore, if

#### 21 CFR Ch. I (4–1–17 Edition)

changes are needed in the radiation safety specifications for the product, or its use, or related radiation control procedures such that the  information in the original application would no longer be correct with respect to radi- ation safety, the applicant shall submit in advance of such changes a  request for an amendment to the  exemption. He also shall submit a request for ex- tension of the exemption, if needed, at least 60 days before the  expiration date. The application for  amendment or extension of an exemption shall in- clude the following information:

- The exemption  number  and  expi- ration  date.

- The amendment or extension re- quested and basis for  the  amendment or extension.

- If the radiation safety specifica- tions for the product or class  of prod- ucts or the product’s or  class of prod- ucts’ use or related radiation control procedures differ from the description provided in the original application, a description of such changes.

- Ruling on an application. (1) The Director, Center for Devices and Radio- logical Health, may grant an exemp- tion including in the written notice of exemption such conditions or terms as may be necessary to protect the public health and safety and shall notify the applicant in writing of his action. The conditions or terms of the exemption may include specifications concerning the manufacture, use, control, and dis- posal of the excess or surplus exempted product of class of products as provided in the Code of Federal Regulations, title 41, subtitle C. Each exemption will be assigned an identifying number.

(2) The Director, Center for Devices and Radiological Health, shall  amend or withdraw an exemption whenever he determines that such action is nec- essary to protect the public health or otherwise is justified by provisions of the act or this subchapter. Such action shall become effective on the date spec- ified in the written notice of the action sent to the applicant, except that it shall become effective immediately when the Director determines that it is necessary to prevent an imminent health hazard.

- Identification of equipment covered by exemption. The manufacturer of any

#### Food and Drug Administration, HHS	§ 1020.10

product for which an exemption is granted shall provide the following identification in the form of a tag or label permanently affixed or inscribed on such product so as to be legible and readily accessible to view when the product is fully assembled for use or in such other manner as may be pre- scribed in the exemption:

CAUTION

This electronic product has been exempted from Food and Drug Administration radi- ation safety performance standards pre- scribed in the Code of Federal Regulations, title 21, chapter I, subchapter J, pursuant to Exemption No. lll, granted on lllllll

[42 FR 44229, Sept. 2, 1977; 42 FR 61257, Dec. 2,

1977, as amended at 44 FR 17657, Mar. 23, 1979;

46 FR 8460, 8958, Jan. 27, 1981; 50 FR 7518, Feb.

22, 1985; 50 FR 13564, Apr. 5, 1985; 53 FR 11254,

Apr. 6, 1988; 59 FR 14365, Mar. 28, 1994; 65 FR

17138, Mar. 31, 2000]

## Subpart B—Alternate Test Procedures

##### § 1010.13   Special test procedures.

The Director, Center for Devices and Radiological Health, may, on the basis of a written application by a manufac- turer, authorize test programs other than those set forth in the standards under this subchapter for an electronic product if he determines that such products are not susceptible to satis- factory testing by the procedures set forth in the standard and that the al- ternative test procedures assure com- pliance with the standard.

[40 FR 32257, July 31, 1975, as amended at 53

FR 11254, Apr. 6, 1988]

## Subpart C—Exportation of Electronic Products

##### § 1010.20   Electronic products intended for export.

The performance standards pre- scribed in this subchapter shall not apply to any  electronic  product  which is intended solely for export if:

- Such product and the outside of any shipping container used in the ex- port of such product are labeled or tagged to show that such product is in- tended for export, and

- Such product meets all the appli- cable requirements of the country to which such product is intended for ex- port.

[40 FR 32257, July 31, 1975]

## PART	1020—PERFORMANCE STANDARDS FOR IONIZING RADI- ATION EMITTING PRODUCTS

Sec.

1020.10   Television receivers.

1020.20    Cold-cathode gas discharge tubes.

- Diagnostic  x-ray  systems  and  their major components.

- Radiographic  equipment.

- Fluoroscopic  equipment.

- Computed   tomography   (CT)   equip- ment.

- Cabinet x-ray systems.

AUTHORITY:   21   U.S.C.   351,   352,   360e–360j,

360hh–360ss, 371, 381.

SOURCE:  38  FR  28632,  Oct.  15,  1973,  unless

otherwise noted.

##### § 1020.10   Television receivers.

  - Applicability. The provisions of this section are applicable to television receivers manufactured subsequent to January 15, 1970.

  - Definitions. (1) External surface means the cabinet or enclosure pro- vided by the manufacturer as part of the receiver. If a cabinet or enclosure is not provided as part of the receiver, the external surface shall be considered to be a hypothetical cabinet, the plane surfaces of which are located at those minimum distances from the chassis sufficient to enclose all components of the receiver except that portion of the neck and socket of the cathode-ray tube which normally extends beyond the plane surfaces of the enclosure.

- Maximum test voltage means 130 root mean square  volts  if  the  receiver is designed to operate from nominal 110 to 120 root mean square volt power sources. If the receiver is designed to operate from a power source having some voltage other than from nominal 110 to 120 root mean square volts, max- imum test voltage means 110 percent of the nominal root mean square voltage specified by the manufacturer for the power source.

- Service controls means all of those controls on a television receiver pro- vided by the manufacturer for purposes

§ 1020.20

of  adjustment  which,  under  normal usage, are not accessible to the user.

- Television receiver means an elec- tronic product designed to receive and display a television picture through broadcast, cable, or closed circuit tele- vision.

- Usable picture means a picture in synchronization and transmitting viewable intelligence.

- User controls means all of those controls on a television receiver, pro- vided by the manufacturer for purposes of adjustment, which on a fully assem- bled receiver under normal usage, are accessible to the user.

  - Requirements—(1) Exposure rate limit. Radiation exposure rates pro- duced by a television receiver shall not exceed 0.5 milliroentgens per hour at a distance of five (5) centimeters from any point on the external surface  of the receiver, as measured in accord- ance with this section.

- Measurements. Compliance  with the exposure rate limit defined in para- graph (c)(1) of this section shall be de- termined by measurements  made with an instrument, the radiation sensitive volume of which shall have a cross sec- tion parallel to the external surface  of the receiver with an area of ten (10) square centimeters and no dimension larger than five (5)  centimeters. Meas- urements made with instruments hav- ing other areas must be corrected for spatial nonuniformity of the radiation field to obtain  the exposure  rate aver- age over a ten (10) square centimeter area.

- Test conditions.  All measurements shall be made with the receiver dis- playing a usable picture and with the power source operated at supply voltages up to the maximum test volt- age of the receiver and, as applicable, under the following specific conditions:

  - On television receivers manufac- tured subsequent to January 15, 1970, measurements shall be made with all user controls adjusted so as to produce maximum x-radiation emissions  from the receiver.

  - On television receivers manufac- tured subsequent to June 1, 1970, meas- urements shall be made with all user controls and all service controls ad- justed  to  combinations  which  result  in

#### 21 CFR Ch. I (4–1–17 Edition)

the  production  of  maximum  x-radi- ation emissions.

  - On television receivers manufac- tured subsequent to June 1, 1971, meas- urements shall be made under the con- ditions described in paragraph (c)(3) (ii) of this section, together with  condi- tions identical to those which  result from that component or circuit failure which maximizes x-radiation emis- sions.

- Critical component warning. The manufacturer shall permanently affix or inscribe a warning label, clearly leg- ible under conditions of service, on all television receivers which could produce radiation exposure rates in ex- cess of the requirements of this section as a result of failure or improper ad- justment or improper replacement of a circuit or shield component. The warn- ing label shall include the specification of operating high voltage and an in- struction for adjusting the high volt- age to the specified value.

##### § 1020.20   Cold-cathode   gas   discharge tubes.

- Applicability. The provisions of this section  are  applicable to  cold- cathode gas discharge tubes designed to demonstrate the effects of a flow of electrons or the production of x-radi- ation as specified herein.

- Definitions. Beam blocking device means a movable or removable portion of any enclosure around a cold-cathode gas discharge tube, which may be opened or closed to permit or prevent the emergence of an exit beam.

Cold-cathode gas discharge tube means an electronic device in which electron flow is produced and sustained by ion- ization of contained gas atoms and ion bombardment of the cathode.

Exit beam means that portion of the radiation which passes through the ap- erture resulting from the opening of the beam blocking device.

Exposure means the sum of the elec- trical charges on all of the ions of one sign produced in air when all electrons liberated by photons in a volume ele- ment of air are completely  stopped  in air divided by the mass of the air in the volume element. The special unit of ex- posure is  the roentgen.  One (1)  roent- gen equals 2.58 10¥4  coulombs/kilo- gram.

#### Food and Drug Administration, HHS	§ 1020.30

- Requirements—(1) Exposure rate limit. (i) Radiation exposure rates pro- duced by cold-cathode gas discharge tubes shall not exceed 10 mR./hr. at a distance of thirty (30) centimeters from any point on the external surface  of the tube, as measured in accordance with this section.

(ii) The divergence of the exit beam from tubes designed primarily to dem- onstrate the effects of x radiation, with the beam blocking device in the open position, shall not exceed (Pi) stera- dians.

(2) Measurements. (i) Compliance with the exposure rate limit defined in para- graph (c)(1)(i) of this section shall be determined by measurements averaged over an area of one hundred (100) square centimeters with no linear di- mension greater than twenty (20) centi- meters.

(ii) Measurements of exposure rates from tubes in enclosures from which the tubes cannot be removed without destroying the function  of  the  tube may  be  made  at  a  distance  of  thirty

- centimeters from any point on the external surface of the enclosure, pro- vided:

  - In the case of enclosures con- taining tubes designed primarily to demonstrate the production of x radi- ation, measurements shall be made with any beam blocking device in the beam blocking position, or

  - In the case of enclosures con- taining tubes designed primarily to demonstrate the effects of a flow of electrons, measurements shall be made with all movable or removable parts of such enclosure in the position which would maximize external exposure lev- els.

- Test conditions. (i) Measurements shall be made under the conditions of use specified in instructions provided by the manufacturer.

(ii) Measurements shall be made with the tube operated under forward and reverse  polarity.

- Instructions, labels, and warnings.

- Manufacturers shall provide, or cause to be provided, with each tube to which this section is applicable, appro- priate safety instructions,  together with instructions for the use of such tube, including the specification of a power source for use with the tube.

- Each enclosure or tube shall have inscribed on or permanently affixed to it, tags or labels, which identify the in- tended polarity of the terminals and:

(a) In the case of tubes designed pri- marily to demonstrate the heat effect, fluorescence effect,  or  magnetic  effect, a warning that application of power in excess of that specified may result  in the production of x-rays in excess of al- lowable limits; and (b) in the case of tubes designed primarily to dem- onstrate the  production  of  x-radiation, a warning that this device produces x- rays when energized.

- The tag or label required by this paragraph shall be located on the tube or enclosure so as to be readily visible and legible when the product is fully assembled for use.

##### § 1020.30   Diagnostic x-ray systems and their major components.

- Applicability. (1) The provisions of this section are applicable to:

- The following components of diag- nostic x-ray systems:

- Tube housing assemblies, x-ray controls, x-ray high-voltage  genera- tors, x-ray tables, cradles, film chang- ers, vertical cassette holders  mounted in a fixed location and cassette holders with front panels, and beam-limiting devices manufactured after August 1, 1974.

- Fluoroscopic imaging assemblies manufactured after August 1, 1974, and before April 26, 1977, or after June 10, 2006.

- Spot-film devices and image in- tensifiers manufactured after April 26, 1977.

- Cephalometric devices manufac- tured after February 25, 1978.

- Image receptor support devices for mammographic x-ray systems man- ufactured after September 5, 1978.

- Image receptors that are elec- trically powered or connected with the x-ray system manufactured on or after June 10, 2006.

- Fluoroscopic air kerma display devices manufactured on or after  June 10, 2006.

- Diagnostic x-ray systems, except computed tomography x-ray systems, incorporating one or more of such com- ponents; however, such x-ray systems shall be required to comply only with

#### § 1020.30

21 CFR Ch. I (4–1–17 Edition)

those   provisions   of   this   section   and

§§ 1020.31 and 1020.32, which relate to the components certified in accordance with paragraph (c) of this section and installed  into  the  systems.

- Computed tomography (CT)  x- ray systems manufactured before No- vember 29, 1984.

- CT gantries manufactured after September 3, 1985.

- The following provisions of this section and § 1020.33 are applicable  to CT x-ray systems manufactured or re- manufactured on or after November 29, 1984:

- Section 1020.30(a);

- Section 1020.30(b) ‘‘Technique fac- tors’’;

- Section 1020.30(b) ‘‘CT,’’ ‘‘Dose,’’ ‘‘Scan,’’	‘‘Scan	time,’’	and ‘‘Tomogram’’;

- Section 1020.30(h)(3)(vi) through (h)(3)(viii);

- Section 1020.30(n);

- Section 1020.33(a) and (b);

- Section 1020.33(c)(1) as it affects

§ 1020.33(c)(2); and

- Section 1020.33(c)(2).

- The provisions of this section and

§ 1020.33 in its entirety, including those provisions in paragraph (a)(2) of this section, are applicable to CT x-ray sys- tems manufactured  or  remanufactured on or after September 3, 1985. The date of manufacture of the CT system is the date  of  manufacture  of  the  CT  gantry.

- Definitions. As used in this section and §§ 1020.31, 1020.32, and 1020.33, the following   definitions   apply:

Accessible surface means the external surface of the enclosure or housing pro- vided by the manufacturer.

Accessory component means:

  - A component used with diagnostic x-ray systems, such as a cradle or film changer, that is not necessary for the compliance of the system with applica- ble provisions of this subchapter but which requires an initial determination of compatibility with the system; or

  - A component necessary for com- pliance of the system with applicable provisions of this subchapter but which may be interchanged with similar com- patible components without affecting the system’s compliance, such as one of a set of interchangeable beam-limiting devices; or

  - A component compatible with all x-ray systems with which it may be used and that does not require compat- ibility or installation  instructions, such as a tabletop cassette holder.

Air kerma means kerma in air (see definition of Kerma).

Air kerma rate (AKR) means  the  air kerma per unit time.

Aluminum equivalent means the thick- ness of aluminum (type 1100 alloy) 1 af- fording the same attenuation, under specified conditions, as the material in question.

Articulated joint means a joint be- tween two separate sections of a table- top which joint provides the capacity for one of the sections to pivot on the line segment along which the sections join.

Assembler means any person engaged in the business of assembling, replac- ing, or installing one or more compo- nents into a diagnostic x-ray system or subsystem. The term includes the owner of an x-ray system or his or her employee or agent who assembles com- ponents into an x-ray system that is subsequently used to provide profes- sional or commercial services.

Attenuation block means a block  or stack of type 1100 aluminum alloy, or aluminum alloy having equivalent at- tenuation, with dimensions 20 centi- meters (cm) or larger by 20 cm or larg- er by 3.8 cm, that is large enough to intercept  the  entire  x-ray  beam.

Automatic exposure control (AEC) means a device which automatically controls one or more technique factors in order to obtain at a preselected loca- tion(s) a required quantity of radi- ation.

Automatic exposure rate control (AERC) means a device which automatically controls one or more technique factors in order to obtain at a preselected loca- tion(s) a required quantity of radiation per unit time.

Beam axis means a line from the source through the centers of the x-ray fields.

1 The  nominal  chemical  composition  of type 1100 aluminum alloy is 99.00 percent minimum aluminum, 0.12 percent copper, as given in ‘‘Aluminum Standards and Data’’ (1969). Copies may be obtained from The Alu- minum  Association,  New  York,  NY.

#### Food and Drug Administration, HHS	§ 1020.30

Beam-limiting device means a device which provides a means to restrict the dimensions of the x-ray field.

C-arm       fluoroscope       means	a fluoroscopic x-ray system in which the image receptor and the x-ray tube housing assembly are connected or co- ordinated to maintain a spatial rela- tionship. Such a system  allows  a change in the direction of the beam axis with respect to the patient  with- out moving the patient.

Cantilevered tabletop means a tabletop designed   such   that   the   unsupported

portion can be extended at least 100 cm beyond the support.

Cassette holder means a device, other than  a  spot-film  device,  that  supports and/or  fixes  the  position  of  an  x-ray film cassette during an x-ray exposure. Cephalometric  device  means  a  device intended for the radiographic visualiza- tion  and  measurement  of  the  dimen-

sions of the human head.

Coefficient of variation means  the ratio of the standard deviation to the mean value of a population of observa- tions. It is estimated using the fol- lowing equation:

where:

s = Estimated standard deviation of the pop- ulation.

X¯ = Mean value of observations in sample.

Xi = ith observation sampled.

n = Number of observations sampled.

Computed tomography (CT) means the production of a tomogram by the ac- quisition and computer processing of x- ray transmission data.

Control panel means that part of the x-ray control upon which are mounted the switches, knobs, pushbuttons, and other hardware necessary for manually setting the technique factors.

Cooling curve means the graphical re- lationship between heat units stored and cooling time.

Cradle means:

- A removable device which sup- ports and may restrain a patient above an x-ray table; or

- A device;

  - Whose patient support structure is interposed between the patient and the image receptor during normal use;

  - Which is equipped with means for patient restraint; and

  - Which is capable of rotation about its long (longitudinal) axis.

CT gantry means tube housing assem- blies, beam-limiting devices, detectors,

and the supporting structures, frames, and covers which hold and/or enclose these components.

Cumulative air kerma means the total air kerma accrued from the beginning of an examination or procedure and in- cludes all contributions from fluoroscopic and radiographic irradia- tion.

Diagnostic source assembly means the tube housing assembly with a beam- limiting  device  attached.

Diagnostic x-ray system means an x- ray system designed for irradiation of any part of the human body for the purpose of diagnosis or visualization.

Dose means the absorbed dose as de- fined by the International Commission on Radiation Units and Measurements. The absorbed dose, D, is the quotient of de by dm, where de is the mean energy imparted to matter of mass dm; thus D

= de/dm, in units of J/kg, where the special name for the unit of absorbed dose is gray (Gy).

Equipment means x-ray equipment.

Exposure (X) means the quotient  of dQ by dm where dQ is the absolute value of the total charge of the ions of one sign produced in air when all the electrons and positrons liberated or created  by  photons  in  air  of  mass  dm

#### § 1020.30

are completely stopped in air; thus X = dQ/dm, in units of C/kg. A second meaning of exposure is the process or condition during which the x-ray tube produces x-ray radiation.

Field emission equipment means equip- ment which uses an x-ray tube  in which electron emission from the cath- ode is due solely to action of  an  elec- tric  field.

Fluoroscopic air kerma display device means a device, subsystem, or compo- nent that provides the display of AKR and  cumulative  air  kerma  required  by

§ 1020.32(k). It includes radiation detec- tors, if any, electronic and computer components, associated software, and data  displays.

Fluoroscopic imaging assembly means a subsystem in which x-ray photons produce a set of fluoroscopic images or radiographic images recorded from the fluoroscopic image receptor. It in- cludes the image receptor(s), electrical interlocks, if any, and structural mate- rial providing linkage between  the image receptor and diagnostic source assembly.

Fluoroscopic irradiation time  means the cumulative duration during an ex- amination or procedure of operator-ap- plied continuous pressure to the device, enabling x-ray tube activation in any fluoroscopic mode of operation.

Fluoroscopy  means  a  technique  for generating x-ray images and presenting them simultaneously and continuously as  visible  images.  This  term  has  the same   meaning   as   the   term   ‘‘radi- oscopy’’ in the standards of the Inter- national Electrotechnical Commission.

General purpose radiographic x-ray sys- tem means any radiographic x-ray sys- tem which, by design, is not limited to radiographic examination of specific anatomical  regions.

Half-value layer (HVL) means the thickness of specified material which attenuates the beam of radiation to an extent such that the AKR is reduced to one-half of its original value. In this definition the contribution of all scat- tered radiation, other than any which might be present initially in the beam concerned, is deemed to be excluded.

Image intensifier means a device, in- stalled in its housing, which instanta- neously converts an x-ray pattern into

#### 21 CFR Ch. I (4–1–17 Edition)

a  corresponding  light  image  of  higher energy  density.

Image receptor means any device, such as a fluorescent screen, radiographic film, x-ray image intensifier tube, solid-state detector, or gaseous detec- tor, which transforms incident x-ray photons either into a visible image or into another form which can be made into a visible image by further trans- formations. In those cases where means are provided to preselect a portion  of the image receptor, the term ‘‘image receptor’’ shall mean the preselected portion of the device.

Image receptor support device means, for mammography x-ray systems, that part of the system designed to support the image receptor during a mammo- graphic examination and to provide a primary protective barrier.

Isocenter means the center of the smallest sphere through  which  the beam axis passes when the equipment moves through a full range of rotations about its common center.

Kerma means the quantity as defined by   the   International   Commission   on Radiation    Units    and    Measurements. The kerma, K, is the quotient of dEtr  by dm, where dEtr  is the sum of the initial kinetic energies of all the charged par- ticles  liberated  by  uncharged  particles in a mass dm of material; thus K = dEtr/ dm, in units of J/kg, where the special name  for  the  unit  of  kerma  is  gray (Gy).   When   the   material   is   air,   the quantity  is  referred  to  as  ‘‘air  kerma.’’ Last-image-hold		(LIH)	radiograph means an image obtained either by re- taining  one  or  more  fluoroscopic  im- ages,  which  may  be  temporally  inte- grated, at the end of a fluoroscopic ex- posure  or  by  initiating  a  separate  and distinct   radiographic   exposure   auto- matically  and  immediately  in  conjunc- tion	with	termination	of	the

fluoroscopic exposure.

Lateral fluoroscope means the x-ray tube and image receptor combination in a biplane system dedicated to the lateral projection. It consists of the lateral x-ray tube housing assembly and the lateral image receptor that are fixed in  position  relative  to the  table with the x-ray beam axis parallel  to the plane of the table.

#### Food and Drug Administration, HHS	§ 1020.30

Leakage radiation means radiation emanating from the diagnostic source assembly except for:

- The useful beam; and

- Radiation  produced  when  the  ex- posure switch or timer is not activated. Leakage  technique  factors  means  the technique  factors  associated  with  the diagnostic  source  assembly  which  are used   in   measuring   leakage   radiation.

They are defined as follows:

- For diagnostic source assemblies intended for capacitor energy storage equipment, the maximum-rated peak tube potential  and  the maximum-rated number of exposures in an hour for op- eration at the maximum-rated peak tube potential with the quantity of charge per exposure being 10 millicoulombs (or 10 mAs) or the min- imum obtainable from the unit, which- ever is larger;

- For diagnostic source assemblies intended for field emission equipment rated for pulsed operation, the max- imum-rated peak tube potential and the maximum-rated number of x-ray pulses in an hour for operation at the maximum-rated peak tube potential; and

- For all other diagnostic source as- semblies, the maximum-rated peak tube potential and the maximum-rated continuous tube current for the max- imum-rated peak tube potential.

Light field means that area of the intersection of the light beam from the beam-limiting device and one of the set of planes parallel to and including the plane of the image receptor, whose pe- rimeter is the locus of points at which the illuminance is one-fourth of the maximum in the intersection.

Line-voltage regulation means the dif- ference between the no-load and the load line potentials expressed as a per- cent  of  the  load  line  potential;  that  is,

Percent	line-voltage	regulation	= 100(Vn ¥ Vi)/Vi

where:

Vn = No-load line potential and Vi = Load line potential.

Maximum line current means the root mean square current in the supply line of an x-ray machine operating at its maximum  rating.

Mode of operation means, for fluoroscopic systems, a distinct meth-

od of fluoroscopy or radiography pro- vided by the manufacturer and selected with a  set of  several technique  factors or other control settings uniquely asso- ciated with the mode. The set of dis- tinct technique factors and control set- tings for the mode may be selected by the operation of a single control. Ex- amples of distinct modes of operation include normal fluoroscopy (analog or digital), high-level control fluoroscopy, cineradiography (analog or digital), digital subtraction angiography, elec- tronic radiography using the fluoroscopic image receptor, and photospot recording. In a specific mode of operation, certain system variables affecting air kerma, AKR, or image quality, such as image magnification, x-ray field size, pulse rate, pulse dura- tion, number of pulses, source-image receptor distance (SID), or optical ap- erture, may be adjustable or may vary; their variation per se does not com- prise a mode of operation  different from the one that has been selected.

Movable tabletop means a tabletop which, when assembled for use, is capa- ble of movement with respect to its supporting structure within the  plane of the tabletop.

Non-image-intensified	fluoroscopy means fluoroscopy using only a fluores- cent screen.

Peak tube potential means the max- imum value of the potential difference across the x-ray tube during an expo- sure.

Primary protective barrier means the material, excluding filters, placed in the useful beam to reduce the radiation exposure for protection purposes.

Pulsed mode means operation of the x- ray system such that the x-ray tube current is pulsed by  the  x-ray  control to produce one or more exposure inter- vals of duration less than one-half sec- ond.

Quick change x-ray tube means an x- ray tube designed for use in its associ- ated tube housing such that:

- The tube cannot be inserted in its housing in a manner that would result in noncompliance of the system with the requirements of paragraphs (k) and

- of this section;

#### § 1020.30

- The focal spot position will not cause noncompliance with the provi- sions of this section or § 1020.31 or 1020.32;

- The shielding within the tube housing cannot be displaced; and

- Any removal and subsequent re- placement of a beam-limiting device during reloading of the tube in the tube housing will not result in noncompli- ance of the x-ray system with the ap- plicable field limitation and alignment requirements of §§ 1020.31 and 1020.32.

Radiation therapy simulation system means a radiographic or fluoroscopic x- ray system intended for localizing the volume to be exposed during radiation therapy and confirming  the  position and size of the therapeutic irradiation field.

Radiography means a technique for generating and recording an x-ray pat- tern for the purpose of providing the user with an image(s) after termi- nation of the exposure.

Rated line voltage means the range of potentials, in volts, of the supply line specified by the manufacturer at which the x-ray machine is designed to oper- ate.

Rated output current means the max- imum allowable load current of the x- ray high-voltage generator.

Rated output voltage means the allow- able peak potential, in volts, at the output terminals of the x-ray high- voltage generator.

Rating means the operating limits specified by the manufacturer.

Recording means producing a retriev- able form of an image resulting from x- ray photons.

Scan means the complete process of collecting x-ray transmission data for the production of a tomogram. Data may be collected simultaneously dur- ing a single scan for the production of one or more tomograms.

Scan time means the period of time between the beginning and end of x-ray transmission data accumulation for a single  scan.

Solid state x-ray imaging device means an assembly, typically in a rectangular panel configuration, that intercepts x- ray photons and converts the photon energy into a modulated electronic sig- nal representative of the x-ray inten- sity  over  the  area  of  the  imaging  de-

#### 21 CFR Ch. I (4–1–17 Edition)

vice. The electronic signal is then used to create an image for display and/or storage.

Source means the focal spot of the x- ray tube.

Source-image receptor distance (SID) means the distance from the source to the center of the input surface of the image receptor.

Source-skin distance (SSD) means the distance from the source to the center of the entrant x-ray field in the plane tangent to the patient skin surface.

Spot-film device means a device in- tended to transport and/or position a radiographic image receptor  between the x-ray source and  fluoroscopic image receptor. It includes a device in- tended to hold a cassette  over  the input end of the fluoroscopic image re- ceptor for the purpose of producing a radiograph.

Stationary tabletop means a tabletop which, when assembled for use, is in- capable of movement with respect  to its supporting structure  within  the plane of the tabletop.

Technique factors means the following conditions of operation:

  - For capacitor energy storage equipment, peak tube potential in kilo- volts (kV) and quantity of charge in milliampere-seconds (mAs);

  - For field emission equipment rated for pulsed operation, peak tube potential in kV and number of x-ray pulses;

  - For CT equipment designed for pulsed operation, peak tube potential in kV, scan time in seconds, and either tube current in milliamperes (mA), x- ray pulse width in seconds, and the number of x-ray pulses per scan, or the product of the tube current, x-ray pulse width, and the number of x-ray pulses in mAs;

  - For CT equipment not designed for pulsed operation, peak tube poten- tial in kV, and either tube current in mA and scan time in seconds, or the product of tube current and exposure time in mAs and the scan time when the scan time and  exposure time are equivalent; and

  - For all other equipment, peak tube potential in kV, and either tube current in mA and exposure time in seconds, or the product of tube current and exposure time in mAs.

#### Food and Drug Administration, HHS	§ 1020.30

Tomogram means the depiction of the x-ray attenuation properties of a sec- tion through a body.

Tube means an x-ray tube, unless otherwise specified.

Tube housing assembly means the tube housing with tube installed. It includes high-voltage and/or filament trans- formers and other appropriate ele- ments when they are contained within the tube housing.

Tube rating chart means the set of curves which specify the rated limits of operation of the tube in terms of the technique  factors.

Useful beam means the radiation which passes through the tube housing port and the aperture of the beam-lim- iting device when the  exposure  switch or timer is activated.

Variable-aperture beam-limiting device means a beam-limiting  device  which has the capacity for stepless  adjust- ment of the x-ray field size at a  given SID.

Visible area means the portion of the input surface of the  image  receptor over which incident x-ray photons are producing a visible image.

X-ray control means a device which controls input power to the x-ray high- voltage generator and/or the x-ray tube. It includes equipment such as timers, phototimers, automatic bright- ness stabilizers, and similar devices, which control the technique factors of an x-ray exposure.

X-ray equipment means an x-ray sys- tem, subsystem, or component thereof. Types of x-ray equipment are as fol- lows:

- Mobile x-ray equipment means x- ray equipment mounted on a perma- nent base with wheels and/or casters for moving while completely assem- bled;

- Portable x-ray equipment means x- ray equipment designed to be hand-car- ried; and

- Stationary x-ray equipment means x- ray equipment which is installed in a fixed location.

X-ray field means that area of the intersection of the useful beam and any one of the set of planes parallel to and including the plane of the image recep- tor, whose perimeter is the locus of points at which the AKR is one-fourth of the maximum in the intersection.

X-ray high-voltage generator means a device which transforms electrical en- ergy from the potential supplied by the x-ray control to the tube operating po- tential. The device may also include means for transforming alternating current to direct current, filament transformers for the x-ray tube(s), high-voltage switches, electrical pro- tective devices, and other appropriate elements.

X-ray subsystem means any combina- tion of two or more components of an x-ray system for which there are re- quirements  specified  in   this   section and §§ 1020.31 and 1020.32.

X-ray system means an assemblage of components for the controlled produc- tion of x-rays. It includes minimally an x-ray high-voltage generator, an x-ray control, a tube housing assembly, a beam-limiting device, and the nec- essary supporting structures. Addi- tional components which function with the system are considered integral parts of the system.

X-ray table means a patient support device with its patient support struc- ture (tabletop) interposed between the patient and the image receptor during radiography and/or fluoroscopy. This includes, but is not limited to, any stretcher equipped with a radiolucent panel and  any  table  equipped  with  a cassette tray (or bucky), cassette tun- nel, fluoroscopic image receptor, or spot-film device beneath the tabletop.

X-ray tube means any electron tube which is designed for the conversion of electrical energy into x-ray energy.

- Manufacturers’ responsibility. Man- ufacturers     of     products     subject     to

§§ 1020.30 through 1020.33 shall certify that each of their products meet all ap- plicable requirements  when  installed into a diagnostic x-ray system accord- ing  to  instructions.  This  certification shall be made under the format speci- fied in § 1010.2 of this chapter. Manufac- turers may  certify  a  combination  of two or more components if they obtain prior authorization in writing from the Director of the Office  of  Communica- tion, Education, and Radiation  Pro- grams of the Center for Devices and Radiological    Health.    Manufacturers shall not be held responsible for non- compliance   of   their   products   if   that

#### § 1020.30

noncompliance is due solely to the im- proper installation or assembly of that product by another person; however, manufacturers are responsible for pro- viding assembly instructions adequate to assure compliance of their compo- nents with the applicable provisions of

§§ 1020.30 through 1020.33.

- Assemblers’ responsibility. An as- sembler who installs one or more com- ponents certified as required by para- graph (c) of this section shall install certified components that  are  of  the type required by § 1020.31, § 1020.32, or

§ 1020.33 and shall assemble, install, ad- just, and test the certified components according to the instructions of their respective manufacturers. Assemblers shall not  be  liable  for  noncompliance  of a certified  component if  the  assembly of  that  component  was  according  to the component manufacturer’s instruc- tion.

- Reports of assembly. All assemblers who install certified components shall file a report of assembly, except as specified in paragraph (d)(2) of this sec- tion. The report will be construed  as the assembler’s certification and iden- tification under §§ 1010.2 and 1010.3 of this chapter. The assembler shall af- firm in the report that the manufactur- er’s instructions were followed in the assembly or that the certified compo- nents as assembled into the  system meet   all   applicable   requirements   of

§§ 1020.30 through 1020.33. All assembler reports must be on a  form  prescribed by the Director, CDRH. Completed  re- ports must be submitted to the Direc- tor, the purchaser, and, where applica- ble, to the State agency responsible for radiation protection within 15 days fol- lowing  completion  of  the  assembly.

- Exceptions to reporting require- ments. Reports of assembly need not be submitted for any of the following:

  - Reloaded or replacement tube housing assemblies that are reinstalled in or newly assembled into an existing x-ray  system;

  - Certified accessory components that have been identified as such to CDRH  in  the  report  required  under

§ 1002.10 of this chapter;

  - Repaired components,  whether or not removed from the system and reinstalled during the course of repair,

#### 21 CFR Ch. I (4–1–17 Edition)

provided the original installation into the system was reported; or

(iv)(A) Components installed tempo- rarily in an x-ray system in place of components removed temporarily for repair, provided the temporarily in- stalled component is identified by a tag or label bearing the following in- formation:

Temporarily Installed Component

This certified component has been assem- bled, installed, adjusted, and tested by me according to the instructions provided by the manufacturer.

Signature Company Name

Street Address, P.O. Box City, State, Zip Code Date of Installation

(B) The replacement of the tempo- rarily installed component by a compo- nent other than the component origi- nally removed for repair shall be re- ported as specified in paragraph (d)(1) of this section.

- Identification of x-ray components. In addition to the identification re- quirements specified in § 1010.3 of this chapter, manufacturers of components subject to this section and §§ 1020.31, 1020.32, and 1020.33, except high-voltage generators contained within tube housings and  beam-limiting  devices that are integral  parts  of  tube housings, shall permanently inscribe or affix thereon the model number and se- rial number of the product so that they are legible and accessible to view. The word ‘‘model’’ or ‘‘type’’  shall  appear as part of the manufacturer’s required identification of certified x-ray compo- nents. Where the certification of a sys- tem or subsystem, consisting of two or more components, has been authorized under paragraph (c) of this section, a single inscription,  tag,  or  label  bearing the model number and serial number may be used to identify the product.

  - Tube housing assemblies. In a simi- lar manner, manufacturers of tube housing assemblies shall also inscribe or affix thereon the name of the manu- facturer, model number, and serial number of the x-ray tube which the tube housing assembly incorporates.

  - Replacement of tubes. Except as specified in paragraph (e)(3) of this sec- tion, the replacement of  an  x-ray  tube in    a    previously    manufactured    tube

#### Food and Drug Administration, HHS	§ 1020.30

housing assembly certified under para- graph (c) of this section constitutes manufacture of a new tube housing as- sembly, and the manufacturer is sub- ject to the provisions of paragraph (e)(1) of this section. The manufacturer shall remove, cover, or deface any pre- viously affixed inscriptions, tags, or la- bels that are no longer applicable.

  - Quick-change x-ray tubes. The re- quirements of paragraph (e)(2) of this section shall not apply to tube housing assemblies designed and designated by their original manufacturer to contain quick change x-ray tubes. The manu- facturer of quick-change x-ray tubes shall include with each replacement tube a label with the tube manufactur- er’s name, the model, and serial num- ber of the x-ray tube. The manufac- turer of the tube shall instruct the as- sembler who installs the new tube to attach the label to the tube housing as- sembly and to remove, cover, or deface the previously affixed  inscriptions, tags, or labels that are described by the tube manufacturer as no longer appli- cable.

- [Reserved]

- Information to be provided to assem- blers. Manufacturers of components listed in paragraph (a)(1) of this section shall provide to assemblers subject to paragraph (d) of this section and, upon request, to others at a cost not to ex- ceed the cost of publication and dis- tribution, instructions for assembly, installation, adjustment, and testing of such components adequate to assure that the products will comply with ap- plicable provisions of this section and

§§ 1020.31, 1020.32, and 1020.33, when as- sembled, installed, adjusted, and  tested as directed. Such instructions shall in- clude specifications of  other  compo- nents compatible with that to be in- stalled when compliance of  the  system or subsystem depends on their compat- ibility. Such specifications may  de- scribe  pertinent  physical   characteris- tics of the components and/or  may  list by manufacturer model number the components which are compatible. For x-ray controls and generators manufac- tured after May 3, 1994, manufacturers shall  provide:

  - A statement of the rated line volt- age and the range of line-voltage regu-

lation for operation at maximum line current;

  - A statement of the maximum line current of the x-ray system based on the maximum input voltage and cur- rent characteristics of the tube hous- ing assembly compatible with rated output voltage and rated output cur- rent characteristics of the x-ray con- trol and associated high-voltage gener- ator. If the rated input voltage and current characteristics of the tube housing assembly are not known by the manufacturer of the x-ray control and associated high-voltage generator, the manufacturer shall provide informa- tion necessary to allow the assembler to determine the maximum line cur- rent for the particular tube housing as- sembly(ies);

  - A statement of the technique fac- tors that constitute the maximum line current condition described in para- graph (g)(2) of this section.

- Information to be provided to users. Manufacturers of x-ray  equipment shall provide to purchasers and, upon request, to others at a cost not to ex- ceed the cost of publication and dis- tribution, manuals or instruction sheets which shall include the fol- lowing technical and safety informa- tion:

  - All x-ray equipment. For x-ray equipment  to  which  this  section  and

§§ 1020.31, 1020.32, and 1020.33 are appli- cable,  there  shall  be  provided:

- Adequate instructions concerning any radiological safety procedures and precautions which may be necessary because of unique features of the equip- ment; and

- A schedule of the maintenance necessary to keep the equipment in compliance    with    this    section    and

§§ 1020.31, 1020.32, and 1020.33.

  - Tube housing assemblies. For each tube housing assembly, there shall be provided:

    - Statements of the leakage tech- nique factors for all combinations of tube housing assemblies and beam-lim- iting devices for which the tube hous- ing assembly manufacturer states com- patibility, the minimum filtration per- manently in the useful beam expressed as millimeters (mm) of aluminum equivalent, and the peak tube potential

#### § 1020.30

at which the aluminum equivalent was obtained;

    - Cooling curves for the anode and tube housing; and

    - Tube rating charts. If the tube is designed to operate  from  different types of x-ray high-voltage generators (such as single-phase self rectified, sin- gle-phase half-wave rectified, single- phase full-wave rectified, 3-phase 6- pulse, 3-phase 12-pulse, constant poten- tial, capacitor energy storage) or under modes of operation such as alternate focal spot sizes or speeds of anode rota- tion which affect its rating, specific identification of the difference in rat- ings shall be noted.

  - X-ray controls and generators. For the x-ray control and associated x-ray high-voltage generator, there shall be provided:

    - A statement of the rated line volt- age and the range of line-voltage regu- lation for operation at maximum line current;

    - A statement of the maximum line current of the x-ray system based on the maximum input voltage and output current characteristics of the tube housing assembly  compatible  with rated output voltage and rated current characteristics of the x-ray control and associated high-voltage generator. If the rated input voltage and current characteristics of the tube housing as- sembly are not known by the manufac- turer of the x-ray control and associ- ated high-voltage generator, the manu- facturer shall provide necessary infor- mation to allow the purchaser to deter- mine the maximum line current for his particular tube housing assembly(ies);

    - A statement of the technique factors that constitute the maximum line current condition described in paragraph (h)(3)(ii) of this section;

    - In the case of battery-powered generators, a specification of the min- imum state of charge necessary for proper operation;

    - Generator rating and duty cycle;

    - A statement of the maximum de- viation from the preindication given by labeled technique factor control set- tings or indicators during any radio- graphic or CT exposure where the equipment is connected to a power sup- ply as described in accordance with this  paragraph.  In  the  case  of  fixed

#### 21 CFR Ch. I (4–1–17 Edition)

technique factors, the maximum devi- ation from the nominal fixed value of each factor shall be stated;

    - A statement of the maximum deviation from the continuous indica- tion of x-ray tube potential and cur- rent during any fluoroscopic exposure when the equipment is connected to a power supply as described in accord- ance with this paragraph; and

    - A statement describing the measurement criteria for all technique factors used in paragraphs (h)(3)(iii), (h)(3)(vi), and (h)(3)(vii) of this section; for example, the beginning and endpoints of exposure time measured with respect to a certain percentage of the voltage waveform.

  - Beam-limiting device. For each variable-aperture beam-limiting de- vice, there shall be provided;

- Leakage technique factors for all combinations of tube housing assem- blies and beam-limiting devices for which the beam-limiting device manu- facturer states compatibility; and

- A statement including the min- imum aluminum equivalent of  that part of the device through which the useful beam passes and including the x- ray tube potential at which the alu- minum equivalent was obtained. When two or more filters are provided as part of the device, the statement shall in- clude the aluminum equivalent of each filter.

  - Imaging system information. For x- ray systems manufactured on or after June 10, 2006, that produce images using the fluoroscopic image receptor, the following information shall be pro- vided in a separate, single section  of the user’s instruction manual or in a separate manual devoted to this infor- mation:

    - For each mode of operation, a de- scription of the mode and detailed in- structions on how the mode is engaged and disengaged. The description of the mode shall identify those technique factors and system controls that are fixed or automatically adjusted by se- lection of the mode of operation, in- cluding the manner in which the auto- matic adjustment is controlled. This information shall include how the oper- ator can recognize which mode of oper- ation has been selected prior to initi- ation of x-ray production.

#### Food and Drug Administration, HHS	§ 1020.30

    - For each mode of operation, a de- scriptive example(s) of any specific clinical procedure(s) or imaging task(s) for which the mode is recommended or designed and how each mode should be used. Such recommendations do not preclude other clinical uses.

  - Displays of values of AKR and cu- mulative air kerma. For fluoroscopic x- ray systems manufactured on or after June 10, 2006, the following shall be provided:

    - A schedule of maintenance for any system  instrumentation  associated with the display of air kerma informa- tion necessary to maintain the displays of AKR and  cumulative  air  kerma within the limits of allowed uncer- tainty specified by § 1020.32(k)(6) and, if the capability for user calibration of the display is provided, adequate in- structions for such calibration;

    - Identification of the distances along the beam axis:

      - From the focal spot to the isocenter, and

      - From the focal spot to the ref- erence location to which displayed val- ues of AKR and cumulative air kerma refer according to § 1020.32(k)(4);

    - A rationale for specification of a reference irradiation location alter- native to 15 cm from the isocenter to- ward the x-ray source along the beam axis when such alternative  specifica- tion        is        made        according        to

§  1020.32(k)(4)(ii).

- [Reserved]

- Warning label. The control panel containing the main power switch shall bear the warning  statement,  legible and accessible to view:

‘‘Warning: This x-ray unit may be dan- gerous to patient and operator unless safe exposure factors, operating instructions and maintenance schedules are observed.’’

- Leakage radiation from the diag- nostic source assembly. The leakage radi- ation from the diagnostic source as- sembly measured at a distance of  1 meter in any direction from the source shall  not  exceed  0.88  milligray   (mGy) air kerma (vice 100 milliroentgen (mR) exposure) in 1 hour when the x-ray tube is operated at the leakage technique factors. If the maximum  rated  peak tube potential of the tube housing as- sembly is greater than the  maximum rated  peak  tube  potential  for  the  diag-

nostic source assembly, positive means shall be provided to limit  the  max- imum x-ray tube potential to that  of the diagnostic source assembly. Com- pliance shall be determined by meas- urements averaged over an area of 100 square cm with no linear dimension greater than 20 cm.

- Radiation from components other than the diagnostic source assembly. The radiation emitted by  a  component other than the diagnostic source as- sembly shall not exceed  an  air  kerma of 18 microGy (vice 2 mR exposure) in 1 hour at 5 cm from any accessible sur- face of the component when it is oper- ated in an assembled x-ray system under any conditions for which it was designed. Compliance shall be deter- mined by measurements averaged over an area of 100 square cm with no linear dimension greater than 20 cm.

- Beam  quality—(1)  Half-value  layer (HVL). The  HVL  of  the  useful  beam  for a  given  x-ray  tube  potential  shall  not be    less    than    the    appropriate    value shown in table 1 in paragraph (m)(1) of this  section  under  the  heading  ‘‘Speci- fied  Dental  Systems,’’  for  any  dental  x- ray    system    designed    for    use    with intraoral   image   receptors   and   manu- factured  after  December  1,  1980;  under the     heading     ‘‘I—Other     X-Ray     Sys- tems,’’  for  any  dental  x-ray  system  de- signed  for  use  with  intraoral  image  re- ceptors   and   manufactured   before   De- cember 1, 1980, and all other x-ray sys- tems  subject  to  this  section  and  manu- factured before June 10, 2006; and under the     heading     ‘‘II—Other     X-Ray     Sys- tems,’’    for    all    x-ray    systems,    except dental  x-ray  systems  designed  for  use with  intraoral  image  receptors,  subject to  this  section  and  manufactured  on  or after June 10, 2006. If it is necessary to determine  such  HVL  at  an  x-ray  tube potential  which  is  not  listed  in  table  1 in  paragraph  (m)(1)  of  this  section,  lin- ear   interpolation   or   extrapolation   may be  made.  Positive  means 2   shall  be  pro- vided  to  ensure  that  at  least  the  min- imum  filtration  needed  to  achieve  the

2 In the case of a system, which is to be op- erated with more than one thickness of fil- tration, this requirement can be met by a fil- ter interlocked with the kilovoltage selector which will prevent x-ray emissions if the minimum required filtration is not in place.

#### § 1020.30

above beam quality requirements is in the useful beam during each exposure. Table 1 follows:

#### 21 CFR Ch. I (4–1–17 Edition)

the useful beam during each exposure. Table 1 follows:

TABLE 1

| X-Ray Tube Voltage (kilovolt peak) | Minimum HVL (mm of aluminum) |
| --- | --- |
| Designed Operating Range | Measured Op- erating Poten- tial | Specified Dental Sys- tems 1 | I—Other X- Ray Sys- tems 2 | II—Other X- Ray Sys- tems 3 |
| Below 51 | 30 | 1.5 | 0.3 | 0.3 |
|  | 40 | 1.5 | 0.4 | 0.4 |
|  | 50 | 1.5 | 0.5 | 0.5 |
| 51 to 70 | 51 | 1.5 | 1.2 | 1.3 |
|  | 60 | 1.5 | 1.3 | 1.5 |
|  | 70 | 1.5 | 1.5 | 1.8 |
| Above 70 | 71 | 2.1 | 2.1 | 2.5 |
|  | 80 | 2.3 | 2.3 | 2.9 |
|  | 90 | 2.5 | 2.5 | 3.2 |
|  | 100 | 2.7 | 2.7 | 3.6 |
|  | 110 | 3.0 | 3.0 | 3.9 |
|  | 120 | 3.2 | 3.2 | 4.3 |
|  | 130 | 3.5 | 3.5 | 4.7 |
|  | 140 | 3.8 | 3.8 | 5.0 |
|  | 150 | 4.1 | 4.1 | 5.4 |

1 Dental x-ray systems designed for use with intraoral image receptors and manufactured after December 1, 1980.

2 Dental x-ray systems designed for use with intraoral image receptors and manufactured before or on December 1, 1980, and all other x-ray systems subject to this section and manufactured before June 10, 2006.

3 All x-ray systems, except dental x-ray systems designed for use with intraoral image receptors, subject to this section and manufactured on or after June 10, 2006.

- Optional filtration. Fluoroscopic systems manufactured on or after June 10, 2006, incorporating an x-ray tube(s) with a continuous output of 1 kilowatt or more and an anode heat storage ca- pacity of 1 million heat units or more shall provide the option of adding x-ray filtration to the diagnostic source as- sembly in addition to the amount need- ed   to   meet   the   HVL   provisions   of

§ 1020.30(m)(1). The selection of this ad- ditional x-ray filtration shall  be  either at the option of  the  user  or  automatic as part of the selected mode of oper- ation. A means of indicating which combination of additional filtration  is in the x-ray beam shall be provided.

- Measuring compliance. For capac- itor energy storage equipment, compli- ance shall be determined with the max- imum selectable quantity of charge per exposure.

- Aluminum equivalent of material be- tween patient and image receptor. Except when used in a CT x-ray system, the aluminum equivalent of each of the items listed in table 2 in paragraph (n) of this section, which are used between the patient and image receptor, may not exceed the indicated limits. Com- pliance shall be determined by x-ray measurements  made  at  a  potential  of

100 kilovolts peak and with an x-ray beam that has an  HVL  specified  in table 1 in paragraph (m)(1) of this sec- tion for the potential. This  require- ment applies to front panel(s) of cas- sette holders and film changers pro- vided by the manufacturer for patient support or for prevention of foreign ob- ject intrusions. It does not apply to screens and their associated mechan- ical support panels or grids. Table 2 fol- lows:

| TABLE 2 |  |
| --- | --- |
| Item | Maximum Aluminum Equivalent (millimeters) |
| 1. Front panel(s) of cassette holders (total of all) | 1.2 |
| 2. Front panel(s) of film changer (total of all) | 1.2 |
| 3. Cradle | 2.3 |
| 4. Tabletop, stationary, without articulated joints | 1.2 |
| 5. Tabletop, movable, without articulated joint(s) (including stationary subtop) | 1.7 |
| 6. Tabletop, with radiolucent panel having one articulated joint | 1.7 |

#### Food and Drug Administration, HHS	§ 1020.31

TABLE 2—Continued

| Item | Maximum Aluminum Equivalent (millimeters) |
| --- | --- |
| Tabletop, with radiolucent panel having two or more articulated joints Tabletop, cantilevered Tabletop, radiation therapy simulator | 2.3 2.3 5.0 |

- Battery charge indicator. On bat- tery-powered generators, visual means shall be provided on the control panel to indicate whether the battery is in a state of charge adequate for proper op- eration.

- [Reserved]

- Modification of certified diagnostic x-ray components and systems. (1) Diag- nostic x-ray components and systems certified in accordance with § 1010.2 of this chapter shall not be modified such that the component or system fails to comply  with  any  applicable   provision of this chapter unless a variance in ac- cordance with § 1010.4 of this chapter or an exemption under section 534(a)(5) or 538(b) of the Federal Food, Drug, and Cosmetic  Act  has  been  granted.

(2) The owner of a diagnostic x-ray system who uses the system in a pro- fessional or commercial capacity may modify the system, provided the modi- fication does not result in the  failure  of the system or component  to  comply with  the  applicable   requirements   of this  section  or  of  § 1020.31,  1020.32,  or

1020.33. The owner who causes such modification need not submit the  re- ports required by subpart B of part 1002 of this chapter, provided the owner records the date and the details of the modification in the system records and maintains this information,  and  pro- vided the modification of the x-ray sys- tem does not result in a failure to com- ply with § 1020.31, 1020.32, or 1020.33.

[71 FR 34028, June 10, 2006, as amended at 72

FR 17401, Apr. 9, 2007]

##### § 1020.31   Radiographic equipment.

The provisions of this  section  apply to equipment for radiography, except equipment for fluoroscopic imaging or for recording images from the fluoroscopic image receptor, or com- puted tomography x-ray systems man- ufactured on or after November 29, 1984.

- Control and indication of technique factors—(1) Visual indication. The tech- nique factors to be used during an ex- posure shall be indicated before the ex- posure begins, except when automatic exposure controls are used, in which case the technique factors which are set prior to the exposure shall be indi- cated. On equipment having fixed tech- nique factors, this requirement may be met by permanent markings. Indica- tion of technique factors shall be visi- ble from the operator’s position except in the case of spot films made by the fluoroscopist.

- Timers. Means shall be provided to terminate the exposure at  a  preset time interval, a preset product of cur- rent and time, a preset number of pulses, or a preset  radiation  exposure to the image receptor.

  - Except during serial radiography, the operator shall be able to terminate the exposure at any time during an ex- posure of greater than one-half second. Except during panoramic dental radi- ography, termination of exposure shall cause automatic resetting of the timer to its initial setting or to zero. It shall not be possible to make an exposure when the timer is set to a zero or off position if either position is provided.

  - During serial radiography, the op- erator shall be able to terminate the x- ray exposure(s) at any time, but means may be provided to permit completion of any single exposure of the series in process.

- Automatic exposure controls. When an automatic exposure control is pro- vided:

  - Indication shall be made on the control panel when this mode of oper- ation is selected;

  - When the x-ray tube potential is equal to or greater than 51 kilovolts peak (kVp), the minimum exposure time for  field  emission  equipment rated for pulsed operation shall be equal  to  or  less  than  a  time  interval

#### § 1020.31

equivalent to two pulses and the min- imum exposure time for all other equipment shall be equal to or less than 1⁄60 second or a time interval re- quired to deliver 5 milliampere-seconds (mAs), whichever is greater;

  - Either the product of peak x-ray tube potential, current, and exposure time shall be limited to  not  more  than 60 kilowatt-seconds (kWs) per exposure or  the  product  of  x-ray  tube  current and exposure time shall be limited  to not more than 600 mAs per exposure, except when the x-ray tube potential is less than 51 kVp, in which case the product of x-ray tube current and expo- sure time shall be limited to not more than 2,000 mAs per exposure; and

  - A visible signal shall indicate when an exposure has been terminated at the limits described in paragraph (a)(3)(iii) of this section, and manual resetting shall be required before fur- ther  automatically  timed  exposures can be made.

- Accuracy. Deviation of technique factors from indicated values shall not exceed the limits given in the informa- tion    provided    in    accordance    with

§ 1020.30(h)(3).

- Reproducibility. The following re- quirements shall apply when the equip- ment is operated on an adequate power supply as specified by the  manufac- turer in accordance with the require- ments of § 1020.30(h)(3):

  - Coefficient of variation. For any specific combination of selected tech- nique factors, the estimated coefficient of variation of the air kerma shall be no greater than 0.05.

  - Measuring compliance. Determina- tion of compliance shall be based on 10 consecutive measurements taken with- in a time period of 1 hour. Equipment manufactured after September 5, 1978, shall be subject to the additional re- quirement that all variable controls for technique factors shall be adjusted to alternate settings and  reset to the test setting after each measurement. The percent line-voltage regulation shall be determined for each measurement. All values for percent line-voltage regula- tion shall be within 1 of the mean value for all measurements. For equip- ment having automatic exposure con- trols, compliance shall be determined with  a  sufficient  thickness  of  attenu-

#### 21 CFR Ch. I (4–1–17 Edition)

ating material in the useful beam such that the technique factors can be ad- justed to provide individual  exposures of a minimum of 12 pulses on field emission equipment rated for pulsed operation or no less than one-tenth second per exposure on all other equip- ment.

- Linearity. The following require- ments apply when the equipment is op- erated on a power supply  as  specified by the manufacturer in accordance with the requirements of § 1020.30(h)(3) for any fixed x-ray tube potential with- in the range of 40 percent to 100 percent of the maximum rated.

  - Equipment having independent se- lection of x-ray tube current (mA). The average ratios of air kerma to the indi- cated milliampere-seconds product (mGy/mAs) obtained at any two con- secutive tube current settings shall not differ by more than 0.10 times their sum. This is: X1 - X20.10(X1 + X2); where X1 and X2 are the average mGy/ mAs values obtained at each of two consecutive mAs selector settings or at two settings differing by no more than a factor of 2 where the mAs selector provides continuous selection.

  - Equipment having selection of x-ray tube current-exposure time product (mAs). For equipment  manufactured  after May 3, 1994, the average ratios of air kerma to the indicated milliampere- seconds product (mGy/mAs) obtained at any two consecutive mAs selector settings  shall  not  differ  by  more  than

0.10  times  their  sum.  This  is:  X1   -  X2

0.10 (X1 + X2); where X1 and X2 are the average mGy/mAs values obtained at each of two consecutive mAs selector settings or at two settings  differing  by no more than a factor of 2 where the mAs selector provides continuous se- lection.

  - Measuring compliance. Determina- tion of compliance will be based on 10 exposures, made within 1  hour,  at  each of the two settings. These two settings may include any two focal spot sizes except where  one  is  equal  to  or  less than 0.45 mm and the other is greater than 0.45 mm. For purposes of this re- quirement, focal spot size is the focal spot size specified by the x-ray tube manufacturer. The percent line-voltage regulation shall be determined for each measurement.    All    values    for    percent

#### Food and Drug Administration, HHS	§ 1020.31

line-voltage regulation at any one combination of technique factors shall be within 1 of the mean value for all measurements at these technique fac- tors.

- Field limitation and alignment for mobile, portable, and stationary general purpose x-ray systems. Except when spot-film devices are in service, mobile, portable, and stationary general pur- pose radiographic x-ray systems shall meet the following requirements:

  - Variable x-ray field limitation. A means for stepless adjustment of the size of the x-ray field shall be provided. Each dimension of the minimum field size at an SID of 100 centimeters (cm) shall be equal to or less than 5 cm.

  - Visual definition. (i) Means for vis- ually defining the perimeter of the  x- ray field shall be provided. The total misalignment of the edges of the vis- ually defined field with the respective edges of the x-ray field along either the length or width of the visually defined field shall  not  exceed  2 percent  of  the distance from the source to the center of the visually defined field when the surface upon which it appears is per- pendicular to the axis of the  x-ray beam.

- When a light localizer is used to define the x-ray field, it  shall  provide an average illuminance of not less than 160 lux (15 footcandles) at 100 cm or at the maximum SID, whichever is less. The average illuminance shall be based on measurements made in the approxi- mate center of each quadrant of the light field. Radiation therapy  simula- tion systems are exempt from this re- quirement.

- The edge of the light field at 100 cm or at the maximum SID, whichever is less, shall have a contrast ratio, cor- rected for ambient lighting, of not less than 4 in the case of beam-limiting de- vices designed for use on stationary equipment, and a contrast ratio of not less than 3 in the case of beam-limiting devices designed for use on mobile and portable equipment.  The contrast  ratio is defined as I1/I2, where I1 is the illumi- nance 3 mm from the edge of the light field toward the center of the field; and I2 is the illuminance 3 mm from the edge of the light field away from the center of the field. Compliance shall be

determined with  a measuring aperture of 1 mm.

- Field indication and alignment on stationary general purpose x-ray equip- ment. Except when  spot-film  devices are in service, stationary general pur- pose x-ray systems shall meet the fol- lowing requirements in addition to those prescribed in paragraph (d) of this  section:

  - Means shall be provided to indi- cate when the axis of the x-ray beam is perpendicular to the plane of the image receptor, to align the center of the  x- ray field with respect to the  center  of the image receptor to within 2 percent of the SID, and to indicate the SID to within 2 percent;

  - The beam-limiting device shall numerically indicate the field size in the plane of the image receptor to which it is adjusted;

  - Indication of field size dimensions and SIDs shall be specified in centi- meters and/or inches and shall be such that aperture adjustments result in x- ray field dimensions in the plane of the image receptor which correspond to those indicated by the beam-limiting device to within 2 percent of the SID when the beam axis is indicated to be perpendicular to the plane of the image receptor; and

  - Compliance measurements will be made at discrete SIDs and  image  recep- tor dimensions in common clinical use (such as SIDs of 100, 150, and 200 cm and/or 36, 40, 48, and 72 inches and nominal image receptor dimensions of 13, 18, 24, 30, 35, 40, and 43 cm and/or 5, 7, 8, 9, 10, 11, 12, 14, and 17 inches) or at any other specific dimensions at which the beam-limiting device or its associ- ated diagnostic x-ray  system  is  unique- ly  designed  to  operate.

- Field limitation on radiographic x- ray equipment other than general purpose radiographic systems—(1) Equipment for use with intraoral image receptors. Radio- graphic equipment designed for use with an intraoral image receptor  shall be provided with means to limit the x- ray beam such that:

- If the minimum source-to-skin distance (SSD) is 18 cm or more, the x- ray field at the minimum SSD shall be containable in a circle having a diame- ter of no more than 7 cm; and

#### § 1020.31

- If the minimum SSD is  less  than 18 cm, the x-ray field at the minimum SSD shall be containable in a circle having a diameter of no more than  6 cm.

- X-ray systems designed for one image receptor size. Radiographic equip- ment designed for only one image re- ceptor size at a fixed SID shall be pro- vided with means to limit the  field  at the plane of the image receptor to di- mensions no greater than those of the image receptor, and to align the center of the x-ray field with the center of the image receptor to within 2 percent  of the SID, or shall be provided with means to both size and align the x-ray field such that the x-ray field at the plane of the image receptor does not extend beyond any edge of the image receptor.

- Systems designed for mammog- raphy—(i) Radiographic systems de- signed only  for  mammography and  gen- eral  purpose   radiography   systems, when special attachments for mam- mography are in service, manufactured on or after November 1, 1977, and before September 30, 1999, shall be provided with means to  limit  the  useful  beam such that the x-ray field at  the  plane  of the image receptor does not extend be- yond any edge of the image receptor  at any designated SID except the edge  of the image  receptor designed to  be adja- cent to the chest wall where the x-ray field may not extend beyond  this  edge by more than 2 percent of the SID. This requirement can be met with a system that performs as prescribed in para- graphs (f)(4)(i), (f)(4)(ii), and (f)(4)(iii) of this section. When the beam-limiting device and image receptor support de- vice are designed to be used to immo- bilize the breast during a  mammo- graphic procedure and the SID  may vary, the SID indication specified in paragraphs (f)(4)(ii) and (f)(4)(iii) of this section shall be the maximum SID for which the beam-limiting device or ap- erture  is  designed.

- Mammographic beam-limiting de- vices manufactured on or after Sep- tember 30, 1999, shall be provided with a means to limit the useful beam such that the x-ray field at the plane of the image receptor does not extend beyond any edge of the image receptor by more than  2  percent  of  the  SID.  This  require-

#### 21 CFR Ch. I (4–1–17 Edition)

ment can be met with a system that performs as prescribed in paragraphs (f)(4)(i), (f)(4)(ii), and (f)(4)(iii) of this section. For systems that  allow changes in the SID, the SID indication specified in paragraphs (f)(4)(ii) and (f)(4)(iii) of this section shall be the maximum SID for which the beam-lim- iting device or aperture is designed.

- Each image receptor support de- vice manufactured on or after Novem- ber 1, 1977, intended for installation on a system designed for mammography shall have clear and permanent mark- ings to indicate the maximum image receptor size for which it is designed.

- Other x-ray systems. Radiographic systems not specifically covered in paragraphs (d), (e), (f)(2), (f)(3), and (h) of this section and systems covered in paragraph (f)(1) of this section, which are also designed for use with extraoral image receptors and when used with an extraoral image receptor, shall be pro- vided with means  to  limit  the  x-ray field in the plane  of  the  image  receptor so that such field does not exceed each dimension of the image receptor  by more than 2 percent of the  SID,  when the axis of the x-ray beam is perpen- dicular to the plane of the image recep- tor. In addition, means shall  be  pro- vided to align the center of the x-ray field with the center of the image re- ceptor to within 2  percent  of  the  SID, or means shall be provided to both size and align the x-ray field such that the x-ray field at the plane of the image re- ceptor does not extend beyond any edge of the image receptor. These require- ments may be met with:

  - A system which performs in ac- cordance with paragraphs (d) and (e) of this section; or when alignment means are also provided, may be met with ei- ther;

  - An assortment of removable, fixed-aperture, beam-limiting devices sufficient to meet the requirement for each combination of image  receptor size and SID for which the unit is de- signed. Each such device shall  have clear and permanent markings to indi- cate the image receptor size and  SID for which it is designed; or

  - A beam-limiting device having multiple fixed apertures sufficient to meet the requirement for each com- bination of image receptor size and SID

#### Food and Drug Administration, HHS	§ 1020.31

for which the unit is designed. Perma- nent, clearly legible markings shall in- dicate the image receptor size and SID for which each aperture is designed and shall indicate which aperture is in posi- tion for use.

- Positive beam limitation (PBL). The requirements of this paragraph shall apply to radiographic systems which contain   PBL.

  - Field size. When a PBL system is provided, it shall prevent x-ray produc- tion when:

    - Either the length or width of the x-ray field in the plane of the image re- ceptor differs from the corresponding image receptor dimension by more than 3 percent of the SID; or

    - The sum of the length and width differences as stated in paragraph (g)(1)(i) of this section without regard to sign exceeds 4 percent of the SID.

    - The beam limiting device is  at an SID for which PBL is not designed for sizing.

  - Conditions for PBL. When provided, the PBL system shall function as de- scribed in paragraph (g)(1) of this sec- tion whenever all the following condi- tions are met:

    - The image receptor is inserted into a permanently mounted cassette holder;

    - The image receptor length and width are less than 50 cm;

    - The x-ray beam axis is within 3 degrees of vertical and the SID is 90 cm to 130 cm inclusive; or the x-ray beam axis is within 3 degrees of horizontal and the SID is 90 cm to 205 cm inclu- sive;

    - The x-ray beam axis is perpen- dicular to the plane of the image recep- tor to within 3 degrees; and

    - Neither tomographic nor stereo- scopic radiography is being performed.

  - Measuring compliance. Compliance with the requirements of paragraph (g)(1) of this section shall be deter- mined when the equipment indicates that the beam axis is perpendicular  to the plane of the image receptor and the provisions of paragraph (g)(2) of this section are met. Compliance shall be determined no sooner than 5 seconds after insertion of the image receptor.

  - Operator initiated undersizing. The PBL system shall be capable of oper- ation  such  that,  at  the  discretion  of

the operator, the size of  the  field  may be made smaller than the size of the image receptor through stepless adjust- ment of the field  size.  Each  dimension of the minimum field size at an SID of 100 cm shall be equal to or less than 5 cm. Return to PBL function as de- scribed in paragraph (g)(1) of this sec- tion shall occur  automatically  upon any change of image receptor size or SID.

  - Override of PBL. A capability may be provided for overriding PBL in case of system  failure  and  for  servicing  the system. This override may be for all SIDs and image receptor sizes. A key shall be required for any override capa- bility that is accessible to the oper- ator. It shall not be possible to remove the key while PBL is overridden. Each such key switch or key shall be clearly and durably labeled as follows:

For X-ray Field Limitation System Failure The override capability is considered acces- sible to the operator if it is referenced in the operator’s manual  or  in  other  material  in- tended for the operator or if its location is such that the operator would consider it part of the operational controls.

- Field limitation and alignment for spot-film devices. The following require- ments shall apply to spot-film devices, except when the spot-film device is provided for use with a radiation ther- apy  simulation  system:

  - Means shall be provided between the source and the patient for adjust- ment of the x-ray field size in the plane of the image receptor to the size of that portion of the image receptor which has been selected on the spot- film selector. Such adjustment shall be accomplished automatically  when  the x-ray field size in the  plane  of  the image receptor is greater than the se- lected portion of the image receptor. If the x-ray field size is less than the size of the selected portion of the image re- ceptor, the field  size shall not open automatically to the size of the  se- lected portion of the image receptor unless the operator has selected that mode of operation.

  - Neither the length nor  the  width of the x-ray field in the plane of the image receptor shall differ from the corresponding dimensions of the se- lected portion of the image receptor by more  than  3  percent  of  the  SID  when

#### § 1020.31

adjusted for full coverage of the  se- lected portion  of  the  image  receptor. The sum, without regard to sign, of the length and width differences shall not exceed 4 percent of the SID. On spot- film devices manufactured after Feb- ruary 25, 1978, if the angle between the plane of the image receptor and beam axis is variable, means shall  be  pro- vided to indicate when the axis of the x-ray  beam  is  perpendicular  to  the plane of the image receptor, and com- pliance shall be determined with  the beam axis indicated to  be  perpen- dicular to the plane of the image recep- tor.

  - The center of the x-ray field in the plane of the image receptor shall be aligned with the center of the selected portion of the image receptor to within 2 percent of the SID.

  - Means shall be provided to reduce the x-ray field size in the plane of the image receptor to a size smaller  than the selected portion of the image re- ceptor such that:

- For spot-film devices used on fixed-SID fluoroscopic systems which are not required to, and do not provide stepless adjustment of the x-ray field, the minimum field size, at the greatest SID, does not exceed 125 square cm; or

- For spot-film devices used on fluoroscopic systems that have a vari- able SID and/or stepless adjustment of the field size, the  minimum  field  size, at the greatest SID, shall be contain- able in a square of 5 cm by 5 cm.

  - A capability may be provided for overriding the automatic x-ray field size adjustment in case of system fail- ure. If it is so provided, a signal visible at the fluoroscopist’s position shall in- dicate whenever the automatic x-ray field size adjustment override is en- gaged. Each such system failure over- ride switch shall be clearly labeled as follows:

For X-ray Field Limitation System Failure

    - Source-skin distance—(1) X-ray sys- tems designed for use with an intraoral image receptor shall be provided with means to limit the source-skin dis- tance to not less than:

- Eighteen cm if operable above 50 kVp; or

- Ten cm if not operable above 50 kVp.

#### 21 CFR Ch. I (4–1–17 Edition)

(2) Mobile and portable x-ray systems other than dental shall  be  provided with means to limit the source-skin distance to not less than 30 cm.

- Beam-on indicators. The x-ray con- trol shall provide visual indication whenever x-rays are produced. In addi- tion, a  signal  audible  to  the  operator shall indicate that the exposure has terminated.

- Multiple tubes. Where two or more radiographic tubes are  controlled  by one exposure switch, the tube or tubes which have been selected shall be clearly indicated  before  initiation  of the exposure. This indication shall be both on the x-ray control and at or near the tube housing assembly which has been selected.

- Radiation from capacitor energy storage equipment. Radiation emitted from the x-ray tube shall not exceed:

- An air kerma of 0.26 microGy (vice

0.03 mR exposure) in 1 minute at 5 cm from any accessible surface of the diag- nostic source assembly, with the beam- limiting device fully  open,  the  system fully charged, and the exposure switch, timer, or any discharge mechanism not activated. Compliance shall be deter- mined by measurements averaged  over an area of 100 square cm, with no linear dimension greater than 20 cm; and

- An air kerma of 0.88 mGy (vice 100 mR exposure) in 1 hour at 100 cm from the x-ray source, with the beam-lim- iting device fully open, when the sys- tem is discharged through  the  x-ray tube either manually or  automatically by use of a discharge switch or deacti- vation of the input power. Compliance shall be determined by  measurements of the maximum air kerma per dis- charge multiplied by the  total  number of discharges in 1 hour (duty cycle). The measurements shall be averaged over an area of 100 square cm with no linear dimension greater than 20 cm.

- Primary protective barrier for mam- mography x-ray systems—(1) For x-ray systems manufactured after  September 5, 1978, and before September 30, 1999, which are designed only for mammog- raphy, the transmission of the primary beam through any image receptor sup- port provided with the system shall be limited such that the air kerma 5 cm from any accessible surface beyond the plane  of  the  image  receptor  supporting

#### Food and Drug Administration, HHS	§ 1020.32

device does not exceed 0.88 microGy (vice 0.1 mR exposure) for each activa- tion of the tube.

  - For mammographic x-ray systems manufactured on or after September 30, 1999:

    - At any SID where exposures can be made, the image receptor support device shall provide a primary protec- tive barrier that intercepts the cross section of the useful beam along every direction except at the chest wall edge.

    - The x-ray system shall not per- mit exposure unless the appropriate barrier is in place to intercept the use- ful beam as required in paragraph (m)(2)(i) of this section.

    - The transmission of the useful beam through the primary protective barrier shall be limited such  that  the air kerma 5 cm from any accessible surface beyond the plane  of  the  pri- mary protective barrier does  not  ex- ceed 0.88 microGy (vice 0.1 mR expo- sure)  for  each  activation  of  the  tube.

  - Compliance with the requirements of paragraphs (m)(1) and (m)(2)(iii) of this section for transmission shall be determined with the x-ray system oper- ated at the  minimum  SID  for  which  it is designed, at the  maximum  rated peak tube potential, at the maximum rated product of x-ray tube current and exposure time (mAs) for the maximum rated peak tube potential, and by measurements averaged over an area of 100 square cm with no linear dimension greater than 20 cm. The sensitive vol- ume of the radiation measuring instru- ment shall not be positioned beyond the edge of the primary protective bar- rier along the chest wall side.

[70 FR 34036, June 10, 2005]

##### § 1020.32   Fluoroscopic equipment.

The provisions of this  section  apply to equipment for fluoroscopic imaging or for recording images from the fluoroscopic image receptor, except computed tomography x-ray systems manufactured on or after November 29, 1984.

- Primary protective barrier—(1) Limi- tation of useful beam. The fluoroscopic imaging assembly shall be provided with a primary protective  barrier which intercepts the entire cross sec- tion of the useful beam at any SID. The x-ray  tube  used  for  fluoroscopy  shall

not produce x-rays unless  the  barrier  is in position to intercept the  entire  use- ful beam. The AKR due to transmission through the barrier with the  attenu- ation block in the useful  beam  com- bined with radiation from  the fluoroscopic image receptor shall not exceed 3.34 10¥3 percent of the en- trance AKR,  at  a  distance  of  10  cm from any accessible surface of the fluoroscopic imaging assembly  beyond the plane of the image receptor. Radi- ation therapy simulation systems  shall be exempt from this requirement pro- vided the systems are intended only for remote control operation and the man- ufacturer sets forth instructions for as- semblers with respect to control loca- tion as part of the information re- quired in § 1020.30(g). Additionally, the manufacturer shall provide to  users, under § 1020.30(h)(1)(i), precautions con- cerning the importance of remote con- trol   operation.

- Measuring compliance. The AKR shall be measured in accordance with paragraph (d) of this section. The AKR due to transmission through the pri- mary barrier combined with radiation from the fluoroscopic image receptor shall be determined by measurements averaged over an area of 100 square cm with no linear dimension greater  than 20 cm. If the source is below the table- top, the measurement shall be made with the input surface of the fluoroscopic imaging assembly posi- tioned 30 cm above the tabletop. If the source is above the tabletop and the SID is variable, the measurement  shall be made with the end of the beam-lim- iting device or spacer as close to the tabletop as it can be placed, provided that it shall not be closer than 30 cm. Movable grids and compression devices shall be removed from the useful beam during the measurement. For all meas- urements, the attenuation block shall be positioned in the useful beam 10 cm from the point of measurement of en- trance AKR and between this point and the input surface of the fluoroscopic imaging  assembly.

- Field limitation—(1) Angulation. For fluoroscopic equipment manufac- tured after February 25, 1978, when the angle between  the  image  receptor and the  beam  axis  of  the  x-ray  beam  is

#### § 1020.32

variable, means shall be provided to in- dicate when the axis of the x-ray beam is perpendicular to the plane of the image receptor. Compliance with para- graphs (b)(4) and (b)(5) of this section shall be determined with the beam axis indicated to be perpendicular to the plane of the image receptor.

  - Further means for limitation. Means shall be provided to permit further lim- itation of the x-ray field to sizes small- er than the limits of paragraphs (b)(4) and (b)(5). Beam-limiting devices man- ufactured after May 22, 1979, and incor- porated in equipment with a variable SID and/or the capability of a visible area of greater than 300 square cm, shall be provided with means for stepless adjustment of the x-ray field. Equipment with a fixed SID and the ca- pability of a visible area of no greater than 300 square cm shall be provided with either stepless adjustment of the x-ray field or with a means to further limit the x-ray field size at the plane of the image receptor to 125 square cm or less. Stepless adjustment shall, at the greatest SID, provide continuous field sizes from the maximum obtainable to a field size containable in a square of 5 cm by 5 cm. This paragraph does not apply to non-image-intensified fluoros- copy.

  - Non-image-intensified fluoroscopy. The x-ray field produced by non-image- intensified  fluoroscopic  equipment shall not extend beyond the entire visi- ble area of the image receptor. Means shall be provided for stepless adjust- ment of field size. The minimum field size, at the greatest SID, shall be con- tainable in a square of 5 cm by 5 cm.

  - Fluoroscopy and radiography using the fluoroscopic imaging assembly with inherently circular image receptors. (i) For fluoroscopic equipment manufac- tured before June 10, 2006, other than radiation therapy simulation systems, the following applies:

    - Neither the length nor the width of the x-ray field in the plane of the image receptor shall exceed that of the visible area of the image receptor by more than 3 percent of the SID. The sum of the excess length and the excess width shall be no greater than 4 per- cent of the SID.

    - For rectangular x-ray fields used with   circular   image   receptors,   the

#### 21 CFR Ch. I (4–1–17 Edition)

error in alignment shall be determined along the length and width dimensions of the x-ray field which pass  through the center of the visible area of the image receptor.

- For fluoroscopic equipment man- ufactured on or after June 10, 2006, other than radiation therapy  simula- tion systems, the maximum area of the x-ray field in the plane of the image re- ceptor shall conform with one of the following requirements:

  - When every linear dimension of the visible area of the image receptor measured through the center of the visible area is less than or equal to 34 cm in any direction, at least 80 percent of the area of the x-ray  field  overlaps the visible area of the image.

  - When any linear dimension of the visible area of the image receptor measured through the center of the visible area is greater than 34 cm  in any direction, the x-ray field measured along the direction of greatest mis- alignment with the visible area of the image receptor does not extend beyond the edge of the visible  area  of  the image receptor by more than 2 cm.

  - Fluoroscopy and radiography using the fluoroscopic imaging assembly with inherently rectangular image receptors. For x-ray systems manufactured on or after June 10, 2006, the following ap- plies:

- Neither the length nor  the  width of the x-ray field in the plane of the image receptor shall exceed that of the visible area of the image receptor by more than 3 percent of the SID. The sum of the excess length and the excess width shall be no greater than 4 per- cent of the SID.

- The error in alignment shall be determined along the length and width dimensions of the x-ray  field  which pass through the center of the visible area of the image receptor.

  - Override capability. If the fluoroscopic x-ray field size is adjusted automatically as the SID or image re- ceptor size is changed, a capability may be provided for overriding the automatic adjustment in case of sys- tem failure. If it is so provided, a signal visible at the fluoroscopist’s position shall indicate whenever the automatic field  adjustment  is  overridden.  Each

#### Food and Drug Administration, HHS	§ 1020.32

such   system   failure   override   switch shall be clearly labeled as follows:

For X-ray Field Limitation System Failure

- Activation of tube. X-ray produc- tion in the fluoroscopic mode shall be controlled by a device which requires continuous pressure by the operator for the entire time of any exposure. When recording serial radiographic images from the fluoroscopic image receptor, the operator shall be able to terminate the x-ray exposure(s) at any time, but means may be provided to permit com- pletion of any single exposure of the se- ries in process.

- Air kerma rates. For fluoroscopic equipment, the following requirements apply:

  - Fluoroscopic equipment  manufac- tured before May 19, 1995—(i) Equipment provided with automatic exposure rate control (AERC) shall not be operable at any combination of tube potential and current that will result in  an  AKR  in excess of 88 mGy per minute (vice 10 R/ min exposure rate) at the measurement point specified in § 1020.32(d)(3),  except as specified in § 1020.32(d)(1)(v).

  - Equipment  provided   without AERC shall not be operable at any combination of tube  potential  and  cur- rent  that  will  result  in  an  AKR  in   ex- cess of 44 mGy per minute (vice 5 R/min exposure  rate)   at   the   measurement point specified in § 1020.32(d)(3),  except as specified in § 1020.32(d)(1)(v).

  - Equipment provided with both an AERC mode and a manual mode shall not be operable at any combina- tion of tube potential and current that will result in an AKR in excess of 88 mGy per minute (vice 10 R/min expo- sure rate) in either mode at the meas- urement	point	specified	in

§ 1020.32(d)(3),    except    as    specified    in

§ 1020.32(d)(1)(v).

  - Equipment may be modified in accordance with § 1020.30(q) to comply with § 1020.32(d)(2). When the equipment is modified, it shall bear a label indi- cating the date of the modification  and the   statement:

Modified  to  comply  with  21  CFR 1020.32(h)(2).

  - Exceptions:

- During recording of fluoroscopic images, or

- When a mode of operation has an optional high-level control, in which case that mode shall not be operable at any combination of tube potential and current that will result in an AKR in excess    of    the    rates    specified    in

§ 1020.32(d)(1)(i),   (d)(1)(ii),   or   (d)(1)(iii)

at  the  measurement  point  specified  in

§ 1020.32(d)(3), unless the high-level con- trol is activated. Special means of acti- vation of high-level controls shall be required. The high-level control shall be operable only when continuous man- ual activation is provided by the oper- ator.  A  continuous  signal  audible  to the fluoroscopist  shall  indicate  that the high-level control is being em- ployed.

  - Fluoroscopic equipment manufac- tured on or after May 19, 1995—(i) Shall be equipped with AERC if operable at any combination of tube potential and current that results in an AKR greater than 44 mGy per minute (vice 5 R/min exposure rate)  at  the  measurement point specified in § 1020.32(d)(3). Provi- sion for manual selection of technique factors  may  be  provided.

- Shall not be operable at any com- bination of tube potential and current that will result in an  AKR  in  excess  of 88 mGy per minute (vice 10 R/min expo- sure rate) at the measurement point specified in § 1020.32(d)(3), except as specified in § 1020.32(d)(2)(iii):

- Exceptions:

- For equipment manufactured prior to June 10, 2006, during the re- cording of images from a fluoroscopic image receptor using photographic film or a video camera when the x-ray source is operated in a pulsed mode.

- For equipment manufactured on or after June 10, 2006, during the re- cording of images from  the fluoroscopic image receptor for the purpose of providing the user with a re- corded image(s)  after  termination  of the exposure. Such recording does not include images resulting from a last- image-hold feature that are not re- corded.

- When a mode of operation has an optional high-level control and the control is activated, in which case the equipment shall not be operable at any combination of tube potential and cur- rent that will result in an AKR in  ex- cess of 176 mGy per minute (vice 20 R/

#### § 1020.32

min exposure rate) at the measurement point specified in § 1020.32(d)(3). Special means of activation of high-level con- trols shall be required. The high-level control shall  be  operable  only when continuous manual activation is pro- vided by the operator. A continuous signal audible to the fluoroscopist shall indicate that the high-level control is being employed.

  - Measuring compliance. Compliance with paragraph (d) of this section shall be determined as follows:

    - If the source is below the x-ray table, the AKR shall be measured  at  1 cm above the tabletop or cradle.

    - If the source is above the x-ray table, the AKR shall be measured at 30 cm above the tabletop with the end of the beam-limiting device or spacer po- sitioned as closely as possible to the point of measurement.

    - In a C-arm type of fluoroscope, the AKR shall be measured at 30 cm from the input surface of the fluoroscopic imaging assembly,  with the source positioned at any available SID, provided that the end of the beam- limiting device or spacer is no closer than 30 cm from the input  surface  of the fluoroscopic imaging assembly.

    - In a C-arm type of fluoroscope having an SID less than 45 cm, the AKR shall be measured at the min- imum  SSD.

    - In a lateral type of fluoroscope, the air kerma rate  shall  be  measured at a point 15 cm from the centerline of the x-ray table and in the  direction  of the x-ray source with the end of the beam-limiting device or spacer posi- tioned as closely as  possible  to  the point of measurement.  If  the  tabletop is movable, it shall be positioned as closely as possible to the lateral x-ray source, with the end of the beam-lim- iting device or spacer no closer than 15 cm to the centerline of the x-ray table.

  - Exemptions. Fluoroscopic radi- ation therapy simulation systems are exempt from the requirements set forth in paragraph (d) of this section.

- [Reserved]

- Indication of potential and current. During	fluoroscopy	and cinefluorography, x-ray tube potential and current shall be continuously indi- cated. Deviation of x-ray tube poten- tial and current from the indicated val-

#### 21 CFR Ch. I (4–1–17 Edition)

ues shall not exceed the maximum de- viation as stated by  the   manufacturer in accordance with § 1020.30(h)(3).

- Source-skin distance. (1) Means shall be provided to limit the source- skin distance to not less than 38 cm on stationary fluoroscopes and to not less than 30 cm on mobile and portable fluoroscopes. In addition, for fluoroscopes intended for specific sur- gical application that would be prohib- ited at the source-skin distances speci- fied in this paragraph, provisions  may be made for operation at shorter source-skin distances but in no case less than 20 cm. When provided, the manufacturer must set forth pre- cautions with respect to the optional means of spacing, in addition to other information as required in § 1020.30(h).

- For stationary, mobile, or  port- able C-arm fluoroscopic systems manu- factured on or after June 10, 2006, hav- ing a maximum source-image receptor distance of less than 45 cm, means shall be provided  to limit the  source-skin distance to not less than 19 cm. Such systems shall be labeled for extremity use only. In addition, for those systems intended for specific surgical applica- tion that would be prohibited at the source-skin distances specified in this paragraph, provisions may be made for operation at shorter source-skin dis- tances but in no case less than 10 cm. When provided, the manufacturer must set forth precautions with respect to the optional means of spacing, in addi- tion to other  information  as  required in § 1020.30(h).

- Fluoroscopic irradiation time, dis- play, and signal. (1)(i) Fluoroscopic equipment  manufactured  before   June 10, 2006, shall be  provided  with  means to  preset  the  cumulative  irradiation time of the fluoroscopic tube. The max- imum cumulative time of the timing device shall not exceed 5 minutes  with- out resetting. A signal audible to the fluoroscopist shall indicate the comple- tion of any preset cumulative irradia- tion-time.   Such   signal   shall    continue to sound while x-rays  are  produced until the timing device is reset. Fluoroscopic equipment may be modi- fied in accordance with § 1020.30(q) to comply     with     the     requirements     of

§ 1020.32(h)(2).   When   the   equipment   is

#### Food and Drug Administration, HHS	§ 1020.32

modified,  it  shall  bear  a  label  indi- cating the statement:

Modified  to  comply  with  21  CFR 1020.32(h)(2).

(ii) As an alternative to the require- ments of this paragraph, radiation therapy simulation systems may be provided with a means to indicate the total cumulative exposure time during which x-rays were produced, and which is capable of being reset between x-ray examinations.

  - For x-ray controls manufactured on or after June 10, 2006, there shall be provided for each fluoroscopic tube:

    - A display of the fluoroscopic irra- diation time at the  fluoroscopist’s working position. This display shall function independently of the audible signal described in § 1020.32(h)(2)(ii). The  following  requirements  apply:

- When the x-ray tube is activated, the fluoroscopic irradiation time in minutes and tenths of minutes shall be continuously displayed  and  updated  at least once every 6 seconds.

- The fluoroscopic irradiation time shall also be displayed within 6 seconds of termination of an exposure and re- main displayed until reset.

- Means shall be provided to reset the display to zero prior to the begin- ning of a new examination or proce- dure.

    - A signal audible to the fluoroscopist shall sound for each pas- sage of  5 minutes  of fluoroscopic irra- diation time during an examination or procedure. The signal shall sound until manually reset or,  if  automatically reset, for at least 2 second.

(i) Mobile and portable fluoroscopes. In addition to the other requirements of this section, mobile and portable fluoroscopes shall provide an image re- ceptor incorporating more than a sim- ple fluorescent screen.

- Display of last-image-hold (LIH). Fluoroscopic equipment manufactured on or after June 10, 2006, shall be equipped with means to display LIH image following  termination  of  the fluoroscopic exposure.

  - For an LIH image obtained by re- taining pretermination fluoroscopic images, if the number of images and method of combining images are se- lectable by the user, the selection shall

be  indicated  prior  to  initiation  of  the fluoroscopic exposure.

  - For an LIH image obtained by ini- tiating a separate radiographic-like ex- posure at the termination of fluoroscopic imaging, the techniques factors for the LIH image shall be se- lectable prior to the fluoroscopic expo- sure, and the combination selected shall be indicated prior to initiation of the fluoroscopic exposure.

  - Means shall be provided to clearly indicate to the user whether a dis- played image is the LIH radiograph or fluoroscopy. Display of the LIH radiograph shall be replaced by the fluoroscopic image concurrently with re-initiation of fluoroscopic exposure, unless separate displays are provided for the LIH radiograph  and fluoroscopic images.

  - The predetermined or selectable options for  producing  the  LIH radiograph shall be described in the in- formation required by § 1020.30(h). The information shall include a  description of any technique factors applicable for the selected option and the impact  of the selectable options on image charac- teristics and the magnitude of  radi- ation   emissions.

- Displays of values of AKR and cu- mulative air kerma. Fluoroscopic equip- ment manufactured on or after June 10, 2006, shall display at the fluoroscopist’s working position the AKR and cumu- lative air kerma. The following re- quirements apply for each x-ray tube used during an examination or proce- dure:

  - When the x-ray tube is activated and the number of images produced per unit time is greater than six images per second, the AKR in mGy/min shall be continuously displayed and updated at least once every second.

  - The cumulative air kerma in units of mGy shall be displayed either within 5 seconds of termination of an exposure or displayed continuously and updated at least once every 5 seconds.

  - The display of the AKR shall be clearly distinguishable from the dis- play of the cumulative air kerma.

  - The AKR and cumulative air kerma shall represent the value for conditions of free-in-air irradiation at one of the following reference locations

#### § 1020.33

specified according to the type of fluo- roscope. The reference location shall be identified and described specifically in the information provided to users ac-

#### 21 CFR Ch. I (4–1–17 Edition)

tomograms produced in a single scan; that is:

7T

cording to § 1020.30(h)(6)(iii).

- For fluoroscopes with x-ray source below the x-ray table, x-ray source above  the  table,  or  of  lateral  type,  the

where:

CTDI 1 / nT 

7T

Dzdz

reference locations shall be the respec- tive	locations	specified	in

§ 1020.32(d)(3)(i), (d)(3)(ii), or (d)(3)(v) for measuring compliance with air-kerma rate    limits.

- For C-arm fluoroscopes, the ref- erence location shall be 15 cm from the isocenter toward the x-ray source along the beam axis. Alternatively, the reference location shall be at a point specified by the manufacturer to rep- resent the location of the intersection of the x-ray beam with the patient’s skin.

  - Means  shall  be  provided  to  reset

z = position along a line perpendicular to the

tomographic plane. D(z) = Dose at position z.

T = Nominal tomographic section thickness. n = Number of tomograms produced in a sin-

gle scan.

This definition assumes that the dose profile is centered around z = 0 and that, for a multiple tomogram system, the scan increment between adjacent scans is nT.

- Contrast scale means the change in linear attenuation coefficient per CT number relative to water; that is:

to zero the display of cumulative air kerma prior to the commencement of a new examination or procedure.

  - The   displayed   AKR   and   cumu-

### Contrast scale 

where:

x w

CTx CTw

lative air kerma shall not deviate from the actual values by more than 35 per- cent over the range of 6 mGy/min and 100 mGy to the maximum indication of AKR and cumulative air kerma, respec- tively. Compliance shall be determined with an irradiation time greater than 3 seconds.

[70 FR 34039, June 10, 2005, as amended at 80

FR 19532, Apr. 13, 2015]

##### § 1020.33   Computed  tomography  (CT) equipment.

    - Applicability. (1) The provisions of this section, except for paragraphs (b), (c)(1), and (c)(2) are applicable as speci- fied herein to CT x-ray systems manu- factured or remanufactured on or after September 3, 1985.

(2) The provisions of paragraphs (b), (c)(1), and (c)(2) are applicable to CT x- ray systems manufactured or remanu- factured on or after November 29, 1984.

    - Definitions. As used in this sec- tion, the following definitions apply:

      - Computed tomography dose index (CTDI) means the integral of the dose profile along a line perpendicular to the tomographic plane divided by the product of the nominal tomographic section  thickness  and  the  number  of

w = Linear attenuation coefficient of water.

x  = Linear attenuation coefficient of mate- rial of interest.

(CT)w = CT number of water.

(CT)x = CT number of material of interest.

- CT  conditions  of  operation   means all selectable parameters governing the operation of a CT x-ray system  includ- ing nominal tomographic  section thickness, filtration, and the technique factors as defined in § 1020.30(b)(36).

- CT number means the number used to represent the x-ray attenuation as- sociated with each elemental area of the CT image.

- [Reserved]

- CT dosimetry phantom means the phantom used for determination of the dose  delivered  by  a  CT  x-ray  system. The phantom shall be a right circular cylinder of polymethl-methacrylate of density 1.190.01 grams per cubic centi- meter. The phantom shall be at least 14 centimeters in length and shall have diameters of 32.0 centimeters for test- ing any CT system designed to image any section of the body (whole body scanners) and 16.0 centimeters for any system designed to image the head (head scanners) or for any whole body scanner operated in the head scanning mode.     The     phantom     shall     provide

Food and Drug Administration, HHS	§ 1020.33

means for the placement of a dosim- eter(s) along its axis of rotation and along a line parallel to the axis of rota- tion 1.0 centimeter from the outer sur- face and within the phantom. Means for the placement of a dosimeter(s) or alignment device at other  locations may be provided for convenience. The means used for placement of a dosim- eter(s) (i.e., hole size) and the type of dosimeter(s) used is at the discretion of the manufacturer. Any effect on the doses measured due to the removal of phantom material to accommodate dosimeters shall be accounted for through appropriate corrections to the reported data or included in the state- ment of maximum deviation for the values obtained using the phantom.

- Dose profile means the dose as a function of position along a line.

- Modulation transfer function means the modulus of the Fourier transform of the impulse response of the system.

- Multiple tomogram system means a CT x-ray system which obtains x-ray transmission data simultaneously dur- ing a single scan to produce more than one tomogram.

- Noise means the standard devi- ation of the fluctuations in CT number expressed as a percent of the attenu- ation coefficient of water. Its estimate (Sn) is calculated using the following expression:

### 100 CS s

Sn

w

where:

CS = Contrast scale.

w = Linear attenuation coefficient of water. s = Estimated standard deviation of the CT numbers of picture elements in a specified

area of the CT image.

- Nominal tomographic section thick- ness means the full-width at half-max- imum of the sensitivity  profile  taken at the center of the cross-sectional vol- ume over which x-ray transmission data are collected.

- Picture element means an ele- mental area of a tomogram.

- Remanufacturing means modi- fying a CT system in such a way that the resulting dose and imaging per- formance become substantially equiva- lent to any CT x-ray system manufac- tured by the original manufacturer on

or  after  November  29,  1984.  Any  ref- erence   in   this   section   to   ‘‘manufac- ture’’,   ‘‘manufacturer’’,   or   ‘‘manufac- turing’’    includes    remanufacture,    re- manufacturer,  or  remanufacturing,  re- spectively.

- Scan increment means the amount of relative displacement of the patient with respect to the CT x-ray system be- tween successive scans measured along the direction of such displacement.

- Scan sequence means a preselected set of two or more scans performed	consecutively	under preselected CT conditions of oper- ations.

- Sensitivity profile means the rel- ative response of the CT x-ray system as a function of position along a line perpendicular to the tomographic plane.

- Single tomogram system means a CT x-ray system which obtains x-ray transmission data during a scan to produce a single tomogram.

- Tomographic plane means that geometric plane which the manufac- turer identifies as corresponding to the output  tomogram.

- Tomographic section means the volume of an object whose x-ray at- tenuation properties are imaged in a tomogram.

    - Information to be provided for users. Each manufacturer of a CT x-ray sys- tem shall provide the following tech- nical and safety information, in addi- tion to that required under § 1020.30(h), to purchasers and, upon request, to others at a cost not to exceed the cost of publication and distribution of such information. This information shall be identified and provided in a separate section of the user’s instruction man- ual or in a separate manual devoted only  to  this  information.

      - Conditions of operation. A state- ment of the CT conditions of operation used to provide the information re- quired by paragraph (c) (2) and (3) of this  section.

      - Dose information. The following dose information obtained by using the CT dosimetry phantom. For any CT  x- ray system designed to image both the head and body, separate dose informa- tion shall be provided for each applica- tion.  All  dose  measurements  shall  be

#### § 1020.33

performed with the CT dosimetry phan- tom placed on the patient couch or support device without additional at- tenuating materials present.

        - The CTDI at the following loca- tions in the dosimetry phantom:

- Along the axis of rotation of the phantom.

- Along a line parallel to the axis of rotation and 1.0 centimeter interior to the surface of the phantom with the phantom positioned so that CTDI is the maximum obtainable at this depth.

- Along lines parallel to the axis of rotation and 1.0 centimeter interior to the surface of the phantom at positions 90, 180, and 270 degrees from the posi- tion in paragraph (c)(2)(i)(b) of this sec- tion. The CT conditions of operation shall be the typical values suggested by the manufacturer for CT of the head or body. The location of the position where the CTDI is maximum as speci- fied in paragraph (c)(2)(i)(b) of this sec- tion shall be given by the manufac- turer with respect to the housing of the scanning mechanism or other readily identifiable feature of the CT x-ray system in such a manner as to permit placement of the dosimetry phantom in this   orientation.

        - The CTDI in the center location of the dosimetry phantom for each se- lectable CT condition of operation that varies either the rate or duration of x- ray exposure. This CTDI shall be pre- sented as a value that is normalized to the CTDI in the center location of the dosimetry phantom from paragraph (c)(2)(i) of this section, with the CTDI of paragraph (c)(2)(i) of this section having a value of one. As each indi- vidual CT condition of operation is changed, all other independent CT con- ditions of operation shall be main- tained at the typical values described in paragraph (c)(2)(i) of this section. These data shall encompass the range of each CT condition of operation stat- ed by the manufacturer as appropriate for CT of the head or body. When more than three selections of a CT condition of operation are available, the normal- ized CTDI  shall be  provided, at  least, for the minimum, maximum, and mid- range value of the CT condition of op- eration.

        - The CTDI at the location coinci- dent with the maximum CTDI at 1 cen-

#### 21 CFR Ch. I (4–1–17 Edition)

timeter interior to the surface of the dosimetry phantom for each selectable peak tube potential. When more than three selections of peak tube potential are  available,  the  normalized  CTDI shall be provided, at least, for the min- imum, maximum, and a  typical  value of peak tube potential. The  CTDI  shall be presented as a value that is normal- ized to the maximum CTDI located at 1 centimeter interior to the surface of the dosimetry phantom from paragraph (c)(2)(i) of this section,  with  the  CTDI of paragraph (c)(2)(i) of this section having a value of one.

        - The dose profile in the center lo- cation of the dosimetry phantom for each selectable nominal tomographic section thickness. When more than three selections of nominal tomo- graphic section thicknesses are avail- able, the information shall be provided, at least, for the minimum, maximum, and midrange value of nominal tomo- graphic section thickness. The dose profile shall be presented on the same graph and to the same scale as the cor- responding sensitivity profile required by paragraph (c)(3)(iv) of this section.

        - A statement of the maximum de- viation from the values given in the in- formation provided according to para- graph (c)(2) (i), (ii), (iii), and (iv) of this section. Deviation of actual values may not exceed these limits.

      - Imaging performance information. The following performance data shall be provided for the CT conditions of op- eration used to provide the information required by paragraph (c)(2)(i) of this section. All other aspects of data col- lection, including the x-ray attenu- ation properties of the material in the tomographic section, shall be  similar to those used to provide the dose infor- mation required by paragraph (c)(2)(i) of this section. For any CT x-ray sys- tem designed to image both the head and body, separate imaging perform- ance information shall be provided for each application.

- A statement of the noise.

- A graphical presentation of the modulation transfer function for the same image processing and display mode as that used in the statement of the noise.

- A statement of the nominal tomographic section thickness(es).

#### Food and Drug Administration, HHS	§ 1020.33

- A graphical presentation of the sensitivity profile, at the location cor- responding  to  the  center  location  of the dosimetry phantom, for each se- lectable nominal tomographic section thickness for which the dose profile is given according to paragraph (c)(2)(iv) of this section.

- A description of the phantom or device and test protocol or procedure used to determine the specifications and a statement of the maximum devi- ation from the specifications provided in accordance with paragraphs (c)(3) (i), (ii), (iii), and (iv) of this section. Deviation of actual values may not ex- ceed these limits.

- Quality assurance. The manufac- turer of any CT x-ray system shall pro- vide the following with each  system. All information required by this sub- section shall be provided in a separate section of the user’s instructional man- ual.

- A phantom(s) capable of providing an indication of contrast scale, noise, nominal tomographic section thick- ness, the spatial resolution  capability of the system for low and high contrast objects, and measuring the mean CT number of water or a reference mate- rial.

- Instructions on the use of the phantom(s) including a schedule of testing appropriate for the system, al- lowable variations for the indicated pa- rameters, and a method to store as records, quality assurance data.

- Representative images obtained with the phantom(s) using the same processing mode and CT conditions of operation as in paragraph (c)(3) of this section for a properly functioning sys- tem of the same model. The represent- ative images shall be of two forms as follows:

- Photographic copies of the images obtained from the image display de- vice.

- Images stored in digital form on a storage medium compatible with the CT x-ray system. The CT x-ray system shall be provided with the means to display these images on the image dis- play device.

- [Reserved]

- Control and indication of conditions of operation—(1) Visual indication. The CT  conditions  of  operation  to  be  used

during a scan or a scan sequence shall be indicated prior to initiation of a scan or a scan sequence. On equipment having all or some of these conditions of operation at fixed values, this re- quirement may be met by permanent markings. Indication of the CT condi- tions of operation shall be visible from any position from which scan initi- ation is possible.

(2) Timers. (i) Means shall be provided to terminate the x-ray exposure auto- matically by either deenergizing the x- ray source or shuttering the x-ray beam in the event of equipment failure affecting data collection. Such termi- nation shall occur within an interval that limits the total scan time to no more than 110 percent of its preset value through the use of either a backup timer or devices which monitor equipment function. A visible signal shall indicate when the x-ray exposure has been terminated through these means and  manual resetting  of the CT conditions of operation shall be re- quired prior to the initiation  of  an- other scan.

(ii) Means shall be provided so that the operator can terminate the x-ray exposure at any time during a scan, or series of scans under x-ray system con- trol, of greater than one-half second duration. Termination of the x-ray ex- posure shall necessitate resetting of the CT conditions of operation prior to the initiation of another scan.

- Tomographic plane indication and alignment. (1) For any single tomogram system, means shall be provided to per- mit visual determination of the tomo- graphic plane or a reference plane off- set from the tomographic plane.

- For any multiple tomogram sys- tem, means shall be provided to permit visual determination  of  the  location  of a reference plane. The relationship of the reference plane to the planes of the tomograms shall be provided  to  the user in addition to other information provided according to § 1020.30(h). This reference plane can be  offset from the location of the tomographic planes.

- The distance between  the indi- cated location of  the  tomographic plane or reference plane and its actual location may not exceed 5 millimeters.

- For any offset alignment system, the manufacturer shall provide specific

#### § 1020.40

instructions with respect to the use of this system for patient positioning, in addition to other information provided according to § 1020.30(h).

- If a mechanism using a light source is used to satisfy the require- ments of paragraphs (g) (1) and (2) of this section, the light source shall allow visual determination of the loca- tion of the tomographic plane or ref- erence plane under ambient light con- ditions of up to 500 lux.

- Beam-on and shutter status indica- tors. (1) Means shall be provided on the control and on or near the  housing  of the scanning mechanism to provide vis- ual indication when and only when x rays are produced and, if applicable, whether the  shutter  is open  or  closed. If the x-ray production period is less than one-half second, the indication of x-ray production shall be actuated for one-half second. Indicators at or near the housing of the scanning mechanism shall be discernible from any point ex- ternal to the patient opening where in- sertion of any part of the human body into the primary beam is possible.

- For systems that allow high volt- age to be applied to the x-ray tube con- tinuously and that  control  the  emis- sion of x-ray with a shutter, the radi- ation emitted may not exceed 0.88 milligray (vice 100 milliroentgen expo- sure) in 1 hour at any point 5  cm  out- side the external surface  of  the  housing of the scanning mechanism when the shutter is closed. Compliance shall be determined by measurements average over an area of 100 square cm with no linear dimension greater than 20 cm.

  - Scan increment accuracy. The devi- ation of indicated scan increment from actual scan increment may not exceed

1 millimeter. Compliance shall be measured as follows: The determina- tion of the  deviation  of  indicated versus actual scan increment shall be based on measurements taken with a mass 100 kilograms or less, on the pa- tient support device. The patient sup- port device  shall  be  incremented  from a typical starting position to the max- imum incrementation distance or 30 centimeters, whichever is less, and then returned to the starting position. Measurement of actual versus indi- cated scan increment may be taken anywhere along this travel.

#### 21 CFR Ch. I (4–1–17 Edition)

(j) CT number mean and standard devi- ation. (1) A method shall be provided to calculate the mean and standard devi- ation of CT numbers for an array of picture elements about any location in the image. The number of elements in this  array  shall  be  under  user  control.

- The manufacturer shall provide specific  instructions  concerning  the use of the method provided for calcula- tion of CT number mean and standard deviation in addition to other informa- tion provided according to § 1020.30(h).

[49 FR 34712, Aug. 31, 1984; 49 FR 37381, Sept.

24, 1984, as amended at 49 FR 47388, Dec. 4,

1984;  56  FR  36098,  Aug.  1,  1991;  67  FR  9587,

Mar. 4, 2002; 70 FR 34042, June 10, 2005]

##### § 1020.40   Cabinet x-ray systems.

  - Applicability. The provisions of this section are applicable to cabinet x- ray systems manufactured or assem- bled on or after April 10, 1975, except that the provisions as applied to x-ray systems designed primarily for the in- spection of carry-on baggage are appli- cable to such systems manufactured or assembled on or after April 25, 1974. The provisions of this section are not applicable to systems which are de- signed exclusively for microscopic ex- amination of material, e.g., x-ray dif- fraction, spectroscopic, and electron microscope equipment or to systems for intentional exposure of humans to x-rays.

  - Definitions. As used in this section the following definitions apply:

    - Access panel means any barrier or panel which is designed to be removed or opened for maintenance or service purposes, requires tools to open, and permits access to the interior of the cabinet.

    - Aperture means any opening in the outside surface of the cabinet, other than a port, which remains open during generation of x radiation.

    - Cabinet x-ray system means an x- ray system with the x-ray tube in- stalled in an enclosure (hereinafter termed cabinet) which, independently of existing architectural structures  ex- cept the floor on which it may  be placed, is intended to contain at least that portion of a material being irradi- ated,  provide  radiation  attenuation, and exclude personnel from its interior

#### Food and Drug Administration, HHS	§ 1020.40

during generation of x radiation. In- cluded are all x-ray systems designed primarily for the inspection of carry-on baggage at airline, railroad, and bus terminals, and in similar facilities. An x-ray tube used within a shielded part of a building, or x-ray equipment which may temporarily or occasionally incor- porate portable shielding is not consid- ered a cabinet x-ray system.

    - Door means any barrier which is designed to be movable or opened for routine operation purposes, does not generally require tools to open, and permits access to the interior of the cabinet. For the purposes of paragraph (c)(4)(i) of this section, inflexible hard- ware rigidly affixed to the door shall be considered part of the door.

    - Exposure means the quotient of dQ by dm where dQ is the absolute value of the total charge of the ions of one sign produced in air when all the elec- trons (negatrons and positrons) liber- ated by photons in a volume element of air having mass dm are completely stopped in air.

    - External surface means the outside surface of the cabinet x-ray system, in- cluding the high-voltage generator, doors, access panels, latches, control knobs, and other permanently mounted hardware and including the plane across any aperture or port.

    - Floor means the underside exter- nal surface of the cabinet.

    - Ground fault means an accidental electrical grounding of an electrical conductor.

    - Port means any opening in the outside surface of the cabinet which is designed to remain open, during gen- eration of x-rays, for the purpose of conveying material to be irradiated into and out of the cabinet, or for par- tial insertion for irradiation of an ob- ject whose dimensions do not permit complete insertion into the cabinet.

    - Primary beam means the x radi- ation emitted directly from the from the target and passing through the window of the x-ray tube.

    - Safety interlock means a device which is intended to prevent the gen- eration of x radiation when access by any part of the human body to the in- terior of the cabinet x-ray system through a door or access panel is pos- sible.

    - X-ray system means an assem- blage of components for the controlled generation of x-rays.

    - X-ray tube means any electron tube which is designed for the conver- sion of electrical energy into x-ray en- ergy.

  - Requirements—(1) Emission limit. (i) Radiation emitted from the cabinet x- ray system shall not exceed an expo- sure of 0.5 milliroentgen in one hour at any point five centimeters outside the external  surface.

  - Compliance with the exposure limit in paragraph (c)(1)(i) of this sec- tion shall be determined by measure- ments averaged over a cross-sectional area of ten square centimeters with no linear dimension greater than 5 centi- meters, with the cabinet x-ray system operated at those combinations of x- ray tube potential, current, beam ori- entation, and conditions of scatter ra- diation which produce the maximum x- ray exposure at the external surface, and with the door(s) and access panel(s) fully closed as well  as  fixed  at  any other position(s) which will allow the generation of x radiation.

- Floors. A cabinet x-ray system shall have a permanent floor. Any sup- port surface to which a cabinet x-ray system is permanently affixed may be deemed the floor of the system.

- Ports and apertures. (i) The inser- tion of any part of the human body through any port into the primary beam shall not be possible.

(ii) The insertion of any part of the human  body  through  any  aperture shall not be possible.

- Safety interlocks. (i) Each door of a cabinet x-ray system shall have a min- imum of  two  safety  interlocks.  One, but not both of the required interlocks shall be such that door opening results in physical disconnection of the energy supply circuit to the high-voltage gen- erator, and  such  disconnection  shall not be dependent upon any moving part other than the door.

- Each access  panel  shall  have  at least one safety interlock.

- Following interruption of x-ray generation by the functioning of any safety interlock, use of a control pro- vided  in  accordance  with  paragraph

#### § 1020.40

(c)(6)(ii) of this section shall be nec- essary for resumption of x-ray genera- tion.

- Failure of any single  component of the cabinet x-ray system shall not cause failure of more than one required safety   interlock.

- Ground fault. A ground fault shall not result in the generation of x-rays.

- Controls and indicators for all cabi- net x-ray systems. For all systems to which this section is applicable there shall be provided:

  - A key-actuated control to  insure that x-ray generation is not possible with the key removed.

  - A control or controls to initiate and terminate the generation of x-rays other than by functioning of a safety interlock or the main power control.

  - Two independent means which indicate when and only when x-rays are being generated, unless the x-ray gen- eration period is less than one-half sec- ond, in which case the indicators shall be activated for one-half second, and which are discernible from any point at which initiation of x-ray generation is possible. Failure of a single component of the cabinet x-ray system shall not cause failure of both indicators to per- form their intended function. One, but not both, of the indicators required by this subdivision may be a milli- ammeter labeled to indicate x-ray tube current. All other indicators shall be legibly  labeled  ‘‘X-RAY  ON’’.

  - Additional means other than milliammeters which indicate when and only  when  x-rays  are  being  gen- erated, unless the x-ray generation pe- riod is less than one-half second in which case the indicators shall be acti- vated for one-half second, as needed to insure that at least one indicator is visible from each door, access panel, and port, and is legibly labeled ‘‘X- RAY  ON’’.

- Additional controls and indicators for cabinet x-ray systems  designed  to admit humans. For cabinet x-ray sys- tems designed to admit humans there shall also be provided:

  - A control within the cabinet for preventing and terminating x-ray gen- eration, which cannot be reset, over- ridden or bypassed from the outside of the  cabinet.

#### 21 CFR Ch. I (4–1–17 Edition)

  - No means by which x-ray genera- tion can be initiated from within the cabinet.

  - Audible and visible warning sig- nals within the cabinet which are actu- ated for at least 10 seconds imme- diately prior to the first initiation of x- ray generation after closing any door designed to admit humans. Failure of any single component of the cabinet x- ray system shall not cause failure of both the audible and visible warning signals.

  - A visible warning signal within the cabinet which remains actuated when and only when x-rays are being generated, unless the x-ray generation period is less than one-half second in which case the indicators shall be acti- vated for one-half second.

  - Signs indicating the meaning of the warning signals provided pursuant to paragraphs (c)(7) (iii) and (iv) of this section and containing instructions for the use of the control provided pursu- ant to paragraph (c)(7)(i) of this sec- tion. These signs shall be legible, ac- cessible to view, and illuminated when the main power control is in the ‘‘on’’ position.

- Warning labels. (i) There shall be permanently affixed or inscribed on the cabinet x-ray system at the location of any controls which can be used to ini- tiate x-ray generation, a clearly legible and visible label bearing the state- ment:

CAUTION: X-RAYS PRODUCED WHEN ENERGIZED

(ii) There shall be permanently af- fixed or inscribed on the cabinet x-ray system adjacent to each port a clearly legible and visible label bearing the statement:

CAUTION: DO NOT INSERT ANY PART OF THE BODY WHEN SYSTEM IS ENER- GIZED—X-RAY HAZARD

- Instructions. (i) Manufacturers of cabinet x-ray systems shall provide for purchasers, and to others upon request at a cost not to exceed the cost of prep- aration and distribution, manuals and instructions which shall include at least the following technical and safety information: Potential, current, and duty cycle ratings of the x-ray genera- tion equipment; adequate instructions

#### Food and Drug Administration, HHS	§ 1030.10

concerning any radiological safety pro- cedures and precautions which may be necessary because of unique features of the system; and a schedule of mainte- nance necessary to keep the system in compliance with this section.

(ii) Manufacturers of cabinet x-ray systems which are intended to be as- sembled or installed by the purchaser shall provide instructions for assembly, installation, adjustment and testing of the cabinet x-ray system adequate to assure that the system is in compli- ance with applicable provisions of this section when assembled, installed, ad- justed and tested as directed.

- Additional requirements for x-ray baggage inspection systems. X-ray sys- tems designed primarily for the inspec- tion of carry-on baggage at airline, railroad, and bus terminals, and at similar  facilities,  shall  be  provided with means, pursuant to paragraphs (c)(10) (i) and (ii) of this section, to in- sure operator presence at the control area in a position which permits sur- veillance of the ports and doors during generation of x-radiation.

- During an exposure or preset suc- cession of exposures of one-half second or greater duration, the means pro- vided shall enable the operator to ter- minate the exposure or preset succes- sion of exposures at any time.

- During an exposure or preset suc- cession of exposures of less than one- half second  duration,  the  means pro- vided may allow completion of the ex- posure in progress but shall enable the operator to prevent additional expo- sures.

  - Modification of a certified system. The modification of a cabinet  x-ray system,   previously   certified   pursuant to § 1010.2 by any person engaged in the business  of  manufacturing,  assembling or  modifying  cabinet  x-ray  systems shall be construed as  manufacturing under the act  if  the  modification  af- fects any aspect of the system’s per- formance for which this section has an applicable requirement. The manufac- turer who performs such modification shall recertify  and  reidentify  the  sys- tem in accordance with  the  provisions of §§ 1010.2 and 1010.3 of this chapter.

[39 FR 12986, Apr. 10, 1974]

## PART	1030—PERFORMANCE STANDARDS FOR MICROWAVE AND RADIO FREQUENCY EMIT- TING PRODUCTS

AUTHORITY: 21 U.S.C. 351, 352, 360, 360e–360j,

360hh–360ss, 371, 381.

##### § 1030.10   Microwave ovens.

- Applicability. The provisions of this standard are applicable to micro- wave ovens manufactured after October 6, 1971.

- Definitions—(1) Microwave oven means a device designed  to  heat,  cook, or dry food through the application of electromagnetic energy at frequencies assigned by the Federal Communica- tions Commission in the normal ISM heating bands ranging from 890 mega- hertz to 6,000 megahertz. As defined in this standard, ‘‘microwave ovens’’ are limited to those manufactured  for  use in homes, restaurants, food vending, or service establishments, on interstate carriers,  and  in  similar  facilities.

- Cavity means that portion of the microwave oven in which food may be heated, cooked, or dried.

- Door means the movable barrier which prevents access to the cavity during operation and whose function is to prevent emission of microwave en- ergy from the passage or opening which provides access to the cavity.

- Safety interlock means a device or system of devices which is intended to prevent generation of microwave en- ergy when  access  to  the  cavity is  pos- sible.

- Service adjustments or service proce- dures means those servicing methods prescribed by the manufacturer for a specific product model.

- Stirrer means that feature of a microwave oven which is intended to provide uniform heating of the load by constantly changing the standing wave pattern  within  the  cavity  or  moving the load.

- External surface means the outside surface of the cabinet or enclosure pro- vided by the manufacturer as part  of the microwave oven, including doors, door handles, latches, and control knobs.

- Equivalent plane-wave power den- sity means the square of the root-mean-

§ 1030.10

square (rms) electric field strength di- vided by the impedance of free space (377 ohms).

- Requirements—(1) Power density limit. The equivalent plane-wave power density existing in the proximity of the external oven surface shall not exceed 1 milliwatt per square centimeter at any point 5 centimeters or more from the external surface of the oven, measured prior to acquisition  by  a  purchaser, and, thereafter, 5 milliwatts per square centimeter at any such point.

  - Safety interlocks. (i) Microwave ovens shall have a minimum of two op- erative safety interlocks. At least one operative safety interlock on a fully assembled microwave oven shall not be operable by any part of  the  human body, or any object with a straight insertable length of 10 centimeters. Such interlock must also be concealed, unless its actuation is prevented when access to the interlock is possible. Any visible actuator or device to prevent actuation of this safety interlock must not be removable without  disassembly of the oven or its door. A magnetically operated interlock is considered to be concealed, or its actuation is consid- ered to be prevented, only if a test magnet held in place on the oven by gravity or its own attraction cannot operate the safety interlock. The test magnet shall be capable of lifting vertically at zero air gap at least 4.5 kilograms, and at 1 centimeter air gap at least 450 grams when the face of the magnet, which is toward the interlock when the magnet is in the test posi- tion, is pulling against one of the large faces of a mild steel armature having dimensions of 80 millimeters by 50 mil- limeters by 8 millimeters.

- Failure of any single mechanical or electrical component of the micro- wave oven shall not cause all safety interlocks to be inoperative.

- Service adjustments or service procedures on the microwave oven shall not cause the safety interlocks to become inoperative or the microwave radiation emission to exceed the power density limits of this  section as a  re- sult of such service adjustments or pro- cedures.

- Microwave radiation emission in excess of the limits specified in para- graph (c)(1) of this section shall not be

#### 21 CFR Ch. I (4–1–17 Edition)

caused by insertion of  an  insulated wire through any opening in the exter- nal surfaces of a fully assembled oven into the cavity, waveguide, or other microwave-energy-containing spaces while the door is closed, provided the wire, when inserted, could consist of two straight segments forming an ob- tuse angle of not less than 170 degrees.

- One (the primary) required safety interlock shall prevent microwave ra- diation emission in excess of the re- quirement of paragraph (c)(1) of this section; the other (secondary) required safety interlock  shall  prevent  micro- wave radiation emission in excess of 5 milliwatts  per  square   centimeter   at any point 5 centimeters or more  from the external  surface  of  the  oven.  The two required safety interlocks shall be designated as primary or secondary in the  service  instructions  for  the  oven.

- A means of monitoring one or both of the required safety interlocks shall be provided which shall cause the oven to become inoperable and remain so until repaired if the required safety interlock(s) should fail to perform re- quired functions as specified in this section. Interlock failures shall not disrupt the monitoring function.

  - Measurement and test conditions. (i) Compliance  with  the  power  density limit in paragraph (c)(1) of this section shall be determined by measurement of the equivalent plane-wave power den- sity made with an instrument which reaches 90 percent of its steady-state reading within 3 seconds, when the sys- tem  is  subjected  to  a  step-function input signal. Tests for compliance shall account for all measurement errors and uncertainties to ensure that the equiv- alent plane-wave  power  density  does not exceed the limit prescribed by paragraph (c)(1) of this section.

- Microwave ovens shall be in com- pliance with the power density limits if the maximum reading obtained at the location of greatest microwave radi- ation emission, taking into account all measurement errors and uncertainties, does not exceed the limit specified in paragraph (c)(1) of this section, when the emission is measured through at least one stirrer cycle. As provided in

§ 1010.13 of this chapter, a manufacturer

#### Food and Drug Administration, HHS	§ 1030.10

may request alternative test proce- dures if, as a result of the stirrer char- acteristics of a microwave oven, such oven is not susceptible to testing  by the procedures described in this para- graph.

- Measurements shall be made with the microwave oven operating  at its maximum output and containing a load of 275 15 milliliters of tap water initially at 20 5 centigrade placed within the cavity at the center of the load-carrying surface provided by the manufacturer. The water  container shall be a low form 600-milliliter beak- er having an inside diameter of ap- proximately 8.5 centimeters and  made of an electrically nonconductive mate- rial such as glass or plastic.

- Measurements shall be made with the door fully closed as well as with the door fixed in any other posi- tion which allows the oven to operate.

  - User instructions. Manufacturers of microwave ovens to which this section is applicable shall provide, or cause to be provided, with each oven, radiation safety instructions which:

    - Occupy a separate section and are an integral part of the regularly sup- plied users’ manual and cookbook, if supplied separately, and are located so as to elicit the attention of the reader.

    - Are as legible and durable as other instructions with the title em- phasized to elicit the attention of the reader by such means as bold-faced type, contrasting color, a heavy-lined border, or by similar means.

    - Contain the following wording:

PRECAUTIONS TO AVOID POSSIBLE EXPOSURE TO EXCESSIVE MICROWAVE ENERGY

- Do not attempt to operate this oven with the door open since open-door operation can result in harmful exposure to microwave energy. It is important not to defeat or tam- per with the safety interlocks.

- Do not place any object between the oven front face and the door or allow soil or cleaner residue to accumulate on sealing surfaces.

- Do not operate the oven if it is dam- aged. It is particularly important that the oven door close properly and that there is no damage to the: (1) Door (bent), (2) hinges and latches (broken or loosened), (3) door seals and sealing surfaces.

- The oven should not be adjusted or re- paired by anyone except properly qualified service personnel.

    - Include additional radiation  safe- ty precautions or instructions  which may be necessary for particular oven designs or models, as determined by the Director, Center for  Devices and Radiological Health or the manufac- turer.

  - Service instructions. Manufacturers of microwave ovens to which this sec- tion is applicable shall provide or cause to be provided to servicing dealers and distributors and to others upon re- quest, for each oven model, adequate instructions for service adjustments and service procedures, and, in addi- tion, radiation safety instructions which:

    - Occupy a separate section and are an integral part of the regularly sup- plied service manual and are located so as to elicit the attention of the reader.

    - Are as legible and durable as other instructions with the title em- phasized so as to elicit the attention of the reader by such means as bold-faced type, contrasting color, a heavy-lined border, or by similar means.

    - Contain the following wording:

PRECAUTIONS TO BE OBSERVED BEFORE AND DURING SERVICING TO AVOID POSSIBLE EX- POSURE TO EXCESSIVE MICROWAVE ENERGY

- Do not operate or allow the oven to be operated with the door open.

- Make the following safety checks on all ovens to be serviced before activating the magnetron or other microwave source, and make repairs as necessary: (1) Interlock op- eration, (2) proper door closing, (3) seal and sealing surfaces (arcing, wear, and other damage), (4) damage to or loosening of hinges and latches, (5) evidence of dropping or abuse.

- Before turning on microwave power for any service test or inspection within the microwave generating compartments, check the magnetron, wave guide or transmission line, and cavity for proper alignment, integ- rity, and connections.

- Any defective or misadjusted compo- nents in the interlock, monitor, door seal, and microwave generation and transmission systems shall be repaired, replaced, or ad- justed by procedures described in this man- ual before the oven is released to the owner.

- A Microwave leakage check to verify compliance with the Federal performance standard should be performed on each oven prior to release to the owner.

    - Include additional radiation safe- ty precautions or instructions which may be necessary for particular oven

#### Pt. 1040

designs or models, as determined by the Director, Center for  Devices and Radiological Health or the manufac- turer.

  - Warning labels. Except as provided in paragraph (c)(6)(iv) of this section, microwave ovens  shall  have the  fol- lowing warning labels:

    - A label, permanently attached to or inscribed on the oven, which shall be legible and readily viewable during normal oven use, and which shall have the title emphasized and be so located as to elicit the attention of the user. The label shall bear the following warning statement:

PRECAUTIONS FOR SAFE USE TO AVOID POS- SIBLE EXPOSURE TO EXCESSIVE MICROWAVE ENERGY

DO NOT Attempt to Operate This Oven With:

- Object Caught in Door.

- Door That Does Not Close Properly.

- Damaged Door, Hinge, Latch, or Sealing Surface.

    - A label, permanently attached to or inscribed on the external surface of the oven, which shall be legible and readily viewable during servicing, and which shall have the word ‘‘CAUTION’’ emphasized and be so located as to elic- it the attention of service personnel. The label shall bear the following warning statement:

CAUTION: This Device  is  to  be  Serviced Only by Properly Qualified Service Per- sonnel. Consult the Service Manual for Prop- er Service Procedures to Assure Continued Compliance with the Federal Performance Standard for Microwave Ovens and for Pre- cautions to be Taken to Avoid Possible Ex- posure to Excessive Microwave Energy.

    - The labels provided in accord- ance with paragraphs (c)(6)(i) and (ii) of this section shall bear only the state- ments specified in that paragraph, ex- cept for additional radiation safety warnings or instructions which may be necessary for particular oven designs or models, as determined by the Direc- tor, Center for Devices and Radio- logical Health or the manufacturer.

    - Upon application by a manufac- turer, the Director, Center for Devices and Radiological Health, Food and Drug Administration, may grant an ex- emption from one or more of the state- ments (radiation safety warnings) spec- ified in paragraph (c)(6)(i) of this sec-

#### 21 CFR Ch. I (4–1–17 Edition)

tion. Such exemption shall be  based upon a determination  by  the  Director that the microwave oven model for which the exemption is sought should continue to comply with paragraphs (c) (1), (2), and (3) of this section under the adverse condition of use addressed by such precautionary statement(s). An original and two copies of applications shall be submitted to the Division of Dockets Management (HFA–305), Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852. Copies of the written portion of the ap- plication,   including   supporting   data and information, and the Director’s ac- tion on the application will be  main- tained by the Branch for public review. The  application  shall  include:

- The specific microwave oven model(s) for which the exemption is sought.

- The specific radiation safety warning(s) from which exemption is sought.

- Data and information which clear- ly establish that one or more of the ra- diation safety warnings in paragraph (c)(6)(i) of this section is not necessary for the specified microwave oven model(s).

- Such other information and a sample of the applicable product if re- quired by regulation or by the Direc- tor, Center for Devices and Radio- logical Health, to evaluate and act on the application.

[38 FR 28640, Oct. 15, 1973, as amended at 40

FR  14752,  Apr.  4,  1975;  40  FR  52007,  Nov.  7,

1975;  46  FR 8461,  Jan. 27,  1981;  48  FR 57482,

Dec. 30, 1983; 50 FR 13566, Apr. 5, 1985; 53 FR

11254, Apr. 6, 1988; 59 FR 14365, Mar. 28, 1994]

## PART	1040—PERFORMANCE STANDARDS FOR LIGHT-EMITTING PRODUCTS

Sec.

- Laser products.

- Specific purpose laser products.

1040.20 Sunlamp products and  ultraviolet lamps intended for use in sunlamp prod- ucts.

- High-intensity  mercury  vapor  dis- charge lamps.

AUTHORITY: 21 U.S.C. 351, 352, 360, 360e–360j,

360hh–360ss, 371, 381.

#### Food and Drug Administration, HHS	§ 1040.10

##### § 1040.10   Laser products.

  - Applicability. The provisions of this section and § 1040.11, as amended, are applicable as specified to all laser products manufactured or assembled after August 1, 1976, except when:

- Such a laser product is either sold to a manufacturer of an electronic product for use as a component (or re- placement) in such electronic product, or

- Sold by or for a manufacturer of an electronic product for use as a com- ponent (or replacement) in such elec- tronic product, provided that such laser product:

  - Is accompanied by a general warn- ing notice that adequate instructions for the safe installation of the laser product are provided in servicing infor- mation available from the complete laser product manufacturer under para- graph (h)(2)(ii) of this section, and should be followed,

  - Is labeled with  a  statement  that it is designated for use solely as a com- ponent of such electronic product and therefore does not comply with the ap- propriate requirements of  this  section and § 1040.11 for complete laser prod- ucts,  and

  - Is not a removable  laser  system as described in paragraph (c)(2) of this section; and

- The manufacturer of such a laser product, if manufactured after August 20, 1986:

  - Registers,  and  provides  a  listing by type of such laser products manu- factured  that  includes  the   product name, model number and  laser  medium or  emitted  wavelength(s),   and   the name and address of the manufacturer. The manufacturer must submit the registration  and   listing   to   the   Food and Drug Administration, Center for Devices and Radiological Health, Di- rector, Office of Compliance, 10903 New Hampshire Ave., Bldg. 66, rm. 3521,  Sil- ver Spring, MD 20993–0002.

  - Maintains and allows  access  to any sales, shipping, or distribution records that identify the purchaser of such a laser product by name and ad- dress, the product by  type,  the  number of units sold, and the date of sale (ship- ment). These records shall be main- tained and made  available  as  specified in § 1002.31.

  - Definitions. As used in this section and § 1040.11, the following definitions apply:

    - Accessible emission level means the magnitude of accessible laser or collat- eral radiation of a specific wavelength and emission duration at a particular point as measured according to para- graph (e) of this section.  Accessible laser or collateral radiation is  radi- ation to which human access is pos- sible, as defined in paragraphs (b) (12), (15), and (22) of this section.

    - Accessible emission limit means the maximum accessible emission level permitted within a particular class as set forth in paragraphs (c), (d), and (e) of this section.

    - Aperture means any opening in the protective housing or other enclosure of a laser product through which laser or collateral radiation is emitted, thereby allowing human access to such radiation.

    - Aperture stop means an opening serving to limit the size and  to  define the shape of the area over which radi- ation is measured.

    - Class I laser product means any laser product that does not permit ac- cess during the operation to levels of laser radiation in excess of the acces- sible emission limits contained in table I of paragraph (d) of this section. 1

    - Class IIa laser product means any laser product that permits human ac- cess during  operation to levels  of visi- ble laser radiation in excess of the ac- cessible emission limits contained in table I, but does not permit human ac- cess during operation to levels of laser radiation in excess of the accessible emission limits contained in  table  II-A of paragraph (d) of this section. 2

    - Class  II  laser  product  means  any laser  product  that  permits  human  ac- cess during operation to levels of visi- ble laser  radiation in  excess of the  ac- cessible   emission   limits   contained   in table  II-A,  but  does  not  permit  human access   during   operation   to   levels   of

1 Class I levels of laser radiation are not considered to be hazardous.

2 Class IIa levels of laser radiation are not considered to be hazardous if viewed for any period of time less than or equal to 1 103 seconds but are considered to be a chronic viewing hazard for any period of time greater than 1 10 3 seconds.

#### § 1040.10

21 CFR Ch. I (4–1–17 Edition)

laser radiation in excess of the acces- sible emission limits contained in table II of paragraph (d) of this section. 3

    - Class IIIa laser product means any laser product that permits human ac- cess during  operation to levels  of visi- ble laser radiation in excess of the ac- cessible emission limits contained in table II, but does not permit human ac- cess during operation to levels of laser radiation in excess of the accessible emission  limits  contained  in  table  III-

A of paragraph (d) of this section. 4

    - Class IIIb laser product means any laser product that permits human ac- cess during operation to levels of laser radiation in excess of the accessible emission limits  of table  III-A, but  does not permit human access during oper- ation to levels of laser radiation in ex- cess of the accessible emission limits contained  in  table  III-B  of  paragraph

(d) of this section. 5

    - Class III laser product means any Class IIIa or Class IIIb laser product.

    - Class IV laser product means any laser that permits human access during operation to levels of laser radiation in excess of the accessible emission limits contained  in  table  III-B  of  paragraph

- of this section. 6

    - Collateral radiation means any electronic product radiation, except laser radiation, emitted by a laser product as a result of the operation of the laser(s) or any component of the laser product that is physically nec- essary for the operation of the laser(s).

    - Demonstration laser product means any  laser  product  manufactured,  de- signed, intended, or promoted for pur- poses of demonstration, entertainment, advertising  display,  or  artistic  com- position.   The   term   ‘‘demonstration laser product’’ does not apply to laser products which are not manufactured,

3 Class II levels of laser radiation are con- sidered to be a chronic viewing hazard.

4 Class IIIa levels of laser radiation are con- sidered to be, depending upon the irradiance, either an acute intrabeam viewing hazard or chronic viewing hazard, and an acute view- ing hazard if viewed directly with optical in- struments.

5 Class IIIb levels of laser radiation are con- sidered to be an acute hazard to the skin and eyes from direct radiation.

6 Class IV levels of laser radiation are con- sidered to be an acute hazard to the skin and eyes from direct and scattered radiation.

designed, intended, or promoted for such purposes, even though they  may be used for those purposes or are in- tended to demonstrate other applica- tions.

    - Emission duration means the tem- poral duration of a pulse, a series of pulses, or continuous operation, ex- pressed in seconds, during which human access to laser or collateral ra- diation could be permitted as a result of operation, maintenance, or  service of a laser product.

    - Human access means the capacity to intercept laser or collateral radi- ation by any part of the human body. For laser products that contain Class IIIb or IV levels of laser radiation, ‘‘human access’’ also means access to laser radiation that can be reflected di- rectly by any single introduced flat surface from the interior of the product through any opening in the protective housing of the product.

    - Integrated radiance means radiant energy per unit area of a radiating sur- face per  unit  solid  angle  of  emission, expressed in joules per square centi- meter per steradian (Jcm¥2  sr¥1).

    - Invisible radiation means laser or collateral  radiation   having   wave- lengths of equal to or greater  than  180 nm but less than or equal to 400 nm or greater than 710 nm but less than or equal to 1.0 106  nm (1 millimeter).

    - Irradiance means the time-aver- aged radiant power incident on an ele- ment of a surface divided by the area of that element, expressed in watts per square centimeter (W cm¥2).

    - Laser means any device that  can be made to produce or amplify electro- magnetic radiation  at  wavelenghts greater than 250 nm but less than or equal to 13,000 nm or, after August 20, 1986, at wavelengths equal to or greater than  180  nm  but  less  than  or  equal  to

1.0 106 nm primarily by the process of controlled  stimulated  emission.

    - Laser energy source means any de- vice intended for use in conjunction with a laser to supply energy for the operation of the laser. General energy sources such as electrical supply mains or batteries shall not be considered to constitute laser energy sources.

#### Food and Drug Administration, HHS	§ 1040.10

    - Laser product means any manu- factured product or assemblage of com- ponents which constitutes, incor- porates, or is intended to incorporate a laser or laser system. A laser or laser system that is intended for use as a component of an electronic product shall itself be considered a laser prod- uct.

    - Laser radiation means all electro- magnetic radiation emitted by a laser product within the spectral range spec- ified in paragraph (b)(19) of this section that is produced as a result of con- trolled stimulated emission or that is detectable with radiation so produced through the appropriate aperture stop and within the appropriate  solid  angle of acceptance, as specified in paragraph

- of this section.

    - Laser system means a laser in combination with an appropriate laser energy source with or without addi- tional incorporated components. See paragraph (c)(2) of this section for an explanation of the term ‘‘removable laser  system.’’

    - Maintenance means performance of those adjustments or procedures specified in user information provided by the manufacturer with the laser product which are to be performed by the user for the purpose of assuring the intended performance of the product. It does not include operation or service as defined in paragraph (b) (27) and (38) of this  section.

    - Maximum output means the max- imum radiant power and, where appli- cable, the maximum radiant energy per pulse of accessible laser radiation emitted by a laser product during oper- ation,  as  determined  under  paragraph

- of this section.

    - Medical laser product means any laser product which is a medical device as defined in 21 U.S.C. 321(h) and is manufactured, designed, intended or promoted for in vivo laser irradiation of any part of the human body for the purpose of: (i) Diagnosis, surgery, or therapy; or (ii) relative positioning of the human body.

    - Operation means the performance of the laser product over the full range of its functions. It does not include maintenance or service as defined in paragraphs (b) (24) and (38) of this sec- tion.

    - Protective housing means those portions of a laser product which are designed to prevent human access to laser or collateral radiation  in  excess of the prescribed accessible emission limits under conditions  specified  in this section and in § 1040.11.

    - Pulse duration means the time in- crement measured between the half- peak-power points at the leading and trailing edges of a pulse.

    - Radiance means time-averaged radiant power per unit area of a radi- ating surface per unit solid angle of emission, expressed in watts per square centimeter per  steradian  (W  cm¥2 sr¥1).

    - Radiant energy means energy emitted, transferred  or  received  in the form of radiation, expressed in joules (J).

    - Radiant exposure means the radi- ant energy incident on an element of a surface divided by the area of the ele- ment, expressed in joules per square centimeter  (Jcm¥2)

    - Radiant power means time-aver- aged power emitted, transferred or re- ceived in the form of radiation, ex- pressed in watts (W).

    - Remote interlock connector means an electrical connector which permits the connection of external remote interlocks.

    - Safety interlock means a device associated with the protective  housing of a laser product to prevent human ac- cess to excessive radiation in accord- ance with paragraph (f)(2) of this sec- tion.

    - Sampling interval means the time interval during which the level of ac- cessible laser or collateral radiation is sampled by a measurement  process. The magnitude of the sampling inter- val in units of seconds  is  represented by the symbol (t).

    - Scanned laser radiation means laser radiation having a time-varying direction, origin or pattern of propaga- tion with respect to a stationary frame of reference.

    - Service means the performance of those procedures or adjustments de- scribed in the manufacturer’s service instructions which may affect any as- pect of the product’s performance for

#### § 1040.10

which this section and § 1040.11 have ap- plicable requirements. It does not  in- clude maintenance or operation as de- fined in paragraphs (b) (24) and (27) of this   section.

    - Surveying, leveling, or alignment laser product means a laser product manufactured, designed, intended or promoted for one or more of the fol- lowing uses:

- Determining and delineating the form, extent, or position of a point, body, or area by taking angular meas- urement.

- Positioning or adjusting parts in proper relation to one another.

- Defining a plane, level, ele- vation, or straight line.

    - Visible radiation means laser or collateral radiation  having  wave- lengths of greater than 400 nm but less than or equal to 710 nm.

    - Warning logotype means a logo- type as illustrated in either figure 1 or figure 2 of paragraph (g) of this section.

    - Wavelength means the propaga- tion wavelength in air of electro- magnetic radiation.

  - Classification of laser products—(1) All laser products. Each laser product shall be classified in Class I, IIa, II, IIIa, IIIb, or IV in accordance with defi- nitions set forth in paragraphs (b) (5) through (11) of this section. The prod- uct classification shall be based on the highest accessible emission level(s) of

#### 21 CFR Ch. I (4–1–17 Edition)

laser radiation to  which  human  access is possible during operation in accord- ance with paragraphs (d), (e), and (f)(1) of this section.

(2) Removable laser systems. Any laser system that is incorporated into  a laser product subject to the require- ments of this section and that is capa- ble, without modification, of producing laser radiation when removed from such laser product, shall itself be con- sidered a laser product and shall be separately subject to the applicable re- quirements in this subchapter for laser products of its class. It shall be classi- fied on the basis of accessible emission of laser radiation when so removed.

  - Accessible emission limits. Acces- sible emission limits for laser radiation in each class are specified in tables I, II-A, II, III-A, and III-B of this para- graph. The factors, k1 and k2 vary with wavelength and emission duration. These factors are given in table IV of this paragraph, with selected numer- ical values in table V of this paragraph. Accessible emission limits for collat- eral radiation are specified in table VI of this paragraph.

NOTES APPLICABLE TO TABLES I, II-A, II, III-

A AND III-B: (1) The factors k1 and k2 are wavelength-dependent correction factors de- termined from table IV.

(2) The variable t in the expressions of emission limits is the magnitude of the sam- pling interval in units of seconds.

#### Food and Drug Administration, HHS	§ 1040.10

§ 1040.10

21 CFR Ch. I (4–1–17 Edition)

Food and Drug Administration, HHS	§ 1040.10

§ 1040.10

21 CFR Ch. I (4–1–17 Edition)

Food and Drug Administration, HHS	§ 1040.10

§ 1040.10

21 CFR Ch. I (4–1–17 Edition)

- Beam of a single wavelength. Laser or collateral radiation of a single wave- length exceeds the accessible emission limits of a class if its accessible emis- sion level is greater than the accessible emission limit of that class within any of the ranges of emission duration specified in tables I, II-A, II, III-A, and III-B of this paragraph.

- Beam of multiple wavelengths in same range. Laser or collateral radi- ation having two or more wavelengths within any one of the wavelength ranges specified in tables I, II-A, II, III- A, and III-B of this paragraph exceeds the accessible emission limits of  a class if the sum of the ratios of the ac- cessible   emission   level   to   the   cor-

responding accessible emission limit at each such wavelength is greater than unity for that combination of emission duration and wavelength distribution which results in the maximum sum.

- Beam with multiple wavelengths in different ranges. Laser or collateral ra- diation having wavelengths within two or more of the wavelength ranges spec- ified in tables I, II-A, II, III-A, and III- B of this paragraph exceeds the acces- sible emission limits of a class if it ex- ceeds the applicable limits within any one of those wavelength ranges. This determination is made for each wave- length range in accordance with para- graph (d) (1) or (2) of this section.

#### Food and Drug Administration, HHS	§ 1040.10

- Class I dual limits. Laser or collat- eral radiation in the  wavelength  range of greater than 400 nm but less than or equal to 1.400 nm exceeds the accessible emission limits of Class I if it exceeds both:

  - The Class I accessible emission limits for radiant energy within any range of emission duration specified in table I of this paragraph, and

  - The Class I accessible emission limits for integrated radiance within any range of emission duration speci- fied in table I of this paragraph.

  - Tests for determination of compli- ance—(1) Tests for certification. Tests on which certification under § 1010.2 is based shall account for all errors and statistical uncertainties in the meas- urement process. Because compliance with the standard is required for the useful life of a product such tests shall also account for increases in emission and degradation in radiation safety with age.

- Test conditions. Except  as  provided in § 1010.13, tests for compliance with each of the applicable requirements of this section and § 1040.11 shall be made during operation,  maintenance,  or service  as  appropriate:

  - Under those conditions and proce- dures which maximize the accessible emission levels, including start-up, sta- bilized emission, and shut-down of the laser product; and

  - With all controls and adjustments listed in the operation, maintenance, and service instructions adjusted in combination to result in the maximum accessible emission level of radiation; and

  - At points in space to which human access is possible in the product configuration which is necessary to de- termine compliance with each require- ment, e.g., if operation may require re- moval of portions of the protective housing and defeat of safety interlocks, measurements shall be made at points accessible in that product configura- tion; and

  - With the measuring instrument detector so positioned and so oriented with respect to the laser product as to result in the maximum detection of ra- diation by the instrument; and

  - For a laser product other than a laser system, with the laser coupled to

that type of laser energy source which is specified as compatible by the laser product manufacturer and which pro- duces the maximum emission level of accessible radiation from that product.

- Measurement parameters. Acces- sible emission levels of laser and col- lateral radiation shall be based upon the following measurements as appro- priate, or their equivalent:

- For laser products intended to be used in a locale  where  the  emitted laser radiation is unlikely to be viewed with optical instruments, the radiant power (W) or radiant energy (J) detect- able through a circular aperture stop having a diameter of 7 millimeters and within a circular solid angle of accept- ance of 1 10¥3 steradian with colli- mating optics of 5 diopters or less. For scanned  laser  radiation,  the   direction of the solid angle of acceptance shall change as needed to maximize detect- able  radiation,  with  an   angular   speed of up to 5 radians/second. A 50 milli- meter diameter aperture stop with the same  collimating  optics  and  accept- ance angle stated above shall  be  used for all  other  laser  products  (except  that a 7 millimeter diameter aperture stop shall be used in the measurement of scanned laser  radiation  emitted  by laser products manufactured on or be- fore August 20, 1986.

- The irradiance (W cm¥2) or radi- ant exposure (J cm¥2 equivalent to the radiant power (W) or radiant energy (J) detectable through a circular aperture stop having a diameter of 7 millimeters and, for irradiance, within a circular solid angle of acceptance of 1 10¥3 steradian with collimating optics of 5 diopters or less, divided by the area of the aperture stop (cm¥2).

- The radiance (W cm¥2 sr¥1) or integrated radiance (J cm¥2 sr¥1) equivalent to the radiant power (W) or radiant energy (J) detectable through a circular aperture stop having a diame- ter of 7 millimeters and within a cir- cular solid angle of acceptance of 1  10¥5 steradian with collimating  optics of 5 diopters or less, divided by that solid angle (sr) and by the area of the aperture stop (cm¥2).

  - Performance requirements—(1) Pro- tective housing. Each laser product shall have a protective housing that pre- vents human access during operation

#### § 1040.10

to laser and collateral radiation that exceed the limits of Class I and table VI, respectively, wherever and when- ever such human access is not nec- essary for the product to perform its intended function. Wherever and when- ever human access to laser radiation levels that exceed the limits of Class I is necessary, these levels shall not ex- ceed the limits of the lowest class nec- essary to perform the intended func- tion(s) of the product.

- Safety interlocks. (i) Each laser product, regardless of its class, shall be provided with at least one safety inter- lock for each portion of the protective housing which is designed to be re- moved or displaced during operation or maintenance, if removal or displace- ment of  the  protective housing  could permit, in the absence of such inter- lock(s), human access to laser or col- lateral radiation in excess of the acces- sible emission limit applicable under paragraph (f)(1) of this section.

  - Each required safety interlock, unless defeated, shall prevent such human access to laser and collateral radiation upon removal or displace- ment of such portion of the protective housing

  - Either multiple safety interlocks or a means to preclude removal or dis- placement of the interlocked portion of the protective housing shall be pro- vided, if failure of a single interlock would allow;

- Human access to a level of laser radiation in excess of the accessible emission limits of Class IIIa; or

- Laser radiation in excess of the accessible emission limits of Class II to be emitted directly through the open- ing created by removal or displacement of the interlocked portion of the pro- tective  housing.

  - Laser products that incorporate safety interlocks designed to  allow safety interlock defeat shall incor- porate a means of visual or aural indi- cation of interlock defeat. During interlock defeat, such indication shall be visible or audible whenever the laser product is energized, with and without the associated portion of the protective housing removed or displaced.

  - Replacement of a removed or dis- placed  portion  of  the  protective  hous-

#### 21 CFR Ch. I (4–1–17 Edition)

ing shall not be possible while required safety interlocks are defeated.

- Remote interlock connector. Each laser system classified as  a  Class  IIIb or IV laser product shall incorporate a readily available remote interlock con- nector having an electrical potential difference of no greater than 130 root- mean-square volts between terminals. When the terminals of the connector are not electrically joined, human ac- cess to all laser  and  collateral  radi- ation from the laser  product  in  excess of the accessible emission limits  of Class I and table VI shall be prevented.

- Key control.  Each laser system classified as a Class IIIb or IV laser product shall incorporate a key-actu- ated master control. The key shall be removable and the laser shall not be operable when the key is removed.

- Laser  radiation  emission  indicator.

- Each laser system classified as a Class II or IIIa laser product shall in- corporate an emission indicator that provides a visible or audible signal dur- ing emission of accessible laser radi- ation in excess of the accessible emis- sion limits of Class I.

- Each laser system classified as a Class IIIb or IV laser product shall in- corporate an emission indicator which provides a visible or audible signal dur- ing emission of accessible laser radi- ation in excess of the accessible emis- sion limits of Class I,  and sufficiently prior to emission of such radiation to allow appropriate action to avoid expo- sure to the laser radiation.

- For laser  systems  manufactured on or before August 20, 1986, if the laser and laser energy source are housed sep- arately and can be operated at a sepa- ration distance of greater than 2 me- ters, both laser and laser energy source shall incorporate an emission  indicator as required in accordance with para- graph (f)(5) (i) or (ii) of this  section. For laser systems manufactured after August 20, 1986, each separately housed laser and operation control of a laser system that regulates the laser or col- lateral radiation emitted by a product during operation shall incorporate an emission indicator as required in ac- cordance with paragraph (f)(5) (i) or (ii) of this section, if the laser or operation control can be operated at a separation distance   greater   than   2   meters   from

#### Food and Drug Administration, HHS	§ 1040.10

any other separately housed portion of the laser product incorporating an emission indicator.

- Any visible signal required by paragraph (f)(5) (i) or (ii) of this section shall be clearly visible through protec- tive eyewear designed specifically for the wavelength(s) of the emitted laser radiation.

- Emission indicators required by paragraph (f)(5) (i) or (ii) of this section shall be located so that viewing does not require human exposure to laser or collateral radiation in excess of the ac- cessible emission limits of Class I and table VI.

- Beam attenuator. (i) Each laser system classified as a  Class  II,  III,  or IV laser product shall be provided with one or more permanently attached means, other than laser energy source switch(es), electrical supply main con- nectors, or the key-actuated master control, capable of preventing access by any part of the human body to all laser and collateral radiation in excess of the accessible emission limits of Class I and table VI.

(ii) If the configuration, design, or function of the laser product would make unnecessary compliance with the requirement in paragraph (f)(6)(i) of this section, the Director, Office of Compliance (HFZ–300), Center for De- vices and Radiological Health, may, upon written application by the manu- facturer, approve alternate means to accomplish the radiation protection provided by the beam attenuator.

- Location of controls. Each Class IIa, II, III, or IV laser product shall have operational and adjustment con- trols located so that human exposure to laser or collateral radiation in ex- cess of the accessible emission limits of Class I and table VI is unnecessary for operation or adjustment of such con- trols.

- Viewing optics. All viewing optics, viewports, and display screens incor- porated into a laser product, regardless of its  class, shall  limit the  levels of laser and collateral radiation acces- sible to the human eye by means of such viewing optics, viewports, or dis- play screens during operation or main- tenance   to   less   than   the   accessible

emission limits of Class I and table VI. For any shutter or variable attenuator incorporated into such viewing optics, viewports, or display screens, a means shall be provided:

- To prevent access by the human eye to laser and collateral radiation in excess of the accessible emission limits of Class I and table VI whenever the shutter is opened or the attenuator varied.

- To preclude, upon failure of such means as required in paragraph (f)(8)(i) of this section, opening the shutter or varying the attenuator when access by the human eye is possible to laser or collateral radiation in excess of the ac- cessible emission limits of Class I and table VI.

- Scanning safeguard. Laser products that emit accessible scanned laser radi- ation shall not, as a result of any fail- ure causing a change in either scan ve- locity or amplitude, permit human ac- cess to laser radiation in excess of:

  - The accessible emission limits of the class of the product, or

  - The accessible emission limits of the class of the scanned laser radiation if the product is Class IIIb  or  IV  and the accessible emission limits of Class IIIa would be exceeded solely as result of such failure.

- Manual reset mechanism. Each laser system manufactured after  Au- gust 20, 1986, and classified as a Class IV laser product shall be provided with a manual reset to enable resumption of laser radiation emission after interrup- tion of emission caused by the use of a remote interlock or after an interrup- tion of emission in excess of 5 seconds duration due to the unexpected loss of main electrical power.

  - Labeling requirements. In  addition to the requirements of §§ 1010.2 and 1010.3, each laser product shall be sub- ject to the applicable labeling require- ments  of  this  paragraph.

    - Class IIa and II designations and warnings. (i) Each Class IIa laser prod- uct shall have affixed a label  bearing the following wording: ‘‘Class IIa Laser Product—Avoid Long-Term Viewing of Direct  Laser  Radiation.’’

#### § 1040.10

(ii) Each Class II laser product shall have affixed a label bearing the  warn- ing logotype A (figure 1 in this para- graph) and including the following wording:

#### 21 CFR Ch. I (4–1–17 Edition)

[Position I on the logotype] ‘‘LASER RADIATION—DO NOT STARE

INTO BEAM’’; and

[Position 3 on the logotype] ‘‘CLASS  II  LASER  PRODUCT’’.

    - Class IIIa and IIIb designations and warnings. (i) Each Class IIIa laser prod- uct with an irradiance less  than  or equal to 2.5 10¥3 W cm2¥ shall have affixed a label bearing the warning log- otype A (figure 1 of paragraph (g)(1)(ii) of this section) and including the fol- lowing wording:

[Position 1 on the logotype]   ‘‘LASER     RADIATION—DO     NOT     STARE

INTO  BEAM  OR  VIEW  DIRECTLY  WITH

OPTICAL INSTRUMENTS’’; and,

[Position 3 on the logotype] ‘‘CLASS  IIIa  LASER  PRODUCT’’.

- Each Class IIIa laser product with an irradiance greater than 2.5 10¥3  W

cm¥2 shall have affixed a label bearing the warning logotype B (figure 2 in this paragraph) and including the following wording:

[Position 1 on the logotype]  ‘‘LASER  RADIATION—AVOID  DIRECT  EYE

EXPOSURE’’; and,

[Position 3 on the logotype] ‘‘CLASS  IIIa  LASER  PRODUCT’’.

#### Food and Drug Administration, HHS	§ 1040.10

- Each Class IIIb laser product shall have affixed a label bearing the warning logotype B (figure 2 of para- graph (g)(2)(ii) of this section) and in- cluding the following wording:

[Position 1 on the logotype] ‘‘LASER    RADIATION—AVOID    DIRECT

EXPOSURE TO BEAM’’; and,

[Position 3 on the logotype] ‘‘CLASS  IIIb  LASER  PRODUCT’’.

    - Class IV designation and warning. Each Class IV laser product shall have affixed a label bearing the warning log- otype B (figure 2 of paragraph (g)(2)(ii) of this section), and including the fol- lowing wording:

[Position 1 on the logotype]

‘‘LASER RADIATION—AVOID EYE OR SKIN EXPOSURE TO DIRECT OR SCAT- TERED RADIATION’’; and,

[Position 3 on the logotype] ‘‘CLASS  IV  LASER  PRODUCT’’.

    - Radiation output information on warning logotype. Each Class II, III, and IV laser product shall state in appro- priate units, at position 2 on the re- quired warning logotype, the maximum output of laser radiation, the pulse du- ration when appropriate, and the laser medium or emitted wavelength(s).

    - Aperture label. Each laser product, except medical laser products and Class IIa laser products, shall have af- fixed, in close proximity to each aper- ture through which is emitted acces- sible laser or collateral radiation in ex- cess of the accessible emission limits of Class I and table VI of paragraph (d) of this section, a label(s) bearing the fol- lowing wording as applicable.

#### § 1040.10

      - ‘‘AVOID EXPOSURE—Laser radi- ation is emitted from this aperture,’’ if the radiation emitted through such ap- erture is laser radiation.

      - ‘‘AVOID EXPOSURE—Hazardous electromagnetic radiation is emitted from this aperture,’’ if the radiation emitted through such aperture is col- lateral radiation described in table VI, item 1.

      - ‘‘AVOID EXPOSURE—Hazardous x-rays are emitted from this aperture,’’ if the radiation emitted through such aperture is collateral radiation de- scribed in table VI, item 2.

    - Labels for noninterlocked protective housings. For each laser product, labels shall be provided for each portion  of the protective housing which has no safety interlock and which  is  designed to be displaced or removed during oper- ation, maintenance, or service, and thereby could permit human access to laser or collateral radiation  in  excess of the limits of Class I and table VI. Such labels shall be visible on the pro- tective housing prior to  displacement or removal of such portion of the pro- tective housing and visible on the prod- uct in close proximity to the opening created by removal or displacement of such portion of the protective housing, and shall include the wording:

      - ‘‘CAUTION—Laser radiation when open. DO NOT STARE INTO BEAM.’’ for Class II accessible laser radiation.

      - ‘‘CAUTION—Laser radiation when open. DO NOT STARE INTO BEAM OR VIEW DIRECTLY WITH OP- TICAL INSTRUMENTS.’’ for Class IIIa accessible laser radiation with an irra- diance less than or equal to 2.5 10¥3 W cm¥2.

      - ‘‘DANGER—Laser radiation when open. AVOID DIRECT EYE EX- POSURE.’’ for Class IIIa accessible laser radiation with an irradiance greater than 2.5 10¥3 W cm¥2.

      - ‘‘DANGER—Laser radiation when open. AVOID DIRECT EXPO- SURE TO BEAM.’’ for Class IIIb acces- sible laser radiation.

      - ‘‘DANGER—Laser radiation when open. AVOID EYE OR SKIN EXPO- SURE TO DIRECT OR SCATTERED RADIATION.’’ for Class IV accessible laser radiation.

      - ‘‘CAUTION—Hazardous electro- magnetic   radiation   when   open.’’   for

#### 21 CFR Ch. I (4–1–17 Edition)

collateral radiation in excess of the ac- cessible emission limits in  table  VI, item 1 of paragraph (d) of this section.

      - ‘‘CAUTION—Hazardous x-rays when open.’’ for collateral radiation in excess of the accessible emission limits in table VI, item 2 of paragraph (d) of this  section.

    - Labels for defeatably interlocked protective housings. For each laser prod- uct, labels shall be provided for each defeatably interlocked (as described in paragraph (f)(2)(iv) of this section) por- tion of the protective housing which is designed to be displaced or removed during operation, maintenance, or service, and which upon interlock de- feat could permit human access  to laser or collateral radiation in excess of the limits of Class I or table VI. Such labels shall be visible on the prod- uct prior to and during interlock defeat and in close proximity to the opening created by the removal or displace- ment of such portion of the protective housing, and shall include the wording:

- ‘‘CAUTION—Laser radiation when open and interlock defeated. DO NOT STARE INTO BEAM.’’ for Class II ac- cessible laser radiation.

- ‘‘CAUTION—Laser radiation when open and interlock defeated. DO NOT STARE INTO BEAM OR VIEW DI- RECTLY WITH OPTICAL INSTRU- MENTS.’’ for Class IIIa accessible laser radiation with an irradiance less than or equal to 2.5 10¥3 W cm¥2.

- ‘‘DANGER—Laser radiation when open and interlock defeated. AVOID DIRECT EYE EXPOSURE.’’ for Class IIIa accessible laser radiation when an irradiance greater than 2.5  10¥3 W cm¥2.

- ‘‘DANGER—Laser radiation when open and interlock defeated. AVOID DIRECT EXPOSURE  TO BEAM.’’ for Class IIIb accessible laser radiation.

- ‘‘DANGER—Laser radiation when open and interlock defeated. AVOID EYE OR SKIN EXPOSURE TO DIRECT OR SCATTERED  RADIATION.’’  for Clas IV accessible laser radiation.

- ‘‘CAUTION—Hazardous electro- magnetic radiation when open and interlock defeated.’’ for collateral radi- ation in excess of the accessible emis- sion limits  in table VI. item  1 of para- graph (d) of this section.

#### Food and Drug Administration, HHS	§ 1040.10

- ‘‘CAUTION—Hazardous x-rays when open and interlock defeated.’’ for collateral radiation in excess of the accesible emission limits in table VI. item 2 of paragraph (d) of this section.

    - Warning for visible and/or invisible radiation. On the labels specified in this paragraph, if the laser or collateral ra- diation referred to is:

      - Invisible  radiation,  the  word  ‘‘in- visible’’ shall appropriately precede the word ‘‘radiation’’; or

      - Visible and invisible radiation, the words ‘‘visible and invisible’’ or ‘‘visible and/or invisible’’ shall appro- priately precede the word ‘‘radiation.’’

      - Visible laser radiation only, the phrase ‘‘laser light’’ may replace the phrase ‘‘laser radiation.’’

    - Positioning of labels. All labels af- fixed to a laser product shall be posi- tioned so as to make unnecessary, dur- ing reading, human exposure to laser radiation in excess of the accessible emission limits of Class I radiation  or the limits of collateral radiation estab- lished to table VI of paragraph (d)  of this  section.

    - Label specifications. Labels re- quired by this section and § 1040.11 shall be permanently affixed to, or inscribed on, the laser  product,  legible,  and clearly visible during operation, main- tenance,  or  service,  as  appropriate.  If the size, configuration, design, or func- tion of the laser product  would  pre- clude compliance with the  require- ments for any required label or would render the required wording of  such label inappropriate or ineffective, the Director, Office of Compliance (HFZ– 300), Center for Devices and Radio- logical Health, on the Director’s own initiative  or  upon  written  application by the manufacturer, may approve al- ternate means of  providing  such label(s) or alternate wording for such label(s)  as  applicable.

  - Informational requirements—(1) User information. Manufacturers of laser products shall provide as an inte- gral part of any user instruction or op- eration manual which is regularly sup- plied with the product, or, if not so supplied, shall cause to be provided with each laser product:

  - Adequate instructions for assem- bly, operation, and maintenance, in- cluding clear warnings concerning pre-

cautions to avoid possible exposure to laser and collateral radiation in excess of the accessible emission limits in ta- bles  I,  II-A,  II,  III-A,  III-B,  and  VI  of

paragraph (d) of this section, and a schedule of maintenance necessary to keep the product in compliance with this section and § 1040.11.

- A statement of the magnitude, in appropriate units, of the pulse dura- tions(s), maximum radiant power and, where applicable, the maximum radi- ant energy per pulse of the accessible laser radiation detectable in each di- rection in excess of the accessible emission limits in table I of paragraph

- of this section determined under paragraph (e) of this section.

- Legible reproductions  (color  op- tional) of all labels and hazard warn- ings required by paragraph (g) of this section and § 1040.11 to be affixed to the laser product or provided with the laser product, including the information re- quired for positions 1, 2, and 3 of the applicable logotype (figure 1 of para- graph (g)(1)(ii) or figure 2 or paragraph (g)(2)(ii) of this section). The cor- responding position of each label  af- fixed to the product shall  be  indicated or, if provided with the product, a statement that such labels could not be affixed to the product  but  were  sup- plied with the product  and  a  statement of the form and manner in which they were supplied shall be provided.

- A listing of all controls, adjust- ments, and procedures  for  operation and maintenance, including the warn- ing ‘‘Caution—use of controls or ad- justments or performance of proce- dures other than those specified herein may result in hazardous radiation ex- posure.’’

- In the case of laser products other than laser systems, a statment of the compatibility requirements for a laser energy source that will assure compli- ance of the laser product with this sec- tion and § 1040.11.

- In the case of laser products clas- sified with a 7 millimeter diameter ap- erture stop as provided in paragraph (e)(3)(i) of this section, if the use of a 50 millimeter diameter aperture  stop would result in a  higher  classification of  the  product,  the  following  warning

#### § 1040.11

shall be included in the user informa- tion: ‘‘CAUTION—The use of optical in- struments with this product will in- crease eye hazard.’’

(2) Purchasing and servicing informa- tion. Manufacturers of laser products shall provide or cause to be provided:

- In all catalogs, specification sheets, and descriptive brochures per- taining to each laser product, a legible reproduction (color optional) of the class designation and warning required by paragraph (g) of this section to be affixed to that product, including the information required for positions 1, 2, and 3 of the applicable logotype (figure

1 of paragraph (g)(1)(ii) or figure 2 of paragraph (g)(2)(ii) of this section).

- To servicing dealers and distribu- tors and to others upon request at a cost not to exceed the cost of prepara- tion and distribution, adequate in- structions for service adjustments and service procedures for each laser prod- uct model, including clear warnings and precautions to be taken to avoid possible exposure to laser and collat- eral radiation in excess of the acces- sible emission limits in  tables  I,  II-A, II, III-A, III-B, and VI of paragraph (d) of this section, and a schedule of main- tenance necessary to keep the product in  compliance  with  this  section  and

§ 1040.11; and in all such service instruc- tions, a listing of those controls and procedures that could be utilized by persons other than the manufacturers or the manufacturer’s agents to in- crease accessible emission levels of ra- diation and a clear description of the location of displaceable portions of the protective housing that could allow human access to laser or collateral ra- diation in excess of the accessible emission limits in tables I, II-A, II, III- A, III-B, and VI of paragraph (d) of this section. The instructions shall include protective procedures for service per- sonnel to avoid exposure to levels of laser and collateral radiation known to be hazardous for each procedure or se- quence of procedures to be accom- plished, and legible reproductions (color optional) of required labels and hazard warnings.

(i) Modification of a certified product. The modification of a laser product, previously certified under  § 1010.2, by any  person  engaged  in  the  business  of

#### 21 CFR Ch. I (4–1–17 Edition)

manufacturing, assembling, or modi- fying laser products shall be construed as manufacturing under the act if the modification affects any aspect of the product’s performance or intended function(s) for which this section and

§ 1040.11 have an applicable require- ment. The manufacturer who performs such modification shall recertify and reidentify  the  product  in  accordance with the provisions of §§ 1010.2. and 1010.3.

(The information collection requirements contained in paragraph (a)(3)(ii) were ap- proved by the Office of Management and Budget under control number 0910–0176)

[50  FR 33688,  Aug. 20,  1985;  50  FR 42156,  Oct.

18, 1985; 65 FR 17138, Mar. 31, 2000, as amended

at 75 FR 20917, Apr. 22, 2010]

##### § 1040.11   Specific  purpose  laser  prod- ucts.

- Medical laser products. Each med- ical laser product shall comply with all of    the    applicable    requirements    of

§ 1040.10 for laser products  of  its  class. In  addition,  the  manufacturer  shall:

  - Incorporate in each Class III or IV medical laser product a means for the measurement of the level of that laser radiation intended  for  irradiation  of the human body. Such means may have an error in measurement of no more than 20 percent when calibrated in ac- cordance with paragraph (a)(2) of this section. Indication of the measurement shall be in International System Units. The requirements of this paragraph do not apply to any laser radiation that is all of the following:

    - Of a level less than the accessible limits of Class IIIa; and

    - Used for relative positioning of the human body; and

    - Not used for irradiation of the human eye for ophthalmic purposes.

  - Supply with each Class III or IV medical laser product instructions specifying a procedure and schedule for calibration of the measurement system required by paragraph (a)(1) of this sec- tion.

  - Affix to each medical laser prod- uct, in close proximity to each aper- ture through which is emitted acces- sible laser radiation in excess of the ac- cessible emission limits of Class I, a label bearing the wording: ‘‘Laser aper- ture.’’

#### Food and Drug Administration, HHS	§ 1040.20

- Surveying, leveling, and alignment laser products. Each surveying, leveling. or alignment laser product shall com- ply with all of the applicable require- ments of § 1040.10 for a Class I, IIa, II or IIIa laser product and shall not permit human access to laser radiation in ex- cess of the accessible emission limits of Class IIIa.

- Demonstration laser products. Each demonstration laser product shall com- ply with all of the applicable require- ments of § 1040.10 for a Class I, IIa, II, or IIIa laser product and shall not permit human access to laser radiation in ex- cess of the accessible emission limits of Class I and, if  applicable,  Class  IIa, Class  II,  or  Class  IIIa.

[50 FR 33702, Aug. 20, 1985]

##### § 1040.20 Sunlamp products and ultra- violet lamps intended for use in sunlamp products.

- Applicability. (1) The provisions of this section, as amended, are applica- ble as specified herein to the following products manufactured on or after Sep- tember 8, 1986.

- Any sunlamp product.

- Any ultraviolet lamp intended for use in any sunlamp product.

(2) Sunlamp products and ultraviolet lamps manufactured on or after May 7, 1980, but before September 8, 1986, are subject to the provisions of this section as published in  the  FEDERAL  REGISTER of November 9, 1979 (44 FR 65357).

- Definitions. As used in this section the following definitions apply:

  - Exposure position means any posi- tion, distance, orientation, or location relative to the radiating surfaces of the sunlamp product at which the user is intended to be exposed to ultraviolet radiation from the product, as rec- ommended by the manufacturer.

  - Intended means the same as ‘‘in- tended uses’’ in § 801.4.

  - Irradiance means the radiant power incident on a surface at a speci- fied location and orientation  relative to the radiating surface divided by the area of the surface, as the area be- comes vanishingly small, expressed in units of watts per square centimeter (W/cm2).

  - Maximum exposure time means the greatest  continuous  exposure  time  in-

terval recommended by the manufac- turer of the product.

  - Maximum timer interval means the greatest time interval setting on the timer of a product.

  - Protective eyewear means any de- vice designed to be worn by users of a product to reduce exposure of the eyes to radiation emitted by the product.

  - Spectral irradiance means the irra- diance resulting  from  radiation  within a wavelength range divided by the wavelength range as the range becomes vanishingly small, expressed in units of watts per square centimeter per nanometer (W/(cm2/nm)).

  - Spectral transmittance means the spectral	irradiance	transmitted through protective eyewear divided by the spectral irradiance incident on the protective eyewear.

  - Sunlamp product means any elec- tronic product designed to incorporate one or more ultraviolet lamps and in- tended for irradiation of any part  of the living human body, by ultraviolet radiation with wavelengths in air be- tween 200 and 400 nanometers, to in- duce skin tanning.

  - Timer means any device incor- porated into a product that terminates radiation emission after a preset time interval.

  - Ultraviolet lamp means any lamp that produces ultraviolet radiation in the wavelength interval of 200 to 400 nanometers in air and that is intended for use in any sunlamp product.

- Performance requirements—(1) Irra- diance ratio limits. For each sunlamp product and ultraviolet lamp,  the  ratio of the  irradiance  within  the  wave- length range of greater than 200 nano- meters through 260 nanometers to the irradiance within the wavelength  range of greater than 260 nanometers through 320 nanometers may not exceed 0.003 at any distance and direction from the product  or  lamp.

- Timer system. (i) Each sunlamp product shall incorporate a timer sys- tem with multiple timer settings ade- quate for the recommended exposure time intervals for different exposure positions and expected results of the products as specified in the label re- quired by paragraph (d) of this section.

- The maximum timer interval(s) may   not   exceed   the   manufacturer’s

#### § 1040.20

recommended maximum exposure time(s) that is indicated on the label required by paragraph (d)(1)(iv) of this section.

- No timer interval may have an error greater than 10 percent of the maximum timer interval of the prod- uct.

- The timer may not automati- cally reset and cause radiation emis- sion to resume for a period greater than the unused portion of the timer cycle, when emission from the sunlamp product has been terminated.

- The timer requirements do not preclude a product from  allowing  a user to reset the timer before the end of the preset time interval.

- Control for termination of radiation emission. Each sunlamp product shall incorporate a control on the product to enable the person being exposed to ter- minate manually radiation emission from the product at any time without disconnecting the electrical plug or re- moving the ultraviolet lamp.

- Protective eyewear. (i) Each sun- lamp product shall be accompanied by the number of sets of protective eyewear that is equal to the maximum number of persons that the instruc- tions provided under paragraph (e)(1)(ii) of this section recommend to be exposed simultaneously to radiation from such product.

(ii) The spectral transmittance to the eye of the protective eyewear required by paragraph (c)(4)(i) of this section shall not exceed a value of 0.001  over the  wavelength  range  of  greater  than

200 nanometers 320 nanometers and a value of 0.01 over the wavelength range of greater than 320 nanometers through 400 nanometers, and shall be sufficient over the wavelength greater than 400 nanometers to enable the user to see clearly  enough  to  reset  the  timer.

- Compatibility of lamps. An ultra- violet lamp may not be capable of in- sertion and operation in either the ‘‘single-contact medium screw’’ or the ‘‘double-contact medium screw’’ lampholders described in American Na- tional Standard C81.10–1976, Specifica- tions for Electric Lamp Bases and Holders—Screw-Shell Types, which is incorporated by reference. Copies are available from the American National Standards    Institute,    1430    Broadway,

#### 21 CFR Ch. I (4–1–17 Edition)

New York, NY 10018, or available for in- spection at the National Archives and Records Administration (NARA). For information on the availability of this material at NARA, call 202–741–6030.

- Label requirements. In addition to the labeling requirements in part 801 and the certification and identification requirements of §§ 1010.2 and 1010.3, each sunlamp product and ultraviolet lamp shall be subject to the labeling requirements prescribed in this para- graph  and  paragraph  (e)  of  this  section.

  - Labels for sunlamp products. Each sunlamp product shall have a label(s) which contains:

    - A warning statement with the words ‘‘DANGER—Ultraviolet radi- ation. Follow instructions. Avoid over- exposure. As with natural sunlight, overexposure can cause eye and skin injury and allergic reactions. Repeated exposure may cause premature aging of the skin and skin cancer. WEAR PRO- TECTIVE EYEWEAR; FAILURE  TO MAY RESULT IN SEVERE BURNS OR LONG-TERM INJURY TO THE EYES. Medications or cosmetics may increase your sensitivity to the ultraviolet radi- ation. Consult physician before using sunlamp if you are using medications or have a history of skin problems or believe yourself especially sensitive to sunlight. If you do not tan in the sun, you are unlikely to tan from the use of this product.’’

    - Recommended exposure posi- tion(s). Any exposure position may be expressed either in terms of a distance specified both in meters and in feet (or in inches) or through the use of mark- ings or other means to indicate clearly the recommended exposure position.

    - Directions for achieving the rec- ommended exposure position(s) and a warning that the use of other positions may result in overexposure.

    - A recommended exposure sched- ule including duration and spacing of sequential exposures and maximum ex- posure time(s) in minutes.

    - A statement of the time it may take before the expected results ap- pear.

#### Food and Drug Administration, HHS	§ 1040.20

    - Designation of the ultraviolet lamp type to be used in the product.

  - Labels for ultraviolet lamps. Each ultraviolet lamp shall have a label which contains:

    - The words ‘‘Sunlamp—DANGER— Ultraviolet radiation. Follow instruc- tions.’’

    - The model identification.

    - The words ‘‘Use ONLY in fixture equipped with a timer.’’

  - Label specifications. (i) Any label prescribed in this paragraph for sun- lamp products shall be permanently af- fixed or inscribed on an exterior sur- face of the product when fully assem- bled for use so as to be legible and readily accessible to view by the person being exposed immediately before the use of the product.

- Any label prescribed in this para- graph for ultraviolet lamps shall be permanently affixed or inscribed on the product so as to be legible and readily accessible to view.

- If the size, configuration,  design, or function of the sunlamp product or ultraviolet lamp would preclude com- pliance with the requirements for any required label or would render the re- quired wording of such label inappro- priate  or  ineffective,  or  would  render the required label unnecessary, the Di- rector, Office of Communication, Edu- cation, and  Radiation  Programs  10903 New Hampshire Ave., Bldg. 66, rm. 4312, Silver Spring, MD 20993–0002, Center for Devices  and   Radiological   Health,   on the center’s own initiative  or  upon written application by  the  manufac- turer, may approve alternate means of providing such label(s), alternate word- ing for such label(s), or deletion, as ap- plicable.

- In lieu of permanently affixing or inscribing tags or labels on the ultra- violet lamp as required by §§ 1010.2(b) and 1010.3(a), the manfacturer of the ul- traviolet lamp may permanently  affix or inscribe such required tags or labels on the lamp packaging uniquely associ- ated with the lamp, if the name of the manufacturer and month and year of manufacture are  permanently  affixed or inscribed  on  the exterior  surface  of the ultraviolet lamp so as to be legible and readily accessible to view. The name of the manufacturer and month and  year  of  manufacture  affixed  or  in-

scribed on the exterior surface of the lamp may be expressed in code or sym- bols, if the manufacturer has pre- viously supplied the Director, Office of Compliance (HFZ–300), Center for De- vices and Radiological Health, with the key to such code or symbols and the lo- cation of the coded information or symbols on the ultraviolet lamp. The label or tag affixed or inscribed on the lamp packaging may provide either the month and year of manufacture with- out abbreviation, or information to allow the date to be readily decoded.

- A label may contain statements or illustrations in addition to those re- quired by this paragraph if the addi- tional statements are not false or mis- leading in any particular; e.g.,  if  they do not diminish the impact of the re- quired statements; and are not prohib- ited by this chapter.

- Instructions to be provided to users. Each manufacturer of a sunlamp prod- uct and ultraviolet lamp shall provide or cause to be provided to purchasers and, upon request, to others at a cost not to exceed the cost of publication and distribution, adequate instructions for use to avoid or to minimize poten- tial injury to the user, including the following technical and safety informa- tion as applicable:

  - Sunlamp products. The users’ in- structions for a sunlamp product shall contain:

    - A reproduction of the label(s) re- quired in paragraph (d)(1) of this sec- tion prominently displayed at the be- ginning  of  the  instructions.

    - A statement of the maximum number of people who may be exposed to the product at the same time and a warning that only that number of pro- tective eyewear has been provided.

    - Instructions for the proper oper- ation of the product including the func- tion, use, and setting of the timer and other controls, and the use of protec- tive eyewear.

    - Instructions for determining the correct exposure time and schedule for persons according to skin type.

    - Instructions for obtaining repairs and recommended replacement compo- nents and accessories which are com- patible    with    the    product,    including

#### § 1040.30

21 CFR Ch. I (4–1–17 Edition)

compatible protective eyewear, ultra- violet lamps, timers, reflectors, and fil- ters, and which will, if installed or used as instructed, result in continued com- pliance with the standard.

  - Ultraviolet lamps. The users’ in- structions for an ultraviolet lamp not accompanying a sunlamp product shall contain:

    - A reproduction of the label(s) re- quired in paragraphs (d)(1)(i) and (2) of this section, prominently displayed at the beginning of the instructions.

    - A warning that the instructions accompanying the sunlamp product should always be followed to avoid or to minimize potential injury.

    - A clear identification by brand and model designation of all lamp mod- els for which replacement lamps are promoted, if applicable.

- Test for determination of compliance. Tests on which certification  pursuant to § 1010.2 is based shall account for all errors and statistical uncertainties in the process and, wherever applicable, for changes in radiation emission or degradation in radiation  safety  with age of the product. Measurements for certification purposes shall be made under those operational conditions, lamp voltage, current, and position as recommended  by  the  manufacturer. For these measurements, the meas- uring instrument shall be positioned at the recommended exposure position and so oriented as to result in the max- imum detection of the radiation by the instrument.

[50 FR 36550, Sept. 6, 1985, as amended at 67

FR 9587, Mar. 4, 2002; 69 FR 18803, Apr. 9, 2004;

75 FR 20917, Apr. 22, 2010]

##### § 1040.30   High-intensity	mercury vapor discharge lamps.

- Applicability. The provisions of this section apply to any high-inten- sity mercury vapor discharge  lamp that is designed, intended, or promoted for illumination purposes and is manu- factured or assembled after March 7, 1980, except as described in paragraph (d)(1)(ii) of this section.

- Definitions. (1) High-intensity mer- cury  vapor  discharge  lamp  means  any

porating a high-pressure arc discharge tube that has a fill consisting pri- marily of mercury and that is con- tained within an outer envelope.

- Advertisement means any catalog, specification sheet, price list, and any other descriptive or commercial bro- chure and literature, including video- tape and film, pertaining to high-inten- sity mercury vapor discharge lamps.

- Packaging means any lamp carton, outer wrapping, or other means of con- tainment that is intended for the stor- age, shipment, or display of a high-in- tensity mercury vapor lamp and is in- tended to identify the contents or rec- ommend its use.

- Outer envelope means the lamp ele- ment, usually glass, surrounding a high-pressure arc discharge tube, that, when  intact,  attenuates  the   emission of  shortwave  ultraviolet  radiation.

- Shortwave ultraviolet radiation means ultraviolet radiation with wave- lengths shorter than 320 nanometers.

- Cumulative operating time means the sum of the times during which elec- tric current passes through the high- pressure arc discharge.

- Self-extinguishing lamp means a high-intensity mercury vapor dis- charge lamp that is intended to comply with the requirements of paragraph (d)(1) of this section as applicable.

- Reference ballast is an inductive re- actor designed to have the operating characteristics as listed in Section 7 in the American National Standard Speci- fications for High-Intensity Discharge Lamp Reference Ballasts (ANSI C82.5– 1977) 1 or its equivalent.

- General requirements for all lamps.

  - Each high-intensity mercury vapor discharge lamp shall:

    - Meet the requirements of either paragraph (d) or paragraph (e) of this section; and

    - Be permanently labeled or marked in such a manner that the name of  the  manufacturer  and  the month and year of manufacture of the lamp can be determined on an intact lamp and after the outer envelope of the lamp is broken or removed. The name of the manufacturer and month

lamp  including  any  ‘‘mercury  vapor’’

and ‘‘metal halide’’ lamp, with the ex- ception of  the tungsten  filament self- ballasted  mercury  vapor  lamp,  incor-

1 Copies are available from American Na- tional Standards Institute, 1430 Broadway, New York, NY 10018.

#### Food and Drug Administration, HHS	§ 1040.30

and year of manufacture may be ex- pressed in code or symbols, provided the manufacturer has previously sup- plied the Director, Center for Devices and Radiological Health, with the key to the code or symbols and the location of the coded information or symbols on the lamp.

  - In lieu of permanently affixing or inscribing tags or labels  on  the  product as required by §§ 1010.2(b) and 1010.3(a) of this  chapter,  the  manufacturer  of any high-intensity mercury vapor dis- charge lamp may permanently affix or inscribe such required tags or labels  on the  lamp  packaging  uniquely   associ- ated  with  the  applicable  lamp.

- Requirements for self-extinguishing lamps—(1) Maximum cumulative oper- ating time. (i) Each self-extinguishing lamp manufactured after March 7, 1980 shall cease operation within a cumu- lative operating time  not to exceed  15 minutes  following  complete  breakage or removal of the outer envelope (with the  exception  of  fragments  extending

50 millimeters or less from the base shell); and

(ii) Each self-extinguishing lamp manufactured after September 7, 1981, shall cease operation within a cumu- lative operating time  not to exceed  15 minutes following breakage or removal of at least 3 square centimeters of con- tiguous surface of the outer envelope.

- Lamp labeling. Each self-extin- guishing lamp shall be clearly marked with the letter ‘‘T’’ on the outer enve- lope and on another part of the lamp in such a manner  that it  is visible  after the outer envelope of the lamp is bro- ken or removed.

- Lamp packaging. Lamp packaging for each self-extinguishing lamp shall clearly and prominently display:

  - The letter ‘‘T’’; and

  - The words ‘‘This lamp should self- extinguish within 15 minutes after the outer envelope is broken or punctured. If such damage occurs,  TURN  OFF AND REMOVE LAMP to avoid possible injury from hazardous shortwave ultra- violet  radiation.’’

- Requirements for lamps that are not self-extinguishing lamps—(1) Lamp label- ing. Any high-intensity mercury vapor discharge lamp that does not comply with paragraph (d)(1) of this section shall  be  clearly  and  legibly  marked

with the letter ‘‘R’’ on the outer enve- lope and on another part of the lamp in such a manner  that it  is visible  after the outer envelope of the lamp is bro- ken or removed.

- Lamp packaging. Lamp packaging for each high-intensity mercury vapor discharge lamp that does not comply with paragraph (d)(1) of this section shall clearly and prominently display:

  - The letter ‘‘R’’; and

  - The words ‘‘WARNING: This lamp can cause serious skin burn and eye in- flammation from shortwave ultraviolet radiation if outer envelope of the lamp is broken or punctured. Do not use where people will remain for more than a few minutes unless adequate shield- ing or other safety precautions are used. Lamps that will automatically extinguish when the outer envelope is broken or punctured are commercially available.’’

- Lamp advertisement. Advertising for any high-intensity mercury vapor discharge lamp that does not comply with paragraph (d)(1) of this section shall prominently display the following wording: ‘‘WARNING: This lamp can cause serious skin burn and eye inflam- mation from shortwave ultraviolet ra- diation if outer envelope of the lamp is broken or punctured. Do not use where people will remain for more than a few minutes unless adequate shielding or other safety precautions are used. Lamps that will automatically extin- guish when the outer  envelope is bro- ken or punctured are commercially available.’’

- Test conditions. Any high-intensity mercury vapor discharge lamp under test for compliance with the require- ments set  forth  in  paragraph  (d)(1)  of this section shall be started and oper- ated under the following conditions as applicable:

- Lamp voltage, current, and ori- entation shall be those indicated or recommended by the manufacturer for operation of the intact lamp.

- The lamp shall be operated on a reference  ballast.

- The lamp shall be started in air that has a temperature of 25 5 C. Heating and movement of the air sur- rounding the lamp shall be that pro- duced by the lamp and ballast alone.

#### Pt. 1050

- If any test is performed in an en- closure, the enclosure shall be not less than 0.227 cubic meter (8 cubic feet).

- Any lamp designed to be operated only in a specific fixture or luminaire that the lamp manufacturer supplies or specifies shall be tested in that fixture or luminaire. Any other lamp shall be tested with no reflector or other sur- rounding material.

[44 FR 52195, Sept. 7, 1979, as amended at 53

FR 11254, Apr. 6, 1988]

## PART	1050—PERFORMANCE STANDARDS FOR SONIC, INFRASONIC, AND ULTRASONIC RADIATION-EMITTING PRODUCTS

AUTHORITY: 21 U.S.C. 351, 352, 360, 360e–360j,

360hh–360ss, 371, 381.

##### § 1050.10   Ultrasonic therapy products.

- Applicability. The provisions of this section are applicable as specified herein to any ultrasonic therapy prod- uct for use in physical therapy manu- factured on or after February 17, 1979.

- Definitions. The following defini- tions apply to words and phrases used in this section:

- Amplitude modulated waveform means a waveform in which the ratio of the temporal-maximum pressure am- plitude spatially averaged over the ef- fective radiating surface to the root- mean-square pressure amplitude spa- tially averaged over the effective radi- ating surface is greater than 1.05.

- Applicator means that portion of a fully assembled ultrasonic therapy product that is designed to emit ultra- sonic radiation and which includes one or more ultrasonic transducers and any associated housing.

- Beam cross-section means the sur- face in any plane consisting of  the points at which the intensity is greater than 5 percent of the spatial-maximum intensity  in  that  plane.

- Beam nonuniformity ratio means the ratio of the temporal-average spa- tial-maximum intensity to the tem- poral-average effective intensity.

- Centroid of a surface means the point whose  coordinates are the  mean values of the coordinates of the points of the surface.

21 CFR Ch. I (4–1–17 Edition)

- Collimating applicator means an ap- plicator that does not meet the defini- tion of a focusing applicator as speci- fied in paragraph (b)(15) of  this section and for which the  ratio  of  the  area  of at least one beam cross-section, whose centroid is 12 centimeters from the centroid of the effective radiating sur- face, to the area of the effective radi- ating  surface  is  less  than  two.

- Continuous-wave waveform means a waveform in which the ratio of the temporal-maximum pressure amplitude spatially averaged over the effective radiating surface to the root-mean- square pressure amplitude spatially averaged over the effective radiating surface is less than or equal to 1.05.

- Diverging applicator means an ap- plicator that does not meet the defini- tion of a collimating applicator or a fo- cusing applicator as specified in para- graphs (b) (6) and (15) of this section.

- Effective intensity means the ratio of the ultrasonic power to the focal area for a focusing applicator. For all other applicators, the effective inten- sity is the ratio of the ultrasonic power to the effective radiating area. Effec- tive intensity is expressed in watts per square centimeter (W cm¥2).

- Effective radiating area means the area consisting of all points of the ef- fective radiating surface at which the intensity is 5 percent or more of the maximum intensity at the effective ra- diating surface, expressed in square centimeters (cm2).

- Effective radiating surface means the surface consisting of all points 5 millimeters from the applicator face.

- Focal area means the area of the focal surface, expressed in square cen- timeters (cm2).

- Focal length means the distance between the centroids of the effective radiating surface and the focal surface, for a focusing applicator, expressed in centimeters (cm).

- Focal surface means the beam cross-section with smallest area of a focusing  applicator.

- Focusing applicator means an ap- plicator in which the  ratio  of  the  area of the beam cross-section with the smallest area to the effective radiating area is less than one-half.

- Generator means that portion of a fully    assembled    ultrasonic    therapy

#### Food and Drug Administration, HHS	§ 1050.10

product that supplies electrical energy to the applicator. The generator may include, but is not limited to, a power supply, ultrasonic frequency oscillator, service controls, operation controls, and a cabinet to house these compo- nents.

- Maximum beam nonuniformity ratio means the maximum value of the beam nonuniformity ratio characteristic of a model of an ultrasonic therapy prod- uct.

- Operation control means any con- trol used during operation of an ultra- sonic therapy product that affects the ultrasonic radiation emitted by the ap- plicator.

- Pressure amplitude means the in- stantaneous value of the modulating waveform, and is p1(t) in the expression for a pressure wave, p(t) = p1(t) p2(t), where p(t) is the instantaneous pres- sure, p1(t) is the modulating envelope, and p2(t) is the relative amplitude of the carrier wave normalized to a peak height of one. All are periodic func- tions of time, t, at any point in space. The period of p1(t) is greater than the period of p2(t).

- Pulse duration means a time in- terval, expressed in seconds, beginning at the first time the pressure ampli- tude exceeds the minimum pressure amplitude plus 10 percent of the dif- ference between the maximum and minimum pressure amplitudes, and ending at the last time the pressure amplitude returns to this value.

- Pulse repetition rate means the repetition frequency of the waveform modulating the ultrasonic carrier wave expressed in pulses per second (pps).

- Service control means any control provided for the purpose of adjustment that is not used during operation and can affect the ultrasonic radiation emitted by the applicator, or can alter the calibration or accuracy of an indi- cator or operation control.

- Ultrasonic frequency means the frequency of the ultrasonic radiation carrier wave, expressed in Hertz (Hz), kilohertz (kHz), or megahertz (MHz).

- Ultrasonic power means the total power emitted in the form of ultrasonic radiation by the applicator averaged over each cycle of the ultrasonic radi- ation carrier wave, expressed in watts.

- Ultrasonic therapy product means:

- Any device intended to generate and emit ultrasonic radiation for therapeutic purposes at ultrasonic fre- quencies above 16 kilohertz (kHz); or

- Any generator or applicator de- signed or specifically  designated  for use in a device as specified in para- graph (b)(25)(i) of this section.

- Ultrasonic transducer means a de- vice used to convert  electrical  energy of ultrasonic frequency into ultrasonic radiation or vice versa.

- Performance requirements. The re- quirements of this paragraph are appli- cable to each ultrasonic therapy prod- uct as defined in paragraph (b)(25) of this section when the generator and ap- plicator are designated or intended for use together, or to each generator when the applicator(s) intended for use with the generator does not contain controls that affect the functioning of the  generator.

  - Ultrasonic power and intensity—(i) Continuous-wave waveform operation. A means shall be incorporated to indicate the magnitudes of the temporal-aver- age ultrasonic power and the temporal- average effective intensity when emis- sion is of continuous-wave waveform. The error in the indication of the tem- poral-average ultrasonic power shall not exceed 20 percent for all emissions greater than 10 percent of the max- imum emission.

(ii) Amplitude-modulated waveform op- eration. A means shall be incorporated to indicate the magnitudes of the tem- poral-maximum ultrasonic power and the temporal-maximum effective inten- sity when the emission is of amplitude- modulated waveform. The sum of the errors in the indications of the tem- poral-maximum ultrasonic power and the ratio of the temporal-maximum ef- fective intensity to the temporal-aver- age effective intensity specified in paragraph (d)(3)(ii) of this section shall not exceed 20 percent for all emissions greater than 10 percent of the max- imum emission.

  - Treatment time. A means shall be incorporated to enable  the duration  of emission of ultrasonic radiation for treatment to be preset and such means shall terminate emission at the end  of the preset time. Means shall also be in- corporated to enable termination of emission  at  any  time.  Means  shall  be

#### § 1050.10

incorporated to indicate the  magnitude of the duration of  emission  (expressed in minutes) to within 0.5 minute of the preset duration of emission for setting less than 5 minutes, to within 10 per- cent of the preset duration of  emission for settings of from 5 minutes to 10 minutes, and to within 1 minute of the preset duration of emission for settings greater  than  10  minutes.

  - Pulse duration and  repetition  rate. A means shall be incorporated for indi- cating the magnitudes of pulse dura- tion and pulse repetition rate of the emitted ultrasonic  radiation,  if  there are operation  controls  for  varying these  quantities.

  - Ultrasonic frequency. A means shall be incorporated for indicating the magnitude of the ultrasonic  frequency of the emitted ultrasonic radiation, if there is an operation control for  vary- ing  this  quantity.

  - Visual indicator. A means  shall be incorporated to provide a clear, dis- tinct, and readily understood visual in- dicator when and only when electrical energy of appropriate ultrasonic fre- quency is being applied to the ultra- sonic transducer(s).

- Labeling requirements. In  addition to the labeling requirements in part 801 and  the  requirements  of  §§ 1010.2  and

- of this chapter, each ultrasonic therapy product shall be subject to the applicable  labeling  requirements  of this  paragraph.

  - Operation controls. Each operation control shall be clearly labeled identi- fying the function controlled and, where appropriate, the units of meas- ure of that function. If a separate con- trol and indicator are associated with the same function, then labeling the appropriate units of measure of that function is required for the indicator but not for the control.

  - Service controls. Each service con- trol that is accessible without displace- ment or removal of any part of the ul- trasonic therapy product shall be clear- ly labeled identifying the function con- trolled and shall include the phrase ‘‘for service adjustment only.’’

  - Generators. (i) Each generator shall bear a label that states:  The brand name, model designation, and unique serial number or other unique identification so that it is individually

#### 21 CFR Ch. I (4–1–17 Edition)

identifiable; ultrasonic frequency (un- less there is an operation control for varying this quantity); and type of waveform (continuous wave or ampli- tude modulated).

(ii) Generators employing amplitude- modulated waveforms shall also bear a label that provides the following infor- mation: Pulse duration and pulse rep- etition rate (unless there are operation controls for varying these quantities), an illustration of the amplitude-modu- lated waveform, and the ratio of the temporal-maximum effective intensity to the temporal-average effective in- tensity. (If this ratio is a function  of any operation control setting, then the range of the ratio shall be specified, and the waveform illustration shall be provided for the maximum value of this ratio.)

  - Applicators. Each applicator shall bear a label that provides the following information:

    - The brand name, model designa- tion, and unique serial number or other unique identification so the applicator is individually identifiable;

    - A designation of the generator(s) for which the applicator is intended; and

    - The ultrasonic frequency, effec- tive radiating area, maximum beam nonuniformity ratio, type of applicator (focusing, collimating, diverging), and for a focusing applicator the focal length and focal area.

  - Label specification. Labels required by this paragraph shall be permanently affixed to or inscribed on the ultra- sonic therapy product; they shall be legible and clearly visible. If the size, configuration, or design of the ultra- sonic therapy product would preclude compliance with the requirements of this paragraph, the Director, Center for Devices and Radiological Health, may approve alternate means of providing such labels).

- Tests for determination of compli- ance—(1) Tests for certification. Tests on which certification pursuant to  § 1010.2 of this chapter is based shall  account for all measurement errors and uncer- tainties.  Such  tests  shall  also  account for increases in emission and degrada- tion in radiation safety that occur with age.

#### Food and Drug Administration, HHS	§ 1050.10

- Test conditions. Except  as  provided in § 1010.13 of this chapter, tests for compliance with each of the applicable requirements of this section shall be made:

  - For all possible combinations of adjustments of the controls listed in the operation instructions.

  - With the ultrasonic radiation emitted into the equivalent of an infi- nite medium of  distilled,  degassed water at 30 C for measurements con- cerning  the  ultrasonic  radiation.

  - With line voltage variations in the range of 10 percent of the rated value specified by the manufacturer.

- Measurement parameters. Measure- ments for determination of the spatial distribution of the ultrasonic radiation field shall be made with a detector hav- ing dimensions of less than one wave- length in water or an equivalent meas- urement technique.

- Informational requirements—(1) Servicing information. The manufacturer of an ultrasonic therapy product shall provide or cause to be provided to serv- icing dealers and distributors, and to others upon request, at a cost not to exceed the cost of preparation and dis- tribution adequate instructions for op- erations, service, and calibration, in- cluding a description of those controls and procedures that could be used to increase radiation emission levels, and a schedule of maintenance necessary to keep equipment in compliance with this section. The instructions shall in- clude adequate safety precautions that may be necessary regarding ultrasonic radiation exposure.

(2) User information. The manufac- turer of an ultrasonic therapy product shall provide as an integral part of any user instruction or operation manual that is regularly supplied with the product, or, if not so supplied, shall cause to be provided with each ultra- sonic therapy product, and to others upon request, at a cost not to exceed the cost of preparation and distribu- tion:

- Adequate instructions concerning assembly, operation, safe use, any safe- ty   procedures   and   precautions   that

may be necessary regarding the use of ultrasonic radiation, and a schedule of maintenance necessary to keep the equipment in compliance with this sec- tion. The operation instructions shall include a discussion of all operation controls, and shall describe the effect of each control.

- Adequate description of the spa- tial distribution of the ultrasonic radi- ation field and the orientation of the field with respect  to  the  applicator. This will include a textual discussion with diagrams, plots, or photographs representative of the beam pattern. If there is more than one ultrasonic transducer in an applicator and their positions are not fixed relative of each other, then the description must speci- fy the  spatial  distribution  of  the  ultra- sonic radiation field emitted by each ultrasonic transducer and present ade- quate examples of the combination field of the ultrasonic tranducers with regard to safe use. The description  of the  ultrasonic  radiation  field  shall state that such description  applies under conditions specified in paragraph (e)(2)(ii) of this section.

- Adequate description, as appro- priate to the product, of the uncertain- ties in magnitude expressed in terms of percentage error, of the ultrasonic fre- quency effective radiating area, and, where applicable, the ratio of the tem- poral-maximum effective intensity  to the temporal-average effective inten- sity, pulse duration, pulse repetition rate, focal area, and focal length. The errors in indications specified in para- graphs (c)(1) and (c)(2) of this section shall be stated in the instruction man- ual.

- A listing of controls, adjust- ments, and procedures for  operation and maintenance, including the warn- ing ‘‘Caution—use of controls or ad- justments or performance of proce- dures other than those specified herein may result in hazardous exposure to ul- trasonic  energy.’’

[43  FR  7166,  Feb.  17,  1978,  as  amended  at  45

FR  16483,  Mar.  14,  1980;  53  FR  11255,  Apr.  6,

1988]

# SUBCHAPTER K—TOBACCO PRODUCTS

## PART 1100—TOBACCO PRODUCTS SUBJECT TO FDA AUTHORITY

Sec.

- Scope.

- Requirements.

- Definitions.

- Exclusion from tobacco regulation.

AUTHORITY: 21 U.S.C. 387a(b), 387f(d) and Pub. L. 111–31.

SOURCE:  81  FR  29102,  May  10,  2016,  unless

otherwise noted.

##### § 1100.1   Scope.

In addition to FDA’s authority over cigarettes, cigarette tobacco, roll- your-own tobacco, and smokeless to- bacco, FDA deems all other products meeting the definition of tobacco prod- uct under section 201(rr) of the Federal Food, Drug, and Cosmetic Act, except accessories of such other tobacco prod- ucts, to be subject to the Federal Food, Drug, and Cosmetic Act.

##### § 1100.2   Requirements.

Cigarettes, cigarette tobacco, roll- your-own tobacco, smokeless tobacco are subject to chapter IX of the Federal Food, Drug, and Cosmetic Act and its implementing regulations. FDA has deemed all other tobacco products, ex- cept accessories of such other tobacco products, subject to chapter IX of the Federal Food, Drug, and Cosmetic Act and its implementing regulations.

##### § 1100.3   Definitions.

For the purposes of this part:

Accessory means any product that is intended or reasonably  expected to be used with or for the human consump- tion of a tobacco product; does not con- tain tobacco and is not made or derived from tobacco; and meets either of the following:

  - Is not intended or reasonably ex- pected to affect or alter the perform- ance, composition, constituents, or characteristics of a tobacco product; or

  - Is intended or reasonably expected to affect or maintain the performance, composition, constituents, or charac- teristics of a tobacco product but

- Solely controls  moisture  and/or temperature of a stored tobacco prod- uct; or

- Solely provides an external heat source to initiate but not maintain combustion of a tobacco product.

Component or part means any soft- ware or assembly of materials intended or reasonably expected:

  - To alter or affect the tobacco product’s performance, composition, constituents, or characteristics; or

  - To be used with or for the human consumption of a tobacco product. Component or part excludes anything that is an accessory of a tobacco prod- uct.

Package  or  packaging  means  a  pack, box,  carton,  or  container  of  any  kind or, if no other container, any wrapping (including  cellophane),  in  which  a  to- bacco  product  is  offered  for  sale,  sold, or otherwise distributed to consumers. Tobacco product. As  stated  in  section 201(rr) of the Federal Food, Drug, and Cosmetic   Act   in   relevant   part,   a   to-

bacco product:

- Means any product made or de- rived from tobacco that is intended for human consumption, including any component, part, or accessory of a to- bacco product (except for raw mate- rials other than tobacco used in manu- facturing a component, part, or acces- sory of a tobacco product); and

- Does not mean an article that is a drug under section 201(g)(1) of the Fed- eral Food, Drug, and Cosmetic Act, a device under section 201(h) of the Fed- eral Food,  Drug,  and  Cosmetic  Act,  or a combination product described in section 503(g) of the Federal Food, Drug, and Cosmetic Act.

##### § 1100.5   Exclusion  from  tobacco  regu- lation.

If a product made or derived from to- bacco that is intended for human con- sumption is intended for use for any of the purposes described in paragraph (a) or (b) of this section, the product is not a tobacco product as defined in section 201(rr) of the Federal Food, Drug, and Cosmetic Act and will be subject to regulation as a drug, device, or com- bination  product.

Food and Drug Administration, HHS	§ 1105.10

  - The product is intended for use in the diagnosis of disease or other condi- tions, or in the cure, mitigation, treat- ment or prevention of disease, includ- ing use in the cure or treatment of nic- otine addiction (e.g., smoking ces- sation), relapse prevention, or relief of nicotine withdrawal symptoms;

  - The product is intended to affect the structure or any function of the body in any way that is different from effects related to nicotine that were commonly and legally claimed in the marketing of cigarettes and smokeless tobacco products prior to March 21, 2000.

EFFECTIVE  DATE  NOTE: At 82  FR 2217,  Jan.

9, 2017, § 1100.5 was added, eff. Feb. 8, 2017. At

82 FR 9501, Feb. 7, 2017, the effective date was

delayed to Mar. 21, 2017. At 82 FR 14319, Mar. 20, 2017, the effective date was further de- layed to Mar. 19, 2018.

## PART 1105—GENERAL

AUTHORITY: 21 U.S.C. 371(a), 387e, 387j, and 387k.

SOURCE:  81  FR  95868,  Dec.  29,  2017,  unless

otherwise noted.

## Subpart A—General Submission Requirements

##### § 1105.10   Refusal   to   accept   a   pre- market submission.

- FDA will refuse to accept for re- view, as soon as practicable, a pre- market tobacco product  application, modified risk tobacco product applica- tion, substantial equivalence applica- tion, or exemption request or subse- quent abbreviated report for the fol- lowing reasons, if applicable:

  - The submission does not  pertain to  a  tobacco  product  as  defined  in  21

U.S.C. 321(rr).

  - The submission is not in English or does not contain complete English translations of any information sub- mitted within.

  - If submitted in an electronic for- mat, the submission is in a format that FDA cannot process, read, review, and archive.

  - The submission does not contain contact information, including the ap- plicant’s name and address.

  - The submission is from a foreign applicant and does not identify an au- thorized U.S. agent, including the agent’s name and address, for the sub- mission.

  - The submission does not contain a required FDA form(s).

  - The submission does not contain the following product-identifying infor- mation: The manufacturer of the to- bacco product; the product name, in- cluding the brand and subbrand; the product category and subcategory; package type and package quantity; and characterizing flavor.

  - The type of submission is not specified.

  - The submission does not contain a signature of a responsible official, au- thorized to represent the applicant, who either resides in or has a place of business in the United States.

  - For premarket tobacco applica- tions, modified risk tobacco product applications, substantial equivalence applications, and exemption requests only: The  submission  does  not  include a valid claim of  categorical  exclusion in accordance with part 25 of this chap- ter, or an environmental assessment.

- If FDA finds that none of the rea- sons in paragraph (a) of this section ex- ists for refusing to accept a premarket submission, FDA may accept the sub- mission for processing and further re- view. FDA will send to  the  submitter an acknowledgement letter stating the submission has been accepted for proc- essing and further review and will pro- vide the premarket submission track- ing number.

- If FDA finds that any of the rea- sons in paragraph (a) of this section exist for refusing to accept the submis- sion, FDA will notify the submitter in writing of the reason(s) and that the submission has not been accepted, un- less insufficient contact information was provided.

## PART 1105—GENERAL

AUTHORITY: 21 U.S.C. 371(a), 387e, 387j, and 387k.

#### § 1105.10

Subpart A—General Submission Requirements

##### § 1105.10   Refusal   to   accept   a   pre- market submission.

- FDA will refuse to accept for re- view, as soon as practicable, a pre- market tobacco product  application, modified risk tobacco product applica- tion, substantial equivalence applica- tion, or exemption request or subse- quent abbreviated report for the fol- lowing reasons, if applicable:

  - The submission does not  pertain to  a  tobacco  product  as  defined  in  21

U.S.C. 321(rr).

  - The submission is not in English or does not contain complete English translations of any information sub- mitted within.

  - If submitted in an electronic for- mat, the submission is in a format that FDA cannot process, read, review, and archive.

  - The submission does not contain contact information, including the ap- plicant’s name and address.

  - The submission is from a foreign applicant and does not identify an au- thorized U.S. agent, including the agent’s name and address, for the sub- mission.

  - The submission does not contain a required FDA form(s).

  - The submission does not contain the following product-identifying infor- mation: The manufacturer of the to- bacco product; the product name, in- cluding the brand and subbrand; the product category and subcategory; package type and package quantity; and characterizing flavor.

  - The type of submission is not specified.

  - The submission does not contain a signature of a responsible official, au- thorized to represent the applicant, who either resides in or has a place of business in the United States.

  - For premarket tobacco applica- tions, modified risk tobacco product applications, substantial equivalence applications, and exemption requests only: The  submission  does  not  include a valid claim of  categorical  exclusion in accordance with part 25 of this chap- ter, or an environmental assessment.

- If FDA finds that none of the rea- sons in paragraph (a) of this section ex-

#### 21 CFR Ch. I (4–1–17 Edition)

ists for refusing to accept a premarket submission, FDA may accept the sub- mission for processing and further re- view. FDA will send to  the  submitter an acknowledgement letter stating the submission has been accepted for proc- essing and further review and will pro- vide the premarket submission track- ing number.

- If FDA finds that any of the rea- sons in paragraph (a) of this section exist for refusing to accept the submis- sion, FDA will notify the submitter in writing of the reason(s) and that the submission has not been accepted, un- less insufficient contact information was provided.

## PART 1107—ESTABLISHMENT REG- ISTRATION, PRODUCT LISTING, AND SUBSTANTIAL EQUIVALENCE REPORTS

AUTHORITY: 21 U.S.C. 387e(j) and 387j.

## Subpart A—Exemptions

##### § 1107.1   Exemptions.

- General requirements. Under sec- tion 905(j)(3) of the Federal Food, Drug, and Cosmetic Act (21 U.S.C. 387e(j)(3)), FDA may exempt from the require- ments  relating  to  the   demonstration that a tobacco product is substantially equivalent within  the meaning  of sec- tion 910  of the  Federal Food,  Drug, and Cosmetic Act (21 U.S.C. 387j), tobacco products  that  are  modified   by   adding or deleting a tobacco additive, or in- creasing or decreasing the quantity  of an existing tobacco additive, if FDA determines   that:

  - Such modification would be a minor modification of a tobacco prod- uct that can be sold under the Federal Food, Drug, and Cosmetic Act (a le- gally marketed tobacco product);

  - A report under section 905(j)(1) in- tended to demonstrate substantial equivalence is not necessary to ensure that permitting the tobacco product to be marketed  would be appropriate  for protection of the public health; and

  - An exemption is otherwise appro- priate.

- Request for an exemption under sec- tion 905(j)(3) of the Federal Food, Drug,

Food and Drug Administration, HHS	Pt. 1140

and Cosmetic Act. A request for an ex- emption from the requirement of dem- onstrating substantial  equivalence may be made only by the manufacturer of a legally marketed tobacco product for a minor modification to that to- bacco product. To request an exemp- tion, the manufacturer must submit the request and all information sup- porting the request in an electronic format that FDA can process, review, and archive. If the manufacturer is un- able to submit an exemption request in an electronic format, the manufacturer may submit a written request to the Center for Tobacco Products explaining in detail why the manufacturer cannot submit the request in an electronic for- mat and requesting an alternative for- mat. Such request must include an ex- planation of why an alternative format is necessary. All submissions, including requests to submit the information in an alternative format, requests for ex- emptions, and all supporting informa- tion must be legible and in the English language. An exemption request must contain:

  - The manufacturer’s address and contact information;

  - Identification of the tobacco prod- uct(s);

  - A detailed explanation of the pur- pose of the modification;

  - A detailed description of the modification, including a statement as to whether the modification involves adding or deleting a tobacco additive, or increasing or decreasing the quan- tity of an existing tobacco additive;

  - A detailed explanation of why the modification is a minor modification of a tobacco product that can be sold under the Federal Food, Drug, and Cos- metic Act;

  - A detailed explanation of why a report under section 905(j)(1) of the Federal Food, Drug, and Cosmetic Act intended to demonstrate substantial equivalence is not necessary to ensure that permitting the tobacco product to be marketed  would be appropriate  for protection of the public health;

  - A certification (i.e., a signed statement by a responsible official of the manufacturer) summarizing the supporting evidence and providing the rationale for the official’s determina- tion that the  modification does not in-

crease the tobacco product’s appeal to or use by minors, toxicity, addictiveness, or abuse liability;

  - Other information justifying an exemption; and

  - An environmental assessment under part 25 of this chapter prepared in accordance with the requirements of

§ 25.40 of this chapter.

- Exemption determination. FDA will review the information submitted and determine whether to grant or deny an exemption request based on whether the criteria in section 905(j)(3) of the Federal Food, Drug, and Cosmetic Act are met. FDA may request additional information if necessary to make a de- termination. FDA will consider the ex- emption request withdrawn if the in- formation is not provided within the requested timeframe.

- Rescission of an exemption. FDA may rescind an exemption if it finds that the exemption is not appropriate for the protection of public health. In general, FDA will rescind an exemption only after notice and opportunity for a hearing under part 16 of this chapter is provided. However, FDA may rescind an exemption prior to notice and op- portunity for a hearing under part 16 of this chapter if the continuance of the exemption presents a serious risk to public health. In that case, FDA will provide the manufacturer an oppor- tunity for a hearing as soon as possible after the rescission.

[76 FR 38974, July 5, 2011]

## Subpart B [Reserved]

PART 1140—CIGARETTES, SMOKE- LESS TOBACCO, AND COVERED TOBACCO PRODUCTS

##### Subpart A—General Provisions

Sec.

- Scope.

- Purpose.

- Definitions.

##### Subpart B—Prohibition of Sale and Distribu- tion to Persons Younger Than 18 Years of Age

1140.10 General responsibilities of manufac- turers,  distributors,  and  retailers.

- Additional responsibilities of manu- facturers.

§ 1140.1

1140.14     Additional  responsibilities  of  retail- ers.

1140.16    Conditions of manufacture, sale, and distribution.

##### Subpart C [Reserved]

Subpart D—Labeling and Advertising

1140.30    Scope  of  permissible  forms  of  label- ing  and  advertising.

1140.32    Format   and   content   requirements for  labeling  and  advertising.

- Sale and distribution of nontobacco items and services, gifts, and sponsorship of events.

AUTHORITY:  21  U.S.C.  301  et  seq.,  Sec.  102,

Pub. L. 111–31, 123 Stat. 1776.

SOURCE:  75  FR  13230,  Mar.  19,  2010,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1140.1   Scope.

  - This part sets out the restrictions under the Federal Food, Drug, and Cos- metic Act on the sale, distribution, and use of cigarettes,  smokeless  tobacco, and covered tobacco products. Section 1140.16(d) sets out restrictions on the distribution of free samples for ciga- rettes, smokeless tobacco, and other tobacco products (as such term is de- fined in section 201 of  the  Federal Food, Drug, and Cosmetic Act).

  - The failure to comply with any applicable provision in this part in the sale, distribution, and use of ciga- rettes, smokeless tobacco, covered to- bacco products, or other tobacco prod- ucts renders the product misbranded under the Federal Food, Drug, and Cos- metic  Act.

  - References in this part to regu- latory sections of the Code of Federal Regulations are to chapter I of title 21, unless otherwise noted.

[81 FR 29102, May 10, 2016]

##### § 1140.2   Purpose.

The purpose of this part is to estab- lish restrictions on the sale, distribu- tion, and use of cigarettes, smokeless tobacco, and covered tobacco products in order to reduce the number of chil- dren and adolescents who use these products, and to reduce the life-threat- ening consequences associated with to- bacco use.

[81 FR 29102, May 10, 2016]

21 CFR Ch. I (4–1–17 Edition)

##### § 1140.3   Definitions.

For the purposes of this part:

Accessory means any product that is intended or reasonably  expected to be used with or for the human consump- tion of a tobacco product; does not con- tain tobacco and is not made or derived from tobacco; and meets either of the following:

  - Is not intended or reasonably ex- pected to affect or alter the perform- ance, composition, constituents, or characteristics of a tobacco product; or

  - Is intended or reasonably expected to affect or maintain the performance, composition, constituents, or charac- teristics of a tobacco product but

    - Solely controls  moisture  and/or temperature of a stored product; or

    - Solely provides an external heat source to initiate but not maintain combustion of a tobacco product.

Cigarette. (1) Means a product that:

- Is a tobacco product and

- Meets the definition of the term ‘‘cigarette’’ in section 3(1) of the Fed- eral Cigarette Labeling and Adver- tising Act; and

(2) Includes tobacco,  in  any  form, that is functional  in  the  product, which, because of its appearance, the type of tobacco used in the filler, or its packaging and labeling, is likely to be offered to, or purchased by, consumers as a cigarette or as roll-your-own to- bacco.

Cigarette tobacco means any product that consists of loose tobacco that is intended for use by consumers in a cig- arette. Unless otherwise stated, the re- quirements applicable to cigarettes under this chapter also apply to ciga- rette  tobacco.

Component or part means any soft- ware or assembly of materials intended or reasonably expected:

- To alter or affect the tobacco product’s performance, composition, constituents, or characteristics; or

- To be used with or for the human consumption of a tobacco product. Component or part excludes anything that is an accessory of a tobacco prod- uct.

Covered tobacco product means any to- bacco product deemed to be subject to the Federal Food, Drug, and Cosmetic Act  under  § 1100.2  of  this  chapter,  but

#### Food and Drug Administration, HHS	§ 1140.14

excludes any component or part that is not made or derived from tobacco.

Distributor means any person who fur- thers   the   distribution   of   a   tobacco product,   whether   domestic   or   im- ported, at any point from the original place of manufacture to the person who sells or distributes the product to indi- viduals for personal consumption. Com- mon  carriers  are  not  considered  dis- tributors for the purposes of this part. Importer means any person who im- ports any tobacco  product that is  in- tended for sale or distribution to con-

sumers in the United States.

Manufacturer means any person, in- cluding any repacker and/or relabeler, who manufactures, fabricates, assem- bles, processes, or labels a finished to- bacco product.

Nicotine means the chemical sub- stance	named	3-(1-Methyl-2- pyrrolidinyl)pyridine or C[10]H[14]N[2], including any salt or complex of nico- tine.

Package  or  packaging  means  a  pack, box,  carton,  or  container  of  any  kind or, if no other container, any wrapping (including  cellophane)  in  which  a  to- bacco  product  is  offered  for  sale,  sold, or otherwise distributed to consumers. Point  of  sale  means  any  location  at which a consumer can purchase or oth- erwise obtain tobacco products for per-

sonal consumption.

Retailer means any person who sells tobacco products to individuals for per- sonal consumption, or who operates a facility where vending machines or self-service displays are permitted under this part.

Roll-your-own tobacco means any to- bacco product that, because of its ap- pearance, type, packaging, or labeling, is suitable for use and likely to be of- fered to, or purchased by, consumers as tobacco for making cigarettes.

Smokeless tobacco means any tobacco product that consists of cut, ground, powdered, or leaf tobacco and that is intended to be placed in the oral or nasal  cavity.

Tobacco product. As stated in section 201(rr) of the Federal Food, Drug, and Cosmetic Act in relevant part,  a  to- bacco product:

  - Means any product made or de- rived from tobacco that is intended for human   consumption,   including   any

component, part, or accessory of a to- bacco product (except for raw mate- rials other than tobacco used in manu- facturing a component, part, or acces- sory of a tobacco product) and

  - Does not mean an article that is a drug under section 201(g)(1) of the Fed- eral Food, Drug, and Cosmetic Act, a device under section 201(h) of the Fed- eral Food,  Drug,  and  Cosmetic  Act,  or a combination product described in section 503(g) of the Federal Food, Drug, and Cosmetic Act.

[81 FR 29102, May 10, 2016]

## Subpart B—Prohibition of Sale and Distribution to Persons Young- er Than 18 Years of Age

##### § 1140.10 General responsibilities of manufacturers, distributors, and re- tailers.

Each manufacturer, distributor, im- porter, and retailer is responsible for ensuring that the cigarettes, smokeless tobacco, or covered tobacco products it manufactures, labels, advertises, pack- ages, distributes, imports, sells, or oth- erwise holds for sale comply with all applicable requirements under this part.

[81 FR 29103, May 10, 2016]

##### § 1140.12 Additional responsibilities of manufacturers.

In addition to the other responsibil- ities under this part, each manufac- turer shall remove from each point of sale all self-service displays, adver- tising, labeling, and other items that the manufacturer owns that do not comply with the requirements under this part.

##### § 1140.14 Additional responsibilities of retailers.

    - In addition to the other require- ments under this part, each cigarette and smokeless tobacco retailer is re- sponsible for ensuring that all sales of cigarettes or smokeless tobacco to any person comply with the following re- quirements:

      - No retailer may sell cigarettes or smokeless tobacco to any person younger than 18 years of age;

(2)(i) Except as otherwise provided in paragraph (a)(2)(ii) of this section and

§ 1140.16

in § 1140.16(c)(2)(i), each retailer must verify by means of photographic identi- fication containing the bearer’s date of birth that no person purchasing the product is younger than 18 years of age;

(ii) No such verification is required for any person over the age of 26;

        - Except  as  otherwise  provided   in

§ 1140.16(c)(2)(ii), a retailer may sell cigarettes or  smokeless  tobacco  only  in a direct, face-to-face exchange without the assistance of any electronic or me- chanical device (such as a vending ma- chine);

        - No retailer may break or other- wise open any cigarette or smokeless tobacco package to sell or distribute individual cigarettes or a number of unpackaged cigarettes that is smaller than the quantity in the minimum cig- arette     package     size     defined     in

§ 1140.16(b), or any quantity of cigarette tobacco or smokeless tobacco that is smaller than the smallest package dis- tributed by the manufacturer for indi- vidual consumer use; and

        - Each retailer must ensure that all self-service displays, advertising, label- ing, and other items, that are located in the retailer’s establishment and that do not comply with the requirements of this part, are removed or are brought into compliance with the re- quirements under this part.

    - Notwithstanding the require- ments in paragraph (a) of this section and in addition to the other require- ments under this part, each retailer of covered tobacco products is responsible for ensuring that all sales of such cov- ered tobacco products to any person comply with the following require- ments:

- No retailer may sell covered to- bacco products to any person younger than 18 years of age;

(2)(i) Except as otherwise provided in paragraph (a)(2)(ii) of this  section  and in § 1140.16(c)(2)(i), each retailer must verify by means of photographic identi- fication containing the bearer’s date of birth that no person purchasing the product is younger than 18 years of age;

(ii) No such verification is required for any person over the age of 26; and

(3) A retailer may not sell covered to- bacco products with the assistance of any electronic or mechanical device (such as a vending machine), except in

#### 21 CFR Ch. I (4–1–17 Edition)

facilities where the retailer  ensures that no person younger than 18 years of age is present, or permitted to enter, at any  time.

[81 FR 29103, May 10, 2016]

##### § 1140.16   Conditions   of   manufacture, sale, and distribution.

- Restriction on product names. A manufacturer shall not use a trade or brand name of a nontobacco product as the trade or brand name for a cigarette or smokeless tobacco product, except for a tobacco product whose trade or brand name was on both a tobacco product and a nontobacco product that were sold in the United States on Janu- ary 1, 1995.

- Minimum cigarette package size. Ex- cept as otherwise provided under this section,  no  manufacturer,  distributor, or retailer may sell or cause to be sold, or distribute or cause to be distributed, any cigarette package that contains fewer than 20 cigarettes.

- Vending  machines,  self-service  dis- plays, mail-order sales, and other ‘‘imper- sonal’’ modes of sale. (1) Except as oth- erwise  provided  under  this  section,  a retailer may sell cigarettes and smoke- less  tobacco  only  in  a  direct,  face-to- face exchange between the retailer and the consumer. Examples of methods of sale   that   are   not   permitted   include vending  machines  and  self-service  dis- plays.

- Exceptions. The following methods of sale are permitted:

  - Mail-order sales, excluding mail- order redemption of coupons and dis- tribution of free samples through the mail; and

  - Vending machines (including vending machines that sell packaged, single cigarettes) and self-service dis- plays that are  located  in  facilities where the retailer ensures that no per- son younger than 18 years of age is present, or permitted to enter, at any time.

(d)(1) Except as provided in para- graph (d)(2) of this section, no manu- facturer, distributor, or retailer may distribute or  cause  to  be  distributed any free samples of cigarettes, smoke- less tobacco, or other tobacco products (as such term is  defined  in  section  201 of the Federal Food, Drug,  and  Cos- metic   Act).

#### Food and Drug Administration, HHS	§ 1140.16

(2)(i) Paragraph (d)(1) of this section does not prohibit a manufacturer, dis- tributor, or retailer from  distributing or causing to be distributed free sam- ples of smokeless tobacco in a qualified adult-only    facility.

    - Paragraph (d)(2) of this section does not affect the authority of a State or local government to prohibit or oth- erwise restrict the distribution of free samples of smokeless tobacco.

    - For purposes of paragraph (d) of this section, the term ‘‘qualified adult- only facility’’ means a facility or re- stricted area that:

      - Requires each person present to provide to a law enforcement officer (whether on or off duty) or to a secu- rity guard licensed by a governmental entity government-issued  identifica- tion showing a photograph and at least the minimum age established by appli- cable law for the purchase of smokeless tobacco;

      - Does not sell, serve, or distribute alcohol;

      - Is not located adjacent to or im- mediately across from (in any direc- tion) a space that is used primarily for youth-oriented	marketing,	pro- motional,  or  other  activities;

      - Is a temporary structure con- structed, designated, and operated as a distinct enclosed area for  the  purpose of distributing free samples of smoke- less tobacco in accordance with this paragraph (d)(2) of this section;

      - Is enclosed by a barrier that:

        - Is constructed of, or covered with, an opaque material (except for en- trances and exits);

        - Extends from no more than 12 inches above the ground or floor (which area at the bottom of the barrier must be covered with material that restricts visibility but may allow airflow) to at least 8 feet above the ground  or  floor (or to the ceiling); and

        - Prevents persons outside the qualified adult-only facility from see- ing into the qualified adult-only facil- ity, unless they make unreasonable ef- forts to do so; and

      - Does not display on its exterior:

        - Any tobacco product  advertising;

        - A brand name other than in con- junction with words for an area or en- closure to identify an adult-only facil- ity; or

        - Any combination of words that would imply to a reasonable observer that the manufacturer, distributor, or retailer has a sponsorship that would violate § 1140.34(c).

    - Distribution of samples of smokeless tobacco under  paragraph (d)(2) of this section permitted to be taken out of the qualified adult-only facility shall be limited to one  package per adult consumer containing no more than 0.53 ounces (15 grams) of smoke- less tobacco. If such package of smoke- less tobacco contains individual por- tions of smokeless tobacco, the indi- vidual portions of smokeless tobacco shall not exceed eight individual por- tions, and the collective weight of such individual    portions    shall    not    exceed

0.53 ounces (15 grams). Any manufac- turer, distributor, or retailer who dis- tributes or causes to be distributed free samples  also   shall   take   reasonable steps  to  ensure  that  the  amounts  in this paragraph (d)(2)(iv) are limited to one such package per adult consumer per  day.

- Notwithstanding paragraph (d)(2) of this section, no manufacturer, dis- tributor, or retailer may distribute or cause to be distributed any free sam- ples of smokeless tobacco:

  - To a sports team or entertainment group; or

  - At any football, basketball, base- ball, soccer, or hockey event or any other sporting or entertainment event determined by the Secretary to be cov- ered by paragraph (d)(3) of this section.

- The Secretary shall implement a program to ensure compliance with paragraph (d) of this section and sub- mit a report to the Congress on such compliance not later than 18 months after the date of enactment of the Family Smoking Prevention and To- bacco  Control  Act.

- Nothing in paragraph (d) of this section shall be construed to authorize any person to distribute or cause to be distributed any sample of a tobacco product to any individual who has not attained the minimum age established by applicable law for the purchase of such product.

#### § 1140.30

(e) Restrictions on labels, labeling, and advertising. No manufacturer, dis- tributor, or retailer may sell or dis- tribute, or cause to be sold or distrib- uted, cigarettes or smokeless tobacco with labels, labeling, or advertising not in compliance  with  subpart  D of  this part, and other applicable require- ments.

## Subpart C [Reserved]

Subpart D—Labeling and Advertising

##### § 1140.30   Scope of permissible forms of labeling and advertising.

(a)(1) A manufacturer, distributor, or retailer may, in accordance with this subpart D, disseminate or cause to be disseminated advertising or labeling which bears a cigarette or smokeless tobacco brand name (alone or in con- junction with any other word) or any other indicia of tobacco product identi- fication, in newspapers; in  magazines; in periodicals or other publications (whether periodic or limited distribu- tion); on billboards, posters, and plac- ards; in nonpoint-of-sale promotional material (including direct mail); in point-of-sale   promotional   material; and in audio or video formats delivered at a point-of-sale.

(2) A manufacturer, distributor, or retailer intending to disseminate, or to cause to be disseminated,  advertising or labeling for cigarettes or smokeless tobacco in a medium that is not listed in paragraph (a)(1) of this section, shall notify the agency 30 days prior to the use of such medium. The notice shall describe the medium and discuss the extent to which the advertising or la- beling may be seen by persons younger than 18 years of age. The manufac- turer, distributor, or retailer shall send this notice to the Office of Compliance, Center for Tobacco Products, Food and Drug Administration, 9200 Corporate Blvd., Rockville, MD, 20850–3229.

- [Reserved]

- This subpart D does not apply to cigarette or smokeless tobacco pack- age labels.

21 CFR Ch. I (4–1–17 Edition)

##### § 1140.32 Format and content require- ments for labeling and advertising.

- Except  as  provided  in  paragraph

- of this section, each manufacturer, distributor, and retailer advertising or causing to be advertised, disseminating or causing to be disseminated, any la- beling or advertising for cigarettes or smokeless tobacco shall use only black text on a white background. This sec- tion does not apply to advertising:

  - In any facility where vending ma- chines and self- service displays are permitted under this part, provided that the advertising is not visible from outside the facility and that it is af- fixed to a wall or fixture in the facility; or

  - Appearing in any publication (whether periodic or limited distribu- tion) that the manufacturer, dis- tributor, or retailer demonstrates is an adult publication. For the purposes of this section, an adult publication is a newspaper, magazine, periodical, or other publication:

- Whose readers younger than 18 years of age constitute 15 percent or less of the total readership as measured by competent and reliable survey evi- dence; and

- That is read by fewer than 2 mil- lion persons younger than 18 years of age as measured by competent and reli- able survey evidence.

- Labeling and advertising in an audio or video format shall be limited as follows:

  - Audio format shall be limited to words only with no music or sound ef- fects.

  - Video formats shall be limited to static black text only on a white back- ground. Any audio with the video shall be limited to words only with no music or sound effects.

##### § 1140.34 Sale and distribution of non- tobacco items and services, gifts, and sponsorship of events.

- No manufacturer and no dis- tributor of imported cigarettes or smokeless tobacco may market,  li- cense, distribute, sell, or cause to be marketed, licensed, distributed, or sold any item (other than cigarettes or smokeless tobacco or roll-your-own paper) or service, which bears the brand  name  (alone  or  in  conjunction

#### Food and Drug Administration, HHS	§ 1141.1

with any other word), logo, symbol, motto, selling message, recognizable color or pattern of colors, or any other indicia of product identification iden- tical or similar to, or identifiable with, those used for any brand of cigarettes or smokeless tobacco.

- No manufacturer, distributor, or retailer may offer or cause to be of- fered any gift or item (other than ciga- rettes or smokeless tobacco) to any person purchasing cigarettes or smoke- less tobacco in consideration of the purchase thereof, or to any person in consideration of furnishing evidence, such as credits, proofs-of-purchase, or coupons, of such a purchase.

- No manufacturer, distributor, or retailer may sponsor or cause to be sponsored any athletic, musical, artis- tic, or other social or cultural event, or any entry or team in any event, in the brand name (alone or in conjunction with any other word), logo, symbol, motto, selling message, recognizable color or pattern of colors, or any other indicia of product identification iden- tical or similar to, or identifiable with, those used for any brand of cigarettes or smokeless tobacco. Nothing in this paragraph prevents  a  manufacturer, distributor, or retailer from sponsoring or causing to be sponsored any ath- letic, musical, artistic, or  other  social or cultural event, or team or entry, in the name of the corporation which manufactures the tobacco product, pro- vided that both the corporate name and the corporation were registered and in use in the United States prior to January 1, 1995, and that the corporate name does not include any brand name (alone or in conjunction with any other word), logo, symbol, motto, selling message, recognizable color or pattern of colors, or any other indicia of prod- uct identification identical or similar to, or identifiable with, those used for any brand of cigarettes or smokeless tobacco.

## PART 1141—CIGARETTE PACKAGE AND ADVERTISING WARNINGS

##### Subpart A—General Provisions

Sec.

1141.1   Scope.

1141.3   Definitions.

##### Subpart B—Cigarette Package and Advertising Warnings

1141.10   Required warnings.

1141.12   Incorporation   by   reference   of   re- quired  warnings.

1141.14    Misbranding of cigarettes.

##### Subpart C—Additional Disclosure Require- ments for Cigarette Packages and Ad- vertising

- Disclosures regarding cessation. AUTHORITY:  15  U.S.C.  1333;  21  U.S.C.  371,

387c, 387f; Secs. 201 and 202, Pub. L. 111–31, 123

Stat. 1776.

SOURCE:  76  FR  36753,  June  22,  2011,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1141.1   Scope.

  - This part sets forth the require- ments for the display of health warn- ings on cigarette packages and in ad- vertisements for cigarettes. FDA may require additional statements to be dis- played on packages and in advertise- ments under the Federal Food, Drug, and Cosmetic Act or other authorities.

  - The requirements of this part do not apply to manufacturers or distribu- tors of cigarettes that do not manufac- ture, package, or import cigarettes for sale or distribution within the United States.

  - A cigarette retailer shall not be considered in violation of this part as it applies to the display of health warn- ings on a cigarette package if the pack- age:

    - Contains a health warning;

    - Is supplied to the retailer by a license- or permit-holding tobacco product manufacturer, importer, or distributor; and

    - Is not altered by the retailer in a way that is material to the require- ments of section 4(a) of the Federal Cigarette Labeling and Advertising Act (15 U.S.C. 1333(a)) or this part, includ- ing by obscuring the warning, by re- ducing its size, by severing it in whole or in part, or by otherwise changing it in a material way.

  - A cigarette retailer shall not be considered in violation  of  this  part  as it applies to the display of health warn- ings in an advertisement for cigarettes if  the  advertisement  is  not  created  by

§ 1141.3

or on behalf of the retailer and the re- tailer is not otherwise responsible for the inclusion of the required warnings. This paragraph shall not relieve a re- tailer of liability if the  retailer  dis- plays, in a location open to the public, an advertisement that does not contain a health warning or that contains a warning that has been altered by the retailer in a way that is  material  to the requirements of section 4(b) of the Federal Cigarette Labeling and Adver- tising Act (15 U.S.C. 1333(b)), this part, or section 4(c) of the Federal Cigarette Labeling and Advertising Act (15 U.S.C. 1333(c)), including by obscuring the warning, by reducing its size, by sev- ering it in whole or in part, or by oth- erwise changing it in a material way.

##### § 1141.3   Definitions.

For the purposes of this part,

Cigarette means:

    - Any roll of tobacco wrapped in paper or in any substance not con- taining tobacco; and

    - Any roll of tobacco wrapped in any substance containing tobacco which, because of its appearance, the type of tobacco used in the filler, or its packaging and labeling, is likely to be offered to, or purchased by, consumers as  a  cigarette  described  in  paragraph

- of this definition.

Commerce means:

  - Commerce between any State, the District of Columbia, the Common- wealth of Puerto Rico, Guam, the U.S. Virgin Islands, American Samoa, Wake Island, Midway Islands, Kingman Reef, or Johnston Island and any place out- side thereof;

  - Commerce between points in any State, the District of Columbia, the Commonwealth of Puerto Rico, Guam, the U.S.  Virgin  Islands,  American Samoa, Wake Island, Midway Islands, Kingman Reef, or Johnston Island, but through any place outside thereof; or

  - Commerce wholly within the Dis- trict of Columbia, Guam, the U.S. Vir- gin Islands, American Samoa, Wake Is- land, Midway Island, Kingman Reef, or Johnston Island.

Distributor means any person who fur- thers the distribution of cigarettes at any point from the original place of manufacture to the person who sells or distributes the product to individuals

#### 21 CFR Ch. I (4–1–17 Edition)

for personal consumption. Common carriers are not considered distributors for the purposes of this part.

Front panel and rear panel mean the two largest sides or surfaces of the package.

Importer means any person who im- ports any cigarette that is intended for sale or distribution to consumers in the United States.

Manufacturer means any person, in- cluding any repacker or relabeler, who manufactures, fabricates, assembles, processes, or labels a finished cigarette product.

Package means a pack,  box,  carton, or container of any kind in which ciga- rettes are offered for sale, sold, or oth- erwise distributed to consumers.

Person means an individual, partner- ship, corporation, or any other busi- ness or legal entity.

Required  warning  means  the  combina- tion    of    one    of    the    textual    warning statements  and  its  accompanying  color graphic,  which  are  set  forth  in  ‘‘Ciga- rette  Required  Warnings,’’  which  is  in- corporated by reference at § 1141.12.

Retailer means any person who sells cigarettes to individuals for personal consumption, or who operates a facil- ity where vending machines or self- service displays of cigarettes are per- mitted.

United States, when used in a geo- graphical sense, includes the several States, the District of Columbia, the Commonwealth of Puerto Rico, Guam, the U.S. Virgin Islands, American Samoa, Wake Island, Midway Islands, Kingman Reef, and Johnston Island. The term ‘‘State’’ includes any polit- ical division of any State.

## Subpart B—Cigarette Package and Advertising Warnings

##### § 1141.10   Required warnings.

- Packages. (1) It shall be unlawful for any person to manufacture, pack- age, sell, offer to sell, distribute, or im- port for sale or distribution within the United States any cigarettes the pack- age of which fails to bear, in accord- ance with section 4 of the Federal Ciga- rette Labeling and Advertising Act (15

U.S.C. 1333) and this part, one of the re- quired warnings on the front and the rear panels.

Food and Drug Administration, HHS	§ 1141.12

- The required warning shall be ob- tained from the electronic images con- tained in ‘‘Cigarette Required Warn- ings,’’ which is incorporated by ref- erence at § 1141.12, and accurately re- produced as specified in ‘‘Cigarette Re- quired Warnings.’’

- The required warning shall appear directly on the package and shall be clearly visible underneath the cello- phane or other clear wrapping.

- The required warning shall be lo- cated in the upper portion of the front and rear panels  of  the  package  and shall comprise at least the top 50 per- cent of these panels; Provided, however, that on cigarette cartons, the required warning shall be located on the left side of the front and rear panels of the carton and shall comprise at least  the left 50 percent of these panels.

- The required warning shall be po- sitioned such that the text of the re- quired warning and the other informa- tion on that panel of the package have the same orientation.

- Advertisements. (1) It  shall be un- lawful for any manufacturer, importer, distributor, or retailer of cigarettes to advertise or cause to be advertised within the United States any cigarette unless its advertising bears, in accord- ance with section 4 of the Federal Ciga- rette  Labeling  and  Advertising  Act  (15

U.S.C. 1333) and this part, one of the re- quired warnings.

- The text in each required warning shall be in the English language, ex- cept that:

  - In the case of an advertisement that appears in a non-English publica- tion, the text in the required warning shall appear in the predominant lan- guage of the publication whether or not the advertisement is in English; and

  - In the case of an advertisement that appears in an English language publication but that is not in English, the text in the required warning shall appear in the same language as that principally used in the advertisement.

- For    English-language    and    Span- ish-language    warnings,    each    required warning   shall   be   obtained   from   the electronic   images   contained   in   ‘‘Ciga- rette  Required  Warnings,’’  which  is  in- corporated by reference at § 1141.12, and

accurately reproduced as specified in ‘‘Cigarette Required Warnings.’’

- For foreign-language warnings, except for Spanish-language warnings, each required warning shall be ob- tained from the electronic images con- tained in ‘‘Cigarette Required Warn- ings,’’ which is incorporated by ref- erence at § 1141.12, and accurately re- produced as specified in ‘‘Cigarette Re- quired Warnings,’’ including the inser- tion of a true  and  accurate  translation of the textual warning. The inserted textual warning must comply with the requirements of section 4(b)(2) of the Federal Cigarette Labeling and Adver- tising Act (15 U.S.C. 1333(b)(2)).

- The required warning shall occupy at least 20 percent of the area of each advertisement, and shall be placed in accordance with the requirements in the Federal Cigarette Labeling and Ad- vertising   Act.

- Irremovable or permanent warnings. The required warnings shall be indeli- bly printed on or permanently affixed to the package or advertisement. Such warnings, for example, must not be printed or placed on a label affixed to a clear outer wrapper that is likely to be removed to access the product within the package.

##### § 1141.12   Incorporation by reference of required warnings.

‘‘Cigarette Required Warnings’’ Edi- tion 1.0 (June 2011), consisting of elec- tronic files,  U.S. Food and  Drug Ad- ministration,   referred   to   at   § 1141.3,

§ 1141.10(a) and (b), and § 1141.16(a), is in- corporated by reference into this sec- tion with the approval of the Director of the Federal Register under 5 U.S.C. 552(a)  and 1  CFR part 51.  To enforce any edition other than that specified in this section, FDA must publish notice of change in the FEDERAL REGISTER and the material must be available to the public. All approved material is avail- able for inspection at the National Ar- chives   and   Records   Administration (NARA). For information on the avail- ability  of  this  material  at  NARA,  call 202–741–6030

#### § 1141.14

. Also, you may ob- tain a copy of the material by con- tacting the Center for Tobacco Prod- ucts, Food and Drug Administration, Office of Health Communication and Education, ATTN: Cigarette Warning File Requests, 9200 Corporate Blvd., Rockville, MD 20850, 1–877–CTP–1373

##### § 1141.14   Misbranding of cigarettes.

  - A cigarette shall be deemed to be misbranded under section 903(a)(1) of the Federal Food, Drug, and Cosmetic Act if its package does not bear one of the required warnings in accordance with section 4 of the Federal Cigarette Labeling and Advertising Act (15 U.S.C. 1333) and this part. A cigarette shall be deemed to be misbranded under section 903(a)(7)(A) of the Federal Food, Drug, and Cosmetic Act if  its  advertising does not bear one of the required warn- ings in accordance with section 4 of the Federal Cigarette Labeling and Adver- tising Act (15 U.S.C. 1333) and this part.

  - A cigarette advertisement or package will be  deemed  to  include  a brief statement   of  relevant  warnings for the purposes of section 903(a)(8) of the  Federal  Food,  Drug,  and  Cosmetic Act if it bears one of the required warnings in accordance with  section  4 of the Federal Cigarette Labeling and Advertising Act  (15  U.S.C.  1333)  and this part. A cigarette distributed or of- fered for sale in any State shall be deemed to be misbranded under section 903(a)(8) of the Federal Food, Drug, and Cosmetic Act unless the manufacturer, packer, or distributor includes in all advertisements and packages issued or caused to be issued by  the  manufac- turer, packer, or distributor with  re- spect to the cigarette one of  the  re- quired warnings in accordance with section 4 of the Federal Cigarette La- beling and Advertising Act (15 U.S.C. 1333) and this part.

#### 21 CFR Ch. I (4–1–17 Edition)

Subpart C—Additional Disclosure Requirements for Cigarette Packages and Advertising

##### § 1141.16 Disclosures regarding ces- sation.

- The   required   warning   shall   in- clude  a  reference  to  a  smoking  ces- sation   assistance   resource   in   accord- ance  with,  and  as  specified  in,  ‘‘Ciga- rette	Required	Warnings’’	(incor- porated by reference at § 1141.12).

- In meeting the smoking cessation needs of an individual caller, the smok- ing cessation assistance resource re- quired to be referenced by paragraph

- of this section must, as appropriate:

  - Provide factual information about the harms to health associated with cigarette smoking and the health bene- fits of quitting smoking;

  - Provide factual information about what smokers can expect when trying to quit;

  - Provide practical advice (problem solving/skills training) about how  to deal with common issues faced by smokers trying to quit;

  - Provide evidence-based advice about how to formulate a plan to quit smoking;

  - Provide evidence-based informa- tion about effective relapse prevention strategies;

  - Provide factual information on smoking cessation treatments, includ- ing FDA-approved cessation medica- tions; and

  - Provide information, advice, and support that is evidence-based, unbi- ased (including with respect to prod- ucts, services, persons, and other enti- ties), and relevant to tobacco ces- sation.

- The smoking cessation resource must:

  - Other than as described in this section, not advertise or promote any particular product or service;

  - Except to meet the particularized needs of an individual caller as deter- mined in the context of individual counseling, not selectively present in- formation about a subset of FDA-ap- proved cessation products or product categories while failing to mention other FDA-approved cessation products or product categories;

#### Food and Drug Administration, HHS	§ 1143.1

  - Not provide or otherwise encour- age the use of any drug or other med- ical product that FDA has not ap- proved for tobacco cessation;

  - Not encourage the use of any non- evidence-based smoking cessation prac- tices;

  - Ensure that staff providing smok- ing cessation information, advice, and support are trained specifically to help smokers quit by delivering  unbiased and evidence-based information, ad- vice, and support; and

  - Maintain appropriate controls to ensure the criteria described in para- graphs (b) and (c) of this section are met.

- If the Secretary of the Depart- ment of Health and Human Services (Secretary) determines that  a part of the smoking cessation assistance re- source referenced by paragraph (a) of this section does not meet the criteria described in paragraphs (b) and (c) of this section, the Secretary shall take appropriate steps to address the non- compliance.

## PART 1143—MINIMUM REQUIRED WARNING STATEMENTS

Sec.

1143.1   Definitions.

1143.3    Required  warning  statement  regard- ing  addictiveness  of  nicotine.

1143.5    Required  warning  statements  for  ci- gars.

1143.7    Language  requirements  for  required warning    statements.

1143.9    Irremovable   or   permanent   required warning    statements.

1143.11    Does  not  apply  to  foreign  distribu- tion.

- Effective date.

AUTHORITY: 21 U.S.C. 387a(b), 387f(d).

SOURCE:  81  FR  29103,  May  10,  2016,  unless

otherwise noted.

##### § 1143.1   Definitions.

For purposes of this part:

Accessory means any product that is intended or reasonably  expected to be used with or for the human consump- tion of a tobacco product; does not con- tain tobacco and is not made or derived from tobacco; and meets either of the following:

  - Is not intended or reasonably ex- pected  to  affect  or  alter  the  perform-

ance,   composition,   constituents,   or characteristics of a tobacco product; or

  - Is intended or reasonably expected to affect or maintain the performance, composition, constituents, or charac- teristics of a tobacco product but

    - Solely controls  moisture  and/or temperature of a stored tobacco prod- uct; or

    - Solely provides an external heat source to initiate but not maintain combustion of a tobacco product

Cigar means a tobacco product that:

- Is not a cigarette and

- Is a roll of tobacco wrapped in leaf tobacco or any substance containing tobacco.

Cigarette tobacco means any product that consists of loose tobacco that is intended for use by consumers in a cig- arette. Unless otherwise stated, the re- quirements applicable to cigarettes under this chapter also apply to ciga- rette  tobacco.

Component or part means any soft- ware or assembly of materials intended or reasonably expected:

- To alter or affect the tobacco product’s performance, composition, constituents, or characteristics; or

- to be used with or for the human consumption of a tobacco product. Component or part excludes anything that is an accessory of a tobacco prod- uct.

Covered tobacco product means any to- bacco product deemed to be subject  to the  Federal  Food,  Drug,  and  Cosmetic Act pursuant to § 1100.2 of this chapter, but excludes any  component  or  part that is not made or derived from to- bacco.

Package  or  packaging  means  a  pack, box,  carton,  or  container  of  any  kind or, if no other container, any wrapping (including  cellophane),  in  which  a  to- bacco  product  is  offered  for  sale,  sold, or otherwise distributed to consumers. Point  of  sale  means  any  location  at which a consumer can purchase or oth- erwise obtain tobacco products for per-

sonal consumption.

Principal display panels means the panels of a package that are most like- ly to be displayed, presented, shown, or examined by the consumer.

Required warning statement means a textual warning statement required to be on packaging and in advertisements

§ 1143.3

for cigarette tobacco, roll-your-own to- bacco, cigars, and other covered to- bacco products.

Retailer means any person who sells tobacco products to individuals for per- sonal consumption, or who operates a facility where vending machines or self-service displays are permitted under this part.

Roll-your-own tobacco means any to- bacco product that, because of its ap- pearance, type, packaging, or labeling, is suitable for use and likely to be of- fered to, or purchased by, consumers as tobacco for making cigarettes.

Tobacco product. As stated in section 201(rr) of the Federal Food, Drug, and Cosmetic Act in relevant part,  a  to- bacco product:

  - Means any product made or de- rived from tobacco that is intended for human consumption, including any component, part, or accessory of a to- bacco product (except for raw mate- rials other than tobacco used in manu- facturing a component, part, or acces- sory of a tobacco product) and

  - Does not mean an article that is a drug under section 201(g)(1) of the Fed- eral Food, Drug, and Cosmetic Act, a device under section 201(h) of the Fed- eral Food,  Drug,  and  Cosmetic  Act,  or a combination product described in section 503(g) of the Federal Food, Drug, and Cosmetic Act.

##### § 1143.3   Required   warning   statement regarding addictiveness of nicotine.

- Packages. (1) For cigarette to- bacco, roll-your-own tobacco, and cov- ered tobacco products other than ci- gars, it is unlawful for any person to manufacture,  package,  sell,  offer   to sell, distribute, or import for sale or distribution within the United States such product unless the tobacco prod- uct package bears the following re- quired warning statement on the pack- age label: ‘‘WARNING: This product contains nicotine. Nicotine is an ad- dictive   chemical.’’

- The required warning statement must appear directly on the package and must be clearly visible underneath any cellophane or other clear wrapping as follows:

  - Be located in a conspicuous and prominent place on the two principal display  panels  of  the  package  and  the

#### 21 CFR Ch. I (4–1–17 Edition)

warning area must comprise at least 30 percent of each of the principal display panels;

  - Be printed in at least 12-point font size and ensures that the required warning statement occupies the great- est possible proportion of the warning area set aside for the required text;

  - Be printed in conspicuous and legible Helvetica bold or Arial bold type (or other sans serif fonts) and in black text on a white background or white text on a black background in a manner that contrasts by typography, layout, or color, with all other printed material on the package;

  - Be capitalized and punctuated as indicated in paragraph (a)(1) of this section; and

  - Be centered in the warning area in which the text is required to be printed and positioned such that the text of the required warning statement and the other information on the prin- cipal display panel have the same ori- entation.

- A retailer of any tobacco product covered by paragraphs (a)(1) and (2) of this section will not be in violation  of this section for packaging that:

  - Contains a health warning;

  - Is supplied to the retailer by the tobacco product manufacturer, im- porter, or distributor, who has the re- quired state, local, or Alcohol and To- bacco Tax and Trade Bureau (TTB)- issued license or permit, if applicable, and

  - Is not altered by the retailer in a way that is material to the require- ments of this section.

- Advertisements. (1) For cigarette tobacco, roll-your-own tobacco, and covered tobacco products other than ci- gars, it is unlawful for any such to- bacco product manufacturer, packager, importer, distributor, or retailer of the tobacco product to advertise  or  cause to be advertised within the United States any tobacco product unless each advertisement bears the required warn- ing statement specified in paragraph (a)(1) of this section.

- For print advertisements and other advertisements with a visual component (including, for example, ad- vertisements on signs, shelf-talkers, Internet Web pages, and electronic mail    correspondence),    the    required

#### Food and Drug Administration, HHS	§ 1143.5

warning statement must appear in the upper portion of the area of the adver- tisement within the trim area as fol- lows:

  - Occupy at least 20 percent of the area of the advertisement;

  - Appear in at least 12-point font size and ensures that the required warning statement occupies the great- est possible proportion of the warning area set aside for the required text;

  - Appear in conspicuous and leg- ible Helvetica bold or Arial bold  type (or other similar sans  serif  fonts)  and in black text on a white background or white text  on  a black  background  in a manner that contrasts by typography, layout, or color, with all other mate- rial on the advertisement;

  - Be capitalized and punctuated as indicated in paragraph (a)(1) of this section;

  - Be centered in the warning area in which the text is required to appear and positioned such that the text  of the required warning statement and the other textual information in the advertisement have the same orienta- tion; and

  - Be surrounded by a rectangular border that is the same color as the text of the required warning statement and that is not less than 3 millimeters (mm) or more than 4 mm.

- This paragraph (b) applies to a re- tailer only if that retailer is respon- sible for or directs the health warning required under the paragraph. How- ever, this paragraph does not relieve a retailer of liability if the retailer dis- plays, in a location open to the public, an advertisement that does not contain a health warning or contains a health warning that has been altered by the retailer in a way that is  material  to the requirements of this section.

- Self-certification. A tobacco prod- uct that would otherwise be required to bear the warning in paragraph (a)(1) of this section but does not contain nico- tine is not required to bear the warning in paragraph (a)(1) of this section on packages or advertisements if the to- bacco product manufacturer has sub- mitted to FDA a confirmation  state- ment certifying to be true and accurate that the product does not contain nico- tine and that the tobacco product man- ufacturer  has  data  to  support  that  as-

sertion. Any product not required to bear the warning in paragraph (a)(1) of this section must include the  state- ment ‘‘This product is made from to- bacco.’’ on all packages and advertise- ments in accordance with the require- ments of this part.

- Small packages. A tobacco product that would otherwise be required to bear the warning in paragraph (a)(1) of this section but is too small or other- wise unable to accommodate a label with sufficient space to bear such in- formation is exempt from compliance with the requirement provided that the information and specifications required under paragraphs (a)(1) and (2) of this section appear on the carton or other outer container or wrapper if the car- ton, outer container, or wrapper has sufficient space to bear the informa- tion, or appear on a tag otherwise firm- ly and permanently affixed to the to- bacco product package. In such cases, the carton, outer container,  wrapper, or tag will serve as the location of the principal display panels.

##### § 1143.5   Required  warning  statements for cigars.

(a) Packages. (1) It is unlawful for any person to manufacture, package, sell, offer to sell, distribute, or import for sale or distribution within the United States any cigar product unless the product package bears one of the fol- lowing required warning statements on the package label:

- WARNING: Cigar smoking can cause cancers of the mouth and throat, even if you do not inhale.

- WARNING: Cigar smoking can cause lung cancer and heart disease.

- WARNING: Cigars are not a safe alternative to cigarettes.

- WARNING: Tobacco smoke in- creases the risk of lung  cancer  and heart disease, even in nonsmokers.

(v)(A) WARNING: Cigar use while pregnant can harm you and your baby.; or

(B) SURGEON GENERAL WARNING: Tobacco Use Increases the Risk of In- fertility, Stillbirth and Low Birth Weight.

- WARNING: This product con- tains nicotine. Nicotine is an addictive chemical.

#### § 1143.5

- Each required warning statement must appear directly on the package and must be clearly visible underneath any cellophane or other clear wrapping as follows:

  - Be located in a conspicuous and prominent place on the two principal display panels of the package and the warning area must comprise at least 30 percent of each of the principal display panels;

  - Appear in at least 12-point font size and ensure that the required warn- ing statement occupies the greatest possible proportion of the warning area set aside for the required text;

  - Be printed in conspicuous and legible Helvetica bold or Arial bold type (or other similar sans serif fonts) and in black text on a white back- ground or white text on a black back- ground in a manner that contrasts by typography, layout, or color, with all other printed material on the package;

  - Be capitalized and punctuated as indicated in paragraph (a)(1) of this section; and

  - Be centered in the warning area in which the text is required to be printed and positioned such that the text of the required warning statement and the other information on that prin- cipal display panel have the same ori- entation.

- No person may manufacture, package, sell, offer to sell,  distribute, or import for sale or distribution with- in the United States any cigar without a required warning statement, except for cigars that are sold individually and not in a product package. For ci- gars that are sold individually and not in a product package, the required warning statements must be posted at the retailer’s point-of-sale in accord- ance with the following:

- All of the warnings in  paragraph

  - of this section must be placed on a sign that is a minimum  of  8.5  x  11 inches, posted on or within 3 inches of each cash register where payment  may be made so that the sign(s) are unob- structed in their  entirety  and  can  be read easily by each consumer making a purchase;

- The sign must be clear, legible, and conspicuous and be printed  in black Helvetica bold or Arial bold type (or    other    similar    sans    serif    fonts)

#### 21 CFR Ch. I (4–1–17 Edition)

against a solid white background in at least 17 point type with appropriate space between the warning statements;

- Be printed in a manner that con- trasts by typography, layout, or color, with all other printed material; and

- Be capitalized and punctuated as indicated in paragraph (a)(1) of this section.

- A retailer of any cigar covered by paragraphs (a)(1) and (2) of this section will not be in violation of  this  section for  packaging  that:

  - Contains a health warning;

  - Is supplied to the retailer by the tobacco product manufacturer, im- porter, or distributor who has the re- quired state, local, or Alcohol and To- bacco Tax and Trade Bureau (TTB)- issued license or permit, if applicable, and

  - Is not altered by the retailer in a way that is material to the require- ments of this section.

  - Advertisements. (1) It is unlawful for any tobacco product manufacturer, packager, importer, distributor, or re- tailer of cigars to advertise or cause to be advertised within the United States any cigar unless each advertisement bears one of the required warning statements specified in paragraph (a)(1) of this section.

- For print advertisements and other advertisements with a visual component (including, for example, ad- vertisements on signs, shelf-talkers, Internet Web pages, and electronic mail correspondence), each required warning statement must appear in the upper portion of the area of the adver- tisement within the trim area as fol- lows:

- Occupy at least 20 percent of the area of the advertisement;

- Appear in at least 12-point font size that ensures that the required warning statement occupies the great- est possible proportion of the warning area set aside for the text required;

- Appear in conspicuous and leg- ible Helvetica bold or Arial bold  type (or other similar sans  serif  fonts)  and in black text on a white background or white text  on  a black  background  in a manner that contrasts by typography, layout, or color, with all other mate- rial on the advertisement;

#### Food and Drug Administration, HHS	§ 1143.13

- Be capitalized and punctuated as indicated in paragraph (a)(1) of this section;

- Be centered in the warning area in which the text is required to appear and positioned such that the text  of the required warning statement and the other textual information in the advertisement have the same orienta- tion; and

- Be surrounded by a rectangular border that is the same color as the text of the required warning statement and that is not less than 3 mm or more than 4 mm.

- This paragraph (b) applies to a re- tailer only if that retailer is respon- sible for or directs the warning state- ments required under the paragraph. However, this paragraph does not re- lieve a retailer of liability if the re- tailer displays, in a location open  to the public, an advertisement that does not contain a health warning or con- tains a  health warning  that  has been altered by the retailer in a way that is material to the requirements of this section.

  - Marketing requirements. (1) Except for cigars sold individually and not in a product package, the warning state- ments required for packages in para- graph (a)(1) of this section must be ran- domly displayed in each  12-month pe- riod, in as equal a number of times as is possible on each brand of cigar sold in product packaging and be randomly distributed in all areas of the United States in which the product is mar- keted in accordance with a plan sub- mitted by the cigar manufacturer, im- porter, distributor, or retailer to, and approved by, the Food and Drug Ad- ministration.

- The warning statements required for advertisements in paragraph (a)(1) of this section must be rotated quar- terly in alternating sequence in each advertisement for each brand  of  cigar in accordance with a plan submitted by the cigar manufacturer, importer, dis- tributor, or retailer to, and approved by,  the  Food  and  Drug  Administration.

- Each person required to randomly display and distribute or rotate warn- ings in accordance with an FDA-ap- proved plan under this part shall sub- mit a proposed warning plan to FDA no later  than  either  12  months  after  May

10, 2016, or 12 months before advertising or commercially marketing  a  product that is subject to such requirement, whichever  is  later.

##### § 1143.7   Language requirements for re- quired warning statements.

The text in each warning statement required in § 1143.3 or § 1143.5 must be in the  English  language,  except  as   fol- lows:

- In the case of an advertisement that appears in a non-English medium, the text in the required warning state- ment must appear in the predominant language of the medium whether or not the advertisement is in English, and;

- In the case of an advertisement that appears in an English language medium but that is not in English, the text in the required warning statement must appear in the same language as that principally used in the advertise- ment.

##### § 1143.9   Irremovable or permanent re- quired warning statements.

The warning statements required by this section must be indelibly printed on or permanently affixed to the pack- age or advertisement. These warnings, for example, must not be printed or placed on a product label affixed to a clear outer wrapper that is likely to be removed to access the product within the package.

##### § 1143.11   Does not apply to foreign dis- tribution.

The provisions of this part do not apply to a manufacturer or distributor of tobacco products that does not man- ufacture, package, or import tobacco products for sale or distribution within the United States.

##### § 1143.13   Effective date.

- Except as stated in  paragraph  (b) of this section, this part will take  ef- fect 24  months  after May  10,  2016. The effective  date  will  be  with  respect   to the date of manufacture,  provided  that, in any case, beginning 30 days after the effective date, a manufacturer may not introduce into the domestic  commerce of the United States any product, irre- spective of the  date  of  manufacture, that is not  in  conformance  with  this part.

#### Pt. 1150

- The requirement to submit a warning plan to FDA under § 1143.5(c)(3) will take effect 12 months after May 10, 2016.

## PART 1150—USER FEES

Sec.

1150.1   Scope.

1150.3   Definitions.

1150.5    Required information.

1150.7   Yearly class allocation.

1150.9    Domestic   manufacturer   or   importer assessment.

1150.11    Notification of assessments.

1150.13    Payment of assessments.

1150.15    Disputes.

1150.17   Penalties.

AUTHORITY: 21 U.S.C. 371, 387a, 387b, 387i, 387s, 21 CFR 1100.1.

SOURCE: 79 FR 39310, July 10, 2014, unless

otherwise noted.

##### § 1150.1   Scope.

This part establishes requirements related to tobacco product user fees under section 919 of the Federal Food, Drug, and Cosmetic Act (21 U.S.C. 387s). The total amount of user fees may not exceed the amount specified for that fiscal year in section 919(b) of the Federal Food, Drug, and Cosmetic Act. All domestic manufacturers and importers of tobacco products are re- quired to pay to FDA their percentage share of the total assessment for a fis- cal year.

##### § 1150.3   Definitions.

The following definitions are applica- ble to this part:

Class of tobacco products means each of the following types of tobacco prod- ucts as defined in 26 U.S.C. 5702 and for which taxes are required to be paid for the removal of such into domestic com- merce: Cigarettes, cigars, snuff, chew- ing tobacco, pipe tobacco, and roll- your-own tobacco.

Domestic manufacturer means a person who is required to obtain a permit from the Alcohol  and  Tobacco  Tax  and Trade Bureau of the Department of the Treasury with respect to the produc- tion of tobacco products under title 27 of the Code of Federal Regulations.

Fiscal year quarter means a quarter in a fiscal year (the fiscal year is October

1   through   September   30).   The   fiscal year  quarters  are  October  1–December

21 CFR Ch. I (4–1–17 Edition)

31,  January  1–March  31,  April  1–June

30, and July 1–September 30.

Importer means a person who is re- quired to obtain a permit from the Al- cohol and Tobacco Tax and Trade Bu- reau of the Department of the Treasury with respect to the importation of to- bacco products under title 27 of the Code of Federal Regulations.

Total assessment means the total amount of user fees (in dollars) author- ized to be assessed and collected for a specific fiscal year under section 919 of the Federal Food, Drug, and Cosmetic Act.

Units of product means:

- The number of sticks for ciga- rettes and cigars, or

- The weight (measured in pounds) for snuff, chewing tobacco, pipe to- bacco, and roll-your-own tobacco.

Units of product removed and not tax exempt means the units of product:

- Removed (as defined by 26 U.S.C. 5702), and

- Not exempt from Federal excise tax under chapter 52 of title 26 of the United States Code at the time of their removal under that chapter or the Har- monized Tariff Schedule of the United States.

Yearly class allocation means the amount of user fees (in dollars) as- sessed for a class of tobacco  products for a particular fiscal year.

[79 FR 39310, July 10, 2014, as amended at 81

FR 28715, May 10, 2016]

##### § 1150.5   Required information.

  - General. Each domestic manufac- turer and importer of tobacco products that are part of a class of tobacco prod- ucts must submit the information de- scribed in this section for such prod- ucts each month, and the information must be received by FDA no later than the 20th day of each month. The infor- mation must be submitted using the form that FDA provides. The informa- tion must be submitted even if the do- mestic manufacturer or importer had no removals subject to tax during the prior month. FDA will use the informa- tion submitted under this section and any other available information, as FDA determines appropriate, to make tobacco product user fee assessments.

#### Food and Drug Administration, HHS	§ 1150.7

  - Contents. Each domestic manufac- turer and importer must submit the following:

    - Identification information. (i) Its name and the mailing address of its principal place of business;

- The name and a telephone num- ber including area code of an office or individual that FDA may contact for further information;

- The email address and postal ad- dress at which it wishes to receive no- tifications FDA sends under this part;

- The Alcohol and  Tobacco  Tax and Trade Bureau (TTB) Permit Num- ber(s); and

- The Employer Identification Number(s) (EIN).

    - Removal information. The units of product, by class, removed and not tax exempt for the prior month and the Federal excise tax it paid, by class, for such removal.

      - This information must be reported for each TTB tobacco permit.

      - If the domestic manufacturer or importer did not remove any amount of tobacco product, it must report that no tobacco product was removed into do- mestic commerce.

    - Certified copies. Certified copies of the returns and forms that relate to:

      - The removal of tobacco products into domestic commerce (as defined by section 5702 of the Internal Revenue Code of 1986); and

      - The payment of the Federal  ex- cise taxes imposed under chapter 52  of the Internal Revenue Code of 1986.

  - First report for cigars. Domestic manufacturers and importers of cigars must submit the information described in this section beginning no later  than the 20th day of August, 2016. Domestic manufacturers and importers of cigars must submit the information described in this section for each of the prior months of fiscal year 2016 as their first monthly submission. The previous sen- tence only applies for  the  first  report in fiscal year 2016.

  - First report for pipe tobacco. Do- mestic manufacturers and importers of pipe tobacco must submit the informa- tion described in this section beginning no later than the 20th day of August, 2016.

[79 FR 39310, July 10, 2014, as amended at 81

FR 28715, May 10, 2016]

##### § 1150.7   Yearly class allocation.

For each fiscal year, FDA will allo- cate the total assessment among the classes of tobacco products.

- Calculation. FDA will calculate the percentage shares for each class as follows:

- Except for cigars, FDA will mul- tiply the units of product removed and not tax exempt for the most recent full calendar year by the 2003 maximum Federal excise tax rate for that class (class dollar figure).

- For cigars, FDA will:

  - Multiply the units of small cigars removed and not tax exempt for the most recent full calendar year by the 2003 maximum Federal excise tax rate for small cigars (small cigar subclass dollar figure).

  - Multiply the units of large cigars removed and not tax exempt for the most recent full calendar year by the 2003 maximum Federal excise tax rate for large cigars (large cigar subclass dollar figure).

  - Add the small cigar subclass dol- lar figure and the large cigar subclass dollar figure (cigar class dollar figure).

- FDA will total the class dollar fig- ures for all tobacco classes for the most recent full calendar year (total dollar figure).

- FDA will divide the class dollar figure by the total dollar figure to de- termine the percentage share for each class.

- FDA will calculate the allocation for each class of tobacco products by multiplying the percentage share for each class by the total assessment.

- Reallocation. For any class of to- bacco products that is not deemed by FDA to be subject to regulation under chapter IX of the Federal Food, Drug, and Cosmetic Act, the amount of user fees that would otherwise be assessed to such class of tobacco products will be reallocated to the classes of tobacco products that are subject to chapter IX of the Federal Food, Drug, and Cos- metic Act in the same manner and based on the same relative percentages otherwise determined under paragraph

- of this section.

[79 FR 39310, July 10, 2014, as amended at 81

FR 28716, May 10, 2016]

#### § 1150.9

##### § 1150.9   Domestic manufacturer or im- porter assessment.

Each quarter, FDA will calculate the assessment owed by each domestic manufacturer or importer for that quarter.

  - Calculation. (1) For each class of tobacco products except cigars, FDA will calculate the percentage share for each domestic manufacturer and im- porter by dividing the Federal excise taxes that it paid for the class for the prior quarter by the total excise taxes that all domestic manufacturers and importers paid for the class for that same quarter.

- For the cigar class, FDA will cal- culate the percentage share for each domestic manufacturer and importer by dividing the Federal excise taxes that it paid for the class for the prior fiscal year by the total excise taxes that all domestic manufacturers and importers paid for the class for the prior fiscal year.

- If the percentage share calculated for a domestic manufacturer or im- porter in this section, as applicable, is less than 0.0001 percent, the share is ex- cluded from the assessment  for  that class  of  tobacco  products.

- Within each class of tobacco prod- ucts, the assessment owed by a domes- tic manufacturer or importer for the quarter is the yearly class allocation, determined as described in § 1150.7, di- vided by four, multiplied by the domes- tic manufacturer’s or importer’s per- centage share, truncated to the fourth decimal place, for that class of tobacco products.

  - Adjustments. Annually, FDA will make any necessary adjustments to in- dividual domestic manufacturer or im- porter assessments if needed to ac- count for any corrections (for example, to include domestic manufacturers or importers that were not included in a relevant assessment calculation).

[79 FR 39310, July 10, 2014, as amended at 81

FR 28716, May 10, 2016]

##### § 1150.11   Notification of assessments.

- Notification. No later than 30 cal- endar days before the end of each fiscal year quarter, FDA will notify each do- mestic manufacturer and importer of the   amount   of   the   quarterly   assess-

#### 21 CFR Ch. I (4–1–17 Edition)

ment imposed on the domestic manu- facturer or importer.

- Content of notification. The notifi- cation under paragraph (a) of this sec- tion will include the following:

  - The amount of the quarterly as- sessment imposed on the domestic manufacturer or importer and the date that payment of the assessment must be received by FDA;

  - Class assessment information, in- cluding each class’ initial percentage share, the reallocation amount (if any) and each class’ percentage share after any such reallocation, and the quar- terly assessment for each class;

  - Domestic manufacturer or im- porter assessment information, includ- ing the domestic manufacturer’s or im- porter’s percentage share of each rel- evant class of tobacco products and in- voice amount;

  - Any adjustments FDA has made under § 1150.9(b);

  - The manner in which assessments are to be remitted to FDA;

  - Information about the accrual of interest if a payment is late; and

  - Information regarding where to send a dispute and when it needs to be sent.

##### § 1150.13   Payment of assessments.

    - Payment of an assessment must be received by FDA no later than  the last day of each fiscal year quarter.

    - Payments must be submitted to FDA in U.S. dollars and in the manner specified in the notification.

    - Except  as  provided  in  paragraph

    - of this section, if an assessment is not received by the last day of the fis- cal year quarter, FDA will begin as- sessing interest on the  unpaid  amount in accordance with 31 U.S.C. 3717.

      - If FDA does not send the notifica- tion described in § 1150.11(a) 30 calendar days before the end of a quarter, no in- terest will be assessed by FDA under paragraph (c) of this section until 30 calendar days have elapsed from the date FDA sent notification of  the amount owed.

      - If a domestic manufacturer or im- porter disputes the amount of an as- sessment, it must still pay the assess- ment in accordance with paragraphs (a) and (b) of this section.

#### Food and Drug Administration, HHS	§ 1150.17

##### § 1150.15   Disputes.

        - A domestic tobacco manufacturer or importer may dispute an FDA as- sessment. The dispute must include the basis for the dispute, and the dispute must be:

          - Submitted in writing;

          - Received by FDA no later than 45 days after the date on the assessment notification;

          - Legible and in English; and

          - Sent to the address found on our Web	site

        - If FDA determines that there was an error related to the assessment and the assessment was too high, FDA will refund the amount assessed in error to the domestic manufacturer or im- porter.

        - FDA will provide a dated, written response, and its response will provide information about how to submit a re- quest for further Agency review.

        - A request for further Agency re- view under § 10.75 of this  chapter  may be submitted. Such a request must be submitted in writing by the domestic manufacturer or importer and received by FDA within 30 days  from  the  date on FDA’s response. The request for fur- ther Agency review must be legible, in English,  and  submitted  to  the  address

found on our Web site.

##### § 1150.17   Penalties.

- Under section 902(4) of the Federal Food,    Drug,    and    Cosmetic    Act    (21

U.S.C. 387b), a tobacco product is deemed adulterated if the domestic manufacturer or importer of the to- bacco product fails to pay a user fee as- sessed to such manufacturer or im- porter by the later of the date the as- sessment is due, 30 days from the date FDA sent notification of the amount owed, or 30 days after final Agency ac- tion on a resolution of any dispute as to the amount of the fee.

- Under section 902(4) of the Federal Food, Drug, and Cosmetic Act, a to- bacco product is deemed adulterated if the domestic manufacturer or importer of the tobacco product fails to  report the information required by § 1150.5 to calculate assessments under this part.

- The failure to report the informa- tion required by § 1150.5 to calculate as- sessments under  this part is  a prohib- ited act under section 301(e) of the Fed- eral Food, Drug, and Cosmetic Act.

- Information	submitted	under

§ 1150.5 is subject to 18 U.S.C. 1001 and other appropriate civil and criminal statutes.

# SUBCHAPTER  L—REGULATIONS  UNDER  CERTAIN  OTHER ACTS ADMINISTERED BY THE FOOD AND DRUG ADMIN- ISTRATION

## PART 1210—REGULATIONS UNDER THE FEDERAL IMPORT MILK ACT

##### Subpart A—General Provisions

Sec.

- Authority.

- Scope of act.

- Definitions.

##### Subpart B—Inspection and Testing

- Availability for examination and in- spection.

- Sanitary inspection of dairy farms.

- Physical examination of cows.

- Tuberculin test.

- Sanitary inspection of plants.

- Pasteurization;	equipment	and methods.

- Method of bacterial count.

- Authority to sample and inspect.

- Scoring.

##### Subpart C—Permit Control

- Application for permit.

- Permit number.

- Form of tag.

- Permits granted on certificates.

- Temporary permits.

- Permits   for   pasteurized   milk   or cream.

- Permits for raw milk or cream.

- Permits waiving clauses 2 and 5, sec- tion 2 of the Federal Import Milk Act.

- Permits  waiving  clause  4,  section  2 of the Federal Import Milk Act.

##### Subpart D—Hearings

- Hearing procedure for permit denial, suspension, and revocation.

- Hearing before prosecution. AUTHORITY: 21 U.S.C. 141–149.

SOURCE:  38  FR  32104,  Nov.  20,  1973,  unless

otherwise noted.

CROSS REFERENCES: For Animal and Plant Health Inspection Service regulations con- cerning tubercular cattle, see 9 CFR parts 50 and 77. For Animal and Plant Health Inspec- tion Service regulations, see 9 CFR chapter

I. For customs regulations concerning im- portation of milk and cream, see 19 CFR 12.7. For regulations of the Agricultural Mar- keting Service (Marketing Agreements and Orders) covering marketing areas for milk, see 7 CFR chapter X.

## Subpart A—General Provisions

##### § 1210.1   Authority.

For the purposes of  the  regulations in  this  part  the  act  (44  Stat.  1101;  21

    - 141–149) ‘‘To regulate the impor- tation of milk and cream into the United States for the purpose of pro- moting the dairy industry of the United States and protecting the pub- lic health’’ shall be known and referred to as ‘‘the Federal Import Milk Act.’’

##### § 1210.2   Scope of act.

The provisions of the act apply to all milk and cream offered for import into the continental United States.

##### § 1210.3   Definitions.

      - Secretary. Secretary means the Secretary of Health and Human Serv- ices.

      - Commissioner. Commissioner means the Commissioner of Food and Drugs.

      - Milk. For the purposes of the act and of the regulations in this part:

Milk is the whole, fresh, clean, lacteal secretion obtained by the complete milking of one or more healthy cows, properly fed and kept, excluding that obtained within 15 days before and 5 days after calving, or such longer pe- riod as may be necessary to render the milk  practically  colostrum  free.

      - Condensed milk. Condensed milk, as the term is used in section 3, para- graph 2, of the Federal Import Milk Act, includes evaporated milk in the manufacture of which sterilization of the milk and cream is a necessary and usual process; it includes sweetened condensed milk only if it is prepared by a process which insures sterilization of the milk and cream. Condensed milk, as the term is used in section 3, para- graph 3, of the Federal Import Milk Act, means sweetened condensed milk.

      - Sweetened condensed milk. Sweet- ened condensed milk conforms to the definition  and  standard  of  identity  for

Food and Drug Administration, HHS	§ 1210.14

such food as set out in § 131.120 of this chapter.

      - Evaporated milk. Evaporated milk conforms to  the  definition  and  stand- ard of identity for  such  food  as  set  out in § 131.130 of this chapter.

      - Cream. Cream is that portion of the milk, rich in  milk  fat,  which  rises to the surface of milk on standing or is separated from it by centrifugal force. (See §§ 131.150 through 131.157 of this chapter).

      - Pasteurization. Pasteurization is the process of heating every particle of milk or cream to at least 143 F., and holding it at such temperature con- tinuously for at least 30 minutes, or to at least 161 F., and holding it at such temperature continuously for at least 15 seconds.

- Shipper. A shipper is anyone, other than a common carrier, who ships, transports, or causes to be shipped or transported  into  the  United  States milk or cream owned by him.

[38 FR 32104, Nov. 20, 1973, as amended at 42

FR 14091, Mar. 15, 1977]

## Subpart B—Inspection and Testing

##### § 1210.10   Availability  for  examination and inspection.

Dairy farms and plants from which milk or cream is shipped or trans- ported into the United States shall be open at all reasonable times to author- ized agents for necessary examinations and inspections. Failure to permit such examinations and inspections may be considered cause for the suspension or revocation of the permit.

##### § 1210.11   Sanitary  inspection  of  dairy farms.

The sanitary conditions of any dairy farm producing milk or cream to be shipped or transported into the United States or to a plant from which milk or cream is to be shipped or transported into the United States must score at least 50 points out of 100 points, accord- ing to the methods for scoring as pro- vided by the score card for sanitary in- spection of dairy farms in the form pre- scribed by the Secretary.

##### § 1210.12   Physical	examination	of cows.

- Physical examination of any and all cows in herds producing milk or cream which is to be shipped or trans- ported into the United States shall be made by an authorized veterinarian of the United States or of any State or municipality thereof or of the country in which such milk or cream is pro- duced to determine whether such cow or cows are in a healthy condition. Such examination shall be made as often as the Secretary may deem nec- essary and, in any event, shall have been made within one year previous to the time of the importation.

- The result of the physical exam- ination shall be set forth in the form prescribed by the Secretary.

##### § 1210.13   Tuberculin test.

- Except as provided in § 1210.27 any and all animals in  herds  producing milk or cream which is to be shipped or transported raw into the United States shall be free from tuberculosis, as de- termined by a tuberculin test  applied by an official  veterinarian  of  the United States or of any State or mu- nicipality thereof or of the country in which such milk or cream is produced. Such test shall be made as often as the Secretary may deem necessary and, in any event, shall have been made within 1 year previous to the time of the im- portation. All animals showing positive or suspicious reactions to the tuber- culin test must be permanently re- moved from the herd.

- The results of the tuberculin test and all facts concerning the disposal of reacting or suspected animals shall be set forth in the form prescribed by the Secretary.

##### § 1210.14   Sanitary inspection of plants.

The sanitary conditions of any plant handling milk or cream any part  of which is to be shipped or transported into the United States shall score at least 50 points out of 100 points accord- ing to the methods for scoring as pro- vided by the score card for sanitary in- spection of such plants in the form pre- scribed by the Secretary.

§ 1210.15

##### § 1210.15   Pasteurization;	equipment and methods.

All dairy farms and plants at which any milk or cream is pasteurized for shipment or transportation into the United States shall employ adequate pasteurization machinery of a type eas- ily cleaned and of sanitary construc- tion capable of holding every portion of the milk or cream at the required tem- perature for the required time. Such pasteurizing machinery shall be prop- erly equipped with accurate time and temperature recording devices, which shall be kept at all times in good work- ing order. The temperature at the time of heating and holding must invariably be recorded on thermograph charts, initialed, numbered, and dated by the official having jurisdiction over such farms and plants. All  thermograph charts shall be held for a period of 2 years unless within that period they have been examined and released by such authorized agents as are des- ignated by  the Secretary.

##### § 1210.16   Method of bacterial count.

The bacterial count of milk and cream refers to the number of viable bacteria as determined by the standard plate method of the American Public Health Association in use at the time of the examination.

##### § 1210.17   Authority  to  sample  and  in- spect.

Inspectors engaged in the enforce- ment of the Federal Import Milk Act are empowered to test for temperature, to take samples of milk or cream, and to use such means as may be necessary for these purposes.

##### § 1210.18   Scoring.

Scoring of sanitary conditions re- quired by §§ 1210.11, 1210.14 shall be done by an official inspector of the  United States or of any State or municipality thereof or of the country in which the dairy  farm  or  plant  is  located.

## Subpart C—Permit Control

##### § 1210.20   Application for permit.

Application for a permit to ship or transport milk or cream into the United States shall be made by the ac-

21 CFR Ch. I (4–1–17 Edition)

tual shipper upon forms prescribed by the Secretary. The request for forms of applications for permits should be ad- dressed to Commissioner of Food and Drugs, Food and Drug Administration, Department of Health and Human Services, 5600 Fishers Lane, Rockville, MD 20857.

##### § 1210.21   Permit number.

Each permit issued under the Federal Import Milk Act, including each tem- porary permit, shall bear an individual number. The right to the use of such number is restricted solely to the per- mittee.

##### § 1210.22   Form of tag.

Each container of milk or cream shipped or transported into the United States by such permittee shall have firmly attached thereto a tag in the following form, bearing the required information in clear and legible type:

Product   ——————————————————

(State  whether  raw  milk,  pasteurized milk, raw cream, or pasteurized cream.)

Permit number   ——————————————

Federal  Import  Milk  Act,  Department of Health and Human Services.

Shipper   ——————————————————

Address of shipper —————————————

Provided, That in case of unit ship- ments consisting of milk only or cream only under one permit number, in lieu of each container being so marked, the vehicle of transportation, if  sealed, may be tagged with the above tag, which should, in addition, show the number of containers and quantity of contents of each.

##### § 1210.23   Permits  granted  on  certifi- cates.

In the discretion of the Secretary, a permit may be granted on a duly cer- tified statement signed by a duly ac- credited official of an authorized de- partment of any foreign government or of any State of the United States  or any municipality thereof. Such state- ment shall be in the form of a  certifi- cate prescribed by the Secretary, and shall have attached thereto, as a part thereof, signed copies of reports pre- scribed by §§ 1210.12, 1230.13, and also by

§§ 1210.11, 1210.14, as applicable. The necessary  inspections  and  examina- tions  upon  which  the  reports  are  based

#### Food and Drug Administration, HHS	Pt. 1230

shall be made by persons who are act- ing under the direct supervision of the certifying   official.

##### § 1210.24   Temporary permits.

A temporary permit will be granted only upon a satisfactory showing that the applicant therefor has been unable to obtain the necessary inspections re- quired by the applicable provisions of section 2 of the Federal Import Milk Act. Temporary permits shall be valid until the Secretary shall provide for in- spection to ascertain that clauses 1, 2, and 3 of section 2 of the Federal Import Milk Act have been complied with.

##### § 1210.25 Permits for pasteurized milk or cream.

Permits to ship or transport pasteur- ized milk or cream into the  United States will be granted only upon com- pliance with the requirements of clauses 1 and 3 of section 2 of the Fed- eral Import Milk Act, §§ 1210.11, 1210.12, 1210.14, as applicable.

##### § 1210.26 Permits for raw milk or cream.

Except as provided in § 1210.27, per- mits to ship or transport raw milk or cream into the United States will be granted only when the milk or cream comes from  dairy  farms  or  plants where pasteurization is not carried on and  then  only  upon  compliance  with the requirements of clauses  1,  2,  and  3 of section 2 of the Federal Import  Milk Act, §§ 1210.11 to 1210.14 as applicable.

##### § 1210.27 Permits waiving  clauses  2 and 5, section 2 of the Federal Im- port Milk Act.

A permit to ship or transport raw milk into the United States will con- tain a waiver of clauses 2 and 5 of sec- tion 2 of the Federal Import Milk Act when the shipper is an operator of a creamery or condensery, or is a pro- ducer shipping or transporting to a creamery or condensery and the cream- ery or condensery is located in the United States within a radius  of  20 miles of the point of production of such milk, and the milk, prior to its sale, use, or disposal, is pasteurized, con- densed, or evaporated.

##### § 1210.28 Permits waiving clause 4, section 2 of the Federal Import Milk Act.

The Secretary, in his discretion, will issue to a shipper who is an operator of a condensery a permit  waiving the  re- quirements of clause 4, of section 2 of the Federal Import Milk Act and allow- ing milk and cream containing not to exceed 1,200,000 bacteria per cubic cen- timeter to be shipped or transported into the United  States  if  the condensery is located within  a  radius of 15 miles of the point of production of the milk and cream and such milk and cream are to be sterilized in the manu- facture of condensed milk.

## Subpart D—Hearings

##### § 1210.30 Hearing procedure for per- mit denial, suspension, and revoca- tion.

Any person who contests denial, sus- pension, or revocation of a permit shall have an opportunity for a regulatory hearing before  the Food  and Drug  Ad- ministration pursuant to subpart F of part 16 of this chapter.

[41  FR  48269,  Nov.  2,  1976,  as  amended  at  42

FR 15676, Mar. 22, 1977]

##### § 1210.31   Hearing before prosecution.

Before violation of the act is referred to the Department of Justice for pros- ecution under section 5 of the Federal Import Milk Act, an opportunity to be heard will be given to the party against whom prosecution is under consider- ation. The hearing will be private and confined to questions of fact. The party notified may  present evidence,  either oral or written, in person or by attor- ney, to show cause why he should not be prosecuted. After a hearing is held, if it appears that the law has been vio- lated, the facts will be reported to the Department of Justice.

[41 FR 48269, Nov. 2, 1976]

## PART 1230—REGULATIONS UNDER THE FEDERAL CAUSTIC POISON ACT

##### Subpart A—General Provisions

Sec.

- Scope of the act.

- Definitions.

§ 1230.2

##### Subpart B—Labeling

- Placement.

- Required wording.

- Manufacturer; distributor.

- Labeling of ‘‘poison’’.

- Directions for treatment.

- Responsibility   for   labeling   direc- tions for treatment.

- Exemption from labeling directions for treatment.

##### Subpart C—Guaranty

- General guaranty.

- Specific guaranty.

##### Subpart D—Administrative Procedures

- Collection of samples.

- Where samples may be collected.

- Analyzing of samples.

- Investigations.

- Analysis.

- Hearings.

- Hearings; when not provided for.

- Publication.

##### Subpart E—Imports

- Required label information.

- Delivery of containers.

- Invoices.

- Enforcement.

- Samples.

- No violation; release.

- Violation.

- Rejected containers.

- Relabeling of containers.

- Penalties.

AUTHORITY: 15 U.S.C. 1261–1276.

CROSS  REFERENCES:  For  regulations  relat- ing to invoices, entry, and assessment of du- ties, see 19 CFR parts 141, 142, 143, 151, 152. For regulations regarding the examination, classification,  and  disposition  of   foods, drugs, devices, cosmetics, insecticides, fun- gicides, and caustic or corrosive substances, see 19 CFR part 12. For regulations relating to consular invoices, and documentation of merchandise, see 22 CFR parts 91 and 92.

SOURCE:  38  FR  32110,  Nov.  20,  1973,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1230.2   Scope of the act.

The  provisions  of  the  act  apply  to any container which has  been  shipped or delivered for shipment  in  interstate or foreign commerce, as defined in sec- tion  2(c)  of  the  act  (44  Stat.  1407;  15

U.S.C. 402) or which has been received

21 CFR Ch. I (4–1–17 Edition)

from shipment in such commerce for sale or exchange, or which is sold or of- fered for sale or held for sale or ex- change in any Territory  or  possession or in the District of Columbia.

##### § 1230.3   Definitions.

  - The word container as used in the regulations in this part means a retail parcel, package, or container suitable for household use and employed exclu- sively to hold any dangerous caustic or corrosive substance defined in the act.

  - The words suitable for household use mean and imply adaptability for ready or convenient handling in places where people dwell.

## Subpart B—Labeling

##### § 1230.10   Placement.

The label or sticker shall be so firmly attached to the container that it will remain thereon while the container is being used, and be so placed as readily to  attract  attention.

##### § 1230.11   Required wording.

- The common name of the dan- gerous caustic or corrosive substance which shall appear on the label or sticker is the name given in section 2(a) of the act (44 Stat. 1406; 15 U.S.C.

402) or any other name commonly em- ployed to designate and identify such substance.

- Preparations within the scope of the act bearing trade or fanciful names shall, in addition, be labeled with the common name of the dangerous caustic or corrosive substance contained there- in and comply with all the other re- quirements of the act and of the regu- lations in this part.

##### § 1230.12   Manufacturer; distributor.

If the name on the label or sticker is other than that of the manufacturer, it shall be qualified by such words as ‘‘packed for,’’  ‘‘packed by,’’  ‘‘sold  by,’’ or ‘‘distributed by,’’ as the case  may be, or by other appropriate expression.

##### § 1230.13   Labeling of ‘‘poison’’.

The following are styles of uncondensed Gothic capital letters 24- point (type face) size:

Food and Drug Administration, HHS	§ 1230.21

When letters of not less than 24-point size are required on a label  in  stating the word ‘‘poison’’ they must not be smaller than those above set forth.

##### § 1230.14   Directions for treatment.

Except as provided in § 1230.16, the container shall bear in all cases  upon the label or sticker thereof, imme- diately following the word ‘‘Poison,’’ directions for treatment in the case of internal personal injury; in addition, if the substance may cause external in- jury, directions for appropriate treat- ment shall be given.  The  directions shall prescribe such treatments for per- sonal injury as are sanctioned by com- petent medical authority, and the ma- terials called for by  such  directions shall be, whenever practicable, such as are  usually  available  in  the  household.

##### § 1230.15   Responsibility for labeling di- rections for treatment.

A person who receives from a manu- facturer or wholesaler any container which under the conditions set forth in section 2(b)(4) of the act and § 1230.16 does not bear at the time of shipment directions for treatment in the case of personal injury must place such direc- tions on  the label or sticker  if he offers such container for general sale or ex- change.

##### § 1230.16   Exemption  from  labeling  di- rections for treatment.

Manufacturers  and  wholesalers  only, at the time of shipment or delivery for shipment, are exempted from placing directions for treatment on the label or sticker of any container for other than household use, but in any event the in- formation required by section 2(b) (1), (2), and (3) of the act (44 Stat. 1407; 15

U.S.C. 402) and the regulations in this part shall be given.

## Subpart C—Guaranty

##### § 1230.20   General guaranty.

In lieu of a particular guaranty for each lot of dangerous caustic or corro- sive substances, a general continuing guaranty may be furnished by the guarantor to actual or prospective pur- chasers. The following are forms of continuing  guaranties:

- Substances for both household use and other than household use:

The undersigned guarantees that the retail parcels, packages, or containers of the dan- gerous caustic or corrosive substance or sub- stances to be sold to lllll are not mis- branded within the meaning of the Federal Caustic Poison Act.

(Date)

(Signature and address of

guarantor)

- Substances for other than house- hold use (this form may  be  issued  only by    a    manufacturer    or    wholesaler) (§§ 1230.15, 1230.16):

The dangerous caustic or corrosive sub- stance or substances in retail parcels, pack- ages, or containers suitable for household use to be sold to lllll are for other than household use, and guaranteed not to be mis- branded within the meaning of the Federal Caustic Poison Act.

(Date)

(Signature and address of manufacturer or wholesaler)

##### § 1230.21   Specific guaranty.

If a guaranty in respect to any spe- cific lot of dangerous caustic or corro- sive substances be given, it shall be in- corporated in or attached to the bill of sale, invoice, or other schedule bearing the date and the name and quantity of the substance sold, and shall not ap- pear on the label or package. The fol- lowing are forms of specific guaranties:

- Substances for both household use and other than household use:

#### § 1230.30

The undersigned guarantees that the retail parcels, packages, or containers of the dan- gerous caustic or corrosive substance or sub- stances listed herein (or specifying the sub- stances) are not misbranded within the meaning  of  the  Federal  Caustic  Poison  Act.

(Signature and address of guarantor)

- Substances for other than house- hold use (this form may  be  issued  only by    a    manufacturer    or    wholesaler (§§ 1230.15, 1230.16):

The dangerous caustic or corrosive sub- stance or substances listed herein (or speci- fying the substances) in retail parcels, pack- ages, or containers suitable for household use are for other than household use and are guaranteed not to be misbranded within the meaning  of  the  Federal  Caustic  Poison  Act.

(Name and address of manufacturer

or wholesaler)

## Subpart D—Administrative Procedures

##### § 1230.30   Collection of samples.

Samples for examination by or under the direction and supervision of the Food and Drug Administration shall be collected by:

- An authorized agent in the em- ploy of the Department of Health and Human Services;

- Any officer of any State, Terri- tory, or possession, or of the District of Columbia, authorized by the Secretary of Health and Human Services.

##### § 1230.31   Where  samples  may  be  col- lected.

Caustic or corrosive substances with- in the scope of this act (44 Stat. 1406; 15

U.S.C. 401–411)  may  be  sampled  wher- ever found.

##### § 1230.32   Analyzing of samples.

Samples collected by an authorized agent shall be analyzed at the labora- tory designated by the Food and Drug Administration. Only such samples as are    collected    in    accordance    with

§§ 1230.30, 1230.31 may be analyzed by or under the direction and supervision of the Food  and  Drug  Administration. Upon request one subdivision of the sample, if available, shall  be  delivered to  the  party  or  parties  interested.

21 CFR Ch. I (4–1–17 Edition)

##### § 1230.33   Investigations.

Authorized agents in the employ of the Department of Health and Human Services may make investigations, in- cluding the inspection of premises where dangerous caustic and corrosive substances subject to the act are man- ufactured, packed, stored, or held for sale or distribution, and make exami- nations of freight and other transpor- tation records.

##### § 1230.34   Analysis.

- The methods of examination or analysis employed shall be those pre- scribed by the Association of Official Agricultural Chemists, when applica- ble, provided, however, that any meth- od of analysis or examination satisfac- tory to the Food and Drug Administra- tion may be employed.

- All percentages stated in the defi- nitions in section 2(a) of the Federal Caustic Poison Act shall be determined by weight.

##### § 1230.35   Hearings.

Whenever it appears from the inspec- tion, analysis, or test of any container that the provisions of section 3 or 6 of the Federal Caustic Poison  Act  (44 Stat. 1407, 1409; 15 U.S.C. 403, 406) have

been violated and criminal proceedings are contemplated, notice shall be given to the party or parties against whom prosecution is under consideration and to other interested parties, and a date shall be fixed at which such party or parties may be heard.  The  hearing shall be held at the office of  the  Food and Drug Administration designated in the notice and shall be private and con- fined to questions of fact. The parties notified may present evidence, either oral or written, in person or by attor- ney, to show cause why the matter should not be referred for prosecution as a violation of the Federal Caustic Poison  Act.

##### § 1230.36   Hearings; when not provided for.

No hearing is provided for when the health, medical, or drug officer or agent of any State, Territory, or pos- session, or of the District of Columbia, acts under the authority contained in section 8 of the Federal Caustic Poison

#### Food and Drug Administration, HHS	§ 1230.43

Act (44 Stat. 1409; 15 U.S.C. 408) in re-

porting a violation direct to the United States  attorney.

##### § 1230.37   Publication.

  - After judgment of the court in any proceeding under the Federal Caus- tic Poison Act, notice shall be given by publication. Such notice shall include the findings of the court and may in- clude the findings of the analyst and such explanatory statements of fact as the Secretary of Health and Human Services may deem appropriate.

  - This publication may be made in the form of a circular, notice, or bul- letin, as the Secretary of Health and Human Services may direct.

  - If an appeal be taken from the judgment of the court before such pub- lication, that fact shall appear.

## Subpart E—Imports

##### § 1230.40   Required label information.

Containers which  are offered  for im- port shall in all cases bear labels or stickers having thereon the  informa- tion required by section 2(b) (1), (2), and (3) of the Federal Caustic Poison Act and the directions for treatment in the case of personal injury, except such directions need not appear on the label or sticker at the time of shipment by a wholesaler or manufacturer for other than household use.

##### § 1230.41   Delivery of containers.

Containers shall not be delivered to the consignee prior to report of exam- ination, unless a bond has been given on the appropriate form for the amount of the  full invoice  value of  such  con- tainers, together with the duty there- on, and the refusal of the consignee to return such containers for any cause to the custody of the District Director of Customs when demanded, for the pur- pose of excluding them from the coun- try or for any other purpose, the con- signee shall pay an amount equal  to the sum named in the bond, and such part of the duty, if any, as may be pay- able, as liquidated damages for failure to return to the District Director of Customs on demand all containers cov- ered by the bond.

##### § 1230.42   Invoices.

As soon as the importer makes entry, the invoices covering containers and the public stores packages shall be made available, with the least possible delay, for inspection by the representa- tive of the district. When no sample is desired the invoice shall be stamped by the district ‘‘No sample desired, Food and Drug Administration, Department of Health and Human Services, per (ini- tials of inspecting officer).’’

##### § 1230.43   Enforcement.

- Enforcement agency. The Federal Caustic Poison Act shall be enforced by the Food and Drug Administration, De- partment of Health and Human Serv- ices.

- Enforcement of provisions. The en- forcement of the provisions of the Fed- eral Caustic Poison Act as they relate to imported dangerous caustic or cor- rosive substances, will, as a general rule, be under the direction of the chief of the local inspection district of the Food and Drug Administration, De- partment of Health and Human Serv- ices, and District Directors of Customs acting as administrative officers in carrying out directions relative to the detention, exportation, and sale, or other disposition of such substances and action under the bond in case of noncompliance with the provisions of the Federal Caustic Poison Act.

- Chief of district as customs officer. The chief of district shall be deemed a customs officer in enforcing import regulations.

- Nonlaboratory ports. (1) At the ports of entry where there is no dis- trict of the Food and Drug Administra- tion, the District Director  of  Customs or deputy, on the day when the first notice of expected shipment of con- tainers is received, either by invoice or entry, shall notify the  chief  of  district in  whose  territory  the  port  is  located.

- On the day of receipt of such no- tice the chief of district shall  mail  to the District Director of Customs appro- priate notice, if no sample is desired. This notice serves as an equivalent to stamping the invoices at district ports with  the  legend  ‘‘No  sample  desired,

§ 1230.44

Food and Drug Administration, De- partment of Health and Human Serv- ices, per (initials of inspecting offi- cer).’’

- If samples are desired, the Chief of district shall immediately notify the District Director of Customs.

- The District Director of  Customs at once shall forward samples, accom- panied by description of shipment.

- When samples are desired from each shipment of containers, the chief of district shall furnish to District Di- rector of Customs and deputies at ports within the district’s territory a list of such containers, indicating the size of sample necessary. Samples should then be sent promptly on arrival of con- tainers without awaiting special re- quest.

- In all other particulars the proce- dure shall be the same at nonlabora- tory ports as  at laboratory  ports, ex- cept that the time consumed in deliv- ery of notices by mail shall be allowed for.

##### § 1230.44   Samples.

On the same day that samples are re- quested by the district, the District Di- rector of Customs or appraiser shall notify the importer that samples  will be taken, that the containers must be held intact  pending  a notice of  the  re- sult of inspection  and  analysis,  and that in case the containers do not com- ply with the requirements of the Fed- eral Caustic Poison Act, they must be returned to the District Director of Customs for disposition. This notifica- tion may be given by the District Di- rector of Customs or appraiser through individual notices to the importer  or by suitable  bulletin  notices  posted daily in the customhouse.

##### § 1230.45   No violation; release.

As soon as examination of the sam- ples is completed, if no violation of the act is detected, the chief of the district shall send a notice of release to the im- porter and a copy of this notice to the District Director of Customs for his in- formation.

##### § 1230.46   Violation.

  - If a violation of the Federal Caus- tic Poison Act is disclosed, the chief of the district shall send to the importer

#### 21 CFR Ch. I (4–1–17 Edition)

due notice of the nature of the viola- tion and of the time and place where evidence may be presented,  showing that the  containers should  not be re- fused admission. At the  same  time similar notice  regarding  detention  of the containers shall be sent to the Dis- trict Director of Customs,  requesting him to refuse delivery thereof or to re- quire their return  to  customs  custody if by any chance the containers were released without the bond  referred  to in § 1230.41. The time allowed the im- porter  for  representations   regarding the shipment may be extended at his request for a reasonable period to per- mit  him  to  secure  such  evidence.

  - If  the  importer  does  not  reply  to the  notice  of  hearing  in  person  or  by letter  within  the  time  allowed  on  the notice,  a  second  notice,  marked  ‘‘sec- ond  and  last  notice,’’  shall  be  sent  at once by the chief of the district, advis- ing him that failure to reply will cause definite   recommendation   to   the   Dis- trict Director of Customs that the con- tainers  be  refused  admission  and  that the   containers   be   exported   within   3 months under customs supervision.

##### § 1230.47   Rejected containers.

- In all cases where the containers are to be refused  admission,  the  chief of the district within 1 day after hear- ing, or, if the importer does not appear or reply within 3 days after second no- tice, shall notify the  District  Director of Customs in duplicate accordingly.

- Not later than 1 day after receipt of this notice the District Director of Customs shall sign and transmit to the importer one of the copies, which shall serve as notification to the importer that the containers must be exported under customs supervision within 3 months from such date, as provided by law; the other notice shall be retained as office record and later returned as a report to the chief of the district. In all cases the importer shall return his no- tice to the District Director of Cus- toms, properly certified as to the infor- mation required, as the form provides.

##### § 1230.48   Relabeling of containers.

- If containers are to be released after relabeling, a notice shall be sent by the chief of district direct to the im- porter, a carbon copy being sent to the

#### Food and Drug Administration, HHS	§ 1230.49

District Director of Customs. This no- tice must state specifically the condi- tions to be performed, so as to bring the performance thereof under the pro- visions of the customs bonds on con- sumption and warehouse entries, these bonds including provisions requiring compliance with all of the require- ments of the Federal Caustic Poison Act and all regulations and instruc- tions issued thereunder. The notice will also state the officer to be notified by the importer when the containers are ready for inspection.

- The importer must return the no- tice to the District Director of Cus- toms or chief of district, as designated, with the certificate thereon filled out, stating that he has complied with the prescribed conditions and that the con- tainers are ready for inspection at the place named.

- This notice will be delivered to the inspection officer, who, after in- spection, will endorse the result there- of on the back of the notice and return the same to the District Director of Customs or to the chief of district, as the case may be.

- When the conditions to be com- plied with are under the supervision of the chief of district, and these condi- tions have been fully met, he shall re- lease the containers to the importer, sending a copy of the notice of release to the District Director of Customs for his information. If the containers have not been properly relabeled within the period allowed, the chief of district shall immediately give notice in dupli- cate to the District Director of Cus- toms of the results of inspection. The District Director of Customs shall sign and immediately transmit one copy of the notice to the importer and proceed in the usual manner.

- If the containers are detained sub- ject to relabeling to be performed under the supervision of the District Director of Customs, the District Di- rector of Customs, as soon as re- labeling is  accomplished,  will  notify the importer that the containers are released.

- If the containers have not been properly relabeled within the period al- lowed, their sale after labeling as re- quired  by  the  act  or  other  disposition

must be effected by the District Direc- tor of Customs.

- When the final action has been taken on containers which have been refused admission, sold, or otherwise disposed of as provided for by the act or which have been relabeled under the supervision of the District Director of Customs, he shall send to the chief of district a notice of such final action, giving the date and disposition.

- When relabeling is allowed the importer must furnish satisfactory evi- dence as to the identity of the con- tainers before release is given. The re- labeling must be done at a stated place and apart from other containers of a similar  nature.

- When containers are shipped to another port for relabeling or expor- tation, they must be shipped under cus- toms carrier’s manifest, in the same manner as shipments in bond.

- District Directors of Customs will perform the inspection service when- ever containers are to be exported, sold, or otherwise disposed of, and in other cases when there is no officer of the district available.

- District Directors of Customs and representatives of the district will con- fer and arrange the apportionment of the inspection service according to local conditions. Officers of the district will, whenever feasible, perform the in- spection service in connection with re- labeling.

##### § 1230.49   Penalties.

- In case of failure to comply  with the  instructions  or  recommendations of the chief of district as to conditions under which containers may be dis- posed of, the District Director of Cus- toms shall notify the  chief  of  district in all cases coming to his attention within 3 days after inspection or  after the expiration of the 3 months allowed by law if no action is taken.

- The chief of district, upon receipt of the above-described notice, and in all cases of failure to meet the condi- tions imposed in order to comply with the provisions of the Federal Caustic Poison Act coming  directly under  his supervision, shall transmit to the Dis- trict Director of Customs such evi- dence as he may have at hand tending

#### Pt. 1240

to indicate the importer’s liability and make a recommendation accordingly.

- The District Director of Customs, within 3 days of the receipt of this rec- ommendation, whether favorable or otherwise, shall notify the importer that, the legal period of 3 months for exportation or relabeling having ex- pired, action will be taken within 30 days to enforce the terms of the bond.

## PART 1240—CONTROL OF COMMUNICABLE DISEASES

##### Subpart A—General Provisions

Sec.

1240.3    General definitions.

1240.10      Effective bactericidal treatment.

##### Subpart B—Administrative Procedures

1240.20    Issuance  and  posting  of  certificates following    inspections.

1240.30    Measures in the event of inadequate local   control.

1240.45   Report of disease.

##### Subpart C [Reserved]

Subpart D—Specific Administrative Decisions Regarding Interstate Shipments

- Molluscan shellfish.

- Mandatory pasteurization for all milk and milk products in final package form intended for direct human con- sumption.

- Turtles intrastate and interstate re- quirements.

1240.65   Psittacine birds.

1240.75    Garbage.

##### Subpart E—Source and Use of Potable Water

1240.80    General  requirements  for  water  for drinking  and  culinary  purposes.

1240.83    Approval of watering points.

1240.86    Protection of pier water system.

1240.90    Approval  of  treatment  aboard  con- veyances.

1240.95   Sanitation of water boats.  AUTHORITY: 42 U.S.C. 216, 243, 264, 271.

CROSS REFERENCES: For Department of Health and Human Services regulations re- lating to foreign quarantine, sanitation measures, and control of communicable dis- eases, see Centers for Disease Control’s re- quirements as set forth in 42 CFR parts 71 and 72.

SOURCE: 40 FR 5620, Feb. 6, 1975, unless oth-

erwise noted.

21 CFR Ch. I (4–1–17 Edition)

Subpart A—General Provisions

##### § 1240.3   General definitions.

As used in this part, terms shall have the following meaning:

- Bactericidal treatment. The applica- tion of a method or substance for the destruction of pathogens and other or- ganisms as set forth in § 1240.10.

- Communicable diseases. Illnesses due to infectious agents or their toxic products, which may be transmitted from a reservoir to a susceptible host either directly as from an infected per- son or animal or indirectly through the agency of an intermediate plant or ani- mal host, vector, or the inanimate en- vironment.

- Communicable period. The period or periods during which  the  etiologic agent may be transferred directly or indirectly from the body of the  in- fected person or animal to the body of another.

- Contamination. The presence of a certain amount of undesirable sub- stance or material, which may contain pathogenic  microorganisms.

- Conveyance. Conveyance means any land or air carrier, or any vessel as defined in paragraph (n) of this section.

- Garbage. (1) The solid animal and vegetable waste, together with the nat- ural moisture content, resulting from the handling, preparation, or consump- tion of foods in houses, restaurants, ho- tels, kitchens, and similar establish- ments, or (2) any other food waste con- taining  pork.

- Incubation period. The period be- tween the implanting of disease orga- nisms in a susceptible person and the appearance of clinical manifestation of the disease.

- Interstate traffic. (1) The move- ment of any conveyance or the trans- portation of persons or property, in- cluding any portion of such movement or transportation which is entirely within a State or possession,

- From a point of origin in any State or possession to a point of des- tination in any other State or posses- sion, or

- Between a point of origin and a point of destination in the same State or possession but through any other State, possession, or contiguous for- eign country.

#### Food and Drug Administration, HHS	§ 1240.10

- Interstate traffic does not include the following:

  - The movement of any conveyance which is solely for the purpose of un- loading persons or property trans- ported from a foreign country, or load- ing persons or property for transpor- tation to a foreign country.

  - The movement of any conveyance which is solely for the purpose of ef- fecting its repair, reconstruction, reha- bilitation, or storage.

- Milk. Milk is the product  defined in § 131.110 of this chapter.

- Milk products. Food products made exclusively or principally from the lac- teal secretion obtained from one or more healthy milk-producing animals, e.g., cows, goats, sheep, and water buf- falo, including, but not limited to, the following: lowfat milk, skim milk, cream, half and half, dry milk, nonfat dry milk, dry cream, condensed or con- centrated milk products, cultured or acidified milk or milk products, kefir, eggnog, yogurt, butter, cheese (where not specifically exempted by regula- tion), whey, condensed or dry whey or whey products, ice cream, ice milk, other frozen dairy desserts and prod- ucts obtained by modifying the chem- ical or physical characteristics of milk, cream, or whey by using enzymes, sol- vents, heat, pressure, cooling, vacuum, genetic engineering, fractionation, or other similar processes, and any such product made by the addition or sub- traction of milkfat or the addition of safe and suitable optional ingredients for the protein, vitamin, or mineral fortification of the product.

- Minimum heat treatment. The caus- ing of all particles in garbage to be heated to a boiling temperature and held at that temperature for a period of not less than 30 minutes.

- Possession. Any of the possessions of the United States, including Puerto Rico and the Virgin Islands.

- Potable water. Water which meets the standards prescribed in the Envi- ronmental Protection Agency’s Pri- mary Drinking  Water  Regulations  as set forth in 40 CFR part 141 and the Food and Drug Administration’s sani- tation requirements as set forth in this part and part 1250 of this chapter.

- State. Any State, the District of Columbia, Puerto Rico and the Virgin Islands.

- Utensil. Includes any kitchenware, tableware, glassware, cutlery, con- tainers, or equipment with  which food or drink comes in contact during stor- age, preparation, or serving.

- Vessel. Any passenger-carrying, cargo, or towing vessel exclusive of:

- Fishing boats including those used for shell-fishing;

- Tugs which operate only locally in specific harbors and adjacent waters;

- Barges without means of self-pro- pulsion;

- Construction-equipment boats and dredges; and

- Sand and gravel dredging and han- dling  boats.

- Watering point. The specific place or water boat from which potable water is loaded on a conveyance.

- Molluscan shellfish. Any edible spe- cies of fresh or frozen oysters, clams, mussels, and scallops or edible portions thereof, except when the product con- sists entirely of the shucked adductor muscle.

- Certification number means a unique combination of letters and numbers assigned by a shellfish control authority to a molluscan shellfish processor.

- Shellfish control authority means a Federal, State, or foreign agency, or sovereign tribal government, legally responsible for the administration of a program that includes activities  such as classification of molluscan shellfish growing areas, enforcement of molluscan shellfish harvesting con- trols, and certification of molluscan shellfish processors.

- Tag means a record of harvesting information attached to a container of shellstock by the harvester or proc- essor.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983; 57 FR 57344, Dec. 4, 1992;

60 FR 65201, Dec. 18, 1995]

##### § 1240.10   Effective   bactericidal   treat- ment.

Whenever, under the provisions  of this part, bactericidal treatment is re- quired, it shall be accomplished by one or more of the following methods:

#### § 1240.20

  - By immersion of the utensil or equipment for at least 2 minutes in clean hot water at a temperature of at least 170 F or for one-half minute in boiling water;

  - By immersion of the utensil or equipment for at least 2 minutes in a lukewarm chlorine bath containing at least 50 ppm of available chlorine if hypochlorites are used or a concentra- tion of equal bactericidal strength if chloramines are used;

  - By exposure of the utensil or equipment in a steam cabinet at a tem- perature of at least 170 F for at least

15  minutes  or  at  a  temperature  of  200

F for at least 5 minutes;

  - By exposure of the utensil or equipment in an oven or hot  air  cabi- net at a temperature of at least 180 F for at least 20 minutes;

  - In the case of utensils or equip- ment so designed or installed  as  to make immersion or exposure imprac- tical, the equipment may be treated for the prescribed periods of time either at the temperatures or with chlorine solu- tions as specified above, (1) with live steam from a hose if the steam can be confined, (2) with boiling  rinse  water, or (3) by spraying or swabbing with chlorine  solution;

  - Any other method determined by the Commissioner of Food and Drugs, upon application of an owner or oper- ator of a conveyance, to be effective to prevent the spread of communicable disease.

[40 FR 5620, Feb. 6, 1975, as amended at 54 FR

24900, June 12, 1989]

## Subpart B—Administrative Procedures

##### § 1240.20   Issuance and posting of cer- tificates following inspections.

The Commissioner of Food and Drugs may issue certificates based upon in- spections provided for in this part and part 1250. Such certificates shall be prominently posted on conveyances.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983]

##### § 1240.30   Measures in the event of in- adequate local control.

Whenever the Commissioner of Food and  Drugs  determines  that  the  meas-

21 CFR Ch. I (4–1–17 Edition)

ures taken by health authorities of any State or possession (including political subdivisions thereof) are insufficient to prevent the spread of any of the com- municable diseases from such State or possession to any other State or pos- session, he may take such measures to prevent such spread of the diseases as he deems reasonably necessary, includ- ing inspection, fumigation, disinfec- tion, sanitation, pest extermination, and destruction of animals or articles believed to be sources of infection.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983]

##### § 1240.45   Report of disease.

The master of any vessel or person in charge of any conveyance engaged in interstate traffic, on which a case or suspected case of a communicable dis- ease develops shall, as soon as prac- ticable, notify the local health  author- ity at the next port of call, station, or stop, and shall take such measures to prevent the spread of the disease as the local  health  authority  directs.

## Subpart C [Reserved]

Subpart D—Specific Administra- tive Decisions Regarding Interstate Shipments

##### § 1240.60   Molluscan shellfish.

- A person shall not offer for trans- portation, or transport, in interstate traffic any molluscan shellfish handled or stored  in  such  an  insanitary  man- ner, or grown in an area so contami- nated, as to render such molluscan shellfish likely to become agents in, and their transportation likely to con- tribute to the spread of communicable disease from one State or possession to another.

- All shellstock shall bear a tag that discloses the date and place they were harvested (by State and site), type and quantity of shellfish, and by whom they were harvested (i.e., the identification number assigned to the harvester by the shellfish control au- thority, where applicable or, if such identification numbers are not as- signed, the name of the harvester  or the name or registration number of the harvester’s vessel). In place of the tag,

Food and Drug Administration, HHS	§ 1240.62

bulk shellstock shipments may be ac- companied by a bill of lading or similar shipping document that contains the same information.

- All containers of shucked molluscan shellfish shall bear a label that identifies the name, address, and certification number of the packer or repacker of the molluscan shellfish.

- Any molluscan shellfish without such a tag, shipping document,  or label, or with a tag, shipping docu- ment, or label that does not  bear  all the information required by paragraphs

- and (c) of this section, shall be sub- ject to seizure or refusal of entry, and destruction.

[40 FR 5620, Feb. 6, 1975, as amended at 60 FR

65202, Dec. 18, 1995]

##### § 1240.61 Mandatory pasteurization for all milk and milk products in final package form intended for direct human consumption.

  - No person shall cause to be deliv- ered into interstate commerce or shall sell, otherwise distribute, or hold for sale or other distribution after ship- ment in interstate commerce  any  milk or milk product in final package  form for direct human consumption unless the product has been pasteurized or is made from dairy ingredients (milk or milk products)  that have all  been  pas- teurized, except where alternative pro- cedures to pasteurization are provided for by regulation, such as in part 133 of this chapter for  curing  of  certain cheese  varieties.

  - Except as provided in paragraphs

  - and (d) of this section, the terms ‘‘pasteurization,’’ ‘‘pasteurized,’’ and similar terms shall mean the process of heating every particle of  milk  and milk product in properly designed and operated equipment to one of the tem- peratures given in the following table and held continuously at or above that temperature for at least the cor- responding specified time:

| Temperature | Time |
| --- | --- |
| 145 F (63 C) 1 ..................................... 161 F (72 C) 1 ..................................... 191 F (89 C) ....................................... | 30 minutes. 15 seconds. 1 second. |

1 If the fat content of the milk product is 10 percent or more, or if it contains added sweeteners, the specified temperature shall be increased by 5 F (3 C).

| Temperature | Time |
| --- | --- |
| 194 F (90 C) ....................................... | 0.5 second. |
| 201 F (94 C) ....................................... | 0.1 second. |
| 204 F (96 C) ....................................... | 0.05 second. |
| 212 F (100 C) ..................................... | 0.01 second. |

- Eggnog shall be heated to at least the following temperature and time specification:

| Temperature | Time |
| --- | --- |
| 155 F (69 C) ....................................... 175 F (80 C) ....................................... 180 F (83 C) ....................................... | 30 minutes. 25 seconds. 15 seconds. |

- Neither paragraph (b) nor (c) of this section shall be construed as bar- ring any other pasteurization process that has been recognized by the Food and Drug Administration to be equally efficient in the destruction of micro- bial organisms of public health signifi- cance.

[52 FR 29514, Aug. 10, 1987, as amended at 57

FR 57344, Dec. 4, 1992]

##### § 1240.62   Turtles  intrastate  and  inter- state requirements.

- Definition. As used in this section the term ‘‘turtles’’ includes all animals commonly known as turtles, tortoises, terrapins, and all other animals of the order Testudinata, class Reptilia,  ex- cept marine species (families Dermachelidae and Chelonidae).

- Sales; general prohibition. Except as otherwise provided in this section, viable turtle  eggs  and  live  turtles  with a carapace length of less than 4 inches shall not be sold, held for sale, or of- fered for any other type of commercial or public distribution.

- Exceptions. The provisions of this section are not applicable to:

  - The sale, holding for sale, and dis- tribution of live turtles and viable tur- tle eggs for bona fide scientific, edu- cational, or exhibitional purposes, other than use as pets.

  - The sale, holding for sale, and dis- tribution of live turtles and viable tur- tle eggs not in connection with a busi- ness.

  - The sale, holding for sale, and dis- tribution of live turtles and viable tur- tle eggs intended for export only, pro- vided that the outside of the shipping package is conspicuously labeled ‘‘For Export Only.’’

#### § 1240.65

  - Marine turtles excluded from this regulation under the provisions of paragraph (a) of this  section  and  eggs of such turtles.

- Petitions. The Commissioner of Food and Drugs, either on his own ini- tiative or on behalf of any interested person who has  submitted  a  petition, may publish a proposal to amend this regulation. Any such petition shall  in- clude an adequate factual basis to sup- port the petition, and will be  published for comment if it contains reasonable grounds for the proposed regulation. A petition requesting such a regulation, which  would  amend  this  regulation, shall be submitted to the Division of Dockets Management, Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852.

[40 FR 22545, May 23, 1975, as amended at 46

FR 8461,  Jan. 27,  1981;  48  FR 11431,  Mar. 18,

1983; 54 FR 24900, June 12, 1989; 59 FR 14366,

Mar. 28, 1994; 66 FR 56035, Nov. 6, 2001; 70 FR

48073, Aug. 18, 2005; 78 FR 44881, July 25, 2013]

##### § 1240.65   Psittacine birds.

- The term psittacine birds shall in- clude all birds commonly known as parrots, Amazons, Mexican double heads, African grays, cocatoos, ma- caws, parakeets, love birds, lories, lorikeets, and all other birds of the psittacine  family.

- No person shall transport, or offer for transportation, in interstate traffic any psittacine bird unless the shipment is accompanied by a permit from the State health department of  the  State of destination where required by such department.

- Whenever the Surgeon General finds that psittacine birds or human beings in any area are infected with psittacosis and there is such danger of transmission of  psittacosis  from  such area as to endanger the public health, he may declare it an area of infection. No person shall thereafter transport, or offer for transportation, in interstate traffic any psittacine bird from such area, except shipments authorized by the Surgeon General for purposes of medical research and accompanied by a permit issued by him, until the Sur- geon General finds that there is no longer any danger of transmission of psittacosis from such area. As used in this  paragraph,  the  term  ‘‘area’’  in-

#### 21 CFR Ch. I (4–1–17 Edition)

cludes,  but  is  not  limited  to,  specific premises or buildings.

##### § 1240.75   Garbage.

- A person shall not transport, re- ceive, or cause to be transported or re- ceived, garbage  in  interstate   traffic and feed such garbage to swine unless, prior to the feeding, such garbage has received minimum heat treatment.

- A person transporting garbage in interstate traffic shall not make, or agree to make, delivery thereof to any person with knowledge of the intent or customary practice of such person to feed to  swine  garbage which  has  not been subjected to minimum heat treat- ment.

## Subpart E—Source and Use of Potable Water

##### § 1240.80 General requirements for water for drinking and culinary purposes.

Only potable water shall be provided for drinking and culinary purposes by any operator of a  conveyance  engaged in interstate traffic, except as provided in § 1250.84(b) of this chapter. Such water shall either have been obtained from watering points approved by the Commissioner of Food and Drugs, or, if treated  aboard  a  conveyance,  shall have been subjected to treatment ap- proved by the Commissioner of Food and Drugs.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983]

##### § 1240.83   Approval of watering points.

- The Commissioner of Food and Drugs shall  approve  any  watering  point if (1) the water supply thereat  meets the standards prescribed in the Envi- ronmental  Protection  Agency’s  Pri- mary  Drinking   Water   Regulations   as set forth in 40 CFR part 141, and (2) the methods of and facilities for delivery of such water to the conveyance and the sanitary conditions surrounding such delivery prevent the introduction, transmission, or spread of commu- nicable   diseases.

- The Commissioner of Food and Drugs may base his approval or dis- approval of a watering point upon in- vestigations   made   by   representatives

Food and Drug Administration, HHS	Pt. 1250

of State departments of health or of the health authorities of contiguous foreign nations.

- If a watering point has not been approved, the Commissioner of Food and Drugs may permit its temporary use under such conditions as, in his judgment, are necessary to prevent the introduction,  transmission,  or  spread of communicable diseases.

- Upon request of the Commissioner of Food and Drugs, operators of con- veyances shall provide information  as to watering points used by them.

water, have been approved by the Com- missioner of Food and Drugs.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983]

## PART 1250—INTERSTATE CONVEYANCE SANITATION

##### Subpart A—General Provisions

Sec.

- Definitions.

##### Subpart B—Food Service Sanitation on Land and Air Conveyances, and Vessels

| [40 FR 5620, Feb. 6, 1975, as amended at 48 FR | 1250.20 | Applicability. |
| --- | --- | --- |
| 11431, Mar. 18, 1983; 48 FR 13978, Apr. 1, 1983] | 1250.21 | Inspection. |

§ 1240.86   Protection of pier water sys- tem.

No vessel engaged in interstate traf- fic shall make a connection between its nonpotable water system and any pier potable water system unless provisions are made to prevent backflow from the vessel to the pier.

##### § 1240.90   Approval	of	treatment aboard conveyances.

  - The treatment of water aboard conveyances shall be approved by the Commissioner of Food and Drugs if the apparatus used is of such design and is so operated as to be capable of pro- ducing and in fact does produce, pota- ble water.

  - The Commissioner of Food and Drugs may base his approval or dis- approval of the treatment of  water upon investigations made by represent- atives of State departments  of  health or of the health authorities of contig- uous foreign nations.

  - Overboard water treated on ves- sels shall be from areas relatively free of contamination and pollution.

[40 FR 5620, Feb. 6, 1975, as amended at 48 FR

11431, Mar. 18, 1983]

##### § 1240.95   Sanitation of water boats.

No vessel engaged in interstate traf- fic shall obtain water for drinking and culinary purposes from any water boat unless the tanks, piping, and other ap- purtenances used by the water boat in the loading, transportation, and deliv- ery   of   such   drinking   and   culinary

1250.22   General requirements.

- Source identification and inspection of food and drink.

- Special food requirements.

- Storage of perishables.

- Source and handling of ice.

1250.30 Construction,  maintenance  and  use of places where food  is  prepared,  served, or  stored.

- Food-handling  operations.

- Utensils and equipment.

- Refrigeration equipment.

- Health of persons handling food.

- Toilet  and  lavatory  facilities  for  use of food-handling employees.

- Garbage equipment and disposition.

##### Subpart C—Equipment and Operation of Land and Air Conveyances

- Applicability.

- Submittal of construction plans.

- Water  systems;  constant  tempera- ture bottles.

- Ice.

- Drinking utensils and toilet articles.

- Food  handling  facilities  on  railroad conveyances.

- Cleanliness of conveyances.

- Toilet and lavatory facilities.

- Railroad   conveyances;   discharge   of wastes.

- Discharge of wastes on highway con- veyances.

- Discharge  of  wastes  on  air  convey- ances.

##### Subpart D—Servicing Areas for Land and Air Conveyances

- Applicability.

- Inspection and approval.

- Submittal of construction plans.

- General requirements.

1250.65  Drainage.

1250.67   Watering equipment.

1250.70   Employee conveniences.

1250.75   Disposal of human wastes.

§ 1250.3

- Garbage disposal.

##### Subpart E—Sanitation Facilities and Conditions on Vessels

- Applicability.

- Inspection.

- Potable water systems.

- Storage of water prior to treatment.

- Water  in  galleys  and  medical  care spaces.

- Drinking fountains and coolers; ice; constant temperature bottles.

- Water for making ice.

- Wash water.

- Swimming pools.

- Toilets and lavatories.

1250.93   Discharge of wastes.

- Insect control.

- Rodent control.

AUTHORITY: 42 U.S.C. 216, 243, 264, 271.

CROSS REFERENCES: For Department of Health and Human Services regulations re- lating to foreign quarantine and control of communicable diseases, see Centers for Dis- ease Control’s requirements as set forth in 42 CFR parts 71 and 72.

SOURCE: 40 FR 5624, Feb. 6, 1975, unless oth-

erwise noted.

## Subpart A—General Provisions

##### § 1250.3   Definitions.

As used in this part, terms shall have the following meaning:

  - Bactericidal treatment. The applica- tion of a method or substance for the destruction of pathogens and other or- ganisms as set forth in § 1240.10 of this chapter.

  - Communicable diseases. Illnesses due to infectious agents or their toxic products, which may be transmitted from a reservoir to a susceptible host either directly as from an infected per- son or animal or indirectly through the agency of an intermediate plant or ani- mal host, vector, or the inanimate en- vironment.

  - Communicable period. The period or periods during which  the  etiologic agent may be transferred directly or indirectly from the body of the  in- fected person or animal to the body of another.

  - Contamination. The presence of a certain amount of undesirable sub- stance or material, which may contain pathogenic  microorganisms.

  - Conveyance. Conveyance means any land or air carrier, or any vessel as

21 CFR Ch. I (4–1–17 Edition)

defined  in  paragraph  (m)  of  this  sec- tion.

  - Existing vessel. Any vessel the con- struction of which was started prior to the effective date of the regulations in this part.

  - Garbage. (1) The solid animal and vegetable waste, together with the nat- ural moisture content, resulting from the handling, preparation, or consump- tion of foods in houses, restaurants, ho- tels, kitchens, and similar establish- ments, or (2) any other food waste con- taining  pork.

  - Interstate traffic. (1) The move- ment of any conveyance or the trans- portation of persons or property, in- cluding any portion of such movement or transportation which is entirely within a State or possession, (i) from a point of origin in any State or posses- sion to a point of destination in any other State or possession, or (ii) be- tween a point of origin and a point of destination in the same State or pos- session but through any other State, possession, or contiguous foreign coun- try.

- Interstate traffic does not include the following:

  - The movement of any conveyance which is solely for the purpose of un- loading persons or property trans- ported from a foreign country, or load- ing persons or property for transpor- tation to a foreign country.

  - The movement of any conveyance which is solely for the purpose of ef- fecting its repair, reconstruction, reha- bilitation, or storage.

- Possession. Any of the possessions of the United States, including Puerto Rico and the Virgin Islands.

- Potable water. Water which meets the standards prescribed in the Envi- ronmental Protection Agency’s Pri- mary Drinking  Water  Regulations  as set forth in 40 CFR part 141 and the Food and Drug Administration’s sani- tation regulations as set forth in this part and part 1240 of this chapter.

- State. Any State, the District of Columbia, Puerto Rico and the Virgin Islands.

- Utensil. Includes any kitchenware, tableware, glassware, cutlery, con- tainers, or  equipment with  which food or drink comes in contact during stor- age, preparation, or serving.

#### Food and Drug Administration, HHS	§ 1250.26

- Vessel. Any passenger-carrying, cargo, or towing vessel exclusive of:

- Fishing boats including those used for shell-fishing;

- Tugs which operate only locally in specific harbors and adjacent waters;

- Barges without means of self-pro- pulsion;

- Construction-equipment boats and dredges; and

- Sand and gravel dredging and han- dling  boats.

- Wash water. Water suitable for do- mestic uses  other  than  for  drinking and culinary purposes, and medical care purposes excluding hydrotherapy.

- Shellfish. Any fresh, frozen, or in- completely cooked oysters, clams, or mussels, either shucked or in the shell, and any fresh, frozen, or incompletely cooked edible products thereof.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

## Subpart B—Food Service Sanita- tion on Land and Air Convey- ances, and Vessels

##### § 1250.20   Applicability.

All conveyances  engaged  in  inter- state traffic shall comply with the re- quirements prescribed in  this  subpart and § 1240.20 of this chapter.

##### § 1250.21   Inspection.

The Commissioner of Food and Drugs may inspect such conveyance to deter- mine compliance with the  require- ments of this subpart and § 1240.20 of this   chapter.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.22   General requirements.

All food and drink served on convey- ances shall be  clean,  wholesome,  and free from spoilage, and shall be  pre- pared, stored, handled, and served in accordance with the requirements pre- scribed in this subpart and § 1240.20 of this   chapter.

##### § 1250.25   Source identification and in- spection of food and drink.

- Operators of conveyances shall identify, when requested by the Com- missioner of Food and Drugs, the ven- dors,    distributors    or    dealers    from

whom they have acquired or are ac- quiring their food supply, including milk, fluid  milk  products,  ice  cream and other frozen desserts, butter, cheese, bottled water, sandwiches and box lunches.

- The Commissioner of Food and Drugs may inspect any source of such food supply in order to determine whether the requirements of the regu- lations in this subpart  and  in  § 1240.20 of this chapter are being met, and may utilize  the  results  of   inspections   of such  sources  made  by   representatives of State health departments or of the health  authorities  of  contiguous   for- eign   nations.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.26   Special food requirements.

Milk, fluid milk products, ice cream and other frozen desserts, butter, cheese, and shellfish served or sold on conveyances shall conform to the fol- lowing requirements:

- Milk and fluid milk products, in- cluding cream, buttermilk, skim milk, milk beverages, and reconstituted milk, shall be pasteurized and obtained from a source of supply approved by the Commissioner of Food and Drugs. The Commissioner of Food and Drugs shall approve any source of supply at or from which milk or fluid milk prod- ucts are produced, processed, and dis- tributed so as to prevent the introduc- tion, transmission, or spread of com- municable diseases. If a source of sup- ply of milk or fluid milk products has not been approved, the Commissioner of Food and Drugs may permit its tem- porary use under such conditions as, in his judgment, are necessary to prevent the introduction, transmission, or spread of communicable diseases. Con- tainers of milk and fluid milk products shall be plainly labeled to show the contents, the word ‘‘pasteurized’’, and the identity of the plant at which the contents were packaged by name and address, provided that a code may be used in lieu of address.

- Ice cream, other frozen desserts, and butter shall be manufactured from milk or milk products that have been pasteurized or subjected to equivalent heat  treatment.

§ 1250.27

- Cheese shall be (1) pasteurized or subjected to equivalent  heat  treat- ment, (2) made from pasteurized milk products or from milk products which have been subjected to equivalent heat treatment, or (3)  cured  for  not  less than 60 days at a temperature not less than 35 F.

- Milk, buttermilk, and milk bev- erages shall be served in or from the original individual containers in which received from the distributor, or from a bulk container equipped with a dis- pensing device so designed, con- structed, installed, and maintained as to prevent the transmission of commu- nicable diseases.

- Shellfish purchased for consump- tion on any conveyance shall originate from a dealer currently listed by the Public Health Service as holding an un- expired and unrevoked certificate issued by a State authority.

- Shucked shellfish shall be pur- chased in the containers in which they are placed at the shucking plant and shall be kept therein until used. The State abbreviation and the certificate number of the packers shall be perma- nently recorded on the container.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.27   Storage of perishables.

All perishable food or drink shall be kept at or below 50 F, except when being prepared or kept hot for serving.

##### § 1250.28   Source and handling of ice.

Ice coming in contact with food or drink and not manufactured on the conveyance shall be obtained from sources approved by competent health authorities. All ice coming in contact with food or drink shall be stored and handled in such manner as to avoid contamination.

##### § 1250.30 Construction,  maintenance and use of places where food is pre- pared, served, or stored.

- All  kitchens,  galleys,  pantries, and other places where food is pre- pared, served, or stored shall be ade- quately lighted and ventilated: Pro- vided, however, That ventilation of cold storage rooms shall not be  required. All such places where food is prepared, served,   or   stored   shall   be   so   con-

#### 21 CFR Ch. I (4–1–17 Edition)

structed and maintained as to be clean and free from flies, rodents, and other vermin.

- Such places shall not be used for sleeping or living quarters.

- Water of satisfactory sanitary quality, under head or pressure, and adequate in amount and temperature, shall be easily accessible to all rooms in which food is prepared and utensils are cleaned.

- All plumbing shall be so designed, installed, and maintained as to prevent contamination of the water supply, food, and food utensils.

##### § 1250.32   Food-handling operations.

- All food-handling operations shall be accomplished so as to minimize the possibility of contaminating food, drink, or utensils.

- The hands of all persons shall be kept clean while engaged in handling food, drink, utensils, or equipment.

##### § 1250.33   Utensils and equipment.

- All utensils and working surfaces used in connection with the prepara- tion, storage, and serving of food or beverages, and the cleaning of food utensils, shall be so constructed as to be easily cleaned and self-draining and shall be maintained in good repair. Adequate facilities shall be provided for the cleaning and bactericidal treat- ment of all multiuse eating and drink- ing utensils and equipment used in the preparation of food and beverages. An indicating thermometer, suitably lo- cated, shall be provided to permit the determination of the hot water tem- perature when and where hot water is used as the bactericidal agent.

- All multiuse eating and drinking utensils shall be thoroughly cleaned in warm water and subjected to an effec- tive bactericidal treatment after each use. All other utensils that come in contact with food and drink shall be similarly treated immediately fol- lowing the day’s operation. All equip- ment shall be kept clean.

- After bactericidal  treatment, utensils shall be stored and handled in such manner as to prevent contamina- tion before reuse.

#### Food and Drug Administration, HHS	§ 1250.42

##### § 1250.34   Refrigeration equipment.

Each refrigerator shall be equipped with a thermometer located in the warmest portion thereof. Waste water drains from ice boxes, refrigerating equipment, and refrigerated  spaces shall be so installed as to prevent back- flow of contaminating liquids.

##### § 1250.35   Health  of  persons  handling food.

  - Any person who is known or sus- pected to be in a communicable period or a carrier of any communicable dis- ease shall not be  permitted  to  engage in the preparation, handling, or serving of water, other beverages, or food.

  - Any person known or suspected to be suffering from gastrointestinal dis- turbance or  who  has  on  the  exposed portion of the body an open lesion or an infected wound shall not be per- mitted to engage in the preparation, handling, or serving of food or bev- erages.

##### § 1250.38   Toilet  and  lavatory  facilities for use of food-handling employees.

- Toilet and lavatory facilities of suitable design and construction shall be provided for use of food-handling employees. Railroad dining car crew lavatory  facilities  are  regulated  under

§ 1250.45.

- Signs directing food-handling em- ployees to wash their hands after each use of toilet facilities  shall  be  posted so as to be readily observable by such employees. Hand washing  facilities shall include soap, sanitary towels and hot and cold running water or warm running water in lieu of hot and cold running water.

- All toilet rooms shall be main- tained in a clean condition.

##### § 1250.39   Garbage  equipment  and  dis- position.

Watertight, readily cleanable non- absorbent containers with close-fitting covers shall be used to receive and store garbage. Garbage and refuse shall be disposed of as frequently as is nec- essary and practicable.

## Subpart C—Equipment and Oper- ation of Land and Air Con- veyances

##### § 1250.40   Applicability.

The sanitary equipment and facilities on land and air conveyances engaged in interstate traffic and the use of such equipment and facilities shall comply with the requirements  prescribed  in this  subpart.

##### § 1250.41 Submittal of construction plans.

Plans for the construction or major reconstruction of sanitary equipment or facilities for such conveyances shall be submitted to the Commissioner of Food and Drugs for review of the con- formity of such plans with the require- ments of this subpart, except that sub- mittal of plans shall not be required for any conveyance under reconstruction if the owner or operator thereof has made arrangements satisfactory to the Com- missioner of Food and Drugs for inspec- tions of such conveyances while under reconstruction for the purpose of deter- mining conformity with those require- ments.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.42 Water systems; constant tem- perature bottles.

- The water system, whether of the pressure or gravity type, shall be com- plete and closed from  the  filling  ends to the discharge taps, except for pro- tected vent openings. The water sys- tem shall be protected against back- flow.

- Filling pipes or connections through which water tanks are sup- plied shall be provided on both sides of all new railway conveyances and on ex- isting conveyances when they undergo heavy repairs. All filling connections shall be easily cleanable and so located and protected as to minimize the haz- ard of contamination of the water sup- ply.

- On all new or reconstructed con- veyances, water coolers shall be an in- tegral part of the closed system.

- Water filters if used on dining cars and other conveyances will be per- mitted only if they are so operated and

§ 1250.43

maintained at all times as to prevent contamination of the water.

- Constant temperature bottles and other containers used for storing or dispensing potable water shall be kept clean at all times and shall be sub- jected to effective bactericidal treat- ment as often as may be necessary to prevent the contamination of water so stored and dispensed.

##### § 1250.43   Ice.

Ice shall not be permitted to come in contact with water in coolers or con- stant  temperature  bottles.

##### § 1250.44   Drinking  utensils  and  toilet articles.

- No cup, glass, or other drinking utensil which may be used by more than one person shall be provided on any conveyance unless such cup, glass, or drinking utensil shall have been thoroughly cleaned and subjected to ef- fective bactericidal treatment  after each individual use.

- Towels, combs, or brushes for common use shall not be provided.

##### § 1250.45   Food  handling  facilities  on railroad conveyances.

- Both kitchens and pantries of cars hereafter constructed or reconstructed shall be equipped with double sinks, one of which shall be of sufficient size and depth to permit complete immer- sion of a basket of dishes during bac- tericidal treatment; in the pantry a dishwashing machine may be sub- stituted for the double sinks. If chemi- cals are used for bactericidal treat- ment, 3-compartment sinks shall be provided.

- A sink shall be provided for wash- ing and handling cracked ice used in food or drink and shall be used for no other purpose.

- Lavatory facilities for the use of the dining car crew shall  be  provided on each dining car. Such facilities shall be conveniently located and used for hand and face washing only: Provided, however, That where the kitchen and pantry on a dining car hereafter con- structed or reconstructed are so parti- tioned or separated as to impede free passage between them lavatory facili- ties shall be provided in both the kitch- en and the pantry.

#### 21 CFR Ch. I (4–1–17 Edition)

- Wherever toilet and lavatory fa- cilities required by paragraph (c) of this section are not on the dining car, a lavatory shall be provided on the din- ing car for the use of employees. The lavatory shall be conveniently located and used only for the purpose for which it is installed.

##### § 1250.49   Cleanliness of conveyances.

Conveyances while in transit shall be kept clean and free of flies and mosqui- toes. A conveyance which becomes in- fected with vermin shall be placed out of service until such time as it shall have been effectively treated for the destruction of the vermin.

##### § 1250.50   Toilet and lavatory facilities.

Where toilet  and  lavatory  facilities are provided on conveyances they shall be so designed as to permit ready cleaning. On conveyances not equipped with retention facilities, toilet hoppers shall be of such  design  and  so  located as to prevent spattering of water filling pipes or hydrants.

##### § 1250.51   Railroad    conveyances;    dis- charge of wastes.

- New railroad conveyances. Human wastes, garbage, waste water, or other polluting materials shall not be dis- charged from any new railroad convey- ance except at servicing areas approved by the Commissioner of Food and Drugs. In lieu of retention pending dis- charge at approved servicing areas, human wastes, garbage, waste  water, or other polluting materials that have been suitably treated to prevent the spread of communicable diseases may be discharged from such conveyances, except at stations. For the purposes of this section, ‘‘new railroad convey- ance’’ means any such conveyance placed into service for the first time after July 1, 1972, and the terms ‘‘waste water or other polluting materials’’ do not include drainage of drinking water taps  or  lavatory  facilities.

- Nonnew railroad conveyances. Human wastes, garbage, waste  water, or other polluting materials shall  not be discharged from any railroad con- veyance, other than passenger convey- ances for which an extension has been granted pursuant to paragraph (f) of this  section,  after  December  31,  1977,

#### Food and Drug Administration, HHS	§ 1250.51

except at servicing areas approved by the Commissioner of Food  and  Drugs. In lieu of retention  pending  discharge at approved servicing areas, human wastes, garbage, waste water, or other polluting materials that have been suitably treated to prevent  the  spread of communicable diseases may be dis- charged from such conveyances, except at stations. The terms ‘‘waste water or other polluting materials’’ do not in- clude drainage of drinking  water  taps or  lavatory  facilities.

- Toilets. When railroad convey- ances, occupied or open to occupancy by travelers, are at a station or serv- icing area, toilets shall be kept locked unless means are provided to prevent contamination of the area or station.

- Submission of annual report. Each railroad company shall submit to the Center for Food Safety and Applied Nu- trition (HFS–627), Food and Drug Ad- ministration, 5001 Campus Dr., College Park, MD 20740, an annual report of ac- complishments made in modifying con- veyances to achieve compliance with paragraph (b) of this section. Annual reports shall be required until a report is submitted showing that 100 percent of the company’s conveyances can com- ply with the requirements of paragraph

(b) of this section; annual  reports  shall be required subsequent  to  such  report if conveyances  not  capable  of  com- plying with the requirements of para- graph (b) of this section are acquired. Every railroad company shall have not less than 10 percent of its nonpassenger conveyances that are in operation  ca- pable of complying with the require- ments of paragraph (b) of  this  section by December 31, 1974, not less than 40 percent by December 31, 1975, and not less than 70 percent by December 31, 1976. All conveyances, other than pas- senger conveyances for which an exten- sion has been granted pursuant to paragraph (f) of this section, in oper- ation after December 31, 1977, shall be capable   of   complying   with   paragraph

- of this section.

  - Requirements of annual report. An- nual reports required by paragraph (d) of this section shall be submitted with- in 60 days of the end of each calendar year. Each report shall contain at least the following information:

    - Company name and  address.

    - Name, title, and address of the company’s chief operating official.

    - Name, title, address, and tele- phone number of the person designated by the company to be directly respon- sible  for  compliance  with  this  section.

    - A statement that all new railroad conveyances placed into service after July 1, 1972 meet the requirements of this section.

    - A complete, factual, narrative statement explaining why retrofitting of noncomplying nonnew conveyances is incomplete, if it is incomplete.

    - A statement of the percentage of conveyances retrofitted with waste dis- charge facilities in  compliance  with this section as of the reporting date and the percentage expected to be com- pleted by December 31st of the fol- lowing year.

    - A tabular report with the fol- lowing vertical columns: equipment type, e.g., locomotive, caboose, pas- senger car, and any others having toi- lets; number of toilets per conveyance; number of each equipment type in op- eration; and number of each to be ret- rofitted by December 31st of each year until 100 percent compliance with this section is achieved.

  - Variances and extensions—(1) Variances. Upon application by a rail- road company, the Director, Center for Food Safety and Applied Nutrition, may grant a variance from the compli- ance schedule prescribed in paragraph

(d) of this section for nonpassenger conveyances when the requested vari- ance is required to prevent substantial disruption of the railroad company’s operations. Such variance shall not af- fect the final deadline of compliance established in paragraph (d) of this sec- tion.

- Extensions. Upon application by a railroad company, the Director, Center for Food Safety and Applied Nutrition, may grant an extension of time for compliance with the requirements of paragraph (b) of this section beyond December 31, 1977, for passenger con- veyances operated by railroad compa- nies when compliance cannot be achieved without substantial disrup- tion of the railroad company’s oper- ations.

#### § 1250.52

- Application for variance or exten- sion. Application for variances or ex- tensions shall be submitted to the Food and Drug  Administration,  Center  for Food Safety and Applied Nutrition, Manager, Interstate Travel Sanitation Sub-Program, HFF–312, 5001 Campus Dr., College Park, MD 20740, and shall include  the  following  information:

  - A detailed  description  of the  pro- posed deviation from the requirements of paragraphs (b) or (d) of this section.

  - A report, current to the date of the request for a variance or extension, containing the information required by paragraph (e) of this section.

- Administration of variances and ex- tensions. (i) Written notification of the granting or refusal of a variance or ex- tension will be provided to the apply- ing railroad company by the Director, Center for Food Safety and Applied Nu- trition. The notification of a granted variance will state the approved devi- ation from the compliance schedule provided for in paragraph (d) of this section. The notification of a granted extension will state the final date for compliance with the provisions of para- graph (b) of this section.

    - A public file of requested variances and extensions, their disposi- tion, and information relating to pend- ing actions will be maintained in the Division of Dockets Management, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852.

    - After notice to the railroad com- pany and opportunity for hearing in ac- cordance with part 16 of this chapter, a variance or extension may be with- drawn prior to its scheduled termi- nation if the Director, Center for Food Safety and Applied Nutrition, deter- mines that such withdrawal is nec- essary to protect the public health.

CROSS REFERENCE: For statutory exemp- tions for ‘‘intercity rail passenger service,’’ see section 306(i) of 45 U.S.C. 546(i).

[40 FR 5624, Feb. 6, 1975, as amended at 40 FR

30110, July 17, 1975; 46 FR 8461, Jan. 27, 1981;

48  FR 11432,  Mar. 18,  1983;  54  FR 24900,  June

12, 1989; 59 FR 14366, Mar. 28, 1994; 61 FR 14481,

Apr.  2,  1996;  66  FR 56035,  Nov. 6,  2001;  81  FR

49897, July 29, 2016]

EFFECTIVE DATE NOTE: For  a  document staying the effectiveness of § 1250.51 (b) and (d), see 42 FR 57122, Nov. 1, 1977.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 1250.52   Discharge of wastes on high- way conveyances.

There shall be no discharge of excre- ment, garbage, or waste water from a highway conveyance except at serv- icing areas approved by the Commis- sioner of Food and Drugs.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.53   Discharge  of  wastes  on  air conveyances.

There shall be no discharge of excre- ment or garbage from any air convey- ance except at servicing areas approved by the Commissioner of Food and Drugs.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

## Subpart D—Servicing Areas for Land and Air Conveyances

##### § 1250.60   Applicability.

Land and air conveyances engaged in interstate traffic shall use only such servicing areas within the  United States as have been approved by the Commissioner of Food and Drugs as being in compliance with the require- ments prescribed in this subpart.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.61   Inspection and approval.

The Commissioner of Food and Drugs may inspect any such areas to deter- mine whether they shall be approved. He may base his approval or dis- approval on investigations made by representatives  of  State  departments of  health.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.62   Submittal	of	construction plans.

Plans for construction or major re- construction of sanitation facilities at servicing areas shall be submitted to the Commissioner of Food and Drugs for review of the conformity of the pro- posed facilities with the  requirements of this subpart.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

Food and Drug Administration, HHS	§ 1250.79

##### § 1250.63   General requirements.

Servicing areas shall be provided with all necessary sanitary facilities so operated and maintained as to prevent the spread of communicable diseases.

##### § 1250.65   Drainage.

All platforms and other places at which water or food supplies are loaded onto or removed from conveyances shall be adequately drained so as to prevent pooling.

##### § 1250.67   Watering equipment.

- General requirements. All servicing area piping systems, hydrants, taps, faucets, hoses, buckets, and other ap- purtenances necessary for delivery of drinking and culinary water to a con- veyance shall be designed, constructed, maintained and operated in such a manner as to prevent contamination of the water.

- Outlets for nonpotable water. Out- lets for nonpotable water shall be pro- vided with fittings different from those provided for outlets for potable water and each nonpotable water outlet shall be posted with permanent signs warn- ing that the water is unfit for drinking.

- Ice. If bulk ice is used for the cool- ing of drinking water or other bev- erages, or for food preservation pur- poses, equipment constructed so as not to become a factor in the transmission of communicable diseases shall be pro- vided for the storage, washing, han- dling, and delivery to conveyances of such bulk ice, and  such  equipment shall be used for no other purposes.

##### § 1250.70   Employee conveniences.

- There shall be adequate toilet, washroom, locker, and other essential sanitary facilities readily  accessible for use of employees adjacent to places or areas where land and air convey- ances are serviced, maintained, and cleaned. These facilities shall be main- tained in a clean and sanitary condi- tion at all times.

- In the case of diners not in a train but with a crew on board, adequate toi- let facilities shall be available to the crew within a reasonable distance but not exceeding 500 feet of such diners.

- Drinking fountains and coolers shall   be   constructed   of   impervious,

nonoxidizing material, and shall be so designed and constructed as to be eas- ily cleaned. The jet of a drinking foun- tain shall be slanting and the orifice of the jet shall be protected by a guard in such a manner as to prevent contami- nation thereof by droppings from the mouth. The orifice of such  a  jet  shall be located a sufficient distance above the rim of the basin to prevent back- flow.

##### § 1250.75   Disposal of human wastes.

- At servicing areas and at stations where land and air conveyances are oc- cupied by passengers the operations shall be so conducted as to avoid con- tamination of such areas and  stations by human wastes.

- Toilet wastes shall be disposed of through sanitary sewers or by other methods assuring sanitary disposal of such wastes. All soil cans and remov- able containers shall be thoroughly cleaned before being returned to use. Equipment for cleaning such con- tainers and for flushing nonremovable containers and  waste  carts shall  be  so designed as to prevent backflow into the water line, and  such  equipment shall be used for no purpose connected with the handling of food, water or ice.

- All persons who have handled soil cans or other containers which have come in contact with human wastes shall be required to wash their hands thoroughly with soap and warm water and to remove any garments which have become soiled with such wastes before engaging in any work connected with the loading, unloading, trans- porting or other handling of food, water or ice.

##### § 1250.79   Garbage disposal.

- Water-tight, readily cleanable, nonabsorbent containers with close-fit- ting covers shall be used to receive and store garbage.

- Can washing and draining facili- ties shall be provided.

- Garbage cans shall be emptied daily and shall be thoroughly washed before being returned for use.

#### § 1250.80

Subpart E—Sanitation Facilities and Conditions on Vessels

##### § 1250.80   Applicability.

The sanitation facilities and the san- itary conditions on vessels engaged in interstate traffic shall comply with the requirements prescribed in this sub- part, provided that no major structural change will be required on existing ves- sels.

##### § 1250.81   Inspection.

The Commissioner of Food and Drugs may inspect such vessels to determine compliance with the requirements of this  subpart.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.82   Potable water systems.

The following conditions must  be met by vessel water systems used  for the storage and distribution of water which   has   met   the   requirements   of

§ 1240.80 of this chapter.

  - The potable water system, includ- ing filling hose and lines, pumps, tanks, and distributing pipes, shall be separate and distinct from other water systems and shall be used for no other purposes.

  - All potable water tanks shall be independent of any tanks holding non- potable water or other liquid. All pota- ble water tanks shall be independent of the shell of the ship unless (1) the bot- tom of the tank is at least 2 feet above the maximum load water line, (2) the seams in the shell are continuously welded, and (3) there are no rivets in that part of the shell which forms a side of a tank. A deck may be used as the top of a tank provided there are no access or inspection openings or rivets therein, and the seams are continu- ously welded. No toilet or  urinal  shall be installed immediately above  that part of the deck which forms the top of a tank. All potable water  tanks  shall be located at a sufficient height  above the bilge to allow for draining and to prevent submergence in bilge water.

  - Each potable water tank shall be provided with a means of drainage and, if it is equipped with a manhole, over- flow, vent, or a device for measuring depth of water, provision shall be made

#### 21 CFR Ch. I (4–1–17 Edition)

to prevent entrance into the tank of any contaminating substance. No deck or sanitary drain or pipe carrying any nonpotable water or liquid shall be per- mitted to pass through the tank.

  - Tanks and piping shall bear clear marks of identification.

  - There shall be no backflow or cross connection between potable water systems and any other systems. Pipes and fittings conveying potable water to any  fixture, apparatus, or equipment shall be installed in such way that backflow will be prevented. Waste pipes from any part of the pota- ble water system, including treatment devices, discharging to a drain, shall be suitably protected against backflow.

  - Water systems shall be cleaned, disinfected, and flushed whenever the Commissioner of Food and Drugs shall find such treatment necessary to pre- vent the introduction, transmission, or spread of communicable diseases.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.83   Storage   of   water   prior   to treatment.

The following requirements with re- spect to the storage of water on vessels prior to treatment must be  met  in order to obtain approval of treatment facilities under § 1240.90 of this chapter.

- The tank, whether independent or formed by the skin of the ship, deck, tank top, or partitions common with other tanks, shall be free of apparent leakage.

- No sanitary drain shall pass through the tank.

- The tank shall be adequately pro- tected against both the backflow  and discharge into it of bilge or highly con- taminated water.

##### § 1250.84   Water in galleys and medical care spaces.

- Potable water, hot and cold, shall be available in the galley and pantry except that, when potable water stor- age is inadequate, nonpotable  water may be piped to the galley for deck washing and in connection with gar- bage disposal. Any tap discharging nonpotable water which is installed for deck washing purposes shall not  be more than 18 inches above the deck and

#### Food and Drug Administration, HHS	§ 1250.95

shall  be  distinctly  marked  ‘‘For  deck washing only’’.

- In the case  of existing vessels on which heat treated wash water  has been used for the washing of utensils prior to the effective date of the regu- lations in this part, such water may continue to be so used provided con- trols are employed to insure the  heat- ing of all water to at least 170 F before discharge from the heater.

- Potable water, hot and cold, shall be available in medical care spaces for hand-washing and for medical care pur- poses excluding hydrotherapy.

##### § 1250.85 Drinking fountains and cool- ers; ice;  constant temperature  bot- tles.

- Drinking fountains and coolers shall be constructed of impervious, nonoxidizing material, and shall be so designed and constructed as to be eas- ily cleaned. The jet of a drinking foun- tain shall be slanting and the orifice of the jet shall be protected by a guard in such a manner as to prevent contami- nation thereof by droppings from the mouth. The orifice of such  a  jet  shall be located a sufficient distance above the rim of the basin to prevent back- flow.

- Ice shall not be permitted to come in contact with water in coolers or con- stant  temperature  bottles.

- Constant temperature bottles and other containers used for storing or dispensing potable water shall be kept clean at all times and shall be sub- jected to effective bactericidal treat- ment after each occupancy of the space served and at intervals not exceeding one week.

##### § 1250.86   Water for making ice.

Only potable water  shall  be  piped into a freezer for making ice for drink- ing and culinary purposes.

##### § 1250.87   Wash water.

Where systems  installed  on  vessels for wash water, as defined in § 1250.3(n), do  not  comply  with   the   requirements of a potable water  system,  prescribed in § 1250.82, they shall be constructed so as to minimize the possibility  of  the water therein being contaminated. The storage tanks shall comply with the re- quirements   of   § 1250.83,   and   the   dis-

tribution system shall not be cross connected to a system carrying water of a lower sanitary quality. All faucets shall be labeled ‘‘Unfit for drinking’’.

##### § 1250.89   Swimming pools.

- Fill and draw swimming pools shall not be installed or used.

- Swimming pools of the recircula- tion type shall be equipped so as to provide complete circulation, replace- ment, and filtration of the water in the pool every six hours or less. Suitable means of chlorination and, if nec- essary, other treatment of the water shall be provided to maintain the resid- ual chlorine in the pool water at not less than 0.4 part per million and the pH (a measure of the hydrogen ion con- centration) not less than 7.0.

- Flowing-through types of salt water pools shall be so operated that complete circulation and replacement of the water in the pool will be effected every 6 hours or less. The water deliv- ery pipe to the pool shall be inde- pendent of all other pipes and shall originate at a point where maximum flushing of the pump and pipe line is ef- fected after leaving polluted waters.

##### § 1250.90   Toilets and lavatories.

Toilet and lavatory equipment and spaces shall be maintained in a clean condition.

##### § 1250.93   Discharge of wastes.

Vessels operating  on  fresh  water lakes or rivers shall not discharge sew- age, or ballast or bilge water, within such areas adjacent to domestic water intakes as are designated by the Com- missioner of Food and Drugs.

CROSS REFERENCE: For Environmental Pro- tection Agency’s regulations for vessel sani- tary discharges as related to authority under the Federal Water Pollution Control Act, as amended (33 U.S.C. 1314 et seq.), see 40 CFR

part 140.

[40 FR 5624, Feb. 6, 1975, as amended at 48 FR

11432, Mar. 18, 1983]

##### § 1250.95   Insect control.

Vessels shall be maintained free of infestation by flies, mosquitoes, fleas, lice, and other insects known to be vec- tors in the transmission of commu- nicable  diseases,  through  the  use  of

#### § 1250.96

screening, insecticides, and other gen- erally accepted methods of insect con- trol.

##### § 1250.96   Rodent control.

Vessels shall be maintained free of rodent infestation through the use of traps, poisons, and other generally ac- cepted methods of rodent control.

## PARTS 1251–1269 [RESERVED]

PART 1270—HUMAN TISSUE INTENDED FOR TRANSPLANTATION

##### Subpart A—General Provisions

Sec.

1270.1   Scope.

1270.3   Definitions.

##### Subpart B—Donor Screening and Testing

1270.21 Determination of donor suitability for human tissue intended for transplan- tation.

##### Subpart C—Procedures and Records

1270.31     Written procedures.

1270.33    Records, general requirements.

1270.35   Specific records.

##### Subpart D—Inspection of Tissue Establishments

- Inspections.

- Human tissue offered for import.

- Retention, recall, and destruction of human tissue.

AUTHORITY: 42 U.S.C. 216, 243, 264, 271.

SOURCE: 62 FR 40444, July 29, 1997, unless

otherwise noted.

## Subpart A—General Provisions

##### § 1270.1   Scope.

  - The regulations in this part apply to human tissue and to establishments or persons engaged in the recovery, screening, testing, processing, storage, or distribution of human tissue.

  - Regulations in this chapter as they apply to drugs, biologics, devices, or other FDA-regulated commodities do not apply to human tissue, except as specified in this part.

  - Regulations in this chapter do not apply to autologous human tissue.

  - Regulations in this chapter do not apply to hospitals or other clinical fa-

21 CFR Ch. I (4–1–17 Edition)

cilities that receive and store human tissue only for transplantation within the same facility.

##### § 1270.3   Definitions.

- Act for  the purpose of  this part means the Public Health Service Act, section 361 (42 U.S.C. 264).

- Blood component  means any part of a single-donor unit of blood sepa- rated by physical or mechanical means.

- Colloid means a protein or poly- saccharide solution that can be used to increase or maintain osmotic (oncotic) pressure in the intravascular compart- ment such as albumin, dextran, hetastarch; or certain blood compo- nents, such as plasma and platelets.

- Contract services are those func- tions pertaining to the recovery, screening, testing, processing, storage, or distribution of human tissue that another establishment agrees to per- form for a tissue establishment.

- Crystalloid means a balanced salt and/or glucose solution used for elec- trolyte replacement or to increase intravascular volume such as saline, Ringer’s lactate solution, or 5 percent dextrose in water.

- Distribution includes any transfer or shipment of human tissue (including importation or exportation), whether or not such transfer or shipment is en- tirely intrastate and whether or not possession of the tissue is taken.

- Donor means a human being, liv- ing or dead, who is the source of tissue for transplantation.

- Donor medical history interview means a documented dialogue with an individual or individuals who would be knowledgeable of the donor’s relevant medical history and social behavior; such as the donor if living, the next of kin, the nearest available relative, a member of the donor’s household, other individual with an affinity relation- ship, and/or the primary treating phy- sician. The relevant social history in- cludes questions to elicit whether or not the donor met certain descriptions or engaged in certain activities or be- haviors considered to place such an in- dividual at increased risk for HIV and hepatitis.

- Establishment means any facility under one management including all locations, that engages in the recovery,

#### Food and Drug Administration, HHS	§ 1270.3

screening, testing, processing, storage, or distribution of human tissue in- tended for transplantation.

- Human tissue, for  the  purpose  of this  part  means  any   tissue   derived from a human body and recovered be- fore May 25, 2005, which:

  - Is intended for transplantation to another human for the diagnosis, cure, mitigation, treatment, or prevention of any condition or disease;

  - Is recovered, processed, stored, or distributed by methods that do not change tissue function or characteris- tics;

  - Is not currently regulated as a human drug, biological product, or medical device;

  - Excludes kidney, liver, heart, lung, pancreas, or any other vascularized human organ; and

  - Excludes semen or other reproduc- tive tissue, human milk, and bone mar- row.

- Importer of record means the per- son, establishment or their representa- tive responsible for making entry of imported goods in accordance with all laws affecting such importation.

- Legislative consent means relating to any of the laws of the various States that allow the medical examiner or coroner to procure corneal tissue in the absence of consent of the donor’s next- of-kin.

- Person includes an individual, partnership, corporation, association, or other legal entity.

- Physical assessment means a lim- ited autopsy or recent antemortem or postmortem physical examination of the donor to assess for any signs of HIV and hepatitis infection or signs sugges- tive of any risk factor for such infec- tions.

- Plasma dilution means a decrease in the concentration of the donor’s plasma proteins and circulating anti- gens or antibodies resulting from the transfusion of blood or blood compo- nents and/or infusion of fluids.

- Processing means any activity per- formed on tissue, other than tissue re- covery, including preparation, preser- vation for storage, and/or removal from storage to assure the quality and/or sterility of human tissue. Processing includes steps to inactivate and remove adventitious  agents.

- Quarantine means the identifica- tion of human tissue as not suitable for transplantation, including human tis- sue that has not yet been characterized as being suitable for transplantation. Quarantine includes the storage of such tissue in an area clearly identified for such use, or other procedures, such as automated designation, for preven- tion of release of such tissue for trans- plantation.

- Reconstituted blood means the extracorporeal resuspension of a blood unit labeled as ‘‘Red Blood Cells’’  by the addition of colloids and/or crystalloids to produce a hematocrit in the normal range.

- Recovery means the obtaining from a donor of tissue that is intended for use in human transplantation.

- Relevant medical records means a collection of documents including a donor medical history interview, a physical assessment of the donor, lab- oratory test results, medical records, existing coroner and autopsy reports, or information obtained from any source or records which may pertain to donor suitability regarding high risk behaviors, clinical signs and symptoms for HIV and hepatitis, and treatments related to medical conditions sugges- tive of such risk.

- Responsible person means a person who is authorized to perform des- ignated functions for which he or she is trained and qualified.

- Storage means holding tissue.

- Summary of records means a con- densed version of the required testing and screening records that contains the identity of the testing laboratory, the listing and interpretation of all re- quired infectious disease tests, and a listing of the documents reviewed as part of the relevant medical records, and the name of the person or estab- lishment determining the suitability of the human tissue for transplantation.

- Vascularized means containing the original blood vessels which are in- tended to carry blood after transplan- tation.

[62 FR 40444, July 29, 1997, as amended at 69

FR 68680, Nov. 24, 2004]

#### § 1270.21

Subpart B—Donor Screening and Testing

##### § 1270.21 Determination of donor suit- ability for human tissue intended for transplantation.

- Donor specimens shall be tested for the following communicable vi- ruses, using Food and Drug Adminis- tration (FDA) licensed donor screening tests in accordance with manufactur- ers’  instructions:

  - Human immunodeficiency virus, Type 1  (e.g.,  FDA licensed  screening test for anti-HIV–1);

  - Human immunodeficiency virus, Type 2  (e.g.,  FDA licensed  screening test for anti-HIV–2);

  - Hepatitis B (e.g., FDA licensed screening test for HBsAg); and

  - Hepatitis C (e.g., FDA licensed screening test for anti-HCV).

- In the case of a neonate, the mother’s specimen is acceptable for testing.

- Such infectious disease testing shall be performed by a laboratory cer- tified under the Clinical Laboratories Improvement Amendments of 1988 (CLIA).

- Human tissue shall be accom- panied by records indicating that the donor’s specimen has been tested and found negative using FDA licensed screening tests for HIV–1, HIV–2, hepa- titis B, and hepatitis C. FDA licensed screening tests labeled for cadaveric specimens must be used when avail- able.

- Human tissue for transplantation shall be accompanied by a summary of records or copies of the original records of the donor’s relevant medical records as defined in § 1270.3(t) which docu- ments freedom  from  risk factors for and clinical evidence of hepatitis B, hepatitis C, or HIV  infection.  There shall be a responsible person des- ignated and identified in the original record and summary of records as hav- ing made the determination that the human tissue is suitable for transplan- tation.

- Determination by the responsible person that a donor of human tissue in- tended for transplantation is suitable shall include ascertainment of the do- nor’s identity, and accurately recorded relevant medical records (as defined in

#### 21 CFR Ch. I (4–1–17 Edition)

§ 1270.3(t)) which documents freedom from risk factors for and clinical evi- dence of hepatitis B, hepatitis C, and HIV   infection.

- For corneal tissue procured under legislative consent where a donor med- ical history screening interview has not occurred, a physical assessment of the donor is required and other avail- able information shall be reviewed. The corneal tissue shall be accompanied by the summary of records documenting that the corneal tissue was determined to be suitable for transplantation in the absence of the donor medical his- tory interview. Corneal tissue procured under legislative consent shall be docu- mented as such in the summary of records.

- Human tissue shall be determined to be not suitable for transplantation if from:

  - A donor whose specimen has test- ed repeatedly reactive on a screening test for HIV, hepatitis  B,  or  hepatitis C;

  - A donor where blood loss is known or suspected to have occurred and transfusion/infusion of more than 2,000 milliliters (mL) of blood (i.e., whole blood, reconstituted blood, or red blood cells), or colloids within 48 hours; or more than 2,000 mL of crystalloids within 1 hour; or any combination thereof prior to the collection  of  a blood specimen from the tissue donor for testing, unless:

    - A pretransfusion or preinfusion blood specimen from the tissue donor is available for infectious disease  testing; or

    - An algorithm is  utilized that evaluates the volumes administered in the 48 hours prior to collecting the blood specimen from the  tissue  donor to ensure that there has not been plas- ma dilution sufficient to affect test re- sults; or

  - A donor who is 12 years of age or less and has been transfused or infused at all, unless:

    - A pretransfusion or preinfusion blood specimen from the tissue donor is available for infectious disease  testing; or

    - An algorithm  is utilized  that evaluates the volumes administered in the 48 hours prior to collecting the blood  specimen  from  the  tissue  donor

#### Food and Drug Administration, HHS	§ 1270.33

to ensure that there has not been plas- ma dilution sufficient to affect test re- sults.

## Subpart C—Procedures and Records

##### § 1270.31   Written procedures.

      - There shall be written procedures prepared and followed for all signifi- cant steps in the infectious disease testing process under § 1270.21 which shall conform to the manufacturers’ in- structions for use contained in the package inserts for the required tests. These procedures shall be readily avail- able to the personnel in the area where the procedures are performed unless impractical. Any deviation from the written procedures shall be recorded and  justified.

      - There shall be written procedures prepared and followed  for  all  signifi- cant steps for obtaining, reviewing, and assessing the relevant medical  records of the donor as provided in § 1270.21. Such procedures shall be readily avail- able to personnel who may perform the procedures. Any deviation from the written procedures shall  be  recorded and   justified.

      - There shall be written procedures prepared and followed for designating and identifying quarantined tissue.

      - There shall be written procedures prepared, validated, and followed for prevention of infectious disease  con- tamination or cross-contamination by tissue during processing.

      - In conformity with this section, any facility may use current standard written procedures such as those in a technical manual prepared by another organization, provided the procedures are consistent with and at least as stringent as the requirements of this part.

##### § 1270.33   Records,	general	require- ments.

- Records shall be maintained con- currently with the performance of each significant step required in this part in the performance of infectious disease screening and testing of donors of human tissue. All records shall be ac- curate, indelible, and legible. The records shall identify the person per- forming the work, the dates of the var-

ious entries, and shall be as detailed as necessary to provide a complete his- tory of the work performed and to re- late the records to the particular tissue involved.

- All human tissue shall be quar- antined until the following criteria for donor suitability are satisfied:

  - All infectious disease  testing under § 1270.21 has been completed, re- viewed by the responsible person, and found to be negative; and

  - Donor screening has been com- pleted, reviewed by the responsible per- son, and determined to assure freedom from risk factors for and clinical evi- dence of HIV infection, hepatitis B, and hepatitis  C.

- All human tissue processed or shipped prior to determination of donor suitability must be under quarantine, accompanied by records assuring iden- tification of the donor and indicating that the tissue has not been deter- mined to be suitable for transplan- tation.

- All  human  tissue  determined  to be suitable for transplantation must be accompanied by a summary  of  records, or copies of such original records, docu- menting  that  all  infectious  disease testing and screening  under  § 1270.21 has been completed, reviewed by the responsible person, and found to be negative, and that the tissue has been determined to be suitable for trans- plantation.

- Human tissue shall be quar- antined until the tissue is either deter- mined to be suitable for transplan- tation or appropriate disposition is ac- complished.

- All persons or establishments that generate records used in determining the suitability of the donor shall retain such records and make them available for authorized inspection or upon re- quest by FDA. The person(s) or estab- lishment(s) making the determination regarding the suitability of the donor shall retain all records,  or  true  copies of such records required under § 1270.21, including all testing and screening records, and shall make them available for authorized inspection or upon re- quest from FDA. Records that can be retrieved from another location by electronic means meet the require- ments of this paragraph.

§ 1270.35

- Records required under this part may be retained electronically, or as original paper records, or as true copies such as photocopies, microfiche, or microfilm, in which case suitable read- er and photocopying equipment shall be readily  available.

- Records shall be retained at least 10 years beyond the date of transplan- tation if known, distribution, disposi- tion, or expiration, of the tissue, whichever is latest.

[62 FR 40444, July 29, 1997, as amended at 63

FR 16685, Apr. 6, 1998]

##### § 1270.35   Specific records.

Records shall be maintained that in- clude, but are not limited to:

- Documentation of results and in- terpretation of all required infectious disease tests;

- Information on the identity and relevant medical records of  the  donor, as required by § 1270.21(e) in English or, if in another language translated to English and accompanied by a state- ment of authenticity by the translator which specifically identifies the trans- lated document;

- Documentation of the receipt and/ or distribution of human tissue; and

- Documentation of the destruction or other disposition of human tissue.

## Subpart D—Inspection of Tissue Establishments

##### § 1270.41   Inspections.

- An establishment covered by these regulations in this part, includ- ing any location performing contract services, shall permit an authorized in- spector of the Food and Drug Adminis- tration (FDA) to make at any reason- able time and in a reasonable manner such inspection of the establishment, its facilities, equipment, processes, products, and records as may be nec- essary to determine compliance with the provisions of this part. Such in- spections may be made with or without notice and will ordinarily be made dur- ing regular business hours.

- The frequency of inspection will be at the agency’s discretion.

- The inspector shall call upon a re- sponsible person of the establishment and may question the personnel of the

21 CFR Ch. I (4–1–17 Edition)

establishment as the inspector deems necessary.

- The inspector may review and copy any records required to be kept pursuant to part 1270.

- The public disclosure of records containing the name or other positive identification of donors or recipients of human tissue will be handled in accord- ance with FDA’s procedures on disclo- sure of information as set forth in 21 CFR part 20 of this chapter.

##### § 1270.42   Human tissue offered for im- port.

- When human tissue is offered for entry, the importer of record must no- tify the director of the district of the Food and Drug Administration having jurisdiction over the port of entry through which the tissue is imported or offered for import, or such officer of the district as the director may des- ignate to act in his or her behalf in ad- ministering and enforcing this part.

- Human tissue offered for import must be quarantined until the human tissue is released by FDA.

##### § 1270.43   Retention,   recall,   and   de- struction of human tissue.

- Upon a finding that human tissue may be in violation of the regulations in this part, an authorized Food and Drug Administration (FDA) represent- ative may:

  - Serve upon the person who dis- tributed the tissue a written order that the tissue be recalled and/or destroyed, as appropriate, and upon persons in possession of the tissue that the tissue shall be retained until it is recalled by the distributor, destroyed, or  disposed of as agreed by FDA, or  the  safety  of the tissue is confirmed; and/or

  - Take possession of and/or destroy the violative tissue.

- The written order will ordinarily provide that the human tissue be re- called and/or destroyed within 5 work- ing days from the date of receipt of the order and will state with particularity the facts that justify the order.

- After receipt of an order under this part, the person in possession of the human  tissue shall  not  distribute or dispose of the tissue in any manner except to recall and/or destroy the tis- sue  consistent  with  the  provisions  of

#### Food and Drug Administration, HHS	Pt. 1271

the order, under the supervision of an authorized  official  of  FDA.

- In lieu of paragraphs (b) and (c) of this section, other arrangements for assuring the proper disposition of the tissue may be agreed upon by the per- son receiving the written order and an authorized official of FDA. Such ar- rangements may  include  providing FDA with records or other written in- formation that adequately assure that the tissue has been recovered, screened, tested, processed, stored, and distrib- uted in conformance with part 1270.

- Within 5 working days of receipt of a written order for retention, recall, and/or destruction of tissue (or within 5 working days of the agency’s posses- sion of such tissue), the recipient of the written order or prior possessor of such tissue shall request a hearing on the matter in accordance with part 16 of this chapter. The order for destruction will be held in abeyance pending reso- lution of the hearing request.

## PART 1271—HUMAN CELLS, TIS- SUES, AND CELLULAR AND TIS- SUE-BASED PRODUCTS

##### Subpart A—General Provisions

1271.27 Will FDA assign me a registration number?

1271.37 Will establishment registrations and HCT/P listings be available  for  inspec- tion, and how do I  request  information on  registrations  and  listings?

##### Subpart C—Donor Eligibility

1271.45 What requirements does this subpart contain?

1271.47  What  procedures  must  I  establish and   maintain?

1271.50 How do I determine whether a donor is  eligible?

1271.55 What records must accompany an HCT/P after the donor-eligibility deter- mination is complete; and what records must  I  maintain?

1271.60 What quarantine and other require- ments apply before the donor-eligibility determination  is  complete?

1271.65 How do I store an HCT/P from a donor determined to be ineligible,  and what uses of the HCT/P are not prohib- ited?

1271.75   How do I screen a donor?

1271.80 What are the  general  requirements for  donor  testing?

1271.85 What donor testing is required for different  types  of  cells  and  tissues?

1271.90 Are there exceptions from the re- quirement of determining  donor  eligi- bility, and what labeling requirements apply?

Sec.

1271.1 What are the purpose and  scope  of this  part?

1271.3 How  does  FDA  define  important terms  in  this  part?

1271.10 Are my HCT/P’s regulated solely under section 361 of the PHS Act and the regulations in this part, and if so what must I do?

1271.15 Are there any exceptions from the requirements  of  this  part?

- If my HCT/P’s do not meet the cri- teria in § 1271.10, and I do not qualify for any of the exceptions in § 1271.15, what regulations   apply?

##### Subpart B—Procedures for Registration and Listing

- When do I register, submit an HCT/P list, and submit updates?

- How do I register and submit an HCT/P list?

- How is a waiver from the electronic format requirements requested?

- What information is required for es- tablishment registration and HCT/P list- ing?

- When must I amend my establish- ment registration?

##### Subpart D—Current Good Tissue Practice

1271.145 Prevention of  the introduction, transmission, or spread of communicable diseases.

1271.150   Current   good   tissue   practice   re- quirements.

1271.155    Exemptions and alternatives.

1271.160    Establishment  and  maintenance  of a  quality  program.

1271.170   Personnel.

1271.180   Procedures.

1271.190   Facilities.

1271.195    Environmental   control   and   moni- toring.

1271.200    Equipment.

1271.210  Supplies and reagents.

1271.215   Recovery.

1271.220   Processing and process controls.

1271.225    Process changes.

1271.230   Process validation.

1271.250    Labeling controls.

1271.260   Storage.

1271.265    Receipt,   predistribution   shipment, and  distribution  of  an  HCT/P.

1271.270   Records.

1271.290  Tracking.

- Complaint file.

§ 1271.1

##### Subpart E—Additional Requirements for Establishments Described in § 1271.10

1271.330   Applicability.

1271.350    Reporting.

1271.370    Labeling.

##### Subpart F—Inspection and Enforcement of Establishments Described in § 1271.10

1271.390   Applicability.

1271.400    Inspections.

1271.420   HCT/Ps offered for import.

1271.440    Orders of retention, recall, destruc- tion,  and  cessation  of  manufacturing.

AUTHORITY: 42 U.S.C. 216, 243, 263a, 264, 271.

SOURCE:  66  FR  5466,  Jan.  19,  2001,  unless

otherwise noted.

## Subpart A—General Provisions

##### § 1271.1   What   are   the   purpose   and scope of this part?

- Purpose. The purpose of this part, in conjunction with §§ 207.9(a)(5), 210.1(c), 210.2, 807.20(d), and 820.1(a) of this chapter, is to create an electronic registration and listing system for es- tablishments that manufacture human cells, tissues, and cellular and tissue- based products (HCT/P’s) and to estab- lish donor-eligibility, current good tis- sue practice, and other procedures to prevent the  introduction, transmission, and spread of communicable diseases by HCT/P’s.

- Scope. (1) If you are an establish- ment that manufactures HCT/P’s that are regulated solely under the  author- ity of  section  361  of the Public  Health Service Act (the PHS Act), this part re- quires you to register and list your HCT/P’s with the Food and Drug Ad- ministration’s (FDA’s) Center for Bio- logics Evaluation and Research and to comply with the other requirements contained in this part, whether or  not the HCT/P enters into interstate com- merce. Those HCT/P’s that are regu- lated solely under the authority of sec- tion 361 of the PHS Act are described in

§ 1271.10.

(2) If you are an establishment that manufactures HCT/P’s that are regu- lated as drugs, devices and/or  biologi- cal products under section 351 of the PHS  Act  and/or  the  Federal  Food, Drug, and Cosmetic Act, §§ 207.9(a)(5) and 807.20(d) of this chapter require you to   register   and   list   your   HCT/P’s   fol-

21 CFR Ch. I (4–1–17 Edition)

lowing the procedures in part 207 (if a drug and/or biological product) of this chapter or part 807 (if a device) of this chapter. Sections 210.1(c), 210.2, 211.1(b), and 820.1(a) of this chapter require you to comply with the donor-eligibility procedures in subpart C of this part and the current good tissue practice procedures in subpart D of this part, in addition to all  other applicable regula- tions.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 69

FR 29829,  May 25,  2004;  81  FR 60223,  Aug. 31,

2016]

##### § 1271.3   How  does  FDA  define  impor- tant terms in this part?

The following definitions apply only to this part:

  - Autologous use means the implan- tation, transplantation, infusion, or transfer of human cells or tissue back into the individual from whom the cells or tissue were recovered.

  - Establishment means a place of business under one management, at one general physical location, that en- gages in the manufacture of human cells, tissues, and cellular and tissue- based products. ‘‘Establishment’’ in- cludes:

    - Any individual, partnership, cor- poration, association, or other legal en- tity engaged in the manufacture of human cells, tissues, and cellular and tissue-based products; and

    - Facilities that engage in contract manufacturing services for a manufac- turer of human cells, tissues, and cel- lular and tissue-based products.

  - Homologous use means the repair, reconstruction, replacement, or sup- plementation of a recipient’s cells or tissues with an HCT/P that performs the same basic function or functions in the recipient as in the donor.

  - Human cells, tissues, or cellular or tissue-based products (HCT/Ps) means ar- ticles containing or consisting of human cells or tissues that are in- tended for implantation, transplan- tation, infusion, or transfer into a human recipient. Examples of HCT/Ps include, but are not limited to, bone, ligament, skin, dura  mater,  heart valve, cornea, hematopoietic stem/pro- genitor cells derived from peripheral and	cord	blood,	manipulated autologous    chondrocytes,    epithelial

#### Food and Drug Administration, HHS	§ 1271.3

cells on a synthetic matrix, and semen or other reproductive tissue. The fol- lowing articles are not considered HCT/ Ps:

    - Vascularized human organs for transplantation;

    - Whole blood or blood components or blood derivative products subject to listing under parts 607 and 207 of this chapter,  respectively;

    - Secreted or extracted human products, such as milk, collagen, and cell factors; except that semen is con- sidered an HCT/P;

    - Minimally manipulated bone mar- row for homologous use and not com- bined with another article (except for water, crystalloids, or a sterilizing, preserving, or storage agent, if the ad- dition of the agent does not raise new clinical safety concerns with respect to the bone marrow);

    - Ancillary products used in the manufacture of HCT/P;

    - Cells, tissues, and organs derived from animals other than humans; and

    - In vitro diagnostic products as de- fined in § 809.3(a) of this chapter.

    - Blood vessels recovered with an organ, as defined in 42 CFR  121.2,  that are intended for use in organ trans- plantation  and  labeled  ‘‘For  use  in organ   transplantation   only.’’

  - Manufacture means, but is not lim- ited to, any or all steps in the recov- ery, processing, storage, labeling, pack- aging, or distribution of any human cell or tissue, and the screening or testing of the cell or tissue donor.

  - Minimal manipulation means:

- For structural tissue, processing that does not alter the original rel- evant characteristics of the tissue re- lating to the tissue’s utility for recon- struction, repair, or replacement; and

- For cells or nonstructural tissues, processing that does not alter the rel- evant biological characteristics of cells or tissues.

  - Transfer means the placement of human reproductive cells or  tissues into a human recipient.

  - Biohazard legend appears on the label as follows and is used to mark HCT/Ps that present a known or sus- pected relevant communicable disease risk.

  - Blood component means a product containing a part of human blood sepa- rated by physical or mechanical means.

  - Colloid means:

    - A protein or polysaccharide solu- tion, such  as  albumin, dextran, or hetastarch, that can be used to in- crease or maintain osmotic (oncotic) pressure in the intravascular compart- ment; or

    - Blood components such as plasma and  platelets.

  - Crystalloid means an isotonic salt and/or glucose solution used for elec- trolyte replacement or to increase intravascular volume, such as saline solution, Ringer’s lactate solution, or 5 percent dextrose in water.

  - Directed reproductive donor means a donor of reproductive cells or tissue (including semen, oocytes,  and  em- bryos to  which  the  donor  contributed the spermatozoa or oocyte) to a spe- cific recipient, and who knows and is known by the recipient  before  dona- tion. The term directed reproductive donor does not include a sexually inti- mate partner under § 1271.90.

  - Donor means a person, living or dead, who is the source of cells or tis- sue for an HCT/P.

  - Donor medical history interview means a documented dialog about the donor’s medical history and relevant social behavior, including activities, behaviors, and descriptions considered to increase the donor’s relevant com- municable disease risk:

- With the donor, if the donor is liv- ing and able to participate in the inter- view, or

- If not, with an individual or indi- viduals able to provide the information sought in the interview (e.g., the do- nor’s next-of-kin, the nearest available relative,   a   member   of   the   donor’s

#### § 1271.3

household, an individual with an affin- ity relationship, and/or the primary treating physician).

  - Physical assessment of a cadaveric donor means a limited autopsy or re- cent antemortem or postmortem phys- ical examination of the donor to assess for signs of a relevant communicable disease and for signs suggestive of any risk factor for a  relevant  commu- nicable disease.

  - Plasma dilution means a decrease in the concentration of the donor’s plasma proteins and circulating anti- gens or antibodies resulting from the transfusion of blood or blood compo- nents and/or infusion of fluids.

  - Quarantine means the storage or identification of an HCT/P, to prevent improper release, in a physically sepa- rate area clearly identified for such use, or through use of other proce- dures, such as automated designation.

  - Relevant communicable disease agent or disease means:

(1)(i) For all human cells and tissues, a communicable disease or disease agent listed as follows:

- Human immunodeficiency virus, types 1 and 2;

- Hepatitis B virus;

- Hepatitis C virus;

- Human transmissible spongiform encephalopathy, including Creutzfeldt-

Jakob disease; and

- Treponema pallidum.

- For viable, leukocyte-rich cells and tissues, a cell-associated disease agent or disease listed as follows:

- Human T-lymphotropic virus, type I; and

- Human T-lymphotropic virus, type II.

- For reproductive cells or tissues, a disease agent or disease of the genito- urinary tract listed as follows:

- Chlamydia trachomatis; and

- Neisseria gonorrhea.

  - A disease agent or disease not list- ed in paragraph (r)(1) of this section:

    - For which there may be a risk of transmission by an HCT/P, either to the recipient of the HCT/P or to those people who may handle or otherwise come in contact with it, such as med- ical personnel, because the disease agent or disease:

      - Is potentially transmissible by an HCT/P and

#### 21 CFR Ch. I (4–1–17 Edition)

      - Either of the following applies:

        - The disease agent or disease has sufficient incidence and/or prevalence to affect the potential donor popu- lation,  or

        - The disease agent or disease may have been released accidentally or in- tentionally in a manner that  could place potential donors at risk of infec- tion;

    - That could be fatal or life-threat- ening, could result in permanent im- pairment of a body function or perma- nent damage to body structure, or could necessitate medical or surgical intervention to preclude permanent impairment of body function or perma- nent damage to a body structure; and

    - For which appropriate screening measures have been developed  and/or an appropriate screening test for donor specimens has been licensed, approved, or cleared for such use by FDA and  is available.

  - Relevant medical records means a collection of documents that includes a current donor medical history inter- view; a current report of the physical assessment of a cadaveric donor or the physical examination of a living donor; and, if available, the following:

    - Laboratory test results  (other than results of testing for relevant communicable disease agents required under this subpart);

    - Medical records;

    - Coroner and autopsy reports; and

    - Records or other information re- ceived from any source pertaining to risk factors  for relevant communicable disease (e.g., social behavior, clinical signs and symptoms of relevant com- municable disease, and treatments re- lated to medical conditions  suggestive of risk for relevant communicable dis- ease).

  - Responsible person means a person who is authorized to perform des- ignated functions for which he or she is trained and qualified.

  - Urgent medical need means that no comparable HCT/P is available and the recipient is likely to suffer death or se- rious morbidity without the HCT/P.

  - Act means the Federal Food, Drug, and Cosmetic Act.

  - PHS Act means the Public Health Service  Act.

#### Food and Drug Administration, HHS	§ 1271.3

- FDA means the Food and Drug Administration.

- Adverse reaction means a noxious and unintended response to any HCT/P for which there is a reasonable possi- bility that the HCT/P caused the re- sponse.

- Available for distribution means that the HCT/P has been determined to meet all release criteria.

(aa) Complaint means any written, oral, or electronic communication about a distributed HCT/P that alleges:

- That an HCT/P has transmitted or may have transmitted a communicable disease to the recipient of  the  HCT/P; or

- Any other problem with an HCT/P relating to the potential for trans- mission of communicable disease, such as the failure to comply with current good tissue practice.

(bb) Distribution means any convey- ance or shipment (including importa- tion and exportation) of an HCT/P that has been determined to meet all re- lease criteria, whether or not such con- veyance or shipment is entirely intra- state. If an entity does not take  phys- ical possession of an HCT/P, the entity is not considered a distributor.

(cc) Establish and maintain means de- fine, document (in writing or electroni- cally), and implement; then follow, re- view, and, as needed, revise on an ongo- ing basis.

(dd) HCT/P deviation means an event:

- That represents a deviation from applicable regulations in this part or from applicable standards or estab- lished specifications that relate to the prevention of communicable disease transmission or HCT/P contamination; or

- That is an unexpected or unfore- seeable event that may relate to the transmission or potential transmission of a communicable disease or may lead to  HCT/P  contamination.

(ee) Importer of record means the per- son, establishment, or its representa- tive responsible for making entry of imported goods in accordance with all laws affecting such importation.

(ff) Processing means any activity performed on an HCT/P, other than re- covery, donor screening, donor testing, storage, labeling, packaging, or dis- tribution,  such  as  testing  for  micro-

organisms, preparation, sterilization, steps to inactivate or remove adven- titious agents, preservation for  stor- age, and removal from storage.

(gg) Quality audit means a docu- mented, independent inspection and re- view of an establishment’s activities related to core CGTP requirements. The purpose of a quality audit is to verify, by examination and evaluation of objective evidence, the degree of compliance with those aspects of the quality program under review.

(hh) Quality program means an orga- nization’s comprehensive system for manufacturing and tracking HCT/Ps in accordance with this part. A quality program is designed to prevent, detect, and correct deficiencies that may lead to circumstances that increase the risk of introduction, transmission, or spread of communicable diseases.

- Recovery means obtaining from a human donor cells or tissues that are intended for use in human implanta- tion, transplantation, infusion, or transfer.

(jj) Storage means holding HCT/Ps for future processing and/or distribution.

(kk) Validation means confirmation by examination and provision of objec- tive evidence that particular require- ments can consistently be fulfilled. Validation of a process, or process vali- dation, means establishing by objective evidence that a process consistently produces a result or HCT/P meeting its predetermined  specifications.

(ll) Verification means  confirmation by examination and provision of objec- tive evidence that specified require- ments have been fulfilled.

(mm) Importer means a company or individual in the United States that is the owner, consignee, or recipient, at the time of entry, of the foreign estab- lishment’s HCT/P that is imported into the United States.

(nn) United States agent means a per- son residing or maintaining a place of business in the United States whom a foreign establishment designates as its agent. This definition excludes mail- boxes, answering machines or services, or  other  places  where  an  individual

#### § 1271.10

acting  as  the  foreign  establishment’s agent is not physically present.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 68

FR  3826,  Jan.  27,  2004;  69  FR  29829,  May  25,

2004; 69 FR 68680, Nov. 24, 2004; 81 FR 60223,

Aug. 31, 2016]

##### § 1271.10 Are my HCT/P’s regulated solely under section 361 of the PHS Act and the regulations in this part, and if so what must I do?

  - An HCT/P is regulated solely under section 361 of the PHS Act and the regulations in this part if it meets all of the following criteria:

- The HCT/P is minimally manipu- lated;

- The HCT/P is intended for homolo- gous use only, as reflected by the label- ing, advertising, or other indications of the manufacturer’s objective intent;

- The manufacture of the HCT/P does not involve the combination of the cells or tissues with another arti- cle, except for water, crystalloids, or a sterilizing, preserving, or  storage agent, provided that the addition of water, crystalloids, or the sterilizing, preserving, or storage agent does not raise new clinical safety concerns with respect to the HCT/P; and

- Either:

  - The HCT/P does not have a sys- temic effect and is not dependent upon the metabolic activity of living cells for its primary function; or

  - The HCT/P has a systemic effect or is dependent upon the metabolic ac- tivity of living cells for its primary function, and:

- Is for autologous use;

- Is for allogeneic use in a first-de- gree or second-degree blood relative; or

- Is for reproductive use.

- If you are a domestic or foreign establishment that manufactures an HCT/P described in paragraph (a) of this section:

  - You must register with FDA;

  - You must submit to FDA a list of each HCT/P manufactured; and

  - You must comply with the other requirements contained in this part.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 69

FR 68681, Nov. 24, 2004]

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 1271.15 Are  there  any  exceptions from the requirements of this part?

- You are not required to comply with the requirements of this part if you are an establishment that uses HCT/P’s solely for nonclinical sci- entific or educational purposes.

- You are not required to comply with the requirements of this part if you are an establishment that removes HCT/P’s from an individual and im- plants such HCT/P’s into the same indi- vidual during the same surgical proce- dure.

- You are not required to comply with the requirements of this part if you are a carrier who accepts, receives, carries, or delivers HCT/P’s in the usual course of business as a carrier.

- You are not required to comply with the requirements of this part if you are an establishment that does not recover, screen, test, process, label, package, or distribute, but only re- ceives or stores HCT/P’s solely for im- plantation, transplantation,  infusion, or transfer within your facility.

- You are not required to comply with the requirements of this part if you are an establishment that only re- covers reproductive cells or tissue and immediately transfers them into a sex- ually intimate partner of the cell or tissue donor.

- You are not required to register or list your HCT/P’s independently,  but you must comply with all other appli- cable requirements in this part, if  you are an individual under contract, agreement, or  other  arrangement  with a registered establishment and engaged solely in recovering cells or tissues and sending the recovered cells  or  tissues to the registered establishment.

##### § 1271.20 If my HCT/P’s do  not  meet the criteria in § 1271.10, and I  do not qualify for any of the excep- tions in § 1271.15, what regulations apply?

If you are an establishment that manufactures an HCT/P that does not meet the criteria set out in § 1271.10(a), and you do not qualify for any of the exceptions in § 1271.15, your HCT/P will be regulated as a drug, device, and/or biological product under the act and/or section 351 of the PHS Act,  and appli- cable  regulations  in  title  21,  chapter  I.

#### Food and Drug Administration, HHS	§ 1271.25

Applicable regulations include, but are not  limited  to,  §§ 207.9(a)(5),  210.1(c),

210.2,  211.1(b),  807.20(d),  and  820.1(a)  of

this chapter, which require you to fol- low the procedures in subparts C and D of this part.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 81

FR 60223, Aug. 31, 2016]

## Subpart B—Procedures for Registration and Listing

##### § 1271.21   When do I register, submit an HCT/P list, and submit updates?

  - You must register and submit a list of every HCT/P that your establish- ment manufactures within 5 days after beginning operations or within 30 days of the effective date of this regulation, whichever is later.

  - You must update your establish- ment registration annually in  Decem- ber, except as required by § 1271.26. You may accomplish your annual  registra- tion in conjunction with updating your HCT/P list under paragraph (c) of this section.

(c)(i)    If    no    change    described    in

§ 1271.25(c) has occurred since you pre- viously  submitted  an  HCT/P  list,  you are  not  required  to  update  your  listing.

(ii)	If	a	change	described	in

§ 1271.25(c) has occurred, you must up- date your HCT/P listing with the new information:

- At the time of the change, or

- Each June or December, which- ever month occurs first after  the change.

[69 FR 68681, Nov. 24, 2004]

##### § 1271.22   How do I register and submit an HCT/P list?

- You must use the electronic reg- istration and listing in ac- cordance with § 1271.25 for:

  - Establishment registration,

  - HCT/P listings, and

  - Updates of registration and HCT/P listing.

- FDA will periodically issue guid- ance on recommended procedures for providing registration and listing in- formation in electronic format (for ex- ample, method of transmission, media, file formats, preparation, and organiza- tion of files).

- You must provide the information under paragraph (a) of this section in accordance with part 11 of this chapter, except      for      the      requirements      in

§ 11.10(b),   (c),   and   (e)   and   the   cor- responding requirements in § 11.30.

[81 FR 60223, Aug. 31, 2016]

##### § 1271.23 How is a waiver from the electronic format requirements re- quested?

- You may request a waiver from the requirement in § 1271.22 that infor- mation must be provided to FDA in electronic format. Submission of a re- quest for waiver does not excuse timely compliance with the registration and listing requirements.  FDA  will  grant  a waiver request if FDA determines that the use of electronic means for submis- sion of registration and listing infor- mation is not reasonable for the reg- istrant  making  the  waiver  request.

- Waiver requests under this sec- tion must be submitted in writing and must include the specific reasons why electronic submission is not reasonable for the  registrant  and a U.S.  telephone number and mailing address where FDA can contact the  registrant.  Waiv- er requests may be sent to the Center for Biologics Evaluation and Research (CBER), Document Control Center (see addresses in § 600.2 of this chapter).

- If FDA grants the waiver request, FDA may limit its duration and will specify terms of the waiver and provide information on how to submit estab- lishment registration, listings, other information, and updates, as applica- ble.

[81 FR 60224, Aug. 31, 2016]

##### § 1271.25 What information is required for establishment registration and HCT/P listing?

- Your establishment registration must include:

- The legal name(s) of the establish- ment;

- Each physical  location,  including the street address, telephone number, email address, and the postal service ZIP code of the establishment;

- The name, address, telephone number, email address, and title of the reporting  official;

§ 1271.26

- A dated signature by the report- ing official affirming that all informa- tion contained in the establishment registration and HCT/P listing form is true and accurate, to the best of his or her knowledge.

- Each foreign establishment must also submit the name, address, tele- phone number, and email address of each importer that is known to the es- tablishment, and the name of each per- son who imports or offers for import such HCT/P to the United States for purposes of importation; and

- Each foreign establishment must also submit the name, address, tele- phone number, and email address of its United States agent.

  - The United States agent must re- side or maintain a place of business in the United States.

  - Upon request from FDA, the United States agent must  assist  FDA in communications with the foreign es- tablishment, respond to questions con- cerning the foreign establishment’s products that are imported or offered for import into the United States, and assist FDA in scheduling inspections of the foreign establishment. If the Agen- cy is unable to contact the foreign es- tablishment directly or expeditiously, FDA may provide information or docu- ments to the United States agent, and such an action is equivalent to pro- viding the same information or docu- ments to the foreign establishment.

  - The foreign establishment or the United States agent must  report changes in the United  States agent’s name, address, telephone number, or email address to FDA within 30 cal- endar days of the change.

- Your HCT/P listing must include all HCT/P’s (including the established name and the proprietary name) that you recover, process, store, label, pack- age, distribute, or for which you per- form donor screening or testing. You must also state whether each HCT/P meets the criteria set out in § 1271.10.

- Your HCT/P listing update must include:

  - A list of each HCT/P that you have begun recovering, processing, storing, labeling, packaging, distrib- uting, or for which you have begun donor screening or testing, that has not been included in any list previously

#### 21 CFR Ch. I (4–1–17 Edition)

submitted. You must provide all of the information required by § 1271.25(b) for each new HCT/P.

  - A list of each HCT/P  formerly  list- ed in accordance with § 1271.21(a) for which you have discontinued recovery, processing,  storage,   labeling,   pack- aging, distribution,  or donor   screening or testing, including for each HCT/P so listed,  the  identity   by   established name and proprietary  name,  and  the date of discontinuance. We request  but do not require that you include the reason for discontinuance with this in- formation.

  - A list of each HCT/P for which a notice of discontinuance was submitted under paragraph (c)(2) of this section and for which you have resumed recov- ery, processing, storage, labeling, pack- aging, distribution, or donor   screening or testing, including the identity by es- tablished name  and  proprietary  name, the date of resumption, and any other information required by § 1271.25(b) not previously    submitted.

  - Any material change in any infor- mation previously submitted. Material changes include any change in registra- tion and  listing  information,  sub- mitted, such  as  whether  the  HCT/P meets the criteria set out in § 1271.10.

- If your HCT/P is described under

§ 1271.20 and is regulated under a BLA, you must submit the information re- quired under part 207 of this chapter using the procedures under  subpart  E of part 207.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 81

FR 60224, Aug. 31, 2016]

##### § 1271.26   When  must  I  amend  my  es- tablishment registration?

If the ownership or location of your establishment changes, or if there is a change in the United States agent’s name, address, telephone number, or email address, you must submit an amendment to registration within 30 calendar days of the change.

[81 FR 60224, Aug. 31, 2017]

##### § 1271.27   Will  FDA  assign  me  a  reg- istration number?

- FDA will assign each location a permanent registration number.

- FDA acceptance of an establish- ment  registration  and  HCT/P  listing

#### Food and Drug Administration, HHS	§ 1271.47

form does not constitute a determina- tion that an establishment is in com- pliance with applicable rules and regu- lations or that the HCT/P is licensed or approved by FDA.

##### § 1271.37 Will establishment registra- tions and HCT/P listings be avail- able for inspection, and how do I request information on registra- tions and listings?

  - Any registration  on  Form  FDA 3356 filed in  paper or electronic  format by each establishment will be available for public inspection through the Cen- ter for Biologics Evaluation and Re- search Human Cell and Tissue Estab- lishment  Registration—Public  Query Web site by using the CBER electronic Web-based application or by going in person to the Food and Drug Adminis- tration, Division  of  Dockets  Manage- ment Public Reading Room (see ad- dress in § 20.120(a) of this chapter). The following information submitted  under the HCT/P requirements  is   illustrative of the type of information that will be available for public  disclosure  when  it is  compiled:

- A list of all HCT/P’s;

- A list of all HCT/P’s manufactured by each establishment;

- A list of all HCT/P’s discontinued; and

- All data or  information that  has already become a matter of public record.

  - You should direct your other re- quests for information  regarding  HCT/ P establishment registrations  and  HCT/ P listings to the Food and Drug Admin- istration, Center for Biologics  Evalua- tion and Research, Office of Commu- nication, Outreach and Development, 10903 New Hampshire Ave., Bldg. 71, Rm. 3103, Silver Spring, MD 20993–0002.

[80 FR 18094, Apr. 3, 2015]

## Subpart C—Donor Eligibility

SOURCE:  69  FR  29830,  May  25,  2004,  unless

otherwise noted.

##### § 1271.45 What requirements does this subpart contain?

- General. This subpart sets out re- quirements for determining donor eli- gibility, including donor screening and testing.  The  requirements  contained  in

this subpart are a component of cur- rent good tissue practice (CGTP) re- quirements. Other CGTP requirements are set out in subpart D of this part.

- Donor-eligibility determination re- quired. A donor-eligibility determina- tion, based on donor screening and testing for relevant communicable dis- ease agents and diseases, is required for all donors of cells or tissue used in HCT/Ps,    except    as    provided    under

§ 1271.90. In the case of an embryo or of cells derived from an embryo, a donor- eligibility  determination  is   required for both the oocyte donor and the semen donor.

- Prohibition on use. An HCT/P must not be implanted,  transplanted,  in- fused, or  transferred  until  the  donor has been determined to be eligible, ex- cept as provided under §§ 1271.60(d), 1271.65(b), and 1271.90 of this subpart.

- Applicability of requirements. If you are an establishment that performs any function described in this subpart, you must comply with the require- ments contained in this subpart that are applicable to that function.

[69 FR 29830, May 25, 2004, as amended at 69

FR 68681, Nov. 24, 2004]

##### § 1271.47   What  procedures  must  I  es- tablish and maintain?

- General. You must establish and maintain procedures for all steps that you perform in testing, screening, de- termining donor eligibility, and com- plying with all other requirements of this subpart. Establish and maintain means define, document (in writing or electronically), and implement; then follow, review, and as needed, revise on an ongoing basis. You must design these procedures to ensure compliance with the requirements of this subpart.

- Review and approval. Before imple- mentation, a responsible person must review and approve all procedures.

- Availability. Procedures must be readily available to the personnel in the area where the operations to which they relate are performed, or in a near- by area if such availability is imprac- tical.

- Departures from procedures. You must record and justify any departure from a procedure relevant to pre- venting  risks  of  communicable  disease

§ 1271.50

transmission at the time of its occur- rence. You must not make available for distribution any HCT/P from a donor whose eligibility is determined under such a departure unless a responsible person has determined that the depar- ture does not increase the risks of com- municable	disease	transmission through the use of the HCT/P.

- Standard procedures. You may adopt current standard procedures, such as those in a technical manual prepared by another organization, pro- vided that you have verified that the procedures are consistent with and at least as stringent as the requirements of this part and appropriate for your operations.

##### § 1271.50 How do I determine whether a donor is eligible?

- Determination based on screening and testing. If you are the establish- ment responsible for making the donor- eligibility determination, you must de- termine whether a donor is eligible based upon the results of donor screen- ing in accordance with § 1271.75 and donor    testing    in    accordance    with

§§ 1271.80 and 1271.85. A responsible per- son, as defined in § 1271.3(t), must deter- mine and document the eligibility of a cell or tissue donor.

- Eligible donor. A donor is eligible under these provisions only if:

  - Donor screening in  accordance with § 1271.75 indicates that the donor:

- Is free from risk factors for, and clinical evidence of, infection due to relevant communicable disease agents and diseases; and

- Is free from communicable dis- ease	risks	associated	with xenotransplantation;  and

  - The results of donor testing for relevant  communicable  disease  agents in accordance with §§ 1271.80 and 1271.85 are negative or nonreactive, except as provided in § 1271.80(d)(1).

##### § 1271.55 What records must accom- pany an HCT/P after the donor-eli- gibility determination is complete; and what records must I retain?

- Accompanying records. Once a donor-eligibility determination has been made, the following must accom- pany the HCT/P at all times:

#### 21 CFR Ch. I (4–1–17 Edition)

- A distinct identification code af- fixed to the HCT/P container, e.g., al- phanumeric, that relates the HCT/P to the donor and to all records pertaining to the HCT/P and, except in the case of autologous donations, directed repro- ductive donations, or donations made by first-degree or second-degree blood relatives, does not include an individ- ual’s name, social security number, or medical record number;

- A statement whether, based on the  results  of  screening  and  testing, the donor has been determined to be el- igible or ineligible; and

- A summary of the records used to make the donor-eligibility  determina- tion.

- Summary of records. The summary of records required by paragraph (a)(3) of this section must contain the fol- lowing information:

  - A statement that the commu- nicable disease testing was performed by a laboratory:

- Certified to perform such testing on human specimens under the Clinical Laboratory Improvement Amendments of 1988 (42 U.S.C. 263a) and 42 CFR part 493; or

- That has met equivalent require- ments as determined by the Centers for Medicare and Medicaid Services in ac- cordance with those provisions;

  - A listing and interpretation of the results of all communicable disease tests performed;

  - The name and address of the es- tablishment that made the donor-eligi- bility determination; and

  - In the case of an HCT/P from a donor who is ineligible based on screen- ing and released under paragraph (b) of

§ 1271.65, a statement noting the rea- son(s) for the determination of ineligi- bility.

- Deletion of personal information. The accompanying records required by this section must not contain the do- nor’s name or other personal informa- tion that might identify the donor.

- Record retention requirements. (1) You must maintain documentation of:

- Results and interpretation of all testing for relevant communicable dis- ease    agents    in    compliance    with

§§ 1271.80 and 1271.85, as well as the name and address of the testing labora- tory  or  laboratories;

#### Food and Drug Administration, HHS	§ 1271.60

- Results and interpretation of all donor screening for communicable dis- eases in compliance with § 1271.75; and

- The donor-eligibility determina- tion, including the name of the respon- sible person who made the determina- tion and the date of the determination.

  - All records must be accurate, in- delible, and legible. Information on the identity and relevant medical  records of the  donor, as  defined in § 1271.3(s), must be in English or, if in another language, must be retained and trans- lated to English and accompanied by a statement of authenticity by the trans- lator that specifically identifies the translated  document.

  - You must retain required records and  make  them  available  for  author- ized inspection by or upon request from FDA.  Records  that  can  be  readily  re- trieved from another location by elec- tronic   means   are   considered   ‘‘re- tained.’’

  - You must retain the records per- taining to a particular HCT/P at least

10 years  after the  date of its  adminis- tration, or if the date of administra- tion is not known, then at least 10 years after the date of the HCT/P’s dis- tribution, disposition, or expiration, whichever is latest.

[69 FR 29830, May 25, 2004, as amended at 70

FR 29952, May 25, 2005]

##### § 1271.60 What quarantine and other requirements apply before the donor-eligibility determination is complete?

- Quarantine. You must keep an HCT/P  in  quarantine,  as  defined  in

§ 1271.3(q), until completion of the donor-eligibility  determination  re- quired by § 1271.50. You must quar- antine semen from anonymous donors until     the     retesting     required     under

§ 1271.85(d) is complete.

- Identification of HCT/Ps in quar- antine. You must clearly identify as quarantined an HCT/P that is in quar- antine pending completion of a donor- eligibility determination. The quar- antined HCT/P must be easily distin- guishable from HCT/Ps that are avail- able for release and distribution.

- Shipping of HCT/Ps in quarantine. If you ship an HCT/P before completion of the donor-eligibility determination, you must keep it in quarantine during

shipment. The HCT/P must be accom- panied by records:

- Identifying the donor (e.g., by a distinct identification code affixed to the HCT/P container);

- Stating that the donor-eligibility determination has not been completed; and

- Stating that the product must not be implanted, transplanted, infused, or transferred until completion of the donor-eligibility determination, except under the terms of paragraph (d) of this section.

- Use in cases of urgent medical need.

- This subpart C does not prohibit the implantation,    transplantation,    infu- sion, or transfer of an HCT/P from a donor for whom the donor-eligibility determination is not  complete  if  there is  a  documented  urgent  medical  need for the HCT/P, as defined in § 1271.3(u).

- If you make an HCT/P available for use under the provisions of para- graph (d)(1) of this section, you must prominently label it ‘‘NOT EVALU- ATED FOR INFECTIOUS SUB- STANCES,’’ and ‘‘ WARNING: Advise patient of communicable disease risks.’’ The following information must accompany the HCT/P:

  - The results of any donor screening required under § 1271.75 that has been completed;

  - The results of any testing re- quired under § 1271.80 or 1271.85 that has been completed; and

  - A list of any screening or testing required   under   § 1271.75,   1271.80   or

1271.85 that has not yet been com- pleted.

- If you are the establishment that manufactured an HCT/P used under the provisions of paragraph (d)(1) of this section, you must document that you notified the physician using the HCT/P that the testing and screening were not complete.

- In the case of an HCT/P used for an urgent medical need under the pro- visions of paragraph (d)(1) of this sec- tion, you must complete the donor-eli- gibility determination  during or  after the use of the HCT/P, and you must in- form the physician of the results of the determination.

#### § 1271.65

##### § 1271.65 How do I store an HCT/P from a donor determined to be in- eligible, and what uses of the HCT/ P are not prohibited?

- Storage. If you are the establish- ment that stores the HCT/P, you must store or identify HCT/Ps from donors who have been determined to be ineli- gible in a physically separate area clearly identified for such use, or fol- low other procedures, such as auto- mated designation, that are  adequate to prevent improper release until de- struction or other disposition of the HCT/P  in  accordance  with  paragraph

- or (c) of this section.

  - Limited uses of HCT/P from ineli- gible donor. (1) An HCT/P from a donor who has been determined to be ineli- gible, based on the results of required testing and/or screening, is not prohib- ited by subpart C of this part from use for implantation, transplantation, in- fusion, or transfer under the following circumstances:

- The HCT/P is for allogeneic use in a first-degree or second-degree blood relative;

- The HCT/P consists of reproduc- tive cells or tissue from a directed re- productive     donor,     as     defined     in

§ 1271.3(l); or

- There is a documented urgent medical need as defined in § 1271.3(u).

- You must prominently label an HCT/P made available for use under the provisions of paragraph (b)(1) of this section with the Biohazard legend shown in § 1271.3(h) with the statement ‘‘WARNING: Advise patient of commu- nicable disease risks,’’ and, in the case of reactive test results, ‘‘WARNING: Reactive test results for (name of dis- ease agent or disease).’’ The  HCT/P must be accompanied by the records re- quired under § 1271.55.

- If you are the establishment that manufactured an HCT/P used under the provisions of paragraph (b)(1) of this section, you must document that you notified the physician using the HCT/P of the results of testing and screening.

  - Nonclinical use. You may make available for nonclinical purposes an HCT/P from a donor who has been de- termined to be ineligible, based on the results of required testing and/or screening, provided that it is labeled:

- ‘‘For Nonclinical Use  Only’’ and

#### 21 CFR Ch. I (4–1–17 Edition)

- With the Biohazard legend  shown in § 1271.3(h).

##### § 1271.75   How do I screen a donor?

- All donors. Except as provided under § 1271.90, if you are the establish- ment that performs donor screening, you must screen a donor of cells or tis- sue by reviewing the donor’s relevant medical records for:

- Risk factors for, and clinical evi- dence of, relevant communicable dis- ease agents and diseases, including:

  - Human immunodeficiency virus;

  - Hepatitis B virus;

  - Hepatitis C virus;

  - Human transmissible spongiform encephalopathy, including Creutzfeldt- Jakob disease;

  - Treponema pallidum; and

- Communicable disease risks asso- ciated with xenotransplantation.

- Donors of viable, leukocyte-rich cells or tissue.  In addition to the relevant communicable disease agents and dis- eases for which screening is required under paragraph (a) of this section, and except as provided under § 1271.90, you must screen the donor of viable, leu- kocyte-rich cells or tissue by reviewing the donor’s  relevant  medical  records for risk factors for and clinical  evi- dence of relevant cell-associated com- municable disease agents and diseases, including    Human    T-lymphotropic virus.

- Donors of reproductive cells or tis- sue. In addition to the relevant commu- nicable disease agents and diseases for which screening is required under para- graphs (a) and (b) of this section, as ap- plicable, and except as provided under

§ 1271.90, you must screen the donor of reproductive cells or tissue by review- ing the donor’s  relevant  medical records for risk factors for and clinical evidence of infection due to relevant communicable diseases of the genito- urinary tract. Such screening must in- clude screening for the communicable disease agents listed in paragraphs (c)(1) and (c)(2) of this section. How- ever, if the reproductive cells or tissues are recovered by a method that ensures freedom from contamination of the cells or tissue by infectious disease or- ganisms that may be present in the genitourinary tract, then screening for the communicable disease agents listed

#### Food and Drug Administration, HHS	§ 1271.80

in paragraphs (c)(1) and (c)(2) of this section is not required. Communicable disease agents of the genitourinary tract for which you must screen in- clude:

  - Chlamydia trachomatis; and

  - Neisseria gonorrhea.

- Ineligible donors. You must deter- mine ineligible a donor who is identi- fied as having either of the following:

  - A risk factor for or clinical evi- dence of any of the relevant commu- nicable disease agents or diseases for which screening is required under para- graphs (a)(1), (b), or (c) of this section; or

  - Any communicable disease risk associated with xenotransplantation.

- Abbreviated procedure for repeat do- nors. If you have performed a complete donor screening procedure on a living donor within the previous 6 months, you may use an abbreviated donor screening procedure on repeat dona- tions. The abbreviated procedure must determine and document any  changes in  the  donor’s  medical  history  since the previous donation that would make the donor ineligible, including relevant social  behavior.

[66  FR 5466,  Jan. 19,  2001,  as  amended  at 71

FR 14798, Mar. 24, 2006]

##### § 1271.80   What are the general require- ments for donor testing?

(a) Testing for relevant communicable diseases is required. To adequately and appropriately reduce  the risk  of trans- mission of relevant communicable dis- eases,  and  except  as  provided  under

§ 1271.90, if you are the establishment that performs donor testing, you must test a donor specimen for evidence of infection due  to communicable  disease agents  in  accordance  with  paragraph

- of this section. You must  test  for those communicable disease agents specified in § 1271.85. In the case of a donor 1 month of age or younger, you must test a specimen from the birth mother instead of a specimen from the donor.

  - Timing of specimen collection. You must collect the donor specimen for testing at the time of recovery of cells or tissue from the donor; or up to 7 days before or after recovery, except:

    - For donors of peripheral blood stem/progenitor cells, bone marrow (if

not excepted under § 1271.3(d)(4)), or oo- cytes, you may collect the donor speci- men for testing up to 30 days before re- covery;  or

    - In the case of a repeat semen donor from whom a specimen has al- ready been collected and tested, and for whom    retesting    is    required    under

§ 1271.85(d), you are not required to col- lect a donor specimen at the time of each  donation.

  - Tests. You must test using appro- priate FDA-licensed, approved, or cleared donor screening tests, in ac- cordance with the manufacturer’s in- structions, to adequately and appro- priately reduce the risk of  trans- mission of relevant communicable dis- ease agents or diseases; however, until such time as appropriate FDA-licensed, approved, or cleared donor screening tests for Chlamydia trachomatis and for Neisseria gonorrhea are available, you must use FDA-licensed, approved, or cleared tests labeled for  the  detection of those organisms in an asymp- tomatic, low-prevalence  population. You must use a test  specifically  la- beled for cadaveric specimens  instead of a more generally labeled test when applicable and when available. Re- quired testing under this section  must be performed by a laboratory that ei- ther is certified to perform such  test- ing on human specimens under the Clinical Laboratory Improvement Amendments of 1988 (42 U.S.C. 263a) and 42 CFR part 493, or has met equiva- lent requirements as determined by the Centers for Medicare and Medicaid Services.

  - Ineligible donors. You must  deter- mine the following donors to be ineli- gible:

    - A donor whose specimen tests re- active on  a screening test  for a commu- nicable  disease  agent  in  accordance with § 1271.85, except for a donor whose specimen tests reactive on a non- treponemal screening test  for  syphilis and negative on a specific treponemal confirmatory    test;

(2)(i) A donor in whom plasma dilu- tion sufficient to affect the results of communicable disease testing is sus- pected, unless:

#### § 1271.85

      - You test a specimen taken from the donor before transfusion or infu- sion and up to 7 days before recovery of cells or tissue; or

      - You use an appropriate algorithm designed to  evaluate volumes  adminis- tered in the 48 hours before specimen collection, and the algorithm  shows that plasma  dilution  sufficient  to  af- fect the results of communicable dis- ease testing has not occurred.

- Clinical situations in which you must suspect plasma dilution sufficient to affect the results of communicable disease testing include but are not lim- ited to the following:

  - Blood loss is known or suspected in a donor over 12 years of age, and the donor has received a transfusion or in- fusion of any of the following, alone or in  combination:

- More than 2,000 milliliters (mL) of blood (e.g., whole blood, red blood cells) or colloids within 48 hours before death or specimen collection, which- ever occurred earlier, or

- More than 2,000 mL of crystalloids within 1 hour before death or specimen collection,  whichever  occurred  earlier.

  - Regardless of the presence or ab- sence of blood loss, the donor is 12 years of age or younger and has re- ceived a transfusion or infusion of any amount of any of  the  following,  alone or in combination:

    - Blood (e.g., whole blood, red blood cells) or colloids within 48 hours before death or specimen collection, which- ever occurred earlier, or

    - Crystalloids within 1 hour before death or specimen collection, which- ever occurred earlier.

[69 FR 29830, May 25, 2004, as amended at 70

FR 29952, May 25, 2005]

##### § 1271.85 What donor testing is re- quired for different types of cells and tissues?

      - All donors. To adequately and ap- propriately reduce the risk of trans- mission of relevant communicable dis- eases,  and  except  as  provided  under

§ 1271.90, you must test a specimen from the donor of cells or tissue,  whether viable or nonviable, for evidence of in- fection due to relevant communicable disease  agents,  including:

        - Human immunodeficiency virus, type 1;

#### 21 CFR Ch. I (4–1–17 Edition)

        - Human immunodeficiency virus, type 2;

        - Hepatitis B virus;

        - Hepatitis C virus; and

        - Treponema pallidum.

      - Donors of viable, leukocyte-rich cells or tissue.  In addition to the relevant communicable disease agents for which testing is required under  paragraph  (a) of this section, and except as provided under § 1271.90,

        - You must test a specimen from the donor of viable, leukocyte-rich cells or tissue to adequately and appro- priately reduce the risk of trans- mission of relevant cell-associated communicable diseases, including:

- Human T-lymphotropic virus, type I; and

- Human T-lymphotropic virus, type II.

        - You must test a specimen from the donor of viable, leukocyte-rich cells or tissue for evidence of infection due to cytomegalovirus (CMV), to ade- quately and appropriately reduce the risk of transmission. You must estab- lish and maintain a standard operating procedure governing the release of an HCT/P from a donor whose specimen tests reactive for CMV.

      - Donors of reproductive cells or tis- sue. In addition to the communicable disease agents for which testing is re- quired under paragraphs (a) and (b) of this section, as applicable,  and  except as provided under § 1271.90, you must test a specimen from the donor of re- productive cells or tissue to adequately and appropriately reduce the risk of transmission of relevant communicable disease agents of the genitourinary tract. Such testing must include test- ing for the communicable disease agents listed in paragraphs (c)(1) and (c)(2) of this section. However, if the re- productive cells or tissues are recov- ered by a method that ensures freedom from contamination of the cells or tis- sue by infectious disease  organisms that may be present in the genito- urinary tract, then testing for the com- municable disease agents listed in paragraphs (c)(1) and (c)(2) of this sec- tion is not required. Communicable disease agents of the genitourinary tract  for  which  you  must  test  include:

        - Chlamydia trachomatis; and

        - Neisseria gonorrhea.

#### Food and Drug Administration, HHS	§ 1271.90

      - Retesting anonymous semen donors. Except as provided under § 1271.90 and except for directed reproductive  donors as defined in § 1271.3(l), at least 6 months after the date of donation of semen from anonymous donors, you must collect a new specimen from the donor and test it for evidence of infec- tion due to the communicable disease agents for which testing is required under paragraphs (a), (b), and (c) of this section.

      - Dura mater. For donors of dura mater, you must perform an adequate assessment designed to detect evidence of	transmissible	spongiform encephalopathy.

##### § 1271.90 Are there  other  exceptions and what labeling requirements apply?

- Donor-eligibility determination not required. You are not required to make a donor-eligibility determination under

§ 1271.50 or to perform donor screening or testing under §§ 1271.75, 1271.80 and

- for:

  - Cells and tissues for autologous use; or

  - Reproductive cells or tissue do- nated by a sexually intimate partner of the recipient for reproductive use; or

  - Cryopreserved cells or tissue for reproductive use, other than embryos, originally excepted under paragraphs (a)(1) or (a)(2) of this section at the time of donation, that are  subse- quently intended for directed donation, provided that:

    - Additional donations are unavail- able, for example, due to the infertility or health of a donor of the cryopreserved reproductive cells or tis- sue; and

    - Appropriate measures are taken to screen  and  test the donor(s)  before transfer to the recipient.

  - A cryopreserved embryo, origi- nally excepted under paragraph (a)(2) of this section at the time of recovery or cryopreservation, that is subse- quently intended for directed or anony- mous donation. When possible, appro- priate measures should be taken to screen and test the semen and oocyte donors before transfer of the embryo to the recipient.

- Exceptions for reproductive use. An embryo  originally  intended  for  repro-

ductive use for a specific individual or couple that  is  subsequently  intended for directed or anonymous donation for reproductive use is excepted from the prohibition on use under § 1271.45(c) even when the applicable donor eligi- bility requirements under subpart C of this part are not met. Nothing in this paragraph creates an exception for de- ficiencies that occurred in making the donor eligibility determination for ei- ther the oocyte donor or the semen donor as required under § 1271.45(b), or for deficiencies in performing donor screening or testing, as required under

§§ 1271.75, 1271.80, and 1271.85.

- Required labeling. As applicable, you must prominently label an HCT/P described in paragraphs (a) and (b) of this section as follows:

- ‘‘FOR AUTOLOGOUS USE ONLY,’’ if it is stored for autologous use.

- ‘‘NOT EVALUATED FOR INFEC- TIOUS SUBSTANCES,’’ unless  you have performed all otherwise applica- ble    screening    and    testing    under

§§ 1271.75, 1271.80, and 1271.85. This para- graph does not apply  to  reproductive cells or tissue labeled  in  accordance with paragraph (c)(6) of this section.

- Unless the HCT/P is for autologous use only, ‘‘WARNING: Ad- vise recipient of communicable disease risks,’’

- When the donor-eligibility deter- mination under § 1271.50(a) is not per- formed or is not completed; or

- If the results of any screening or testing performed indicate:

  - The presence of relevant commu- nicable disease agents and/or

  - Risk factors for or clinical evi- dence of  relevant  communicable  dis- ease agents or diseases.

- With the Biohazard legend  shown in § 1271.3(h), if the results of any screening  or  testing  performed  indi- cate:

  - The presence of relevant commu- nicable disease agents and/or

  - Risk factors for or clinical evi- dence of  relevant  communicable  dis- ease agents or diseases.

- ‘‘WARNING: Reactive test results for (name of disease agent or disease),’’ in the case of reactive test results.

- ‘‘Advise recipient that screening and  testing  of  the  donor(s)  were  not

#### § 1271.145

performed at the time of recovery or cryopreservation of the reproductive cells or tissue, but have been performed subsequently,’’ for paragraphs (a)(3) or (a)(4) of this section.

[69 FR 29830, May 25, 2004, as amended at 70

FR 29952,  May 25,  2005;  81  FR 40517,  June 22,

2016]

## Subpart D—Current Good Tissue Practice

SOURCE:  69  FR  68681,  Nov.  24,  2004,  unless

otherwise noted.

##### § 1271.145 Prevention of the introduc- tion, transmission, or spread of communicable diseases.

You must recover, process, store, label, package, and distribute HCT/Ps, and screen and test cell and tissue do- nors, in a way that prevents the intro- duction, transmission, or spread of communicable diseases.

##### § 1271.150   Current good tissue practice requirements.

- General. This subpart D and sub- part C of this part set forth current good tissue practice (CGTP) require- ments. You must follow CGTP require- ments to prevent the introduction, transmission, or spread of commu- nicable diseases by HCT/Ps (e.g., by en- suring that the HCT/Ps do not contain communicable disease agents, that they are not contaminated, and that they do not become contaminated dur- ing manufacturing). Communicable diseases include, but are not limited to, those transmitted by viruses, bacteria, fungi, parasites, and transmissible spongiform encephalopathy agents. CGTP requirements govern the meth- ods used in, and the facilities and con- trols used for, the manufacture of HCT/ Ps, including but not limited to all steps in recovery, donor screening, donor testing, processing, storage, la- beling, packaging, and distribution. The CGTP provisions specifically gov- erning determinations of donor eligi- bility, including donor screening and testing, are set out separately in sub- part C of this part.

- Core CGTP requirements. The fol- lowing are core CGTP requirements:

  - Requirements  relating  to  facili- ties in § 1271.190(a) and (b);

21 CFR Ch. I (4–1–17 Edition)

  - Requirements relating to environ- mental control in § 1271.195(a);

  - Requirements relating to equip- ment in § 1271.200(a);

  - Requirements relating to supplies and reagents in § 1271.210(a) and (b);

  - Requirements relating to  recovery in § 1271.215;

  - Requirements relating to proc- essing and process controls in § 1271.220;

  - Requirements relating to labeling controls in § 1271.250(a) and (b);

  - Requirements  relating  to  storage in § 1271.260 (a) through (d);

  - Requirements relating to receipt, predistribution shipment, and distribu- tion of an HCT/P in § 1271.265(a) through (d); and

  - Requirements relating to donor eligibility determinations, donor screening,     and     donor     testing     in

§§ 1271.50, 1271.75, 1271.80, and 1271.85.

- Compliance with applicable require- ments—(1)  Manufacturing  arrangements

- If you are an establishment that en- gages in only some operations subject to the regulations in this subpart and subpart C of this part, and not others, then you need only comply with those requirements applicable to the oper- ations that you perform.

- If you engage another establish- ment (e.g., a laboratory to perform communicable disease testing, or an ir- radiation facility to perform terminal sterilization), under a contract, agree- ment, or other arrangement, to per- form any step in manufacture for you, that establishment is responsible for complying with requirements applica- ble to that manufacturing step.

- Before entering into a contract, agreement, or other arrangement with another establishment to perform any step in manufacture for you, you must ensure that the establishment complies with applicable CGTP requirements. If, during the course of this contract, agreement, or other arrangement, you become aware of information sug- gesting that the establishment may no longer be in compliance with such re- quirements, you must take reasonable steps to ensure the establishment com- plies with those requirements. If you determine  that  the  establishment  is

#### Food and Drug Administration, HHS	§ 1271.155

not in compliance with those require- ments, you must terminate your con- tract, agreement, or other arrange- ment with the establishment.

- If you are the establishment that determines that an HCT/P meets all re- lease criteria and makes the HCT/P available for distribution, whether or not you are the actual distributor, you are responsible for reviewing manufac- turing and tracking records  to deter- mine that the HCT/P has been manu- factured and tracked in compliance with the requirements of this subpart and subpart C of this part and any other applicable requirements.

- With the exception of §§ 1271.150(c) and 1271.155 of this subpart, the regula- tions in this subpart are not being im- plemented for reproductive HCT/Ps de- scribed in § 1271.10 and regulated solely under section 361 of the Public Health Service Act and the regulations in this part, or for the establishments that manufacture    them.

- Compliance with parts 210, 211, and 820 of this chapter. With respect to HCT/ Ps that are drugs (subject to review under an application submitted under section 505 of the Federal Food, Drug, and Cosmetic Act or under a biological product license application under sec- tion 351 of the Public Health Service Act) or that are devices (subject to pre- market review  or  notification  under the device provisions of the act  or under a biological product license ap- plication under section 351 of the Pub- lic Health Service Act), the procedures contained in this subpart and in sub- part C of this part and the current good manufacturing practice regulations in parts 210 and 211 of this chapter and the quality system regulations in  part  820 of this chapter supplement, and do not supersede, each other unless the regu- lations explicitly provide otherwise. In the event that a regulation in part 1271 of this chapter is in conflict with a re- quirement in parts 210, 211, or 820 of this chapter, the regulations more spe- cifically applicable to the product in question will  supersede the  more gen- eral.

- Where appropriate. When a require- ment   is   qualified   by   ‘‘where   appro- priate,’’   it   is   deemed   to   be   ‘‘appro- priate’’  unless  you  can  document  jus- tification  otherwise.  A  requirement  is

‘‘appropriate’’ if nonimplementation of the requirement could reasonably be expected to result in the HCT/P not meeting its specified requirements re- lated to prevention of introduction, transmission, or spread of commu- nicable diseases, or in your inability to carry out any necessary corrective ac- tion.

##### § 1271.155   Exemptions	and	alter- natives.

- General. You may request an ex- emption from or alternative to any re- quirement in subpart C or D of this part.

- Request for exemption or alternative. Submit your request under this section to the Director of the appropriate Cen- ter (the Director), e.g., the Center for Biologics Evaluation and Research or the Center for Devices and Radio- logical Health. The request must be ac- companied by supporting documenta- tion, including all relevant valid sci- entific data, and must contain either:

- Information justifying the re- quested exemption from the require- ment, or

- A description of a proposed alter- native method of meeting the require- ment.

- Criteria for granting an exemption or alternative. The Director may grant an exemption or alternative if he or she finds that such action is consistent with the goals of protecting the public health and/or preventing the introduc- tion, transmission, or spread of com- municable diseases and that:

  - The information submitted justi- fies an exemption; or

  - The proposed alternative satisfies the purpose of the requirement.

- Form of request. You must  ordi- narily make your request for an ex- emption or alternative in writing (hard copy or electronically). However, if cir- cumstances make it  difficult  (e.g., there is inadequate time) to submit your request in writing, you may make the request orally, and the Director may orally grant an exemption or al- ternative. You must follow your oral request with an immediate written re- quest, to which the Director will re- spond in writing.

- Operation under exemption or alter- native.  You  must  not  begin  operating

#### § 1271.160

under the terms of a requested exemp- tion or alternative until the exemption or alternative has been granted. You may apply for an extension of an ex- emption or alternative beyond its expi- ration date, if any.

- Documentation. If you operate under the terms of an exemption or al- ternative, you must maintain docu- mentation  of:

- FDA’s grant of the exemption or alternative, and

- The date on which you began op- erating under the terms of the exemp- tion or alternative.

- Issuance of an exemption or alter- native by the Director. In a public health emergency, the Director may issue an exemption from, or alternative to, any requirement in part 1271. The Director may issue  an  exemption  or alternative under this section if the exemption or alternative is necessary to assure that certain HCT/Ps will be available in a specified location to respond to an un- anticipated immediate need for those HCT/Ps.

##### § 1271.160   Establishment  and  mainte- nance of a quality program.

- General. If you are an establish- ment that performs any step in the manufacture of HCT/Ps, you must es- tablish and maintain a quality pro- gram intended to prevent the introduc- tion, transmission, or spread of com- municable diseases through the manu- facture and use of HCT/Ps. The quality program must be appropriate for the specific HCT/Ps manufactured and the manufacturing steps performed. The quality program must address all core CGTP        requirements        listed        in

§ 1271.150(b).

- Functions. Functions of the qual- ity program must include:

  - Establishing and maintaining ap- propriate procedures relating to core CGTP requirements, and ensuring com- pliance    with    the    requirements    of

§ 1271.180 with respect to such proce- dures, including review, approval, and revision;

  - Ensuring that procedures exist for receiving, investigating, evaluating, and documenting information relating to core CGTP requirements, including complaints, and for sharing any infor- mation pertaining to the possible con-

#### 21 CFR Ch. I (4–1–17 Edition)

tamination of the HCT/P or the poten- tial for transmission of a commu- nicable disease by the HCT/P with the following:

- Other establishments that are known to have recovered HCT/Ps from the same donor;

- Other establishments that are known to have performed manufac- turing steps with respect to the same HCT/P; and

- Relating to consignees, in the case of such information received after the HCT/P is made available for dis- tribution, shipped to the consignee, or administered to the recipient, proce- dures must include provisions for as- sessing risk and appropriate followup, and evaluating the effect this informa- tion has on the HCT/P and for the noti- fication of all entities to whom the af- fected HCT/P was distributed, the quar- antine and recall of the HCT/P, and/or reporting to FDA, as necessary.

  - Ensuring that appropriate correc- tive actions relating to core CGTP re- quirements, including reaudits of defi- ciencies, are taken and documented, as necessary. You must verify corrective actions to ensure that such actions are effective and are in compliance with CGTP. Where appropriate, corrective actions must include both short-term action to address the immediate prob- lem and long-term action to prevent the problem’s recurrence. Documenta- tion of corrective actions must include, where appropriate:

    - Identification of the HCT/P af- fected and a description of its disposi- tion;

    - The nature of the problem requir- ing corrective action;

    - A description of the corrective action taken; and

    - The date(s) of the corrective ac- tion.

  - Ensuring the proper training and education of personnel involved in ac- tivities related to core CGTP require- ments;

  - Establishing and maintaining ap- propriate monitoring systems as nec- essary to comply with the require- ments of this subpart (e.g., environ- mental  monitoring);

  - Investigating and documenting HCT/P deviations and trends of HCT/P

#### Food and Drug Administration, HHS	§ 1271.190

deviations relating to core CGTP re- quirements and making reports if re- quired under § 1271.350(b) or other appli- cable regulations. Each investigation must include a review  and  evaluation of the HCT/P deviation, the efforts made to determine the cause, and the implementation of corrective action(s) to address the HCT/P deviation and prevent recurrence.

- Audits. You must periodically per- form for management review a quality audit, as defined in § 1271.3(gg), of ac- tivities related to core CGTP require- ments.

- Computers. You must validate the performance of computer software for the intended use, and the performance of any changes to that software for the intended use, if you rely upon the soft- ware to comply with core CGTP re- quirements and if the software either is custom software or is commercially available software that has been cus- tomized or programmed (including software programmed to perform a user defined calculation or table) to perform a function related to core CGTP re- quirements. You must verify the per- formance of all other software for the intended use if you rely upon it to com- ply with core CGTP requirements. You must approve and document these ac- tivities and results before implementa- tion.

##### § 1271.170   Personnel.

- General. You must have personnel sufficient to  ensure  compliance  with the requirements of this part.

- Competent performance of functions. You must have personnel with the nec- essary education, experience, and training to ensure competent  perform- ance of their assigned functions. Per- sonnel must perform only those activi- ties for which they are qualified and authorized.

- Training. You must train all per- sonnel, and retrain as necessary, to perform their assigned responsibilities adequately.

##### § 1271.180   Procedures.

- General. You must establish and maintain procedures appropriate to meet core CGTP requirements for all steps that you perform in the manufac- ture of HCT/Ps. You must design these

procedures to prevent circumstances that increase the risk of the introduc- tion, transmission, or spread of com- municable diseases through the use of HCT/Ps.

- Review and approval. Before imple- mentation, a responsible person must review and approve these procedures.

- Availability. These procedures must be readily available to the per- sonnel in the area where the operations to which they relate are performed, or in a nearby area if such availability is impractical.

- Standard procedures. If you adopt current standard procedures from an- other organization, you must verify that the procedures meet the require- ments of this part and are appropriate for your operations.

##### § 1271.190   Facilities.

- General. Any facility used in the manufacture of HCT/Ps must be of suit- able size, construction, and location to prevent contamination of HCT/Ps with communicable disease agents and to ensure orderly handling of HCT/Ps without mix-ups. You must maintain the facility in a good state of repair. You must provide lighting, ventilation, plumbing, drainage, and access to sinks and toilets that are adequate to pre- vent the introduction, transmission, or spread of communicable disease.

- Facility cleaning and sanitation. (1) You must maintain any facility used in the manufacture of HCT/Ps in a clean, sanitary, and orderly manner, to pre- vent the introduction, transmission, or spread of communicable disease.

- You must dispose of sewage, trash, and other refuse in a timely, safe, and sanitary manner.

- Operations. You must divide a fa- cility used in the manufacture of HCT/ Ps into separate or defined areas of adequate size for each operation that takes place in the facility, or you must establish and maintain other control systems to prevent improper labeling, mix-ups, contamination, cross-con- tamination, and accidental exposure of HCT/Ps to communicable disease agents.

- Procedures and records. (1) You must establish and maintain proce- dures for facility cleaning and sanita- tion  for  the  purpose  of  preventing  the

#### § 1271.195

introduction, transmission, or spread of communicable disease. These proce- dures must assign responsibility for sanitation and must describe in suffi- cient detail the cleaning methods to be used and the schedule for cleaning the facility.

  - You must document, and main- tain records of, all cleaning and sanita- tion activities performed to prevent contamination of HCT/Ps. You must re- tain such records 3 years after their creation.

##### § 1271.195   Environmental  control  and monitoring.

    - Environmental control. Where envi- ronmental conditions could reasonably be expected to cause contamination or cross-contamination of HCT/Ps or equipment, or accidental exposure of HCT/Ps to communicable disease agents, you must adequately control environmental conditions and provide proper conditions for operations. Where appropriate, you must provide for the following  control  activities  or  systems:

      - Temperature and humidity con- trols;

      - Ventilation and air filtration;

      - Cleaning and disinfecting of rooms and equipment to ensure aseptic processing operations; and

      - Maintenance of equipment used to control conditions necessary for asep- tic processing operations.

    - Inspections. You must inspect each environmental control system periodi- cally to verify that the system, includ- ing necessary equipment, is adequate and functioning properly. You must take appropriate corrective action as necessary.

    - Environmental monitoring. You must monitor environmental condi- tions where environmental conditions could reasonably  be expected to  cause contamination  or  cross-contamination of HCT/Ps or equipment,  or accidental exposure of HCT/Ps to communicable disease agents. Where appropriate, you must provide environmental moni- toring for microorganisms.

    - Records. You must document, and maintain records of, environmental control and monitoring activities.

#### 21 CFR Ch. I (4–1–17 Edition)

##### § 1271.200   Equipment.

- General. To prevent the introduc- tion, transmission, or spread of com- municable diseases, equipment used in the manufacture of HCT/Ps must be of appropriate design for its use and must be suitably located and installed to fa- cilitate operations, including cleaning and maintenance. Any automated, me- chanical, electronic, or other equip- ment used for inspection,  measuring, or testing in accordance with this part must be capable of producing valid re- sults. You must clean, sanitize, and maintain equipment according to es- tablished schedules.

- Procedures and schedules. You must establish and maintain proce- dures for cleaning, sanitizing, and maintaining equipment to prevent malfunctions, contamination or cross- contamination, accidental exposure of HCT/Ps to communicable disease agents, and other events that could reasonably be expected to result in the introduction, transmission, or spread of communicable diseases.

- Calibration of equipment. Where ap- propriate, you must routinely calibrate according to established procedures and schedules all automated, mechan- ical, electronic, or other equipment used for inspection, measuring, and testing in accordance with this part.

- Inspections. You must routinely inspect equipment for cleanliness, sani- tation, and  calibration, and  to ensure adherence to applicable equipment maintenance  schedules.

- Records. You must document and maintain records of all equipment maintenance, cleaning, sanitizing, cali- bration, and other activities performed in accordance with  this section. You must display records of recent mainte- nance,  cleaning,  sanitizing,   calibra- tion, and other activities on or near each piece of equipment, or make the records readily available to the indi- viduals responsible for performing these activities and to the personnel using the equipment. You must main- tain records of the use of each piece of equipment, including the identification of each HCT/P manufactured with that equipment.

#### Food and Drug Administration, HHS	§ 1271.230

##### § 1271.210   Supplies and reagents.

  - Verification. You must not use supplies and reagents until they have been verified to meet specifications de- signed to prevent circumstances that increase the risk of the introduction, transmission, or spread of commu- nicable diseases. Verification may be accomplished by the establishment that uses the supply or reagent, or by the vendor of the supply or reagent.

  - Reagents. Reagents used in proc- essing and  preservation  of  HCT/Ps must be sterile, where appropriate.

  - In-house reagents. You must vali- date and/or verify the processes used for production of in-house reagents.

  - Records. You must maintain the following records pertaining to supplies and reagents:

    - Records of the receipt of each sup- ply or reagent, including the type, quantity, manufacturer, lot number, date of receipt, and expiration date;

    - Records of the verification of each supply or reagent, including test re- sults or, in the case of vendor verification, a certificate of analysis from the vendor; and

    - Records of the lot of supply or re- agent used in the manufacture of each HCT/P.

##### § 1271.215   Recovery.

If you are an establishment that re- covers HCT/Ps, you must recover each HCT/P in a way that does not cause contamination or cross-contamination during recovery, or otherwise increase the risk of the introduction, trans- mission, or spread of communicable disease through the use of the HCT/P.

##### § 1271.220   Processing and process con- trols.

      - General. If you are an establish- ment that processes HCT/Ps, you must process each HCT/P in a way that does not cause contamination or cross-con- tamination during processing, and that prevents the introduction, trans- mission, or spread of communicable disease through the use of the HCT/P.

      - Pooling. Human cells or tissue from two or more donors must not be pooled (placed  in physical  contact  or mixed in a single receptacle) during manufacturing.

      - In-process control and testing. You must ensure that specified require- ments, consistent with paragraph (a) of this section, for in-process controls are met, and that each in-process HCT/P is controlled until the required  inspec- tion and tests or other verification ac- tivities have been completed, or nec- essary approvals are received and docu- mented. Sampling of in-process HCT/Ps must be representative of the material to be evaluated.

      - Dura mater. (1) When there is a published validated process that re- duces the risk of transmissible spongiform encephalopathy, you must use this process for dura mater (or an equivalent process that you have vali- dated), unless following this process adversely affects the clinical utility of the dura mater.

- When you use a published vali- dated process, you must verify such a process in your establishment.

##### § 1271.225   Process changes.

Any change to a process must be verified  or  validated  in  accordance with § 1271.230, to ensure that the change does not create an adverse im- pact elsewhere in the operation, and must be approved before implementa- tion by a responsible person with ap- propriate knowledge and background. You must communicate approved changes to the appropriate personnel in a timely manner.

##### § 1271.230   Process validation.

  - General. Where the results of proc- essing described in § 1271.220 cannot be fully verified by subsequent inspection and tests, you must validate  and  ap- prove the process according to estab- lished procedures. The validation ac- tivities and results must be docu- mented, including the date and signa- ture of the individual(s) approving the validation.

  - Written representation. Any writ- ten representation that your  proc- essing methods reduce the risk of transmission of communicable disease by an HCT/P, including but not limited to, a representation of sterility or pathogen inactivation of an HCT/P, must be based on a fully verified or validated process.

#### § 1271.250

  - Changes. When changes to a vali- dated process subject to paragraph (a) of this section occur, you must review and evaluate the process and perform revalidation where appropriate. You must  document  these  activities.

##### § 1271.250   Labeling controls.

- General. You must establish and maintain procedures to control the la- beling of HCT/Ps. You must design these procedures to ensure proper HCT/ P identification and to prevent mix- ups.

- Verification. Procedures must in- clude verification of label accuracy, legibility, and integrity.

- Labeling requirements. Procedures must ensure that each HCT/P is labeled in accordance with all applicable label- ing   requirements,   including   those   in

§§ 1271.55,       1271.60,       1271.65,       1271.90,

- , and 1271.370, and that each HCT/P made available  for  distribution is accompanied  by  documentation  of the donor eligibility determination as required under § 1271.55.

##### § 1271.260   Storage.

  - Control of storage areas. You must control your storage areas and stock rooms to prevent:

- Mix-ups, contamination, and cross-contamination of HCT/Ps, sup- plies, and reagents, and

- An HCT/P  from being  improperly made available for distribution.

  - Temperature. You must store HCT/ Ps at an appropriate temperature.

  - Expiration date. Where appro- priate, you must assign an expiration date to each HCT/P based on the fol- lowing factors:

- HCT/P type;

- Processing, including the method of preservation;

- Storage conditions; and

- Packaging.

  - Corrective action. You must take and document corrective action when- ever proper storage conditions are not met.

  - Acceptable temperature limits. You must establish acceptable temperature limits for storage of HCT/Ps at each step of the manufacturing process to inhibit the growth of infectious agents. You must maintain and record storage temperatures for HCT/Ps. You must pe-

#### 21 CFR Ch. I (4–1–17 Edition)

riodically review recorded tempera- tures to ensure that temperatures have been within acceptable limits.

##### § 1271.265 Receipt, predistribution shipment, and distribution of an HCT/P.

- Receipt. You must evaluate each incoming HCT/P for the presence and significance of microorganisms and in- spect for damage and contamination. You must determine whether to accept, reject, or place in quarantine each in- coming HCT/P, based upon pre-estab- lished criteria designed to prevent communicable disease transmission.

- Predistribution shipment. If you ship an HCT/P within your establish- ment or between establishments (e.g., procurer to processor) and the HCT/P is not available for distribution as de- scribed in paragraph (c) of this section, you must first determine and  docu- ment whether pre-established criteria designed to prevent communicable dis- ease transmission have been met, and you must ship the HCT/P in quar- antine.

- Availability for distribution. (1) Be- fore making an HCT/P available for distribution, you must review manu- facturing and tracking records per- taining to the HCT/P, and, on the basis of that record review, you must verify and document that the release criteria have been met. A  responsible person must document and date the deter- mination that an HCT/P is available for distribution.

- You must not make available for distribution an HCT/P that is in quar- antine, is contaminated, is recovered from a donor who has been determined to be ineligible or for whom a donor- eligibility determination has not been completed (except as provided under

§§ 1271.60, 1271.65, and 1271.90), or that otherwise does not meet  release  cri- teria designed  to  prevent  commu- nicable   disease   transmission.

- You must not make available for distribution any HCT/P manufactured under a departure from a procedure rel- evant to preventing risks of commu- nicable disease transmission, unless a responsible person has determined that the departure does  not  increase  the risk of communicable disease through the use of the HCT/P. You must record

#### Food and Drug Administration, HHS	§ 1271.290

and justify any departure from a proce- dure at the time of its occurrence.

- Packaging and shipping. Packaging and shipping containers must be de- signed and constructed to protect the HCT/P from contamination. For each type of HCT/P, you must establish ap- propriate shipping conditions to be maintained  during  transit.

- Procedures. You must establish and maintain procedures, including re- lease criteria, for the activities in paragraphs (a) through (d) of this sec- tion. You must document these activi- ties. Documentation must include:

  - Identification of the HCT/P and the establishment that supplied the HCT/P;

  - Activities performed and the re- sults of each activity;

  - Date(s) of activity;

  - Quantity of HCT/P subject to the activity; and

  - Disposition of the HCT/P (e.g., identity of consignee).

- Return to inventory. You must es- tablish and maintain procedures to de- termine if an HCT/P that is returned to your establishment is suitable to be re- turned to inventory.

##### § 1271.270   Records.

- General. You must maintain records concurrently with the perform- ance of each step required in this sub- part and subpart C of this part. Any re- quirement in this part that an action be documented involves the creation of a record, which is subject to the re- quirements of this section. All records must be accurate, indelible, and leg- ible. The records must identify the per- son performing the work and the dates of the various entries, and must be as detailed as necessary to provide a com- plete history of the work performed and to relate the records to the par- ticular HCT/P involved.

- Records management system. You must establish and maintain a records management system relating to core CGTP requirements. Under this sys- tem, records pertaining to a particular HCT/P must be maintained in such a way as to facilitate review of the HCT/ Ps history before making it available for distribution and, if necessary, sub- sequent to the HCT/Ps  release  as  part of  a  followup  evaluation  or  investiga-

tion. Records pertinent to the manu- facture of HCT/Ps (e.g., labeling and packaging procedures, and equipment logs) must also be maintained and or- ganized under the records management system. If records are maintained in more than one location, then  the records management system must be designed to ensure prompt identifica- tion, location, and retrieval of all records.

- Methods of retention. You may maintain records required under this subpart electronically, as original paper records, or as true copies such as photocopies, microfiche, or microfilm. Equipment that is necessary to make the records available and legible, such as computer and reader equipment, must be readily available. Records stored in electronic systems must be backed up.

- Length of retention. You must re- tain all records for 10 years after their creation, unless stated  otherwise  in this  part.  However,  you  must  retain the records pertaining to a particular HCT/P at least 10 years after the date of its administration, or if the date of administration is not known, then at least 10 years after the date of the HCT/Ps distribution, disposition, or ex- piration, whichever  is  latest.  You  must retain records for archived  specimens of dura mater for 10 years after the ap- propriate disposition of the specimens.

- Contracts and agreements. You must maintain the name and address and a list of the responsibilities of any establishment that performs a manu- facturing step for you. This informa- tion must be available during an in- spection conducted under § 1271.400.

##### § 1271.290   Tracking.

- General. If you perform any  step in the manufacture of an HCT/P  in which you handle the HCT/P, you must track each such HCT/P in accordance with this section, to facilitate the in- vestigation of actual or suspected transmission of communicable disease and take appropriate and timely cor- rective  action.

- System of HCT/P tracking. (1) You must establish and maintain a system of HCT/P tracking that enables the tracking of all HCT/Ps from:

#### § 1271.320

  - The donor to the  consignee or final disposition; and

  - The consignee or final disposition to the donor.

(2) Alternatively, if you are an estab- lishment that performs some but  not all of the steps in  the  manufacture  of an HCT/P in which you handle the HCT/ P, you may participate in a system of HCT/P tracking established and main- tained by another establishment re- sponsible for other steps in the manu- facture of the same HCT/P, provided that  the  tracking  system  complies with all the requirements of this sec- tion.

- Distinct identification code. As part of your tracking system, you must en- sure: That each HCT/P that you manu- facture is assigned and labeled with a distinct identification code, e.g., alpha- numeric, that relates the HCT/P to the donor and to all records pertaining to the HCT/P; and that labeling includes information designed to facilitate ef- fective tracking, using the distinct identification code, from the donor to the recipient and from the recipient to the   donor.   Except   as   described   in

§ 1271.55(a)(1), you must create such a code specifically for  tracking,  and  it may not include an individual’s name, social security number, or medical record number. You may adopt a dis- tinct identification code assigned by another establishment engaged in the manufacturing process, or you may as- sign a new code. If you assign a new code to an HCT/P, you must establish and maintain procedures for relating the new code to the old code.

- Tracking from consignee to donor. As part of your tracking system, you must establish and maintain a method for recording the distinct identifica- tion code and type of each HCT/P dis- tributed to a consignee to enable tracking from the consignee to the donor.

- Tracking from donor to consignee or final disposition. As part of your track- ing system, you must establish and maintain a method for documenting the disposition of each of your HCT/Ps, to enable tracking from the donor to the consignee or final disposition. The information you maintain must permit the prompt identification of the con- signee of the HCT/P, if any.

#### 21 CFR Ch. I (4–1–17 Edition)

- Consignees. At or before the time of distribution of an HCT/P to a con- signee, you must inform the consignee in writing of the requirements in this section and  of  the  tracking  system that you have established and are maintaining to comply with these re- quirements.

- Requirements specific to dura mater donors. You must archive appropriate specimens from each donor of dura mater, under appropriate storage con- ditions, and for the appropriate dura- tion, to enable testing of the archived material for evidence of transmissible spongiform encephalopathy, and to en- able appropriate disposition of any af- fected nonadministered dura mater tis- sue, if necessary.

[69 FR 68681, Nov. 24, 2004, as amended at 70

FR 29952, May 25, 2005]

##### § 1271.320   Complaint file.

- Procedures. You must establish and maintain procedures for the re- view, evaluation, and documentation of complaints as defined in § 1271.3(aa), re- lating to core current good tissue prac- tice (CGTP) requirements, and the in- vestigation of complaints as appro- priate.

- Complaint file. You must maintain a record of complaints that you receive in a file designated for complaints. The complaint file must contain sufficient information about each complaint for proper review and evaluation of the complaint (including the distinct iden- tification code of the HCT/P that is the subject of the complaint) and for deter- mining whether the complaint is an isolated event or represents a trend. You must make the complaint file available for review and copying upon request from FDA.

- Review and evaluation of com- plaints. You must review and evaluate each complaint relating to core CGTP requirements to determine if  the com- plaint is related to an HCT/P deviation or to an adverse reaction, and to deter- mine if a report under § 1271.350 or an- other applicable regulation is required. As soon as practical, you must review, evaluate, and investigate each com- plaint that represents an event re- quired to be reported to FDA, as de- scribed in § 1271.350. You must review and  evaluate  a  complaint  relating  to

#### Food and Drug Administration, HHS	§ 1271.350

core CGTP requirements that does not represent an event required to be re- ported to determine whether an inves- tigation is necessary; an investigation may include referring a copy of the complaint to another establishment that performed manufacturing steps pertinent to the complaint. When no investigation is made, you must main- tain a record that includes the reason no investigation was made, and the name of the individual(s) responsible for the decision not to investigate.

## Subpart E—Additional Require- ments for Establishments De- scribed in § 1271.10

SOURCE:  69  FR  68686,  Nov.  24,  2004,  unless

otherwise noted.

##### § 1271.330   Applicability.

The provisions set forth in this sub- part are being implemented for non- reproductive     HCT/Ps     described     in

§ 1271.10 and regulated solely under sec- tion 361 of the Public Health Service Act and the regulations in this part, and for the establishments that manu- facture those HCT/Ps. HCT/Ps that are drugs or devices regulated under the act, or are biological products  regu- lated under section 351 of the Public Health Service Act, are not subject  to the regulations set forth in this sub- part.

##### § 1271.350   Reporting.

  - Adverse reaction reports. (1) You must investigate any adverse reaction involving a communicable disease re- lated to an HCT/P that you made avail- able for distribution. You  must  report to FDA an adverse reaction involving a communicable disease if it:

- Is fatal;

- Is life-threatening;

- Results in permanent impair- ment of a body function or permanent damage to body structure; or

- Necessitates medical or surgical intervention, including hospitalization.

- You must submit each report on a Form FDA–3500A to the address in paragraph (a)(5) of this section within

15 calendar days of initial receipt of the  information.

- You must, as soon as practical, in- vestigate all adverse reactions that are the subject of these 15-day reports and must submit followup reports within 15 calendar days of the receipt of new in- formation or as requested by FDA. If additional information is not obtain- able, a followup report may be required that describes briefly the  steps  taken to seek additional  information and the reasons why it could not be obtained.

- You may obtain copies of the re- porting form (FDA–3500A) from the Center for Biologics Evaluation and Research (see address in paragraph (a)(5) of this section). Electronic Form FDA–3500A may be obtained .

- You must submit two  copies  of each report described in  this  paragraph to the  Food  and  Drug  Administration, Center for Biologics Evaluation and Research, Document Control Center, 10903 New Hampshire Ave., Bldg. 71, Rm. G112, Silver Spring, MD 20993–0002. FDA  may  waive  the  requirement  for the second copy in appropriate  cir- cumstances.

  - Reports of HCT/P deviations. (1) You must investigate all HCT/P devi- ations related to a distributed HCT/P for which you performed a manufac- turing  step.

- You must report any such HCT/P deviation relating to the core CGTP re- quirements, if the HCT/P deviation oc- curred in your facility or in a facility that performed a manufacturing step for you under contract, agreement, or other arrangement. Each report must contain a description of the HCT/P de- viation, information relevant to the event and the manufacture of the HCT/ P involved, and information on all fol- low-up actions that have been or will be taken in response to the HCT/P devi- ation (e.g., recalls).

- You must report each such HCT/P deviation that relates to a core CGTP requirement on Form FDA 3486 within 45 days of the discovery of the event ei- ther electronically using the Center for Biologics Evaluation and Research electronic Web-based application or by mail to the Food and Drug Administra- tion, Center for Biologics  Evaluation and Research, Document Control Cen- ter,  10903  New  Hampshire  Ave.,  Bldg.

§ 1271.370

71,  Rm.  G112,  Silver  Spring,  MD  20993– 0002.

[69 FR 68686, Nov. 24, 2004, as amended at 80

FR 18095, Apr. 3, 2015]

##### § 1271.370   Labeling.

The following requirements apply in addition to §§ 1271.55, 1271.60, 1271.65,

and 1271.90:

  - You must label each HCT/P made available for distribution clearly and accurately.

  - The following information must appear on the HCT/P label:

    - Distinct identification code af- fixed to the HCT/P container, and as- signed in accordance with § 1271.290(c);

    - Description of the type of HCT/P;

    - Expiration date, if any; and

    - Warnings	required	under

§ 1271.60(d)(2),	§ 1271.65(b)(2),	or

§ 1271.90(c), if applicable and physically possible. If it is not physically possible to include these warnings on the label, the warnings must, instead, accompany the HCT/P.

  - The following information must either appear on the HCT/P label or ac- company the HCT/P:

- Name and address of the establish- ment that determines that the HCT/P meets release criteria and makes the HCT/P available for distribution;

- Storage temperature;

- Other warnings, where appro- priate; and

- Instructions for use when related to the prevention of the introduction, transmission, or spread of commu- nicable diseases.

[69 FR 68686, Nov. 24, 2004, as amended at 70

FR 29952,  May 25,  2005;  81  FR 40518,  June 22,

2016]

## Subpart F—Inspection and En- forcement of Establishments Described in § 1271.10

SOURCE:  69  FR  68687,  Nov.  24,  2004,  unless

otherwise noted.

##### § 1271.390   Applicability.

The provisions set forth in this sub- part are applicable only to HCT/Ps de- scribed in § 1271.10 and regulated solely under section 361 of the Public Health Service Act and the regulations in this part,   and   to   the   establishments   that

21 CFR Ch. I (4–1–17 Edition)

manufacture those  HCT/Ps.  HCT/Ps that are drugs or devices regulated under the act, or are biological prod- ucts regulated under section 351 of the Public Health Service Act, are not sub- ject to the regulations set forth in this subpart.

##### § 1271.400   Inspections.

- If you are an establishment that manufactures   HCT/Ps   described   in

§ 1271.10, whether or not under contract, you must permit the Food and Drug Administration (FDA) to inspect any manufacturing location at any reason- able time and in  a  reasonable  manner to determine compliance with applica- ble provisions of this part. The inspec- tion will be conducted as necessary  in the judgment of the FDA and may in- clude your establishment, facilities, equipment, finished and unfinished ma- terials, containers, processes, HCT/Ps, procedures, labeling, records, files, pa- pers, and controls required to be main- tained under the part. The inspection may be made with or without prior no- tification and will ordinarily be made during regular business hours.

- The frequency of inspection will be at the agency’s discretion.

- FDA will call upon the most re- sponsible person available at the time of the inspection of the establishment and may question the personnel of the establishment as necessary to deter- mine compliance with the provisions of this part.

- FDA’s representatives may take samples, may review and copy any records required to be kept under this part, and may use other appropriate means to record evidence of observa- tions during inspections conducted under this subpart.

- The public disclosure of records containing the name or other positive identification of donors or recipients of HCT/Ps will be handled in accordance with FDA’s procedures on disclosure of information as set forth in parts 20 and 21 of this chapter.

##### § 1271.420   HCT/Ps offered for import.

(a) Except as provided in paragraphs

- and (d) of this section, when an HCT/P is offered for import, the im- porter of record must notify, either be- fore or at the time of importation, the

#### Food and Drug Administration, HHS	§ 1271.440

director of the district of the Food and Drug Administration (FDA) having ju- risdiction over the port of entry through which the HCT/P is imported or offered for import, or such officer of the district as the director may des- ignate to act in his or her behalf in ad- ministering and enforcing this part, and must provide sufficient informa- tion, including information submitted in the Automated Commercial Envi- ronment (ACE) system or any other electronic data interchange system au- thorized by the U.S. Customs and Bor- der Protection Agency as required in part 1, subpart D  of this chapter,  for FDA to make an admissibility deci- sion.

  - Except as provided in  paragraphs

  - and (d) of this section, an HCT/P of- fered for import must be held intact by the importer or consignee, under condi- tions necessary to prevent trans- mission of communicable disease, until an admissibility decision is made by FDA. The HCT/P may be transported under quarantine to the consignee, while the FDA district reviews the doc- umentation accompanying the HCT/P. When FDA makes a decision regarding the admissibility of the HCT/P, FDA will notify the importer of record.

    - This section does not apply to re- productive HCT/Ps regulated solely under section 361 of the Public Health Service Act and the regulations in this part, and donated by a sexually inti- mate partner of the recipient for repro- ductive use.

    - This section does not apply to pe- ripheral blood stem/progenitor cells regulated solely under section 361 of the Public Health Service Act and the regulations in this part, except that paragraphs (a) and (b) of this section apply when circumstances occur under which such imported peripheral blood stem/progenitor cells may present an unreasonable risk of communicable disease transmission which  indicates the need to review the information ref- erenced in paragraph (a) of this sec- tion.

[69 FR 68687, Nov. 24, 2004, as amended at 81

FR 85873, Nov. 29, 2016]

##### § 1271.440 Orders of retention, recall, destruction, and cessation of manu- facturing.

- Upon an agency finding that there are reasonable grounds to believe that an HCT/P is a violative HCT/P because it was manufactured in violation of the regulations in this part and, therefore, the conditions of manufacture of the HCT/P do not provide adequate protec- tions against risks of communicable disease transmission; or the HCT/P is infected or contaminated so as to be a source of dangerous infection to hu- mans; or an establishment is in viola- tion of the regulations in this part and, therefore, does not provide adequate protections against the risks of com- municable disease  transmission,  the Food and Drug Administration (FDA) may take one or more of the following actions:

  - Serve upon the person who dis- tributed the HCT/P  a  written  order that the HCT/P be recalled and/or de- stroyed, as appropriate, and upon per- sons in possession of the HCT/P  that the HCT/P must be retained until it is recalled by the distributor,  destroyed, or disposed of as agreed by FDA, or the safety of the HCT/P is confirmed;

  - Take possession of and/or destroy the violative HCT/P; or

  - Serve upon the establishment an order to cease manufacturing until compliance with the regulations of this part has been achieved. When FDA de- termines there are reasonable grounds to believe there is a danger to health, such order will be effective imme- diately. In other situations, such order will be effective after one of the fol- lowing events, whichever is later:

    - Passage of 5 working days from the establishment’s receipt of  the order; or

    - If the establishment requests a hearing in accordance with paragraph

    - of this section and part 16 of this chapter, a decision in, and in accord- ance with, those proceedings.

      - A written order issued under para- graph (a) of this section will state with particularity the facts that justify the order.

#### § 1271.440

(c)(1) A written order issued under paragraph (a)(1) of this section will or- dinarily provide that the HCT/P be re- called and/or destroyed within 5 work- ing days from the date of receipt of the order. After  receipt of  an  order issued under paragraph (a)(1) of this section, the establishment in possession of the HCT/P must not distribute  or  dispose of the HCT/P in any manner except to recall and/or destroy the HCT/P con- sistent with  the  provisions  of  the order, under the supervision of FDA.

(2) In lieu of paragraph (c)(1) of this section, other arrangements for assur- ing the proper disposition of the HCT/P may be agreed upon by the person re- ceiving the written order and  FDA. Such  arrangements  may  include, among others, providing FDA with records or other written information that adequately ensure that the HCT/P has been recovered, processed, stored, and distributed in  conformance  with this part, and that, except as provided under §§ 1271.60, 1271.65, and 1271.90, the donor of the cells or tissue for the HCT/ P has been determined to be eligible.

        - A written order issued under para- graph (a)(3) of this section will specify the regulations with which you must achieve compliance and will ordinarily specify the particular operations cov- ered  by  the  order.  After  receipt  of  an

#### 21 CFR Ch. I (4–1–17 Edition)

order that is in effect and issued under paragraph (a)(3) of this section, you must not resume operations without prior written authorization of FDA.

        - The recipient of an order issued under this section may request a hear- ing in accordance with part 16 of this chapter. To request a hearing, the re- cipient of the written order or prior possessor of such HCT/P must make the request within 5 working days of receipt of a written order for retention, recall, destruction, and/or cessation (or within 5 working days of the agency’s possession of an HCT/P under para- graph (a)(2) of this section), in accord- ance with part 16 of this chapter. An order of  destruction  will be  held  in abeyance pending resolution of the hearing request. Upon request under part 16 of this chapter, FDA will pro- vide an opportunity for an expedited hearing for an  order  of  cessation  that is not stayed by the Commissioner of Food and Drugs.

        - FDA will not issue an order for the destruction of reproductive tissue under paragraph (a)(1) of this section, nor will it carry out such destruction itself under paragraph (a)(2) of this sec- tion.

## PARTS 1272–1299 [RESERVED]