# Department of Environmental Protection Process for Assessing Data Usability

DEP-EA 001/07

Florida Department of Environmental Protection 2600 Blair Stone Road

Tallahassee, FL 32399-2400

Water Quality Standards Program

Aquatic Ecology and Quality Assurance Section

Month Year

Process for Assessing Data Usability DEP-EA 001/07

Table of Contents

1.	Introduction	3

1.1.	The Data Assessment Process	3

1.2.	Evaluating DQI Failures.	3

2.	Data Quality Indicators (DQI)	3

3.	Laboratory Control Sample or Spike (LCS)	4

3.1.	    General Requirements	4

3.2.	Evaluation of LCS Recovery	5

4.	Matrix Spikes (MS)	5

4.1.	    General Requirements	6

4.2.	    Evaluation of MS Recovery	6

5.	Surrogate Spikes	7

6.	Duplicate Measurements: LCS Duplicates or Replicates (LCSD), Matrix Spike Duplicates (MSD), Matrix Duplicates (MD) and Sample Duplicates (SD)	7

6.1.	    General Requirements	7

6.2.	    Evaluation of Duplicates	8

7.	Calibrations	8

7.1.	    Initial Calibration	8

7.2.	    Calibration Verification (CV)	9

8.	Method Blanks or Other Analytical Blanks	10

9.	Field Quality Control Blanks (Trip Blanks, Field Blanks or Equipment Blanks)	11

10.	Holding Times	11

11.	Quality Control Check Samples	12

11.1.    Biochemical Oxygen Demand (BOD) Analyses	12

11.2.    Chlorophyll Analyses	12

11.3.	Matrix-Specific Quality Control Sample Evaluation.	13

12.	Sample Preservation and Filtration Checks	14

13.	Evaluation of the Reported MDL	14

14.	Evaluation of the Reported PQL	14

15.	Evaluation of Reversals (Parts vs. Whole Comparison)	14

16.	General Principles of Data Quality Assessment	15

17.	Procedure for Data Usability Determinations	16

18.	Summary of the Data Usability Assessment Process	17

Process for Assessing Data Usability DEP-EA 001/07

- Introduction - The intent of this document is to outline the process to be used by the Department when evaluating Data Quality Indicators (DQIs) and determining the usability of analytical data.

  - The Data Assessment Process - Determining if data are usable for a particular purpose is a complex task, requiring a reasonable and balanced evaluation of many factors. The procedural components of a usability assessment must be performed by auditors with sufficient scientific expertise in environmental data verification and validation, and include, but are not limited to:

- Understanding the purpose for auditing the data (project or program Data Quality Objectives [DQOs] provide the context for the audit);

- Identifying the set of data to be audited, the types of analytes or parameters in the data set and the reported values (e.g., concentration) for the analytical results for all components;

- Determining the relationship between each analytical result, the associated decision or action level (e.g., water quality standard or clean up target) and the laboratory’s stated practical quantitation limit (PQL);

- Evaluating the documented calibration, quality control and other supporting data against designated DQIs;

- Establishing the pattern, frequency, and magnitude of any failures or other deficiencies associated with the results;

- Determining the extent to which the audited data set fulfills the DQOs of the project or program;

- Evaluating corroborative data (e.g., performance tests, data from other laboratories);

- Providing usability recommendations to the program data users.

  - Evaluating DQI Failures - Exceeding the acceptance criteria for one or more DQIs does not necessarily mean the data are unusable. The factors mentioned in 1.1 above must be systematically evaluated before a usability decision can be made.

    - The purpose for which the analytical data were collected can vary widely and may include such diverse activities as: initial screening or scoping studies, permit compliance monitoring, assessing waters to determine if they meet water quality standards, developing Total Maximum Daily Loads, or determining whether a permitted waste facility has met “clean closure” site contamination assessment requirements.

    - A DQI failure that is acceptable for a screening study may not be acceptable for declaring a site free from contamination.

    - In practical terms, it is not possible in this document to discuss each potential scenario that might be associated with a data usability recommendation for a particular purpose. Therefore, the factors affecting this decision (outlined in 1.1 and 1.2, above) and the thought process applied to the data usability assessment are addressed by the general principles and examples listed below.

- Data Quality Indicators (DQI) - The following DQIs (Sections 3 – 16) are targets used in the context of assessing data usability. They evaluate bias, precision, accuracy, comparability, completeness, and representativeness. Not all DQIs are applicable to the specific assessment project, and DQIs differ among analytical groups.

Example DQI: To be considered usable, a dataset submittal to DEP consisting only of high sample concentrations (with no values near the PQL) need not be associated with records demonstrating evaluation or verification of the PQL.

  - The application of DQIs to sample results presumes the random occurrence of non- routine QC failures to be expected with all analyses. The frequency of DQI failures is considered in the overall data usability assessment in order to determine those instances where routine or systematic failures indicate a significant data usability problem.

  - Use of DQIs with Other Published Criteria - Except as discussed otherwise in this document, sample results are evaluated for usability based on the procedural requirements and performance criteria established by the reported analytical methods, applicable project DQOs, applicable regulations, applicable The NELAC Institute (TNI) Quality Systems standards EL-V1-ISO-2016-Rev. 2.1, Environmental Laboratory Sector, Vol. 1, Module 1 – 7 Management and Technical Requirements for Laboratories Performing Environmental Analysis (2016), hereafter 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C., and applicable DEP SOPs for field and laboratory activities (DEP-SOP-001/01, DEP-SOP-002/01, and DEP-SOP-003/11, incorporated by reference in 62-160.800(1), F.A.C.).

  - Use of Records to Assess Data Usability - As applicable to the data usability assessment process, any record associated with a reported sample result or set of sample results may be audited, per Rules 62-160.240 & .340, F.A.C. Both original (“raw”) and reduced or summarized versions of data records are inspected in order to determine acceptance of the procedures and performance criteria that are used to generate and evaluate the sample data and associated quality control activities.

  - The presence or absence of critical, archival records that support the sample results is considered when determining whether sufficient documentation is available to assess the usability of the data.

- Laboratory Control Sample or Spike (LCS) – An LCS is used to evaluate the accuracy (analyte percent recovery) of the method without the effects that a sample matrix might contribute. It is associated with a set of up to 20 environmental samples of the same quality system matrix and is used to evaluate the overall performance of the test procedure. The LCS must be a discrete sample with each batch of 20. Every preparation batch must include an LCS and when possible, a matrix spike (MS). A method may require analysis of both in each batch. They each provide important information regarding the analytical process and the sample matrix impact on analysis.

# General Requirements

    - An LCS result is linked or associated with all applicable reported sample results for the same preparation batch of 20 or fewer environmental samples of the same quality system matrix. An LCS is not required for microbiological analysis or other analyses for which no spiking solution is available.

    - The LCS contains all of the analytes of interest for the method or project, or the mix of analytes in the LCS is rotated at a routine frequency as suggested in Module 4 of the 2016 TNI Standard.

    - If a preparation step (digestion or extraction) is required for the analysis, the LCS is subjected to the same preparation steps as other samples in the preparation batch.

    - If no preparation is required, a separate LCS must be analyzed for each set of 20 or fewer environmental samples of the same quality system matrix.

    - The level of the spike in the LCS should be above the PQL but less than or equal to the concentration of the mid-level standard.  Ideally the concentration in the LCS should represent the anticipated concentration in the associated samples set.

    - The LCS is analyzed or reanalyzed within the analytical sequence (the set of samples that are chronologically bracketed by verifications)

    - All reanalyses are documented with a valid explanation of the QC failure that prompted the reanalyses.

    - The concentration of an LCS must fall within the calibration range for the instrument as established by the concentrations of the standards used for the initial calibration and continuing calibration verifications.

    - Where applicable, the LCS is prepared at the method-specified concentration or at a concentration appropriate for the project DQO.

    - An LCS is not required for microbiological analysis or other analyses for which no spiking solution is available.

  - Evaluation of LCS Recovery –

    - The LCS result is compared with the expected spike concentration to calculate the percent recovery. See Appendix A for the percent recovery calculation.

    - Where applicable, the control limits for the LCS are established by the laboratory for a specified matrix at ± 3X the standard deviation of the mean recovery of the cumulative LCS analyses at a specified concentration. The mean recovery is calculated after outliers are eliminated from the recovery data set. Outliers may be identified using an appropriate statistical test. If trend monitoring of individual LCS recoveries indicates the presence of a systematic bias determine possible reasons for the bias using root cause analysis and implement applicable corrective actions.  Recalculate the mean recovery if warranted.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - Where applicable, criteria established by rule, reported method or project DQOs are used to evaluate the LCS recoveries.

    - If no other criteria are applicable, the following control limits are used to evaluate LCS recoveries:

80% - 120% (water) for Nitrate, Nitrite, Nitrate + Nitrite, Ammonium, TKN, Total Phosphorus, Orthophosphate

85%-115% (water) for Metals

Laboratory Established Control Limits for:

All organic analytes for the LCS in any matrix

All other analytes amenable to spiking for the LCS in water

All analytes amenable to spiking for the LCS in non-aqueous matrices

    - When applicable, sample results reported with failed LCS recoveries are qualified with the “J” data qualifier and accompanied with a comment explaining the reason for the qualifier and an estimation of the bias (e.g. if an LCS failed high, the reported sample concentrations are expected to be less than the reported value).  Analytes or methods technically not amenable to spiking for recovery determinations do not have LCSs and are therefore not evaluated.

- Matrix Spikes (MS) – A MS is a sample that is spiked or fortified with known amounts of target analytes and then taken through all sample preparation and analytical steps of the procedure. The MS result is compared to the known value of the spike to calculate percent recovery. See Appendix A for the percent recovery calculation. Method-specified criteria are used to evaluate MS recoveries, where applicable. If an LCS was not analyzed for the sample preparation batch being evaluated, the MS data may be used to evaluate batch  control for all associated samples in the batch provided the acceptance criteria for the LCS are used to evaluate the results. A MS is not conducted for analyses for which spiking solutions are not available.

# General Requirements

    - The MS contains all of the analytes of interest for the method or project, or the mix of analytes in the MS is rotated at a routine frequency as suggested in Module 4 of the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C..

    - If a preparation step (digestion or extraction) is required for the analysis, the MS is subjected to the same preparation steps as other samples in the preparation batch.

    - The MS is analyzed or reanalyzed within a verified calibration bracket (analytical sequence).

    - All reanalyses are documented with a valid explanation of the QC failure that prompted the reanalyses.

    - The concentration of an MS is required to fall within the calibration range for the instrument as established by the concentrations of the standards used for the initial calibration and continuing calibration verifications.

    - Where applicable, the MS is spiked at the method-specified concentration or at a concentration appropriate for the project DQO. Spike levels for a matrix spike may be specified by the method and may range from three to ten times the reporting limit or PQL. If the laboratory can anticipate expected environmental sample concentrations, spike levels should be higher than target analyte concentrations in the spiked sample. For samples that typically have low concentrations of target analytes (ex. trace metals, pesticides, SVOCs in water), the spike levels are typically three to four times the PQL. For other target analytes where concentrations in environmental samples can be relatively high compared to the MDL or PQL (e.g., TKN, total phosphorus) the spike levels can be as high as ten times the PQL.

# Evaluation of MS Recovery

    - The MS spike concentration is compared with the un-spiked (parent) sample concentration. The suitability of the spike concentration is evaluated on a case-by-case basis according to method specifications or project DQOs.

    - MS recovery data is evaluated if specified by applicable DQOs, but the recovery of the MS is only associated with the parent sample unless no LCS is associated with the preparation batch of 20 environmental samples or less. In this case, the MS recovery is associated with all of the samples in the preparation batch of 20 environmental samples or less for evaluation of the acceptability of the sample batch.

    - Where applicable, the control limits for the MS are established by the laboratory for a specified matrix at ± 3X the standard deviation of the mean recovery of the cumulative MS analyses for the associated matrix at a specified concentration. The mean recovery is calculated after outliers are eliminated from the recovery data set. Outliers may be identified using an appropriate statistical test. If trend monitoring of individual MS recoveries indicates the presence of a systematic bias, determine possible reasons for the bias using root cause analysis and implement applicable corrective actions.  Recalculate the mean recovery if warranted.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - If the MS is used to evaluate all associated batch samples, the applicable LCS criteria are used to evaluate the MS recovery.

    - Where applicable, criteria established by rule, reported method or project DQOs are used to evaluate the MS recoveries.

    - If no other criteria are applicable, the following control limits are used to evaluate the MS recoveries:

80% - 120% (water) for Nitrate, Nitrite, Nitrate + Nitrite, Ammonium, TKN, Total Phosphorus, Orthophosphate

70%-130% (water) for Metals

Laboratory Established Control Limits for:

All organic analytes at any concentration

All other analytes amenable to spiking at any concentration All analytes in non-aqueous matrices

    - When applicable, the parent sample result of the failed matrix spike is qualified with a “J” and a comment explaining the error that occurred. MS recoveries may fail due to matrix interference in the sample’s matrix.

- Surrogate Spikes – Surrogate spikes are samples fortified at known concentration(s) with one or more compounds having similar chemical characteristics to the target analytes, but which are not normally found in environmental samples. To assure this, deuterated analogs of some of the target analytes are frequently selected as surrogates. Surrogates are typically spiked at three to five times the reporting limit (PQL) and their recoveries give an indication of method accuracy and are a good check for gross inaccuracy (i.e., incorrect volumes, concentration problems, etc.) in all the samples. Surrogate spikes are evaluated according to method-specified requirements.

  - The concentration of a surrogate spike must fall within the initial calibration range established for the instrument.

  - When applicable, sample results reported with failed surrogate spike recoveries are qualified with a “J” and a comment explaining the error that occurred.

  - Analytes or methods technically not amenable to spiking for recovery determinations are not evaluated.

- Duplicate Measurements: LCS Duplicates or Replicates (LCSD), Matrix Spike Duplicates (MSD) and Matrix Duplicates (MD) - in order to assess the true method precision in the presence of matrix effects, MS/MSD are preferred. The sample duplicate (SD) is also evaluated for precision.

# General Requirements

    - When applicable, replicate rather than duplicate LCSD/MSD/MD data may be evaluated for precision. The analysis of duplicate LCSs is routinely required for some organic analyses such as volatiles (VOCs).

    - An LCSD result is linked or associated with all applicable reported sample results for the same preparation batch or of 20 environmental samples or less and represents the precision of the method.  The MSD/MD is linked only to the parent sample.

    - The LCSD/MSD must be spiked with the same analytes as the LCS/MS at the same concentrations. For some analyses, the spiked analytes can be rotated at a prescribed frequency defined in the method.

    - If a preparation step (digestion or extraction) is required for the analysis, the LCSD/MSD/MD are subjected to the same preparation steps as other samples in the preparation batch.

    - The LCSD/MSD/MD is analyzed or reanalyzed within a verified calibration bracket.

    - All reanalyses are documented with a valid explanation of the QC failure that prompted the reanalyses.

    - The final spiked concentration of the LCSD/MSD should be within the calibration range for the instrument as established by the concentrations of the standards used for the initial calibration.

    - Where applicable to the evaluation, the MSD spike concentration is compared with the un-spiked, parent-sample concentration. The suitability of the spike concentration is evaluated on a case-by-case basis according to method specifications or project DQOs.

# Evaluation of Duplicates

    - The LCS/MS/S result is typically compared with the corresponding LCSD/MSD/SD result with the Relative Percent Difference (RPD) or Percent Relative Standard Deviation (RSD), calculated as described in Appendix A.

    - Where applicable, the control limits for the duplicates are established by the laboratory for a specified matrix at ≤3X the standard deviation of the mean precision (absolute or relative) of the cumulative duplicate pair analysis for the associated matrix at a specified concentration. The mean precision is calculated, after outliers are eliminated from the precision data set. Outliers may be identified using an appropriate statistical test.  If trend monitoring for individual LCSD/MSD recoveries indicates the presence of a systematic bias, determine the reason for the bias and implement applicable corrective action.  Recalculate the mean precision if warranted. Duplicate recoveries should be consistent with sections 3 for LCS and 4 for MS recovery.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - Where applicable, criteria established by rule, reported method or project DQOs are used to evaluate duplicate precision.

    - If no other criteria are applicable, precision data for the duplicates are evaluated using the laboratory control limits, with a target precision of ≤20% Relative Percent Difference (RPD) or Percent Relative Standard Deviation (RSD) for water samples and ≤40% RPD or RSD for non-aqueous and solid-matrix samples

    - When applicable, sample results reported with an LCSD precision failure are qualified with a “J” and a comment explaining the error that occurred. Only the sample associated with MSD/SD precision failure requires qualification.

    - Analytes or methods technically not amenable to precision determinations are not evaluated.

- Calibrations - Except as discussed elsewhere in this document, all calibrations are evaluated according to method-specified requirements and those of the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C.

# Initial Calibration

    - The number of standard concentrations used to perform the initial calibration specified by the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C., should be followed unless the method or project DQOs require or allow a different number.

    - Where applicable, the concentration values specified by the method or project DQOs are used to perform the initial calibration of the instrument or technique.

    - If not specified by the method or project objectives and when applicable to the analytical technology, at least five standard concentrations and a blank are used for a linear calibration curve.

    - The acceptance of the initial calibration is evaluated using the method-specified criteria or the criteria specified by the project DQOs.

    - If not specified by the method or project DQOs and when applicable to the analytical technology, the following expectations could apply to data use evaluations:

      - The correlation coefficient for a regression must be ≥0.995 when applied to the linear calibration curve. The lowest non-zero calibration standard is typically at the reporting limit or PQL.

      - Higher-order calibration curve regressions should have a coefficient of determination of ≥0.99, using a minimum of six non-zero calibration points for a second order curve and seven non-zero calibration points for a third order curve. The calibration points used for the curve may include a calibration blank. Thus, a second order (quadratic) curve would have a minimum of seven calibration points, one of which is the blank.

      - Calibrations must include evaluating relative error and applicable acceptance criteria. Relative error is a measure of the difference between the nominal and fitted values of the calibration standards. For calibrations evaluated using an average response factor, the determination of the RSD of response factors is the measure of the relative error. Details may be found in Section 1.7, Module 4 of the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C.

      - The laboratory must have a written standard operating procedure (SOP) for dropping calibration points from the initial calibration curve. The SOP must be consistent with the requirements in Module 4 of the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C.  The laboratory must be able to justify dropping calibration points, and must ensure that the required number of calibration points remain after the removal of calibration point(s).  In some cases, the MDL and PQL or the upper range may need to be adjusted.

    - The initial calibration range can be extended by an accepted calibration verification of a standard that is above or below the highest or lowest standard, respectively.  This verification must be made before another initial calibration is performed.

    - Applicable sample results whose analyzed concentrations fall outside of the calibration range established by the standard concentrations used for the initial calibration (excluding the calibration blank) and calibration verifications are not usable without further evaluation or qualification. Samples with concentrations higher than the highest calibration standard must be diluted and reanalyzed.

# Calibration Verification (CV)

    - All Initial Calibrations must be verified using a verification standard from a second independent source as described in Module 4 of the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C. Subsequent CVs utilized in analytical runs may or may not be made up from a second source.

    - An acceptable calibration verification is analyzed at the beginning and end of the analytical batch (verified calibration bracket). CVs are typically spiked at levels close to the middle of the calibration range.

    - As applicable or when specified by rule or project DQOs, the calibration verification standard is analyzed at the method-specified frequency. Unless otherwise specified in the method, a CV must be prepared and analyzed at the beginning and end of the analytical batch. For some organic methods it must also be run every 12 hours of analysis time. Some methods require the analysis of a CV every 10 samples.

    - As applicable or when specified by rule or project DQOs, the calibration verification result is evaluated using the method-specified criteria.

    - In the absence of acceptance criteria specified by rule, project DQOs or method, and where applicable, the control limits for the CV are established by the laboratory for a specified matrix at ± 3X the standard deviation of the mean recovery of the cumulative CV analyses at a specified concentration. For this purpose, recovery is defined as the result obtained for the standard divided by the expected concentration, expressed as a percent. The mean recovery is calculated after outliers are eliminated from the recovery data set. Outliers may be identified using an appropriate statistical test. If trend monitoring of individual CV recoveries indicates the presence of a systematic bias, perform a root cause analysis to determine possible reasons for the bias and implement applicable corrective actions.  Recalculate the mean recovery if warranted.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - If a continuing calibration verification did not meet the acceptance criterion, but is immediately reanalyzed with acceptable results, the samples analyzed before the failed verification result are valid, provided that the verification previous to the failed verification and following the reanalyzed verification are also acceptable and comprise a verified calibration bracket for the affected samples.

    - Sample results that are not analyzed within a verified analytical batch or calibration bracket must be reanalyzed or qualified with a “J” as estimated values.

    - Alternatively, for purposes of assessing the usability of the sample data where verifications have failed and where the reported sample values are relevant to the regulatory action level, compliance limit or other project DQO, the failed continuing calibration verifications are evaluated for potential high or low bias according to the applicable requirements for calibration verifications in the 2016 TNI Standard, incorporated by reference in paragraph 62-160.800(3)(b), F.A.C.  (Quality Systems General Requirements).

- Method Blanks or Other Analytical Blanks - All method blanks and other types of analytical blanks used to assess potential analytical contamination are evaluated for the presence of the analytes of interest for the project or method.

  - A blank result is linked or associated with all applicable reported sample results that are associated with the blank.

  - If a preparation step (such as digestion or extraction) is required for the analysis, a method blank must be subjected to the same preparation steps as other samples in the preparation batch.

  - The blank is analyzed or reanalyzed within a verified calibration bracket.

  - All reanalyses are documented with a valid explanation of the error that prompted the reanalyses.

  - Unless specified otherwise by rule, an applicable analytical method or project DQOs, all analytes in a blank should be less than the applicable method detection limit.

  - A method blank is acceptable if it is less than the MDL or if the concentration of analytes detected in the blank (above the reported MDL) is less than 10% of the analyte concentrations in the individual samples of the associated preparation batch (for the affected analytes). If the method blank concentration (above the reported MDL) is greater than 10% of the concentration in the associated samples, those sample results must include a “V” qualifier.

  - BOD method blanks are evaluated against the method criterion.

  - Bacteriological method blanks should have no colonies. Similarly the Most Probable Number (MPN) type methods blanks should be less than 1.0.

- Field Quality Control Blanks (Trip Blanks, Field Blanks or Equipment Blanks) - Field QC Blanks associated with the samples for a specific sampling event are evaluated for contamination as indicated by the presence of the analytes of interest for the project or method.

  - If a preparation step (such as digestion or extraction) is required for the analysis method, the field-generated blank is subjected to the same preparation steps as other samples in the preparation batch.

  - The blank is analyzed or reanalyzed within a verified calibration bracket

  - Unless specified otherwise by rule, an applicable analytical method or project DQOs, it is expected that the concentration of analytes detected in the field blank (above the reported MDL) will be less than 10% of the concentration in the associated field sample (for the affected analytes). If the field QC blank concentration (above the reported MDL) is greater than 10% of the concentration in the associated field samples, those associated sample results must include a “G” qualifier.

- Holding Times – Maximum holding times are evaluated against the times specified in the tables in the DEP SOP FS 1000. Expectations for reporting of holding times are described below. If the maximum holding time is exceeded, results must include a “Q” qualifier. Evaluate usability of these qualified data based on the magnitude of the exceedance relative to the maximum holding time.

  - Sample collection time represents the beginning of the holding time. For those analytes where there is no separate holding time specified for the extract, digestate or other processed sample, the holding period ends when the sample processing begins. Sample collection times, preparation times and analysis times should be recorded with a resolution of 1 minute. Differences between any two times must therefore be evaluated to the minute. Example: A sample is collected on 2/28/2024 at 9:30 am. It is prepared on 3/02/2024 at 10:25 am and analyzed on 3/03/2024 at 1:52 pm (13:52).

  - If holding times are specified in “hours,” each hour of specified holding time is 60 minutes. Calculate the maximum holding time for the sample using 1-minute increments for each hour the sample may be held past the sample collection time before processing must begin, accounting for the time zone in which the sample was collected. For all DEP programs, holding times specified in hours are evaluated to the minute. Example: The analytical result for a sample with an 8-hour holding time collected at 9:20 a.m. must be qualified if processing begins after 5:20 p.m. (17:20).

  - If holding times are specified in “days,” calculate the maximum holding time for the sample using the 24-hour increment (24 60-minutes increments) for each day the sample may be held past the sample collection time before processing must begin, accounting for the time zone in which the sample was collected. For some DEP programs, holding times specified in days are evaluated to the minute. Example: The analytical result for a sample with a 7-day holding time collected at 9:12 am on April 1st must be qualified if processing begins after 9:12 am on April 8th, the seventh day after collection.

  - Clarifications

    - For microbiology, two holding times are applicable. The preparation holding time is evaluated as the duration between the sample collection date and time and the date and time the processed sample begins incubation. The analysis date/time is the time incubation ends and must be the same time the sample was read by the analyst.

    - For BOD or CBOD, the holding time is evaluated as the duration between the sample collection date and time and the date and time of the initial DO measurement for the test.

    - For toxicity testing, the holding time is evaluated as the duration between the sample collection date and time and the date and time of introduction of the last individual test organism into a test sample.

- Quality Control Check Samples – The following criteria are used to evaluate specific quality control check samples for the indicated test.

# Biochemical Oxygen Demand (BOD) Analyses

    - For each preparation batch of 20 samples or less, glucose-glutamic acid (GGA) samples are analyzed to obtain values within the method-specified acceptance control limits.

    - Alternatively, per approved revisions to SM 5210 B, other compounds are used to control the method, when applicable to a specific wastewater effluent.

    - In order of preference, replicate check sample data are evaluated rather than sample replicate data, if both types of data are generated for the associated samples.

    - Precision data for the replicate check sample analyses are evaluated using the laboratory control limits, with a target precision of ≤20% RPD or RSD expected.

    - Where applicable, the control limits for the replicate check sample analyses are established by the laboratory at ≤3X the standard deviation of the mean precision (absolute or relative) of the cumulative check sample analyses. The mean precision is calculated after outliers are eliminated from the precision data set. Outliers may be identified using an appropriate statistical test. The mean precision is recalculated if trend monitoring of individual check sample results indicates the presence of a systematic error that requires corrective action.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - When applicable, sample results reported with failed check sample precision are qualified with a “J” and a comment explaining the error that occurred.

# Chlorophyll Analyses

    - For each preparation batch of 20 samples or less, or at least once every three months, a chlorophyll extract solution in acetone is analyzed to evaluate recovery at the concentration for the extract.

    - A vendor-assayed source of chlorophyll is used to prepare the check sample extract.

    - The concentration of the check sample extract is such that when extrapolated to the volume of a nominal whole-water concentration, the extract appropriately represents the estimated reporting limit for the whole water sample.

    - The check-sample results at the extract concentration are evaluated using control limits established by the laboratory.

    - The control limits for the extract check-sample solution are established by the laboratory at ± 3X the standard deviation of the mean result of the cumulative check- sample analyses expressed as a percentage of the expected value of the check-sample concentration. The mean result is calculated after outliers are eliminated from the check-sample data set. Outliers may be identified using an appropriate statistical test. If trend monitoring of individual check-sample results indicates the presence of a systematic bias, determine the reason for the bias and implement applicable corrective actions.  Recalculate the mean if warranted.

    - Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

  - Matrix-Specific Quality Control Sample Evaluation – A quality control sample may be prepared or purchased to evaluate interferences due to the sample matrix, such as the analyses for trace metal concentrations in seawater.

    - When applicable to the project DQOs and when technically feasible, known or suspected interferences in the matrix specific to the project analytes and samples are evaluated using quality control (QC) samples of known matrix composition and analyte concentration.

    - The QC sample result is typically linked or associated with all applicable reported field sample results for the same preparation batch or analytical sequence of 20 samples or less.

    - The composition of the matrix for the QC sample is controlled for the analytes and interferences of interest and closely matches or approximates the matrix of the associated field samples.

    - A naturally derived or artificially formulated matrix is used to prepare the QC sample.

    - The QC sample is prepared by the analyzing laboratory using a second source standard for the analyte of interest. The second-source standard is selected as indicated in TNI Quality Systems Standard for initial calibrations.

    - Alternatively, the QC sample is obtained from an external source such as a commercial vendor of QC-check samples or Certified Reference Materials.

    - Where applicable, the QC sample is prepared at the method-specified concentration or at a concentration selected according to project DQOs.

    - Where applicable, the concentration of the QC sample is evaluated against any DQOs for minimum quantitation level applicable to the data use and does not exceed 2X the minimum quantitation level.

    - If a preparation step (digestion or extraction) is required for the analysis, the QC sample is subjected to the same preparation steps as other samples in the preparation batch.

    - The QC sample is analyzed or reanalyzed within a verified calibration bracket.

    - All reanalyses are documented with a valid explanation of the QC failure that prompted the reanalyses.

    - The concentration of the QC sample is analyzed to be within the calibration range for the instrument as established by the concentrations of the standards used for the initial calibration and continuing calibration verifications.

    - As applicable, the recovery for the QC sample is evaluated using method- specified criteria, vendor-specified criteria, or control limits established by the laboratory.

    - Where applicable, the control limits for the QC sample are established by the laboratory for a specified matrix at ± 3X the standard deviation of the mean recovery of the cumulative QC sample analyses at a specified concentration. The mean recovery is calculated after outliers are eliminated from the recovery data set. Outliers may be identified using an appropriate statistical test. If trend monitoring of individual QC sample recoveries indicates the presence of a systematic bias, determine possible reasons for the bias using root cause analysis and implement applicable corrective action.  Recalculate the mean precision if warranted. Alternatively, control limits are established by the laboratory using other technically justifiable and scientifically sound procedures in accordance with project DQOs.

    - When applicable, sample results reported with failed QC sample recoveries are qualified with a “J” and a comment explaining the QC failure that occurred.

# Sample Preservation and Filtration Checks

  - The field records demonstrate positive indication that the samples are properly preserved and/or filtered (including thermal preservation).

  - The laboratory records demonstrate that proper thermal preservation was checked upon receipt at the laboratory, per the requirements in the TNI Quality Systems Standards.

  - The laboratory records demonstrate that samples that have been collected on the same day and hand-delivered are received at the laboratory in ice if the temperature check fails the applicable rule or method requirement.

  - When required, the laboratory records demonstrate positive indication that the samples are checked for proper chemical preservation prior to or during sample preparation or analysis.

# Evaluation of the Reported MDL

  - If no analyte is detected in the evaluated sample, the assessment of sample data usability includes establishing that the laboratory determined, evaluated and verified the reported MDL for the analyte according to the requirements in the applicable TNI Quality Systems Standards, regulatory requirement or reported method.

  - The MDL determination, evaluation and verification are considered when establishing the usability of the sample data where the reported MDL is relevant to the regulatory action level, compliance limit or other project DQO.

  - For the purposes of data usability evaluation, the DEP-defined MDL is equivalent to the TNI-defined Limit of Detection (LOD).

  - Data users should consult established target MDLs and PQLs developed by DEP for certain program rules (e.g., Rules 62-4, 62-777, & 62-550, F.A.C.), as applicable.

  - The concept of a statistically derived MDL is not appropriate for some analytes (see Appendix B-c.), but DEP still requires a reporting limit in those cases.

# Evaluation of the Reported PQL

  - If the concentration of the evaluated sample is below the concentration value of the reported PQL or is below the concentration value of the lowest initial calibration standard or continuing calibration verification standard associated with the sample, the assessment of sample data usability includes establishing that the laboratory determined, evaluated and verified the reported PQL for the analyte according to the requirements in the applicable TNI Quality Systems Standards.

  - The PQL determination, evaluation and verification are considered when establishing the usability of the sample data where the reported PQL is relevant to the regulatory action level, compliance limit or other project DQO.

  - For the purposes of data usability evaluation, the DEP-defined PQL is equivalent to the TNI-defined Limit of Quantitation (LOQ).

  - Note that there is greater uncertainty when results are between the PQL and MDL. Results are qualified with an “I” data qualifier.

# Evaluation of Reversals (Parts vs. Whole Comparison)

  - Where applicable, sample results are evaluated to determine if the sum of reported parts or fractions for the associated sample analyte results exceed 120% of the corresponding reported or calculated whole.

  - Evaluation of reversals with respect to sample data usability is assessed on a case-by-case basis with more extensive validation of the specifically affected analyses conducted as needed.

Examples of reversals include the following:

| Parameter Part | Parameter Whole |
| --- | --- |
| Total ammonia | TKN |
| Orthophosphate | Total phosphorus |
| Nitrate | Measured Total nitrite/nitrate |
| Nitrite | Measured Total nitrite/nitrate |
| Sum of nitrite and nitrate | Measured Total nitrite/nitrate |
| Sum of nitrite, nitrate and TKN | Measured Total nitrogen |
| Sum of measured total nitrite/nitrate and TKN | Measured Total nitrogen |
| Filtered sample results (e.g., dissolved metals) | Unfiltered sample results (e.g., total metals) |
| Methyl mercury | Total mercury |
| DOC | TOC |
| Chlorophyll a, corrected for phaeophytin | Chlorophyll a, not corrected |

- General Principles of Data Quality Assessment - The following principles will be used when evaluating sample data for data usability determinations using specified DQIs. The guidelines discussed in this section are only applicable if specific data quality assessment directives have not been provided in any other Department Rule or reporting format. The examples given below are illustrative and simplistic in nature.

  - Supporting data that fail the target acceptance criteria for specific DQIs will be evaluated against the affected sample result and the magnitude of the failure.

Example:

Sample result (exceeding the water quality standard by nearly 209 ug/L), and the fact that it exceeded the high-level standard by only 5%, the data is usable for the project.

A wastewater effluent sample, which was collected to determine compliance with a 10 ug/L arsenic permit limit, was found to have 9.5 ug/L of arsenic.

While the Laboratory Control Sample acceptance criteria previously established by the laboratory was 75 – 125%, the LCS recovery associated ith this sample was only 65%. These data are unusable to demonstrate compliance, due to the close proximity of the sample to the action level, and the low LCS recovery (well outside of the targeted acceptance criteria)

  - The magnitude of the sample result is considered when evaluating the consequences of the failed or absent DQIs.

Example 1: An Everglades surface water sample, analyzed for compliance with a 10 ug/L total phosphorus criterion, was found to contain 21 ug/L of total phosphorus. The sample was properly preserved, but due to a laboratory mistake, the original sample had to be re-analyzed. The second analysis produced a total phosphorus result of 17 ug/L, but now the holding time was exceeded by 1 week. The value was qualified by the laboratory with a “Q” for exceeding the holding time. Despite a reduction in the phosphorus concentration that could have been associated with exceeding the holding time, the analytical result still exceeded the action level, and the datum is usable.

For a waste facility study, benzene was found in groundwater at a concentration of 5.7 ug/L. However, the reported method detection limit (MDL) for this particular benzene analysis was 1 ug/L, with PQL of 4 ug/L. The Chapter 62-777, F.A.C., guidance document for routinely achievable PQLs indicates the target PQL for benzene is 1 ug/L. In this case, despite the elevated detection limit (indicating use of a less sensitive method), the sample result was above the reported PQL, and this result is usable.

Example 2: A surface water sample was analyzed by membrane filtration for fecal coliform, to determine if the 800 Colonies/100 mL water quality standard was exceeded. The analytical result was calculated to be 1,200 Colonies/100 mL. Method blanks were run simultaneously with the surface water samples by filtering 100 mL of sterilized dilution water, which resulted in the presence of 2 – 5 target colonies in the blank samples. The laboratory did not use “V” (analyte detected in blank) to qualify the result. This result is usable despite the failure of the lab to properly qualify the samples, the high magnitude of the result exceedance, compared with the minor contamination in the blanks.

Example 3: Biochemical Oxygen Demand analysis conducted on a surface water sample yielded a result of 5.5 mg/L. Dissolved oxygen depletion in the associated method blank was 0.3 mg/L, exceeding method criteria by 0.1 mg/L. The results were properly qualified with “V”. The data is usable based on the magnitude of the exceedance compared with the quality control failure.

  - The evaluation of laboratory performance or the determination of usability for a data set will be based on a preponderance of Data Quality Indicator failures pointing to specific, systematic problems with the laboratory operation or the data set.

Example scenario 1: The analytical data generated by a County surface water monitoring program were being audited to determine if the data could support Impaired Waters Rule listing decisions for the Total Maximum Daily Load (TMDL) program. Fifteen chlorophyll a samples, which had been analyzed over a 5 year period, were evaluated. Twice in this five-year period, due to laboratory oversights, the 48 hour holding time for filtration was not met. Records demonstrated that the 48 hour holding time was the laboratory’s target and that a quality system for meeting this objective was evident. Upon reviewing additional chlorophyll a quality control records, it was determined these holding time exceedances were isolated instances and that the five year period of data is usable.

Example scenario 2: Upon auditing the laboratory’s quality control information associated with data submitted to the TMDL program, it was discovered that the majority of samples were analyzed beyond the accepted holding time, and that none of these results were qualified with a “Q.” When the laboratory manager was asked about the discrepancy, they indicated that the qualifiers were purposely suppressed so that the TMDL program would use all the data. Because these admissions indicated that the entire quality system was compromised the data were deemed not usable.

  - Project management goals are taken into consideration when making data usability assessments. Specific DQOs are established for certain analytical activities within a project, and these objectives will be evaluated when making a usability statement.

# Procedure for Data Usability Determinations

  - Data auditors will review and evaluate the following information:

- The purpose for auditing the data;

- The reported values of the analytical results;

- The relationship between each analytical result and an associated decision or action level;

- The documented calibration, quality control and other supporting data compared with designated DQIs in Chapter 62-160, F.A.C., or other specified DQOs;

- The pattern, frequency, and magnitude of any failures or other deficiencies associated with the results;

- The extent to which the audited data set fulfills the DQOs of the project or program.

  - Based on the evaluation, the auditors will determine how the data can be used by the relevant Department programs.

Example 1: A cleanup goal established for a waste remediation site stated that all analytical work must be sensitive enough to meet the Cleanup Target Levels (CTL) in Rule 62 – 777 FAC. During the project, a previously unknown area of contamination was discovered, and the Project Manager ordered some grab samples to delineate the aerial extent and magnitude of the newly found contamination. The lab used a method with a higher PQL (that did not meet the original cleanup objectives) but use of this method allowed a rapid turn- around time, enabling an effective adaptive management approach to better address the newly discovered contamination. These data are usable based on this set of objectives (delineating the new area of concern).

Example 2: Four monitoring wells were sampled for trichloroethane (TCE) at a groundwater remediation site previously known to have TCE concentrations ranging from 20-30 ug/L. During this particular sampling, TCE measured at the four wells was found to be below 1 ug/L, and there were no laboratory quality control failures of any kind. However, sampling records indicated that large volumes of groundwater were purged with a centrifugal pump in a short time period, and that dissolved oxygen levels in the wells exceed 7 mg/L. Because the evidence indicates that improper sampling occurred, resulting in excessive aeration and de-gassing of volatile compounds, these data were determined to be unusable

# Summary of the Data Usability Assessment Process

Based on the above inputs, the auditor will extrapolate audit findings to determine the overall performance of a laboratory for a period of record and determine the usability of the data in question for a Department purpose. Usability assessments will evaluate the Data Quality Indicator results for the subject data set relative to DEP program or project objectives, and follow the principles characterized in this guidance document to draw an “overall conclusion” concerning the usability of the data set. This conclusion will address individual samples or will express a general assessment based on examination of a representative sample set over a pre-determined time period, depending on the project objectives and the sample set being evaluated. This assessment will include, where applicable, the evaluation of data trends relative to laboratory corrective actions or laboratory events. The resulting usability determination would be characterized over certain segments of the audited period. Recommendations concerning usability of the data will be communicated with appropriate Department staff.

APPENDIX A

- Formulas for Calculating Precision and Accuracy

Use the following formulas for calculating the precision and accuracy of test measurements and the associated acceptance ranges:

  - Precision

Calculate the precision of replicate samples using one of the following three formulas:

    - Percent Relative Standard Deviation:

%RSD=S÷X×100

Where:		X = Mean (average) of the data points

s = Standard deviation calculated as:

s = {[S(X-Xi)2]¸(n-1)}0.5

X  = mean value of measured concentrations (µg/L)

Xi = value of each measured concentration (µg/L)

n  = number of determinations

    - Relative Percent Difference:

RPD=A-BA+B2× 100

Where:		A = concentration in sample A

B = concentration in sample B

  - Accuracy (as % Recovery)

Determine the accuracy (as % recovery) by calculating the % recovery of a known amount of analyte from the fortified (spiked) blank or sample as follows:

% R= Spiked Concentration-Unspiked ConcentraionCalculated Fortification Spiked Concentration×100

__________________________________________________

- APPENDIX B

- Method Detection Limits and Practical Quantitation Limits

  - Method Detection Limit (MDL)

An MDL is analyte-and matrix-specific and is laboratory-dependent. For an MDL study, all sample processing steps of the analytical method must be included for each analyte and matrix proposed for the method scope and applicability. The MDL may be determined following the procedures specified below unless otherwise required by a published method for which the laboratory is seeking approval as a modified or alternative method, or the MDL may otherwise be determined by any technically justifiable and scientifically sound procedure. A specific MDL procedure must be used if required by the Department to meet DQOs for a specific program activity or data use for which the method is proposed (e.g., if required in a Department rule).

NOTE: The MDL is not equivalent to the Instrument Detection Limit (IDL) which cannot be used in place of the MDL. The IDL is determined using multiple analyses of standards and is useful in determining an experimental concentration level to use when fortifying samples for the MDL determination. The MDL is determined by processing samples through the entire sample preparation and analytical procedure (not just analysis).

The laboratory must determine the MDL using the protocol specified in the published laboratory method, if applicable. If the protocol for determining detection limits is not specified, one of the following three protocols may be used:

- EPA - "Definition and Procedure for the Determination of the Method Detection Limit - Revision 2" (2016), 40 CFR Part 136, Appendix B, 7/1/24, incorporated by reference in 62-160.800(5)(a), F.A.C.;

- IUPAC- “Nomenclature in Evaluation of Analytical Methods including Detection and Quantification Capabilities”, Pure & Appl. Chem., Vol. 67, No. 10, pp. 1699-1723, 1995, incorporated by reference in 62-160.800(5)(b), F.A.C.;

- Hubaux and Vos- “Decision and Detection Limits for Linear Calibration Curves”, Analytical Chemistry, Vol. 42, No. 8, July 1970, pp. 849-855, incorporated by reference in 62-160.800(5)(c), F.A.C.

The method endorsed by the International Union of Pure and Applied Chemistry (IUPAC) is derived from the method published by Lloyd Currie and assumes a constant error model within a small concentration region. This method sets the MDL at a critical value intended to exclude 99% of the analytical noise population from reportable levels. The Appendix B to Part 136, Title 40 method for determining the MDL is designed for use with a variety of physical and chemical methods. It incorporates all sample processing steps used by the laboratory for an analytical method, including all steps of the method for samples analyzed over a prescribed period of time and distributed across all instruments to which the MDL will be applied. The laboratory calculates the MDL by determining the standard deviation of a set of spiked samples (MDLs) and method blanks (MDLb) and assigns either an initial MDL or an ongoing verification of the MDL, as applicable. The method published by Hubaux and Vos, is based on a variable error model and the effect of concentration on the resulting noise distribution is considered in determining the detection limit. While this technique is more robust than that of other models, considerably more effort is required to develop method detection limits.

NOTE: Any methods that support compliance monitoring and reporting for EPA's National Pollutant Discharge Elimination System (NPDES) program, or, for compliance with monitoring requirements in rules authorized under the Safe Drinking Water Act must use the 40 CFR method for determining the MDL. See applicable Department rules for this requirement. If a Department rule refers to a later revision of the 40 CFR Part 136 MDL method, follow the requirements in the later version.

  - Practical Quantitation Limit (PQL)

The Practical Quantitation Limit (PQL) is the lowest level that can be reliably achieved during routine laboratory operating conditions within specified limits of precision and accuracy. Typically, the PQL is 3-5 times the MDL, and it represents a practical and routinely achievable detection level with a relatively good certainty than any reported value is reliable. The PQL for the proposed method must be determined by a defined procedure and/or criteria. For example, the PQL may be chosen to be the lowest calibration standard used for a calibration curve. Alternatively, the PQL may be indicated as the concentration at which the method has been demonstrated to achieve a specified range of precision and accuracy (e.g., %R= 70% - 130%; RPD ≤20%). See section 3.3 above for applicable requirements for determining the PQL. The procedure and/or criteria used to define the PQL must be included in the laboratory standard operating procedure developed for the method, or, in the laboratory quality manual.

  - Analytes Without Statistically Derived MDLs

The concept of a statistically derived MDL as described above is not appropriate for some analytes. DEP still requires a reporting limit in those cases, and requires non-detect results to be reported as the reporting limit with the “U” qualifier; however, the reported MDL may be equal to the reported PQL. Examples of these analytes include: Biochemical Oxygen Demand (BOD); Carbonaceous Biochemical Oxygen Demand (CBOD); Total Chlorine; Chlorophyll; Fecal and Total Coliforms; Color; Dissolved Oxygen (DO); E. Coli; Enterococci; Hardness (by calculation); Hydrogen ion (pH); Hydrogen sulfide; Nitrate (by calculation); Total Nitrogen (by calculation); Odor; Radiochemistry analytes; Residues (TDS, TSS); Specific conductance; Temperature; Titrimetric methods; Turbidity; and Un-ionized ammonia.