## Field Measurement of Temperature

##### Use this SOP in conjunction with the following DEP SOPs:

###### FT 1000 General Field Testing and Measurement

###### FQ 1000 Field Quality Control Requirements

- FS 1000 General Sampling Procedures

###### FD 1000 Documentation Procedures

##### Equipment and Supplies

##### Field Instruments:  Use any of the following instrument types for performing field measurements:

##### Digital thermistor (thermocouple type) and meter typical of field instruments

##### Glass bulb, mercury-filled thermometer (not recommended for field ruggedness)

##### Glass bulb, alcohol-filled thermometer with protective case

##### Bi-metal strip/dial-type thermometer

##### Advanced silicon chip temperature sensor and digital meter

##### Field instruments must be capable of measuring temperature in 0.1oC increments.

##### Standard Thermometer:  NIST-traceable Celsius certified thermometer with scale marks for every 0.1oC increment, a range of 0oC to 100oC (or a range bracketing expected sample temperatures) and correction chart supplied with certification.  The standard thermometer must have a valid certification for the period of measurement.

##### Recordkeeping and Documentation Supplies:

###### Field notebook or forms

###### Indelible pens

##### Calibration and Use

##### General Concerns

##### Select a temperature measuring device meeting the requirements of section 1.1 above.

##### Dial-type and thermocouple-type devices with meters are preferred over the glass thermometers for fieldwork because of their durability and ease of reading.

##### Transport glass thermometers in protective cases.

##### Inspect glass thermometers for liquid separation.  Do not use a thermometer if the liquid has separated.

##### Record the temperature reading to the level of resolution stated by the sensor manufacturer. Most instruments with digital display will provide more decimal figures than are significant.

##### Calibration

##### Follow the calibration activities specified in FT 1000, section 2.2.

##### Verify all thermistor (meter) devices and field thermometers against the NIST-traceable standard thermometer at several temperatures in the expected sample measurement range, using any correction factor indicated by the certificate supplied with the NIST-traceable thermometer.

##### 2.2.2.1.	See the US Geological Survey, Temperature: U.S. Geological Survey Techniques and Methods, Book 9, chapter A6.1, 14p., 2024, incorporated by reference in paragraph 62-160.800(2)(g), F.A.C., National Field Manual for the Collection of Water-Quality Data, Book 9, Chapter A6, Field Measurements, Section 6.1, Temperature, Techniques of Water-Resources Investigations, Version 2, 3/2006 for additional guidance about making temperature comparisons with the standard thermometer.

##### 2.2.2.2.	Make note of the calibration in the calibration records.  See section 4 below.

##### 2.2.2.3.	The field measurement device may be used with a linear correction factor provided that the observed temperature difference with the standard thermometer is documented at incremental temperatures over the range of expected sample temperatures.

##### 2.2.2.4.	Use the resulting correction factor when making temperature measurements of samples with the field measurement device.

##### 2.2.2.5.	Prominently display the correction factor on the field measurement device, with the date last verified.  A calibration correction curve or plot may also be used.

##### To be acceptable, a calibration verification must be within +/- 0.5oC of the corrected reading of the NIST-traceable thermometer.

##### Properly dispose of glass-bulb thermometers that do not meet the above calibration acceptance criteria.

##### Continuing Calibration Verifications:

##### Determine the maximum time between continuing calibration verifications for the specific field temperature measurement device based on instrument stability.

##### Verify the field measurement device, in the “read” or “run” mode, against the standard NIST-traceable thermometer as in section 2.2.2 above.

##### Refer to additional calibration requirements in FT 1000, section 2.2.  Sample readings must be chronologically and quantitatively bracketed by verifications.

##### More frequent calibration verifications may be required for discharge permit compliance measurements or other regulatory requirements.

##### Measuring Sample Temperature

##### Insert or place the thermometer or sensor in situ at a measuring location representative of the sampling source.

##### Do not measure temperature from an intermediate container containing static sample

##### Measure groundwater samples in situ with a downhole probe or in a flow-through container.  Do not measure bailed or pumped samples in an intermediate container containing static sample.

##### Allow the thermometer or temperature sensor to equilibrate to ambient in situ temperature.

##### Record the temperature to the level of resolution stated by the sensor manufacturer after the reading stabilizes and remains constant.

##### Preventive Maintenance:  Refer to FT 1000, section 3.

##### Documentation

##### Standards Documentation:  Document information about the NIST-traceable standard thermometer in the calibration record, including:

##### Unique identification for the thermometer

##### Vendor certificate of calibration, including any correction factor

##### Vendor’s expiration date for the certificate of calibration

##### Field Instrument Calibration Documentation:  Document acceptable calibration and calibration verification for each instrument unit and field test or analysis, linking this record with affected sample measurements.

##### Retain vendor certifications of all factory-calibrated instrumentation.

##### Designate the identity of specific instrumentation in the documentation with a unique description or code for each instrument unit used. Record manufacturer name, model number, and identifying number such as a serial number for each instrument unit.

##### Record the time and date of all initial calibrations and all calibration verifications.

##### Record the instrument reading (value in appropriate measurement units) of all calibration verifications to the level of resolution stated by the sensor manufacturer.

##### Record the name of the analyst(s) performing the calibration.

##### Document the following information about initial calibration and calibration verifications and link to information recorded according to section 4.1 above:

##### Details of the method used to compare the field measurement device to the NIST-traceable standard thermometer.

##### Results of each calibration verification, including the expected reading (per the NIST-traceable standard thermometer)

##### The actual reading of the field measurement device, using any established correction factors and correct units.

##### Retain manufacturers’ instrument specifications.

##### Document whether successful initial calibration occurred.

##### Document whether each calibration verification passed or failed.

##### Document any corrective actions taken to correct instrument performance (such as a new correction factor) according to records requirements of FD 3000.

##### Document date and time of any corrective action.

##### Note any incidence of discontinuation of use of the instrument due to calibration failure.

##### Describe or cite the specific calibration or verification procedure performed (DEP SOP or internal SOP).

##### Record all field-testing measurement data, to include the following:

###### Project name

###### Date and time of measurement or test (including time zone, if applicable)

###### Source and location of the measurement or test sample (e.g., monitoring well identification number, outfall number, station number or other description)

  - Latitude and longitude of sampling source location (if required)

###### Analyte or parameter measured

###### Measurement or test sample value to the level of resolution stated by the sensor manufacturer

###### Reporting units

- “J” qualifier code and explanatory comment if the sample measurement is not chronologically and quantitatively bracketed by acceptable verifications

###### Initials or name of analyst performing the measurement

###### Unique identification of the specific instrument unit(s) used for the test(s)